## Readme
General Notes:
These worksheets in this workbook were derived from the old "Examp9010.xls" workbook. They have
been updated for use with 9095 XLPRO and provide automatic kVp measurment capture and
analysis. In some cases where the kVp data might be erroneous (e.g. a hvl data point with added filtration)
the kVp values are hidden by coloring the font white.
Font colors have been used to help identify different types of data. In general, all user entry data is in a medium-blue color.
All text that should not need to be changed in is black. Dose data is dark-blue, time data is medium-green, and kVp data is dark-red
| EXAMPLE |  | Measured
|  | SET | Dose | Time | HV
|  | mA | mR | ms | kVp
|  | 25 | 6.68 | 44.3 | 58.5
To protect formatting and foumulas, all cells that should not need to be changed are locked and protected.
Those cells that contain formulas have also been hidden. You may of course unlock them and edit as you see fit.
Before unlocking or making changes, be sure you have retained a copy of the original workbook. To remove the protection,
Select "Unprotect Sheet" from the Tools\Protection menu.
The worksheets have been formatted to print a report and exclude comments and extraneous data. To view how a
sheet will look when printed, select "File" and then "Print Preview".
For additional information see the Radcal "XLPROManual.pdf" file either on the install CD or
in your 9000PXL directory.
Note on E13: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.

## Worksheet Information
Sample Worksheets in "9096 Sample Template.xls"
The three sample worksheets in the workbook 9096 Sample Template.xls delivered with
XLPRO have been designed to demonstrate the ease of use of XLPRO with Excel and
a formatted worksheet. These worksheets capture dose and rate & kVp from a 9096 and
implement typical QA measurements and data-collection requirements for a radiology department.
The degree of macro automation used in these worksheets is minimal, which makes it very easy
for users to modify the sheets to fit their individual data-collection needs.
Most of the usefulness of these worksheets comes from the formulas and graphs that are in the
worksheets themselves. The user places the cursor at the cell into which a measurement is to be
inserted and then initiates a measurement through a control-key sequence. This control-key
sequence initiates communication with the 9096 through the XLPRO control interface.
Once XLPRO inserts the measurement into the cell, control is returned to the user either automatically
or by pressing ESC. The user then ensures that the cursor is in the correct cell for the next measurement
and initiates a new measurement. This sequence continues until all the required measurements are completed.
For each section of these worksheets, there is a setup area to input machine parameters, a data collection
area for either manual or automatic data collection, and a comment area to outline the
results of the test and the action to be taken. All three of these worksheets come with the
protection property of Excel turned on so that only the data-entry areas of the worksheet are
unprotected and modifiable by the user. You may turn the protection property off so that you can
modify other areas easily. No password is required.
You should keep in mind when modifying these worksheets that there are items such as the
Settings Area and intermediate calculation areas which are located outside of the main worksheet.
These areas must not be broken by inserting rows into the middle of them.
Macros
The workbook contains the macros listed below. The Auto_Open subroutine initializes
combination CONTROL-SHIFT keys to run the subroutines that they are associated with. These
subroutines first determine the cell at the top-left corner of the Settings Area and use these coordinates to
record the initial settings for Scale and Cell Update. Then they modify the settings area cells as needed to
execute the command, and execute the appropriate 9000XLMacro.xls macro (e.g.MacroRunMaxRate),
which does not exit if no argument is passed to the macro. The subroutines end after restoring the
Settings Area to its initial configuration and advancing the selection by one row. The Auto_Close
subroutine allows you to delete the Auto_Activate commands when the workbook closes.
FOR EXAMPLE
Sub Auto_Open()
Application.OnKey "+^o", "RateOverwrite"
Application.OnKey "+^f", "FluorokVpOverwrite"
End Sub
Sub RateOverwrite()
On Error GoTo Quit 'Trap no Settings Area or its write protection.
'Find the coordinates of the top-left cell of the Settings Area.
r1 = ActiveSheet.Names("_RCTL").RefersToRange.Row
c1 = ActiveSheet.Names("_RCTL").RefersToRange.Column
'Store the current values from the Settings Area.
incr = Cells(r1 + 10, c1 + 2)
units = Cells(r1 + 6, c1 + 2)
'Modify the Settings Area.
Cells(r1 + 10, c1 + 2) = "Overwrite"
Cells(r1 + 6, c1 + 2) = ""
'Run the MaxRate macro.
Application.Run "9000XLMacro.xls!MacroRunMaxRate"
'Restore the Settings Area to its old values and advance the selection by one row.
Cells(r1 + 10, c1 + 2) = incr
Cells(r1 + 6, c1 + 2) = units
ActiveCell.Offset(1, 0).Select
Quit:
If (Err) Then Call MsgBox("No recorded Settings Area or it is protected.", _
vbCritical, "Local Error")
On Error GoTo 0
End Sub
Sub Auto_Close()
If (Right(ActiveWorkbook.Name, 3) = "xlt") Then Exit Sub
i = 0
On Error Resume Next
For j = 1 To ActiveWorkbook.Sheets.Count 'Look for Auto_Activates.
ws = ActiveWorkbook.Sheets(j).Name
If (ws <> "Settings") Then i = Sheets(j).Names("Auto_Activate").Index
Next j
On Error GoTo 0
If (i = 0) Then Exit Sub 'Avoid query if no Auto_Activates are present.
If (MsgBox("Do you want to keep auto activation?", vbYesNo) <> vbNo) Then Exit Sub
On Error Resume Next
For j = 1 To ActiveWorkbook.Sheets.Count 'Delete Auto_Activates.
ws = ActiveWorkbook.Sheets(j).Name
If (ws <> "Settings") Then Sheets(j).Names("Auto_Activate").Delete
Next j
On Error GoTo 0
End Sub
Sheet named Radiographic
The worksheet Radiographic is set up for a strictly radiographic room or station. The red comment indicators,
in the yellow cells indicate comments that describe the steps required for executing the Measure Pulse
| macro (CONTROL-SHIFT P). |  |  | Example | Note: these comments will not be visible when you print.
A. Facility Information
Facility information contains data areas for equipment location, x-ray equipment type and
equipment operational ratings. There are no formulas associated with this section of the
worksheet.
B. kVp Accuracy and Exposure Linearity
kVp Accuracy and Exposure Linearity provides automatic input of the Dose Time & kVp
data from the radiation monitor. Formulas for this section are for kV error
calculation & Pass/Fail criteria ($24:G37), linear-regression calculation of the best-fit
straight line to the measured dose data, and the error of the measured dose data
relative to the best-fit dose data (H24:I37). Linear-regression calculations are performed just to
the right of the graph in an area of the worksheet (T21:AA37) which is not in the print area.
Formulas in this area must be modified by the user if the full data area is not used.
C. mA and Exposure Linearity
mA and Exposure Linearity allows for the automatic input of exposure data from the radiation
monitor. The Set mA data area is unprotected, and users may change these numbers to reflect the
settings on their x-ray machines. Formulas for this section are for the linear-regression calculation
of the best-fit straight line to the measured dose data (F57:F63), the error of the measured dose
data relative to the best-fit dose data (G57:G63), and average dose/mAs (B65). Again, the linefitting
calculations are performed just to the right of the graph in an area of the worksheet
(X56:AB64) that is not in the print area. Formulas in this area must be modified by the user if the
full data area is not used.
D. Timer and Exposure Linearity
Timer and Exposure Linearity allows for the automatic input of exposure data from the radiation
monitor. The Set Time data area is unprotected, and users may change these numbers
to reflect the settings on their x-ray machines. Formulas for this section are for the calculation
of timer error (F79:F94), the linear regression calculation of the best-fit straight
line to the measured dose data (G79:G94), and the error of the measured dose data relative to
the best-fit dose data (H79:H94). Again, the linefitting calculations are performed just to the right
of the graph in an area of the worksheet (X79:AB87) that is not in the print area.
Formulas in this area must be modified by the user if the full data area is not used.
E. Exposure Reproducibility
Exposure Reproducibility is a summary area where the user may enter information about the x-ray
techniques used, a summary of the results, and comments about the action required.
F. Half Value Layer
Formulas for this section are for the calculation of mR/s (E134:F139), the normalization of the mR/s values
(G134:G139), and the calculation of the HVL (D141) done in cells T132:X139. The HVL is calculated
using the mR/s values. The formulas for mR/s calculation use the time in the cell
just to the left of the mR/s cell. Again, the two-point log interpolation for
determining the calculated HVL is done just to the right of this section in a non-printing area of
the worksheet. If the user wishes to do fewer measurements in determining HVL, then the
interpolation formula must be changed also.
G. Phototimer
Phototimer allows capture of data for three different cell areas of the phototimer area.
The last four sections (H., I., J. and K.) are text-entry areas for users to insert their observations
and comments about various non-radiation related aspects of the x-ray machine performance.
Sheet named Special
The "Special" worksheet is set up for a fluoroscopic or radiographic room. The red indicators
in the yellow cells indicate comments that describe the steps required for executing the MeasureFluorokVp
| macro (CONTROL-SHIFT F). |  |  | Example | Note: these comments will not be visible when you print.
A. Facility Information
Facility information contains data areas for equipment location, x-ray equipment type and
equipment operational ratings. There are no formulas associated with this section of the
worksheet.
B. Fluoroscopy
The fluoroscopy section has data-entry areas for machine-setup parameters, for dose-rate or
exposure information, and result and action comments. There are no formulas associated with this
section of the worksheet.
1) Maximum Table Top Exposure Rate
2) Automatic Dose Control
3) Automatic Brightness Control
4) Resolution
5) Timer
6) kV Accuracy
7) Half Value Layer
8) Spot Films
9) Exposure Reproducibility
10) Mechanical
C. Radiographic Overhead X-Ray Tube
See Radiographic Sheet discussion above
Sheet named Special
The "Mobile" worksheet is set up for a mobile x-ray machine. The red comment indicators,
in the yellow cells indicate comments that describe the steps required for executing the Measure kVp
| macro (CONTROL-SHIFT K). |  |  | Example | Note: these comments will not be visible when you print.
This sheet follows the format and measurement principles found in the Radiographic sheet
discussed above.
Sheet named Settings
This sheet is found in all 9000XL workbooks and contains the basic settings information needed to start
making measurments. For further information regarding the Settings sheet see the Radcal "XLPROManual.pdf"
file either on the install CD or in your 9000PXL directory.
Note on F89: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on F147: 1. Place cursor in this cell and press "Control-Shift f" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make a fluoro exposure. 3. Wait for kV and Rate values to settle to stable numbers. 4. Press Escape to end.
Note on F175: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.

## Radiographic
2026 ANNUAL CALIBRATION REPORT | [25 empty] | 9000XL ™ Version 4.10
RADIOLOGY DEPARTMENT | [25 empty] | Operation Settings Area
[7 empty] | DATE: | 2026-07-29 14:40
Units
A. | FACILITY |  |  | Almost anywhere USA | [6 empty] | Room No: | B243 |  | Other: | [11 empty] | Radiation: | R
|  | State Reg. No: |  | CA423-619 |  |  |  | UNIT: | Radiographic | [16 empty] | /Time: | /min
[26 empty] | Scale: | milli
| EQUIPMENT |  |  | Manufacturer | [5 empty] | Model |  |  |  | Serial No. | [11 empty] | Pulse Rate: | 0.25-9-6-5
|  |  | CONSOLE: | General Electric |  |  |  |  | DXR-750 |  |  |  | A417176 | [12 empty] | Mode | D
|  |  | TUBE: | Machlett |  |  |  |  | Dynamax 69 B |  |  |  | 99828288 | [12 empty] | Units Display: | Omit
|  |  | RATINGS: |  |  | Maximum tube kV: | 150 | kVp. |  |  | Maximum tube mA: | 700 | mA. | Phase: | 3 | [10 empty] | Cell Update: | Incremental
[6 empty] | Inherent filtration: | 1.2 | mm Al |  |  | Total Filtration | 3.1 | mm Al | [12 empty] | Auto Zero: | Disabled
[26 empty] | Capture Mode: | Auto
[26 empty] | Screen Tracking: | Yes
B. | kVp ACCURACY AND EXPOSURE LINEARITY | [24 empty] | kVp | 9000
|  | Focal Spot: | lg |  | Field Size: | Max | [19 empty] | Config: | DC+T+kV
|  | 60 | mA | 100 | ms | 6 | mAs | [18 empty] | Scale: | 9096
| Radcal Accu-kV diagnostic sensor placed on table . | [24 empty] | Start delay (ms): | 0
|  | @ | 20.5 | inches from target, | [21 empty] | Width (ms): | 4914.6
| Radcal Chamber at: |  |  | 15.25 | inches from target | [20 empty] | Filtration (mm-Al): | 2.5
|  | Measured |  |  | kV Accuracy |  | Linearity | [10 empty] | Exposure time: | 0.1 | sec | [5 empty] | Dental threshold: | 0
| Set | Dose | Time | HV | Error | Pass/ | mR | % | [8 empty] | Ln Values |  | Reg. |  |  | Regression Output: |  |  |  | Device Index | 0
| kVp | mR | ms | kVp | kV | Fail | regrs. | Error | [8 empty] | kVp | Dose | Dose |  |  |  | b | a |  | Monitor Configuration
| 50 | 198 | 100 | 49.1 | -0.9 |  | 192.6169 | -2.795% | [8 empty] | 3.8939 | 5.2883 | 192.6169 |  |  | Coeff Value | 2.079 | -2.8348 |  | Serial No: | 00000
| 55 | 225 | 100 | 52 | -3 |  | 217.0242 | -3.675% | [8 empty] | 3.9512 | 5.4161 | 217.0242 |  |  | Std Error | 0.029373 | 0.12877 |  | ROM Ver: | 6.02
| 60 | 260 | 100 | 58 | -2 |  | 272.3367 | 4.53% | [8 empty] | 4.0604 | 5.5607 | 272.3367 |  |  | R squared | 0.99761 | 0.032073 |  | Detector: | 5.5 cc
| 65 | 291 | 100 | 61 | -4 | FAIL | 302.4414 | 3.783% | [8 empty] | 4.1109 | 5.6733 | 302.4414 | [6 empty] | Chamb Kq: | D
| 70 | 381 | 101 | 69 | -1 |  | 390.7602 | 2.498% | [8 empty] | 4.2341 | 5.9428 | 390.7602 |  | y = a + (b * x) |  |  |  |  | Atm Corr: | Automatic
| 75 | 421 | 101 | 72 | -3 |  | 426.9119 | 1.385% | [8 empty] | 4.2767 | 6.0426 | 426.9119 |  | a= | -2.8348 |  |  |  | Temp (°C): | 22
| 80 | 525 | 101 | 79 | -1 |  | 517.7412 | -1.402% | [8 empty] | 4.3694 | 6.2634 | 517.7412 |  | b= | 2.079 |  |  |  | Press (kPa): | 101.3
| 85 | 560 | 100 | 83 | -2 |  | 573.7337 | 2.394% | [8 empty] | 4.4188 | 6.3279 | 573.7337
| 90 | 677 | 101 | 88 | -2 |  | 647.9295 | -4.487% | [8 empty] | 4.4773 | 6.5177 | 647.9295
| 95 | 731 | 102 | 91 | -4 | FAIL | 694.698 | -5.226% | [8 empty] | 4.5109 | 6.5944 | 694.698
| 100 | 832 | 100 | 99 | -1 |  | 827.7064 | -0.5187% | [8 empty] | 4.5951 | 6.7238 | 827.7064
| 110 | 1030 | 101 | 110 | 0 |  | 1030.406 | 0.0394% | [8 empty] | 4.7005 | 6.9373 | 1030.406
| 120 | 1210 | 102 | 120 | 0 |  | 1234.7318 | 2.003% | [8 empty] | 4.7875 | 7.0984 | 1234.7318
| 130 | 1400 | 103 | 128 | -2 |  | 1412.0359 | 0.8524% | [8 empty] | 4.852 | 7.2442 | 1412.0359
|  |  |  | Max.: | -4 | kV | Max: | -5.226%
|  |  |  | Limit ± | 3 | kV
Comments:
|  |  |  | kV Accuracy: | Service recommended for the 65 & 95 kV stations.
GE's specifications for kV accuracy need to be checked.
|  |  |  | Exposure Linearity: | Exposures linear to within 6% at all kV stations tested
Acceptable results
C. | mA AND EXPOSURE LINEARITY
|  | Technique: | 60 | kV | 50 | ms
| 21 | in. target to chamber distance
Measured
| SET | Dose | Time | HV | Regression
| mA | mR | ms | kVp | mR | Error | [16 empty] | Regression Output:
| 25 | 6.68 | 44.3 | 58.5 | 8.4324 | -20.78% | [14 empty] | 6.7459 |  |  | b | a
| 50 | 13.2 | 45 | 59 | 17.3399 | -23.88% | [14 empty] | 6.936 |  | Coeff Value | 0.3563 | -0.47518
| 100 | 33.7 | 47.5 | 59.1 | 35.155 | -4.139% | [14 empty] | 7.031 |  | Std Error | 0.015556 | 4.7879
| 200 | 70.9 | 49.2 | 59.4 | 70.7853 | 0.1621% | [14 empty] | 7.0785 |  | R squared | 0.99056 | 7.9673
| 300 | 111 | 48.9 | 59.4 | 106.4155 | 4.308% | [14 empty] | 7.0944
| 400 | 155 | 48.6 | 59.5 | 142.0457 | 9.12% | [14 empty] | 7.1023 | y = a + (b * x)
| 600 | 203 | 48.3 | 59.4 | 213.3062 | -4.832% | [14 empty] | 7.1102 | a= | -0.47518
| Regression Avg.: |  |  |  | Max. Error: | -23.88% | [15 empty] | b= | 0.3563
| 7.014 | mR/mAs @ 21 inches
Comments:
|  |  | Results: | Exposures should be linear to within +/-10%
At 25 and 50 mA the exposures are too low
|  |  | Action: | SERVICE REQUIRED
D. | TIMER AND EXPOSURE LINEARITY
|  | Technique: | 75 | kV | 300 | mA
| 21 | in. target to chamber distance
Measured
| SET | Dose | Time | HV | Timer | Regression
| Time | mR | ms | kVp | Error | mR | Error
| 17 | 78 | 18 | 74.5 | 5.882% | 79.2031 | -1.519% | [14 empty] | Regression Output:
| 25 | 89.8 | 24 | 74 | -4% | 106.4557 | -15.65% | [15 empty] | b | a
| 33 | 118 | 35 | 73.7 | 6.061% | 133.7083 | -11.75% | [14 empty] | Coeff Value | 3.4066 | 21.2914
| 50 | 178 | 52 | 73.6 | 4% | 191.62 | -7.108% | [14 empty] | Std Error | 0.019324 | 7.8785
| 67 | 233 | 65 | 74.2 | -2.985% | 249.5317 | -6.625% | [14 empty] | R squared | 0.99955 | 22.2629
| 100 | 352 | 102 | 74.17 | 2% | 361.9485 | -2.749%
| 125 | 448 | 128 | 74.14 | 2.4% | 447.1128 | 0.1984% | [13 empty] | y = a + (b * x)
| 150 | 535 | 155 | 74.11 | 3.333% | 532.277 | 0.5116% | [13 empty] | a= | 21.2914
| 200 | 704 | 205 | 74.08 | 2.5% | 702.6056 | 0.1985% | [13 empty] | b= | 3.4066
| 250 | 913 | 255 | 74.05 | 2% | 872.9341 | 4.59%
| 300 | 1060 | 306 | 74.02 | 2% | 1043.2627 | 1.604%
| 400 | 1400 | 408 | 73.99 | 2% | 1383.9197 | 1.162%
| 500 | 1760 | 508 | 73.96 | 1.6% | 1724.5768 | 2.054%
| 600 | 2080 | 610 | 73.93 | 1.667% | 2065.2339 | 0.715%
| 800 | 2700 | 815 | 73.9 | 1.875% | 2746.548 | -1.695%
| 1000 | 3420 | 1008 | 73.87 | 0.8% | 3427.8622 | -0.2294%
Comments:
|  |  | Results: | Exposures should be linear to within +/-10%
At .0167 and .025 seconds the exposures are too low
|  |  | Action: | SERVICE REQUIRED
E. | EXPOSURE REPRODUCIBILITY
|  |  |  | Dose | Time | kVp | mR/mAs
|  |  | PASS/FAIL Criteria: | 5% | 5% | 5% | 5%
|  |  | mA: | 200 | ms: | 100 |  | mA: | 200 | ms: | 100 |  | mA: | 200 | ms: | 100
|  |  | kVp: | 60 | mAs: | 20 |  | kVp: | 60 | mAs: | 20 |  | kVp: | 60 | mAs: | 20
|  |  | Measured |  |  |  |  | Measured |  |  |  |  | Measured
|  |  | Dose | Time | HV | mR/ |  | Dose | Time | HV | mR/ |  | Dose | Time | HV | mR/
|  |  | mR | ms | kVp | mAs |  | mR | ms | kVp | mAs |  | mR | ms | kVp | mAs
|  | X1 | 100 | 100 | 59.4 | 5 | X1 | 100 | 100 | 59.4 | 5 | X1 | 100 | 100 | 59.4 | 5
|  | X2 | 99 | 98 | 60 | 5.051 | X2 | 99 | 98 | 60 | 5.051 | X2 | 99 | 98 | 60 | 5.051
|  | X3 | 98 | 98 | 60.1 | 5 | X3 | 98 | 98 | 60.1 | 5 | X3 | 98 | 98 | 60.1 | 5
|  | X4 | 99 | 98.5 | 59.8 | 5.0254 | X4 | 99 | 98.5 | 59.8 | 5.0254 | X4 | 99 | 98.5 | 59.8 | 5.0254
|  | X5 | 94 | 99.6 | 60.3 | 4.7189 | X5 | 94 | 99.6 | 60.3 | 4.7189 | X5 | 94 | 99.6 | 60.3 | 4.7189
|  | Avg | 98 | 98.82 | 59.92 | 4.9591 | Avg | 98 | 98.82 | 59.92 | 4.9591 | Avg | 98 | 98.82 | 59.92 | 4.9591
|  | CV | 2.393% | 0.9395% | 0.5708% | 2.741% | CV | 2.393% | 0.9395% | 0.5708% | 2.741% | CV | 2.393% | 0.9395% | 0.5708% | 2.741%
|  | Pass/Fail | Pass | Pass | Pass | Pass |  | Pass | Pass | Pass | Pass |  | Pass | Pass | Pass | Pass
|  |  | RESULTS: | Reproducibility was within the required +/- 5%.
|  |  | COMMENT: | Acceptable results
F. | HALF VALUE LAYER
Technique
| kV | mA | ms | added filtration |  |  | 74.5 kV (75 kV setting) 200 mA, 2.0 mmAl Added
| 75 | 200 | 100 | 2 | mmAl |  | Calculated HVL @75.0 kV setting is 2.5 mmAl
| 15.25 | in. target to chamber distance
Measured
| mm | Dose | Time | HV | Calculated | [12 empty] | HVL (2 point Log interpolation)
| Al | mR | ms | kVp | mR/s | Atten | [13 empty] | (a)Just > | (b)Just < |  | 0.5 line data
| 0 | 218 | 100.1 | 74.5 | 2177.8222 | 1 | [10 empty] | 1 |  | Target Value | 0.5 |  |  | 0 | 0.5
| 1 | 158 | 99.8 | 74.8 | 1583.1663 | 0.72695 | [10 empty] | 0.72695 |  | Row | 3 | 4 |  | 5 | 0.5
| 2 | 122 | 99.9 | 75 | 1221.2212 | 0.56075 | [10 empty] | 0.56075 |  | Exposure Rate | 1221.2212 | 970
| 3 | 97 | 100 | 74.1 | 970 | 0.4454 | [10 empty] | 0.4454 |  | Thickness (mm) | 2 | 3
| 4 | 78.1 | 99.9 | 84.9 | 781.7818 | 0.35897 | [10 empty] | 0.35897
| 5 | 65.1 | 99.9 | 75 | 651.6517 | 0.29922 | [10 empty] | 0.29922 |  | Calc. HVL | 2.4979 | mmAl
|  | Calculated HVL: | 2.4979 | mmAl
Comments:
|  |  | Results: | Acceptable
|  |  | Action: | None Required
G. | PHOTOTIMER
Technique
| kV | mA | ms | cassette |  |  | 72 | inch SFD |  | 67.5 | inch SCD, Lead Apron
| 121 | 200 | 500 | 14 x 17 | in. |  |  | Results: | Chest Stand + Lead Apron
|  | Center Cell |  |  |  |  | Right Cell |  |  |  |  | Left Cell
|  | Measured |  |  |  |  | Measured |  |  |  |  | Measured
|  | Dose | Time | HV | Dose |  | Dose | Time | HV | Dose |  | Dose | Time | HV | Dose
|  | mR | ms | kVp | Ratio |  | mR | ms | kVp | Ratio |  | mR | ms | kVp | Ratio
| NORM | 44 | 77 | 121.1 | N/A |  | 31 | 52 | 121 | N/A |  | 48 | 85 | 120.8 | N/A
| 1.25 | 52 | 94 | 120.8 | 1.1818 |  | 39 | 68 | 120.8667 | 1.2581 |  | 56 | 101 | 121 | 1.1667
| 1.5 | 65 | 118 | 121 | 1.4773 |  | 47 | 77 | 120.8167 | 1.5161 |  | 72 | 134 | 120.8667 | 1.5
| 2 | 73 | 134 | 120.8667 | 1.6591 |  | 56 | 101 | 120.7667 | 1.8065 |  | 85 | 159 | 121 | 1.7708
| 0.75 | 30 | 51 | 120.8167 | 0.68182 |  | 22 | 35 | 120.7167 | 0.70968 |  | 39 | 68 | 120.8667 | 0.8125
| 0.5 | 22 | 35 | 120.7667 | 0.5 |  | 17 | 27 | 120.8667 | 0.54839 |  | 26 | 43 | 120.8167 | 0.54167
| 0.25 | 13 | 18 | 120.7167 | 0.29545 |  | 9 | 10 | 120.8167 | 0.29032 |  | 13 | 18 | 120.7667 | 0.27083
Comments:
|  |  | Results: | On both normal and 3/4 density settings, Exposures are not balanced
for the three cells. There should be better agreement in % variation
among the three cells.
|  |  | Action: | SERVICE RECOMMENDED
H. | LIGHT FIELD/X-RAY FIELD ALIGNMENT
|  | TABLE: | At 37 inch SFD - Alignment was not within 2% of SFD. The x-ray field is 1/2 inch smaller
and 1/4 inch larger than the light field (cross table), and shifted ~ 1/4 inch longitudinally.
|  | CHEST STAND: | At 72 inch SFD, the alignment is poor: 3/4 inch of the x-ray field is cut off (or shifted)
on the left side (facing the chest unit) and ~3/4 inch too large on the right side.
|  | Comments: | SERVICE REQUIRED
I. | CENTERING OF X-RAY FIELD TO CASSETTE IN BUCKY TRAY
|  | TABLE: | At 40 inch SFD, Center of x-ray field and cassette are within 2% of SFD.
|  | Comments: | Acceptable Results
|  | WALL: | Center of x-ray field and cassette were not within 2% of SFD
|  | Comments: | SERVICE REQUIRED
J. | AUTOMATIC COLLIMATION
SERVICE REQUIRED, this system was not operational.
K. | MECHANICAL
| 1 | Ready light on control panel did not light
| 2 | Chest stand cassette size indicator light did not light.
| 3 | When using low technique (~ 50 - 55 kV, 5 mAs) the RESET indicator came on.
| 4 | Several times, during testing no exposure occurred when the exposure switch was engaged.
L. | IN AIR EXPOSURE FROM CLINICAL PROTOCOLS
Measured
|  |  |  | Generator Settings |  |  | Dose | Time | HV
|  | STUDY |  | KV | mA | ms | mR | ms | kVp
|  | ABDOMEN |  | 75 | 200 | 300 | 704 | 305 | 74.8
|  |  |  | 105 | 200 | 50 | 242 | 48 | 105.1
|  |  |  | 68 | 300 | 100 | 278 | 97 | 68.7
CHEST
|  |  | a) Erect | 75 | 300 | 25 | 89 | 22 | 74.3
|  |  |  | 85 | 400 | 100 | 667 | 98 | 84.2
|  |  |  | 88 | 400 | 100 | 712 | 7 | 87.2
|  |  | b) Supine | 63 | 300 | 25 | 64 | 22 | 62.8
|  |  |  | 65 | 300 | 17 | 42 | 13 | 64.9
|  |  |  | 76 | 300 | 33 | 124 | 31 | 75.8
|  |  | LAT | 95 | 600 | 25 | 3050 | 256 | 93.7
|  |  |  | 110 | 300 | 33 | 278 | 32 | 109.6
c) PHOTOTIMED - on 3/4 Density
|  |  |  | 80 | 300 | 100 | 433 | 97 | 79.7
|  |  |  | 100 | 200 | 500 | 2240 | 506 | 100.1
HIP & PELVIS
|  |  |  | 69 | 100 | 200 | 183 | 203 | 69.2
|  |  |  | 70 | 300 | 200 | 635 | 197 | 69.9
|  |  |  | 74 | 200 | 250 | 608 | 255 | 74.2
|  |  |  | 90 | 300 | 250 | 1400 | 256 | 89.9
|  | L-SPINE |  | 75 | 200 | 250 | 585 | 255 | 75
|  |  |  | 80 | 200 | 250 | 719 | 255 | 80.1
|  |  |  | 85 | 200 | 800 | 2330 | 797 | 85
|  |  |  | 90 | 200 | 800 | 2810 | 797 | 89.8
OBSTRUCTION SERIES
|  |  |  | 78 | 300 | 150 | 604 | 147 | 77.3
|  |  |  | 80 | 300 | 400 | 1680 | 397 | 79.2
|  |  |  | 85 | 300 | 400 | 1770 | 397 | 84.4
CALIBRATION PERFORMED BY:
CERTIFIED RADIOLOGICAL PHYSICIST
Note on C24: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on C57: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on C79: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on C134: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on C157: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on D111: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on H157: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on M157: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.

## Special
2026 ANNUAL CALIBRATION REPORT | [22 empty] | 9000XL ™ Version 4.10
RADIOLOGY DEPARTMENT | [22 empty] | Operation Settings Area
[6 empty] | DATE: | 2026-07-29 14:40
Units
A. | FACILITY |  |  | Almost anywhere USA | [6 empty] | Room No: | 453B |  | Other: | [8 empty] | Radiation: |  | R
|  | State Reg. No: |  | No. 705544 (Fluro) |  |  |  | UNIT: | Special Procedures | [13 empty] | /Time: |  | /min
|  |  |  | No 703190 (Radiographic) | [18 empty] | Scale:
| EQUIPMENT |  |  | Manufacturer | [5 empty] | Model |  |  |  | Serial No. | [8 empty] | Pulse Rate: |  | 15-9-6-5
|  |  | CONSOLE: | General Electric |  |  |  |  | DXR-750 |  |  |  | A417176 | [9 empty] | Mode |  | D
|  |  | TUBE: | Machlett |  |  |  |  | Dynamax 69 B |  |  |  | 99828288 | [9 empty] | Units Display: |  | Omit
|  |  | RATINGS: |  |  | Maximum tube kV: | 165 | kVp. |  |  | Maximum tube mA: | 100 | mA. | Phase: | 1 | [7 empty] | Cell Update: |  | Overwrite
[6 empty] | Inherent filtration: |  | mm Al |  |  | Total Filtration |  | mm Al | [9 empty] | Auto Zero: |  | Disabled
[23 empty] | Capture Mode: |  | Auto
B. | FLUOROSCOPY | [21 empty] | Screen Tracking: |  | Yes
| 1) | Maximum Table Top Exposure Rate | [20 empty] | kVp |  | 9000F
|  | R/min | kV: | [19 empty] | Config: |  | DC+kV
|  | 8.6 | 120.1 | [19 empty] | Scale: |  | 9096
|  | mA: | None seen | [19 empty] | Start delay (ms): |  | 0
[23 empty] | Width (ms): |  | 4914.6
|  | COMMENT: |  | Acceptable. This complies with federal and state regulations | [18 empty] | Filtration (mm-Al): |  | 2.5
[23 empty] | Dental threshold: |  | 0
[23 empty] | Device Index |  | 0
| 2) | Automatic Dose Control | [20 empty] | Monitor Configuration
|  | Technique: | [20 empty] | Serial No: |  | 00000
|  | kV | mA |  | Auto mode will override manual mode if phantom's thickness and density | [17 empty] | ROM Ver: |  | 6.02
|  | 70 | MAX |  | drive system beyond the selected kV (via dial on image intensifier) | [17 empty] | Detector: |  | 5.5 cc
[23 empty] | Chamb Kq: |  | D
[5 empty] | Table Top Exposures | [17 empty] | Atm Corr: |  | Automatic
|  |  | Phantom |  | kV |  | R/min | kVp | [14 empty] | Temp (°C): |  | 22
|  |  | Copper |  | 75 |  | 1.65 | 74.3 | [14 empty] | Press (kPa): |  | 101.3
[5 empty] | 80 |  | 0.73 | 80.1
[5 empty] | 90 |  | 0.46 | 91
|  |  | 0.75" Al + Cu |  | 72 |  | 7.11 | 70.4
[5 empty] | 80 |  | 2.97 | 80.4
[5 empty] | 120 |  | 0.63 | 121.3
|  |  | 1.5" Al + Cu |  | 86.5 |  | 6.955 | 80.8
[5 empty] | 121 |  | 1.6 | 120.2
|  |  | Lead + Cu + Al |  | 121 |  | 8 | 119.7
|  | COMMENT: |  | This unit is performing as per GE's Design. It will operate on the kV selected
(ie 70 kV) unless the patient/phantom's thickness or density drives the system
to maximum mA first. Then it will increase the kV. The above data illustrates
how patient dose is lowered at higher kV's. Conversely, as kV increases
contrast decreases.
| 3) | Automatic Brightness Control
|  | Technique: |  | Different thicknesses of different materials were placed on the table in the beam
The ion chamber was placed as close as possible to the outside entrance to image
intensifier assembly
[5 empty] | Phantom Material |  |  | R/min | kV
[5 empty] | 0.75 " Aluminum |  |  | 0.0065 | 70
[5 empty] | 1.5 " Aluminum |  |  | 0.006 | 70
[5 empty] | Copper |  |  | 0.005 | 70
[5 empty] | Cu + 0.75"Al |  |  | 0.005 | 70
[5 empty] | 1.5" Al + Cu |  |  | 0.006 | 85
|  | COMMENT: |  | Acceptable Results
| 4) | Resolution
|  |  | Technique: | Line pair phantom was taped to the outside entrance of the image intensifier.
|  |  | Results: | 1.1 line pairs per millimeter
|  | COMMENT: |  | Acceptable Results. This is not as good as in Room 1
| 5) | Timer
|  |  | Results: | Has audible indication at the completion of nominal 5 minutes of Fluoro time.
|  |  | COMMENT: | Acceptable Results.
| 6) | kV Accuracy
|  | Technique: |  | Radcal AccukV sensor was placed on the table at 90 degree angle with the
anode/cathode of the x-ray tube.
[6 empty] | Radcal |  | Apparent
|  |  | SET | kV (TV) |  | R/min | kVp | kV Error
|  |  | 75 | 75 |  |  | 76.1 | 1.1
|  |  | 80 | 80 |  |  | 80 | 0
|  |  | 85 | 85,85 |  |  | 82.5 | -2.5
|  |  | 90 | 89 |  |  | 93 | 3
|  |  | 100 | 99,100 |  |  | 95 | -5
|  |  | 110 |  |  |  | 103 | -7
|  |  | 120 | 121 |  |  | 115 | -5
|  |  | COMMENT: | Service Required for the high kV's
NOTE: The kV meter used to measure the kV drove the automatic system to a
minimum kV of 75.
| 7) | Half Value Layer
|  |  | Result: | 2.9 mm of Al at 75 kV
|  |  | COMMENT: | Acceptable Results
| 8) | Spot Films
|  | Technique: |  | Routine clinical settings of 110 kV, 1.5 seconds back-up for phototimer, phantom.
a) Spacing
No everlap noted on the 4/1 and 2/1 films.
b) Film Density
Variation in density corresponded with exposure variations and phantoms.
See patient exposures below.
c) Patient Exposures
|  |  |  | Technique: | Automatic Mode, Copper, Copper plus Aluminum phantoms.
70 kV on Fluoro, 95 kV, 400 mA, 0.5 sec backup.
|  |  |  | Results: | [5 empty] | Dose | Time
|  |  |  | Phantom |  |  | Spacing |  | Grid | mR | ms
[6 empty] | Copper + 1.5" Al |  | 1/1 | No | 1330 | 77
[9 empty] | Yes | 1670 | 93
[6 empty] | Copper + 3/4" Al |  | 2/1 | No | 917 | 56
[9 empty] | Yes | 874 | 56
[6 empty] | Copper + 1.5" Al |  | 2/1 | No | 2010 | 124
[9 empty] | Yes | 1920 | 119
[6 empty] | Copper + 1.5" Al |  | 4/1 | No | 2670 | 1610
[9 empty] | Yes | 2310 | 1450
[6 empty] | Copper +3/4" Al |  | 4/1 | No | 954 | 64
[6 empty] | Copper Alone |  | 4/1 | Yes | 309 | 26
| 9) | Exposure Reproducibility
|  |  | Result: | Exposures were not within the required 5% for the Fluoro mode using Copper
+ 1.5 in Aluminum phantom. The fluoro kV was 70 kV selected via the kV dial on
the image intensifier assembly. Table top exposures varried as follows:
7.51, 6.7, 6.92, 6.7, 6.38, 6.95, 6.91 R/min
|  |  | COMMENT: | Service Needed
| 10) | Mechanical
|  |  | NOTE: | This system appears to have only the 9 inch fluoro field of view
in spite of 6 and 4.5 inch FOV indicators on the equipment.
C. | RADIOGRAPHIC OVERHEAD X-RAY TUBE
| 1). | kVp Accuracy and Exposure Linearity
|  | Focal Spot: | lg |  | Field Size: | Max
Radcal Accu-kV diagnostic sensor placed on table .
| @ | 20.5 | inches from target, |  |  | 6 | mAs
| Radcal Chamber at: |  |  | 15.25 | inches from target
|  | Measured |  |  | kVp Accuracy |  | Linearity | [9 empty] | Exposure time: | 0.999 | sec
| Set | Dose | Time | HV | Error | Pass/ | mR | % | [8 empty] | Ln Values |  | Reg. |  |  | Regression Output:
| kVp | mR | ms | kVp | kV | Fail | regrs. | Error | [8 empty] | kVp | Dose | Dose |  |  |  | b | a
| 50 | 185 | 99.9 | 49 | -1 |  | 181.9213 | -1.692% | [8 empty] | 3.8918 | 5.2204 | 181.9213 |  |  | Coeff Value | 2.1598 | -3.2019
| 55 | 195 | 99.5 | 52 | -3 |  | 206.8338 | 5.721% | [8 empty] | 3.9512 | 5.273 | 206.8338 |  |  | Std Error | 0.029362 | 0.12873
| 60 | 260 | 99.7 | 58 | -2 |  | 261.8475 | 0.7056% | [8 empty] | 4.0604 | 5.5607 | 261.8475 |  |  | R squared | 0.99779 | 0.032187
| 65 | 291 | 100 | 61 | -4 | FAIL | 291.979 | 0.3353% | [8 empty] | 4.1109 | 5.6733 | 291.979
| 70 | 381 | 101 | 69 | -1 |  | 381.0146 | 0.00384% | [8 empty] | 4.2341 | 5.9428 | 381.0146 |  | y = a + (b * x)
| 75 | 421 | 101 | 72 | -3 |  | 417.6974 | -0.7907% | [8 empty] | 4.2767 | 6.0426 | 417.6974 |  | a= | -3.2019
| 80 | 525 | 101 | 79 | -1 |  | 510.375 | -2.866% | [8 empty] | 4.3694 | 6.2634 | 510.375 |  | b= | 2.1598
| 85 | 560 | 100 | 83 | -2 |  | 567.8307 | 1.379% | [8 empty] | 4.4188 | 6.3279 | 567.8307
| 90 | 677 | 101 | 88 | -2 |  | 644.2987 | -5.075% | [8 empty] | 4.4773 | 6.5177 | 644.2987
| 95 | 731 | 102 | 91 | -4 | FAIL | 692.6772 | -5.533% | [8 empty] | 4.5109 | 6.5944 | 692.6772
| 100 | 832 | 100 | 99 | -1 |  | 830.9322 | -0.1285% | [8 empty] | 4.5951 | 6.7238 | 830.9322
| 110 | 1030 | 101 | 110 | 0 |  | 1043.2582 | 1.271% | [8 empty] | 4.7005 | 6.9373 | 1043.2582
| 120 | 1210 | 102 | 120 | 0 |  | 1258.9454 | 3.888% | [8 empty] | 4.7875 | 7.0984 | 1258.9454
| 130 | 1440 | 103 | 129 | -1 |  | 1471.7781 | 2.159% | [8 empty] | 4.8598 | 7.2724 | 1471.7781
|  |  |  | Max.: | -4 | kV | Max: | 5.721%
Comments:
|  |  |  | kV Accuracy: | Service recommended for the 65 & 95 kV stations.
GE's specifications for kV accuracy need to be checked.
|  |  |  | Exposure Linearity: | Exposures linear to within 6% at all kV stations tested
Acceptable results
| 2). | mA and Exposure Linearity
|  | Technique: | 60 | kV | 50 | ms
| 21 | in. target to chamber distance
Measured
| SET | Dose | Time | HV | Regression
| mA | mR | ms | kVp | mR | Error | [15 empty] | Regression Output:
| 25 | 6.68 | 55 | 58.5 | 8.4324 | -20.78% | [13 empty] | 6.7459 |  |  | b | a
| 50 | 13.2 | 52 | 59 | 17.3399 | -23.88% | [13 empty] | 6.936 |  | Coeff Value | 0.3563 | -0.47518
| 100 | 33.7 | 54 | 59.1 | 35.155 | -4.139% | [13 empty] | 7.031 |  | Std Error | 0.015556 | 4.7879
| 200 | 70.9 | 53 | 59.4 | 70.7853 | 0.1621% | [13 empty] | 7.0785 |  | R squared | 0.99056 | 7.9673
| 300 | 111 | 52 | 59.4 | 106.4155 | 4.308% | [13 empty] | 7.0944
| 400 | 155 | 53 | 59.5 | 142.0457 | 9.12% | [13 empty] | 7.1023 | y = a + (b * x)
| 600 | 203 | 55 | 59.4 | 213.3062 | -4.832% | [13 empty] | 7.1102 | a= | -0.47518
[5 empty] | Max. Error: | -23.88% | [14 empty] | b= | 0.3563
Regression Avg.:
| 7.014 | mR/mAs @ 21 inches
Comments:
|  |  | Results: | Exposures should be linear to within +/-10%
At 25 and 50 mA the exposures are too low
|  |  | Action: | SERVICE REQUIRED
| 3). | Timer and Exposure Linearity
|  | Technique: | 75 | kV | 300 | mA
| 21 | in. target to chamber distance
Measured
| SET | Dose | Time | HV | Timer | Regression
| Time | mR | ms | kVp | Error | mR | Error
| 17 | 59.5 | 18 | 74.5 | 5.882% | 77.019 | -22.75% | [12 empty] | Regression Output:
| 25 | 89.8 | 22 | 74 | -12% | 104.3019 | -13.9% | [13 empty] | b | a
| 33 | 118 | 35 | 73.7 | 6.061% | 131.5847 | -10.32% | [12 empty] | Coeff Value | 3.4104 | 19.043
| 50 | 178 | 52 | 73.6 | 4% | 189.5608 | -6.099% | [12 empty] | Std Error | 0.0198 | 8.0727
| 67 | 233 | 65 | 74.2 | -2.985% | 247.5368 | -5.873% | [12 empty] | R squared | 0.99953 | 22.8117
| 100 | 352 | 102 | 74.17 | 2% | 360.0785 | -2.244%
| 125 | 448 | 128 | 74.14 | 2.4% | 445.3374 | 0.5979% | [11 empty] | y = a + (b * x)
| 150 | 535 | 155 | 74.11 | 3.333% | 530.5963 | 0.8299% | [11 empty] | a= | 19.043
| 200 | 704 | 205 | 74.08 | 2.5% | 701.1141 | 0.4116% | [11 empty] | b= | 3.4104
| 250 | 913 | 255 | 74.05 | 2% | 871.6319 | 4.746%
| 300 | 1060 | 306 | 74.02 | 2% | 1042.1497 | 1.713%
| 400 | 1400 | 408 | 73.99 | 2% | 1383.1853 | 1.216%
| 500 | 1760 | 508 | 73.96 | 1.6% | 1724.2208 | 2.075%
| 600 | 2080 | 610 | 73.93 | 1.667% | 2065.2564 | 0.7139%
| 800 | 2700 | 815 | 73.9 | 1.875% | 2747.3276 | -1.723%
| 1000 | 3420 | 1008 | 73.87 | 0.8% | 3429.3987 | -0.2741%
Comments:
|  |  | Results: | Exposures should be linear to within +/-10%
At .0167 and .025 seconds the exposures are too low
|  |  | Action: | SERVICE REQUIRED
| 4). | Exposure Reproducibility
|  |  | TECHNIQUES USED = | 55, 65, 75, 85, 95 kV, 200 Ma, 0.1s
75 kV, 300 Ma, 4/5 s
|  |  | RESULTS= | Reproducibility was within the required +/- 5%.
Exposures varied by 8%
|  |  | COMMENT: | Acceptable results
| 5). | Half Value Layer
Technique
| kV | mA | ms | added filtration |  |  | 74.5 kV (75 kV setting) 200 mA, 2.0 mmAl Added
| 75 | 200 | 100 | 2 | mmAl |  | Calculated HVL @75.0 kV setting is 2.5 mmAl
| 15.25 | in. target to chamber distance
Measured
| mm | Dose | Time | HV | Calculated | [11 empty] | HVL (2 point Log interpolation)
| Al | mR | ms | kVp | mR/s | Atten | [12 empty] | (a)Just > | (b)Just < |  | 0.5 line data
| 0 | 218 | 100.1 | 74.5 | 2177.8222 | 1 | [9 empty] | 1 |  | Target Value | 0.5 |  |  | 0 | 0.5
| 1 | 158 | 99.8 | 74.8 | 1583.1663 | 0.72695 | [9 empty] | 0.72695 |  | Row | 3 | 4 |  | 5 | 0.5
| 2 | 122 | 99.9 | 75 | 1221.2212 | 0.56075 | [9 empty] | 0.56075 |  | Exposure Rate | 1221.2212 | 970
| 3 | 97 | 100 | 74.1 | 970 | 0.4454 | [9 empty] | 0.4454 |  | Thickness (mm) | 2 | 3
| 4 | 78.1 | 99.9 | 84.9 | 781.7818 | 0.35897 | [9 empty] | 0.35897
| 5 | 65.1 | 99.9 | 75 | 651.6517 | 0.29922 | [9 empty] | 0.29922 |  | Calc. HVL | 2.4979 | mmAl
|  | Calculated HVL: | 2.4979 | mmAl
Comments:
|  |  | Results: | Acceptable
|  |  | Action: | None Required
| 6). | Phototimer
Technique
| kV | mA | ms | cassette |  |  | 72 | inch SFD |  | 67.5 | inch SCD, Lead Apron
| 121 | 200 | 500 | 14 x 17 | in. |  |  | Results: | Chest Stand + Lead Apron
|  | Center Cell |  |  |  |  | Right Cell |  |  |  |  | Left Cell
|  | Measured |  |  |  |  | Measured |  |  |  |  | Measured
|  | Dose | Time | HV | Dose |  | Dose | Time | HV | Dose |  | Dose | Time | HV | Dose
|  | mR | ms | kVp | Ratio |  | mR | ms | kVp | Ratio |  | mR | ms | kVp | Ratio
| NORM | 44 | 77 | 121.1 | N/A |  | 31 | 52 | 121 | N/A |  | 48 | 85 | 120.8 | N/A
| 1.25 | 52 | 94 | 120.8 | 1.1818 |  | 39 | 68 | 120.8667 | 1.2581 |  | 56 | 101 | 121 | 1.1667
| 1.5 | 65 | 118 | 121 | 1.4773 |  | 47 | 77 | 120.8167 | 1.5161 |  | 72 | 134 | 120.8667 | 1.5
| 2 | 73 | 134 | 120.8667 | 1.6591 |  | 56 | 101 | 120.7667 | 1.8065 |  | 85 | 159 | 121 | 1.7708
| 0.75 | 30 | 51 | 120.8167 | 0.68182 |  | 22 | 35 | 120.7167 | 0.70968 |  | 39 | 68 | 120.8667 | 0.8125
| 0.5 | 22 | 35 | 120.7667 | 0.5 |  | 17 | 27 | 120.8667 | 0.54839 |  | 26 | 43 | 120.8167 | 0.54167
| 0.25 | 13 | 18 | 120.7167 | 0.29545 |  | 9 | 10 | 120.8167 | 0.29032 |  | 13 | 18 | 120.7667 | 0.27083
|  |  | Results: | On both normal and 3/4 density settings, Exposures are not balanced
for the three cells. There should be better agreement in % variation
among the three cells.
|  |  | Action: | SERVICE RECOMMENDED
| 7). | Light Field/X-ray Field Alignment
|  | TABLE: | At 37 inch SFD - Alignment was not within 2% of SFD. The x-ray field is 1/2 inch smaller
and 1/4 inch larger than the light field (cross table), and shifted ~ 1/4 inch longitudinally.
|  | CHEST STAND: | At 72 inch SFD, the alignment is poor: 3/4 inch of the x-ray field is cut off (or shifted)
on the left side (facing the chest unit) and ~ 3/4 inch too large on the right side.
|  | Comments: | SERVICE REQUIRED
| 8). | Centering of X-ray field to Cassette in Bucky Tray
|  | TABLE: | At 40 inch SFD, Center of x-ray field and cassette are within 2% of SFD.
|  | Comments: | Acceptable Results
|  | WALL: | Center of x-ray field and cassette were not within 2% of SFD
|  | Comments: | SERVICE REQUIRED
| 9). | Automatic Collimation
SERVICE REQUIRED, this system was not operational.
| 10). | Mechanical
| 1 | Ready light on control panel did not light
| 2 | Chest stand cassette size indicator light did not light.
| 3 | When using low technique (~ 50 - 55 kV, 5 mAs) the RESET indicator came on.
| 4 | Several times, during testing no exposure occurred when the exposure switch was engaged.
| 11). | In air exposure from clinical protocols
Measured
|  |  |  | Generator Settings |  |  | Dose | Time | HV
|  | STUDY |  | KV | mA | ms | mR | ms | kVp
|  | ABDOMEN |  | 75 | 200 | 300 | 704 | 305 | 74.8
|  |  |  | 105 | 200 | 50 | 242 | 48 | 105.1
|  |  |  | 68 | 300 | 100 | 278 | 97 | 68.7
CHEST
|  |  | a) Erect | 75 | 300 | 25 | 89 | 22 | 74.3
|  |  |  | 85 | 400 | 100 | 667 | 98 | 84.2
|  |  |  | 88 | 400 | 100 | 712 | 7 | 87.2
|  |  | b) Supine | 63 | 300 | 25 | 64 | 22 | 62.8
|  |  |  | 65 | 300 | 17 | 42 | 13 | 64.9
|  |  |  | 76 | 300 | 33 | 124 | 31 | 75.8
|  |  | LAT | 95 | 600 | 25 | 3050 | 256 | 93.7
|  |  |  | 110 | 300 | 33 | 278 | 32 | 109.6
c) PHOTOTIMED - on 3/4 Density
|  |  |  | 80 | 300 | 100 | 433 | 97 | 79.7
|  |  |  | 100 | 200 | 500 | 2240 | 506 | 100.1
HIP & PELVIS
|  |  |  | 69 | 100 | 200 | 183 | 203 | 69.2
|  |  |  | 70 | 300 | 200 | 635 | 197 | 69.9
|  |  |  | 74 | 200 | 250 | 608 | 255 | 74.2
|  |  |  | 90 | 300 | 250 | 1400 | 256 | 89.9
|  | L-SPINE |  | 75 | 200 | 250 | 585 | 255 | 75
|  |  |  | 80 | 200 | 250 | 719 | 255 | 80.1
|  |  |  | 85 | 200 | 800 | 2330 | 797 | 85
|  |  |  | 90 | 200 | 800 | 2810 | 797 | 89.8
OBSTRUCTION SERIES
|  |  |  | 78 | 300 | 150 | 604 | 147 | 77.3
|  |  |  | 80 | 300 | 400 | 1680 | 397 | 79.2
|  |  |  | 85 | 300 | 400 | 1770 | 397 | 84.4
CALIBRATION PERFORMED BY:
CERTIFIED RADIOLOGICAL PHYSICIST
Note on C17: 1. Place cursor in this cell and press "Control-Shift f" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make a fluoro exposure. 3. Wait for kV and Rate values to settle to stable numbers. 4. Press Escape to end.
Note on C159: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on C187: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on C209: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on C249: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on C271: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on G87: To Make a Fluoro Rate & kVp Measurement 1. Place cursor in this cell and press "Control-Shift-f" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make a fluoro exposure. 3. Wait for kV and Rate values to settle to stable numbers. 4. Press Escape to end. 5. Move cursor to next cell
Note on H30: To Make a Fluoro Rate & kVp Measurement 1. Place cursor in this cell and press "Control-Shift-f" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make a fluoro exposure. 3. Wait for kV and Rate values to settle to stable numbers. 4. Press Escape to end. 5. Move cursor to next cell
Note on H271: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on H325: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on I55: To Make a Fluoro Rate & kVp Measurement 1. Place cursor in this cell and press "Control-Shift-f" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make a fluoro exposure. 3. Wait for kV and Rate values to settle to stable numbers. 4. Press Escape to end. 5. Move cursor to next cell
Note on K121: To make a dose & time measurement (BE SURE THE kVp SENSOR IS IN THE BEAM): 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Press Escape to end.
Note on M271: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.

## Mobile
2026 ANNUAL CALIBRATION REPORT | [24 empty] | XLPRO ™ Version 4.10
RADIOLOGY DEPARTMENT | [24 empty] | Operation Settings Area
[7 empty] | DATE: | 2026-07-29 14:40
Units
A. | FACILITY |  |  | Almost anywhere USA | [6 empty] | Room No: | 223 |  | Other: | [10 empty] | Radiation: | R
|  | State Reg. No: |  | CA 324-916 |  |  |  | UNIT: | Radiographic Mobile | [15 empty] | /Time: | /min
[25 empty] | Scale: | milli
| EQUIPMENT |  |  | Manufacturer | [5 empty] | Model |  |  |  | Serial No. | [10 empty] | Pulse Rate: | 15-9-6-5
|  |  | CONSOLE: | General Electric |  |  |  |  | AMX-4 |  |  |  | 371294 | [11 empty] | Mode | D
|  |  | TUBE: | General Electric |  |  |  |  | GE MX75 |  |  |  | MX9184 | [11 empty] | Units Display: | Omit
|  |  | RATINGS: |  |  | Maximum tube kV: | 75 | kVp. Maximum tube mA: |  |  |  | 40 | mA. | Phase: | 1 | [9 empty] | Cell Update: | Incremental
[25 empty] | Auto Zero: | Disabled
[25 empty] | Capture Mode: | Auto
B. | kVp ACCURACY AND EXPOSURE LINEARITY | [23 empty] | Screen Tracking: | Yes
|  | Focal Spot: | lg |  | Field Size: | Max | [18 empty] | kVp | 9000
| Radcal Accu-kV diagnostic sensor placed on table . | [23 empty] | Config: | DC+T+kV
| @ | 20.5 | inches from target, |  |  | 6 | mAs | [17 empty] | Scale: | 9096
| Radcal Chamber at: |  |  | 15 | inches from target | [19 empty] | Start delay (ms): | 0
| Radcal Chamber at: |  |  | 15.25 | inches from target | [19 empty] | Width (ms): | 5000
|  | Measured |  |  | kV Accuracy |  | Linearity | [9 empty] | Exposure time: | 0.0999 | sec | [5 empty] | Filtration (mm-Al): | 2.5
| Set | Dose | Time | HV | Error | Pass/ | mR | % | [7 empty] | Ln Values |  | Reg. |  |  | Regression Output: |  |  |  | Dental threshold: | 0
| kVp | mR | ms | kVp | kV | Fail | regrs. | Error | [7 empty] | kVp | Dose | Dose |  |  |  | b | a |  | Device Index | 0
| 50 | 81 | 99.9 | 50.5 | 0.5 |  | 83.8508 | 3.4% | [7 empty] | 3.922 | 4.3944 | 83.8508 |  |  | Coeff Value | 2.1689 | -4.0772 |  | Monitor Configuration
| 55 | 100 | 99.5 | 55 | 0 |  | 100.9043 | 0.8962% | [7 empty] | 4.0073 | 4.6052 | 100.9043 |  |  | Std Error | 0.037879 | 0.16778 |  | Serial No: | 00008
| 60 | 125 | 99.7 | 61 | 1 |  | 126.3098 | 1.037% | [7 empty] | 4.1109 | 4.8283 | 126.3098 |  |  | R squared | 0.99635 | 0.039822 |  | ROM Ver: | 4.01
| 65 | 157 | 100 | 68 | 3 |  | 159.868 | 1.794% | [7 empty] | 4.2195 | 5.0562 | 159.868 | [6 empty] | Detector: | 6 cc
| 70 | 181 | 101 | 73 | 3 |  | 186.4629 | 2.93% | [7 empty] | 4.2905 | 5.1985 | 186.4629 |  | y = a + (b * x) |  |  |  |  | Chamb Kq: | 1.02
| 75 | 218 | 101 | 77 | 2 |  | 209.3342 | -4.14% | [7 empty] | 4.3438 | 5.3845 | 209.3342 |  | a= | -4.0772 |  |  |  | Atm Corr: | Automatic
| 80 | 248 | 101 | 83 | 3 |  | 246.3301 | -0.6779% | [7 empty] | 4.4188 | 5.5134 | 246.3301 |  | b= | 2.1689 |  |  |  | Temp (°C): | 24.1
| 85 | 284 | 100 | 86.5 | 1.5 |  | 269.4155 | -5.413% | [7 empty] | 4.4601 | 5.649 | 269.4155 | [6 empty] | Press (kPa): | 100
| 90 | 325 | 101 | 91 | 1 |  | 300.7407 | -8.067% | [7 empty] | 4.5109 | 5.7838 | 300.7407
| 95 | 330 | 102 | 95 | 0 |  | 330.1499 | 0.04541% | [7 empty] | 4.5539 | 5.7991 | 330.1499
| 100 | 404 | 100 | 103 | 3 |  | 393.4301 | -2.687% | [7 empty] | 4.6347 | 6.0014 | 393.4301
| 110 | 481 | 101 | 113 | 3 |  | 480.9999 | -2.893e-05% | [7 empty] | 4.7274 | 6.1759 | 480.9999
| 120 | 572 | 102 | 125 | 5 | FAIL | 598.7 | 4.46% | [7 empty] | 4.8283 | 6.3491 | 598.7
| 124 | 606 | 103 | 129 | 5 | FAIL | 641.0303 | 5.465% | [7 empty] | 4.8598 | 6.4069 | 641.0303
|  |  |  | Max.: | 5 | kV | Max: | -8.067%
|  |  |  | Limit ± | 3 | kV
Comments:
|  |  |  | kV Accuracy: | Results acceptable below 120kV.
SERVICE REQUIRED FOR kV STATIONS ABOVE 110kV
|  |  |  | Exposure Linearity: | Exposures linear to within 6% at all kV stations tested
Very Good Results, which are better than the other Capmobile
C. | mAs AND EXPOSURE LINEARITY
|  | Technique: | 70 | kV | 50 | ms
| 20 | in. target to chamber distance
Measured
| SET | Data | Time | HV | Regression |  | mR/
| mAs | mR | ms | kVp | mR | Error | mAs | [13 empty] | Regression Output:
| 1 | 33.4 | 55 | 70.3 | 52.6483 | -36.56% | 52.6483 | [14 empty] | b | a
| 2 | 34.6 | 52 | 70.2 | 63.9551 | -45.9% | 31.9776 | [13 empty] | Coeff Value | 11.3069 | 41.3414
| 3 | 51.2 | 54 | 70.1 | 75.262 | -31.97% | 25.0873 | [13 empty] | Std Error | 0.7334 | 10.9405
| 4 | 65.7 | 53 | 70.2 | 86.5688 | -24.11% | 21.6422 | [13 empty] | R squared | 0.94814 | 27.1034
| 5 | 82 | 52 | 70.4 | 97.8757 | -16.22% | 19.5751
| 6 | 112 | 53 | 79.2 | 109.1825 | 2.581% | 18.1971 | [12 empty] | y = a + (b * x)
| 7 | 128 | 55 | 70 | 120.4894 | 6.233% | 17.2128 | [12 empty] | a= | 41.3414
| 8 | 157 | 55 | 70.3 | 131.7962 | 19.12% | 16.4745 | [12 empty] | b= | 11.3069
| 10 | 170 | 55 | 70 | 154.4099 | 10.1% | 15.441
| 12 | 202 | 55 | 70.1 | 177.0237 | 14.11% | 14.752
| 14 | 228 | 55 | 70 | 199.6374 | 14.21% | 14.2598
| 16 | 260 | 55 | 69.9 | 222.2511 | 16.98% | 13.8907
| 22 | 311 | 55 | 70 | 290.0922 | 7.207% | 13.186
| 27 | 346 | 55 | 69.9 | 346.6265 | -0.1807% | 12.838
| 35 | 384 | 55 | 70 | 437.0813 | -12.14% | 12.488
[5 empty] | Max. Error: | -45.9%
Regression Avg.:
| 19.978 | mR/mAs @ 20 inches
Comments:
|  |  | Results: | These results are not as good as a the other Capmobile.
Even though this is a capacitor discharge unit.
Exposures should be linear to within +/- 10% from 1 to 10 or 14 mAs
|  |  | Action: | SERVICE REQUIRED
D. | EXPOSURE REPRODUCIBILITY
|  |  | TECHNIQUES USED= | 55, 70,85,100 kV, 6 mAs
75 kV, 8 mAs
70 kV, 3, 5, 10 mAs
|  |  | RESULTS= | Reproducibility was within the required +/- 5%.
Exposures varied by 8%
|  |  | Comment: | Acceptable results
E. | HALF VALUE LAYER
Technique
| kV | mAs | ms | added filtration |  |  | 75.7 kV (75 kV setting) 8 mA, 2.0 mmAl Added
| 75 | 8 | 100 | 2 | mmAl |  | Calculated HVL @75.0 kV setting is 2.7 mmAl
| 15.25 | in. target to chamber distance
HVL (2 point Log interpolation)
|  | Measured | [17 empty] | (a)Just > | (b)Just < |  | 0.5 line data
| mm | Dose | Time | HV | Calculated | [11 empty] | 1 |  | Target Value | 0.5 |  |  | 0 | 0.5
| Al | mR | ms | kVp | mR/s | Atten | [10 empty] | 0.74854 |  | Row | 3 | 4 |  | 5 | 0.5
| 0 | 171 | 100 | 75.7 | 171 | 1 | [10 empty] | 0.5848 |  | Exposure Rate | 100 | 81
| 1 | 128 | 100 | 75.9 | 128 | 0.74854 | [10 empty] | 0.47368 |  | Thickness (mm) | 2 | 3
| 2 | 100 | 100 | 76.3 | 100 | 0.5848 | [10 empty] | 0.39708
| 3 | 81 | 100 | 76.5 | 81 | 0.47368 | [10 empty] | 0.3345 |  | Calc. HVL | 2.7434 | mmAl
| 4 | 67.9 | 100 | 76.8 | 67.9 | 0.39708
| 5 | 57.2 | 100 | 76.9 | 57.2 | 0.3345
|  | Calculated HVL: | 2.7434 | mmAl
Comments:
|  |  | Results: | Acceptable
|  |  | Action: | None Required
F. | LIGHT FIELD/X-RAY FIELD ALIGNMENT
|  | Findings: | At 27 inch SFD - Alignment was not within 2% of SFD. The x-ray field is 1/2 inch smaller
than the light field, as well as shifted/
|  | Comments: | While this is in compliance, it is at the limit. Service recommended to prevent
it from being out of compliance.
G. | MECHANICAL
| 1 | Locks and light fields were working porperly.
| 2 | Audible beam on indicator needs to be louder.
H. | IN AIR EXPOSURE FROM CLINICAL PROTOCOLS
|  | Measurement Distance from Source: |  |  |  |  | 15 | inches
Measured
|  |  |  | Generator Settings |  |  | Dose | Time | HV
|  | STUDY |  | KV | mAs |  | mR | ms | kVp
|  | ABDOMEN |  | 80 | 22 |  | 756 | 274 | 78.6
|  | & C-Spine |  | 80 | 22 |  | 756 | 274 | 79
CHEST
|  |  |  | 70 | 5 |  | 142 | 14 | 69.8
|  |  |  | 75 | 8 |  | 172 | 980 | 74 | at 20 inches
|  | COLANG. |  | 90 | 35 |  | 1520 | 500 | 89.5
|  | HIP |  | 108 | 27 |  | 1720 | 461 | 107.5
|  |  |  | 85 | 27 |  | 999 | 360 | 84
|  | L-SPINE |  | 120 | 35 |  | 2590 | 618 | 119
|  |  |  | 120 | 50 |  | 3230 | 900 | 118.5
|  | PELVIS |  | 75 | 22 |  | 629 | 242 | 74.5
CALIBRATION PERFORMED BY:
CERTIFIED RADIOLOGICAL PHYSICIST
Note on C23: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on C55: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on C102: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on H135: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.

## Settings
XLPRO ™ Version 4.10 |  |  |  | DEFAULT
Operation Settings Area |  |  | 3500 | 3500 | Max. Freq. (Hz)
|  |  | 20 | 20 | Initial Delay (ms)
Units |  |  | 7 | 7 | End Delay (ms)
Radiation: |  | R | 100 | 100 | Width (ms)
/Time: |  | /min | 7 | 7 | Pulses in window
Scale: |  |  | 3 | 3 | Initial Interval (ms)
Pulse Rate: |  | 15-9-6-5 | 20 | 20 | Noise threshold (kV)
Mode |  | D | 0.9 | 0.9 | Upper threshold
Units Display: |  | Rescale | 0.7 | 0.7 | Lower threshold
Cell Update: |  | Incremental | 3 | 3 | FT-resolution control
Auto Zero: |  | Disabled |  | INSTRUMENT PARAMETERS
Capture Mode: |  | Auto | 0.09838 | 0.09838 | 9000 Sample Interval (ms)
Screen Tracking: |  | Yes | 38.08 | 38.08 | 9000 Input Filter (us) |  |  | 4081/4082
kVp |  |  | 37.3 | 37.3 | 37.3 | 37.3 | Sample Interval (us)
Config: | kVSDmA | DC+T+kV | 2.9764 | 2.9764 | 2.9764 | 2.9764 | Input Filter -3dB (kHz)
Scale: |  | 9096 | 16.2 | 16.2 | 16.2 | 16.2 | Output Filter (us)
Start delay (ms): |  | 0 | 4 | 4 | 2.6 | 2.6 | Averaging Intervals
Width (ms): |  | 5000 | 0.75 | 0.75 | Width fraction of kVp
Filtration (mm-Al): |  | 2.5 | 0.17 | 0.17 | Duplicate width (ms)
Dental threshold: |  | 0 | 0.0768 | 0.0768 | 9095 Sample Interval (ms)
Device Index |  | 0
Monitor Configuration |  |  |  |  | COEFFICIENTS FOR PRACTICAL kV
Serial No: |  | N/A | Power | Values
ROM Ver: |  | N/A | Min. kV | 20 | 36 | 150 | 161
Detector: |  | N/A | 0 | -18.6452 | -0.017472 | -0.017472 | 0
Chamb Kq: |  | N/A | 1 | 0.85068 | 1.0308e-05 | 1.0308e-05 | 0
Atm Corr: |  | Automatic | 2 | -0.019681 | 2.3082e-05 | 2.3082e-05 | 0
Temp (°C): |  | 22 | 3 | 0.00025663 | -1.662e-07 | -1.662e-07 | 0
Press (kPa): |  | 101.3 | 4 | -2.1424e-06 | 4.3106e-10 | 4.3106e-10 | 0
|  |  | 5 | 0 | 0 | 0 | 0
|  |  | EXP? | 1 | 0 | 0 | 0
NON-LINEAR CORRECTION FACTORS
|  |  | 1030 | 2700 | 0.12
|  |  | 0.343 | -0.569 | 0.2809 | -0.0331
|  |  | 2.9804 | -5.4582 | 4.6956 | -1.4557