## Data Form
[7 empty] | RADIOGRAPHIC CALIBRATION REPORT | [7 empty] | XLPRO ™ Version 4.06
[11 empty] | DATE: | 2026-07-30 02:06 |  |  | Operation Settings Area
[11 empty] | Tech: | JR
Units
A. |  | FACILITY: | [6 empty] | Room No: |  | Bldg: |  |  |  | Radiation: | R
|  | State Reg. No: | [5 empty] | UNIT: | [5 empty] | Additional Information. | /Time: | /min
[8 empty] | System ID : | [5 empty] | 1. Look for cells with comments like this: (e.g. E24) | Scale:
| EQUIPMENT |  |  | Manufacturer |  |  |  | Model |  |  | Serial No. |  |  | 2. This spreadsheet has been setup | Pulse Rate: | 15
|  |  | GENERATOR: | Electromed |  |  | EDEC 80 | [6 empty] | with protected cells to prevent inadvertent | Mode
|  |  | TUBE: | Varian |  |  | A192 | [6 empty] | modification of formulas. To remove the protection go to | Units Display: | Omit
|  |  | HOUSING: | [10 empty] | the Tools menu, Protection and select Unprotect Sheet. | Cell Update: | Incremental
|  |  | COLLIMATOR: | N/A | [10 empty] | Auto Zero: | Disabled
|  |  | FILTRATION: | 2.5 | mm Al |  | Anode angle: | 12 | deg | Phase: | cp |  |  |  | Capture Mode: | Auto
|  |  | RATINGS: | Maximum kV: |  | 150 | kVp. Max. tube mA: |  |  | 800 | mA. |  |  |  | Screen Tracking: | Yes
[5 empty] | Small FS | 0.6 | mm |  | Large FS | 1.2 | mm |  |  |  | kVp | FT
B. | kVp ACCURACY AND RADIATION OUTPUT | [9 empty] | NOTES: |  |  |  | Config: | Time+kV
|  | Focal Spot: | auto |  | 100 | mA | 500 | ms |  | 1. | color | denotes manual data entry |  |  | Scale: | 4083
| Radcal 40X5-W kV sensor placed on table . | [5 empty] | 50 | mAs, |  | 2. | other | color | is Radcal 4083 data |  | Start delay (ms): | 0
| @ | 60 | cm from target, | [6 empty] | 3. | Apparent Error values will become RED if the the specification value is exceeded. |  |  |  | Width (ms): | 5000
[15 empty] | Filtration (mm-Al): | 2.5
|  |  |  | Specification: | ms ± | 5% | kVp ± | 5 | [6 empty] | Dental threshold:
| kV | FS | Measured |  | Apparent Error | [9 empty] | Device Index | 0
| Set | lg/sm | msec | kVp | ms |  | kV | [7 empty] | Monitor Configuration
| 40 | sm | 513.9 | 40.12 | 13.9 | 2.78% | 0.12 | 0.3% | [6 empty] | Serial No: | N/A
| 45 | sm | 495.3 | 45.01 | -4.7 | -0.94% | 0.01 | 0.02222% | [6 empty] | ROM Ver: | N/A
| 50 | sm | 502 | 50.02 | 2 | 0.4% | 0.02 | 0.04% | [6 empty] | Detector: | N/A
| 55 | sm | 494.2 | 55.05 | -5.8 | -1.16% | 0.05 | 0.09091% |  | 0 | 0 | 5 | -5 |  | Chamb Kq: | N/A
| 60 | sm | 500.4 | 60.12 | 0.4 | 0.08% | 0.12 | 0.2% |  | 200 | 0 | 5 | -5 |  | Atm Corr: | Automatic
| 65 | sm | 504.3 | 65.19 | 4.3 | 0.86% | 0.19 | 0.2923% | [6 empty] | Temp (°C): | 22
| 70 | sm | 509.8 | 70.62 | 9.8 | 1.96% | 0.62 | 0.8857% | [6 empty] | Press (kPa): | 101.3
| 75 | sm | 516.6 | 76.04 | 16.6 | 3.32% | 1.04 | 1.387% | [6 empty] | Press (kPa): | 101.3
| 80 | sm | 514.5 | 81.44 | 14.5 | 2.9% | 1.44 | 1.8%
| 85 | sm | 511.4 | 86.87 | 11.4 | 2.28% | 1.87 | 2.2%
| 90 | sm | 509.6 | 91.83 | 9.6 | 1.92% | 1.83 | 2.033%
| 95 | sm | 509.1 | 97.27 | 9.1 | 1.82% | 2.27 | 2.389%
| 100 | sm | 508.9 | 102.4 | 8.9 | 1.78% | 2.4 | 2.4%
| 110 | sm | 507.7 | 112.3 | 7.7 | 1.54% | 2.3 | 2.091%
| 120 | lg | 471.5 | 122.3 | -28.5 | -5.7% | 2.3 | 1.917%
| 130 | lg | 468.2 | 132.3 | -31.8 | -6.36% | 2.3 | 1.769%
| 140 | lg | 465.3 | 142.3 | -34.7 | -6.94% | 2.3 | 1.643%
| 150 | lg | 462.9 | 152 | -37.1 | -7.42% | 2 | 1.333%
Comments:
|  |  |  | kV and Timer Accuracy: | Time error exceeds 5% above 120 kV when unit switches to the
large FS.
|  |  |  | Action: | Service required
B. | Waveform Record
|  | Focal Spot: | Auto |  | 80 | kV | 500 | mA | 100 | ms
| Radcal 40X5-W kV sensor placed on table . | [5 empty] | 50 | mAs,
| @ | 60 | cm from target,
|  |  | 0 | 80
|  |  | 200 | 80
Comments:
|  |  |  | kV and Timer Accuracy: | Excellent rise & fall
|  |  |  | Action: | None
B. | kV& Time variation with mA
|  | Focal Spot: | Auto |  | 100 | ms
Radcal 40X5-W kV sensor placed on table .
| @ | 60 | cm from target,
|  |  | Specification: | ms ± | 5% | kVp ± | 5
| SET | 100 | 40 | Apparent Error
| mA | ms | kVp | ms |  | kV
| 10 | 101 | 41.7 | 1 | 1% | 1.7 | 4.25%
| 15 | 99.8 | 41.49 | -0.2 | -0.2% | 1.49 | 3.725% |  | 0 | 0 | 5 | -5
| 25 | 103.6 | 40.87 | 3.6 | 3.6% | 0.87 | 2.175% |  | 600 | 0 | 5 | -5
| 40 | 106.9 | 40.72 | 6.9 | 6.9% | 0.72 | 1.8%
| 75 | 100.9 | 40.28 | 0.9 | 0.9% | 0.28 | 0.7%
| 100 | 102.7 | 40.24 | 2.7 | 2.7% | 0.24 | 0.6%
| 150 | 100.4 | 40.12 | 0.4 | 0.4% | 0.12 | 0.3%
| 200 | 101.1 | 39.96 | 1.1 | 1.1% | -0.04 | -0.1%
| 300 | 101.2 | 39.85 | 1.2 | 1.2% | -0.15 | -0.375%
| 400 | 98.6 | 39.75 | -1.4 | -1.4% | -0.25 | -0.625%
| 500 | 98.1 | 39.68 | -1.9 | -1.9% | -0.32 | -0.8%
|  |  |  | Average: | 1.3% | 0.42364 | kV
|  |  |  | Standard Deviation | 2.439% | 0.68925 | kV
|  | Focal Spot: | Auto
Radcal 40X5-W kV sensor placed on table .
| @ | 60 | cm from target,
|  |  | Specification: | ms ± | 5% | kVp ± | 5
| SET | 100 | 80 | Apparent Error
| mA | ms | kVp | ms |  | kV
| 10 | 95 | 82.01 | -5 | -5% | 2.01 | 2.513%
| 15 | 94.5 | 81.93 | -5.5 | -5.5% | 1.93 | 2.413% |  | 0 | 0 | 5 | -5
| 25 | 99.4 | 81.78 | -0.6 | -0.6% | 1.78 | 2.225% |  | 600 | 0 | 5 | -5
| 40 | 103.7 | 81.77 | 3.7 | 3.7% | 1.77 | 2.212%
| 75 | 98.3 | 81.67 | -1.7 | -1.7% | 1.67 | 2.087%
| 100 | 104.1 | 81.67 | 4.1 | 4.1% | 1.67 | 2.087%
| 150 | 103.1 | 81.55 | 3.1 | 3.1% | 1.55 | 1.938%
| 200 | 98.1 | 81.46 | -1.9 | -1.9% | 1.46 | 1.825%
| 300 | 98.9 | 81.49 | -1.1 | -1.1% | 1.49 | 1.862%
| 400 | 99.1 | 81.38 | -0.9 | -0.9% | 1.38 | 1.725%
| 500 | 99.5 | 81.29 | -0.5 | -0.5% | 1.29 | 1.613%
|  |  |  | Average: | -0.5727% | 1.6364 | kV
|  |  |  | Standard Deviation | 3.176% | 0.22633 | kV
Comments:
|  |  |  | kV and Timer Accuracy: | Time error at limits for low mA @ 80 kV
|  |  |  | Action: | Watch for future change
Note on C53: 1. Go to Waveform Data Sheet 2. Goto Cell F5. 3. Enter Info 2. Goto Cell E7 and capture waveform
Note on D98: 1. Set x-ray machine 2. Ctrl-Shift-K 3. Expose (data will enter sheet automatically) 4. Change x-ray setting & Expose 5. Repeat for all settings 6. Press Esc to end.
Note on D119: 1. Set x-ray machine 2. Ctrl-Shift-K 3. Expose (data will enter sheet automatically) 4. Change x-ray setting & Expose 5. Repeat for all settings 6. Press Esc to end.
Note on E24: Place cursor here 1. Set x-ray machine 2. Ctrl-Shift-K 3. Expose (data will enter sheet automatically) 4. Change x-ray setting & Expose 5. Repeat for all settings 6. Press Esc to end.
Note on U7: These comments will provide useful instructions and will not print.

## Waveform Data Sheet
XLPRO ™ Version 4.06
Operation Settings Area |  |  |  | Sample Waveform Capture Sheet
Units
Radiation: |  | R |  | Generator Setting |  |  |  | FS | Filtration mm Al |  |  | Generator Setting |  |  |  | FS | Filtration mm Al
/Time: |  | /min |  | kVp | ms | mA | FS | Dist (cm) | Inherent | Total |  | kVp | ms | mA | FS | Dist (cm) | Inherent | Total
Scale: |  |  | SET: | 80 | 100 | 500 | Auto | 60 | 1.2 | 2.5 | SET: | 80 | 100 | 10 | Auto | 60 | 1.2 | 2.5
Pulse Rate: |  | 15
Mode |  |  | 99.7 | ms | 81.28 | kVp |  |  |  |  | 96 | ms | 82.04 | kVp
Units Display: |  | Rescale | Peak= | 81.275 | kVp | PPV= | 81.2 | kV | Control= | 32832 | Peak= | 82.04 | kVp | PPV= | 81.8 | kV | Control= | 32832
Cell Update: |  | Incremental | Rel Dose= | N/A | /s | Practical= | 81.245 | kV |  |  | Rel Dose= | N/A | /s | Practical= | 81.88 | kV
Auto Zero: |  | Disabled | Pk Freq= | 117.6197 | Hz | Rate= | 3.0096 | Hz | Avg. mA= | N/A | Pk Freq= | 667.5715 | Hz | Rate= | 3.1267 | Hz | Avg. mA= | 0
Capture Mode: |  | Auto | Width= | 99.6804 | ms | Pulses= | 1 |  | mAs= | N/A | Width= | 95.9471 | ms | Pulses= | 1 |  | mAs= | 0
Screen Tracking: |  | Yes | ms | kV |  |  |  | LINES | Hz | kV | ms | kV |  |  |  | LINES | Hz | kV
kVp |  | FT | -0.29867 | 0 |  |  |  | -0.29867 | 0 | -2.3201e-15 | -0.29867 | 0 |  |  |  | -0.29867 | 0 | 4.7063e-15
Config: | kVS | Time+kV | -0.224 | 0 |  |  |  | 100.4271 | 3.2697 | 0.015947 | -0.224 | 0 |  |  |  | 96.4698 | 3.2697 | 0.0071022
Scale: |  | 4083 | -0.14933 | 0 |  |  |  | 81.275 | 6.5395 | 0.026946 | -0.14933 | 0 |  |  |  | 82.04 | 6.5395 | 0.021736
Start delay (ms): |  | 0 | -0.074667 | 0 |  |  |  | 81.275 | 9.8092 | 0.029861 | -0.074667 | 0 |  |  |  | 82.04 | 9.8092 | 0.039724
Width (ms): |  | 4778 | 0 | 37.76 |  |  |  | 81.245 | 13.0789 | 0.0245 | 0 | 37.76 |  |  |  | 81.88 | 13.0789 | 0.054657
Filtration (mm-Al): |  | 2.5 | 0.074667 | 37.76 |  |  |  | 81.245 | 16.3486 | 0.013722 | 0.074667 | 37.76 |  |  |  | 81.88 | 16.3486 | 0.061583
Dental threshold: |  | 0 | 0.14933 | 45.825 |  |  |  | 0 | 19.6184 | 0.0048445 | 0.14933 | 49.9 |  |  |  | 0 | 19.6184 | 0.058625
Device Index |  | 0 | 0.224 | 58.73 |  |  |  | 0 | 22.8881 | 0.010172 | 0.224 | 67.555 |  |  |  | 0 | 22.8881 | 0.046938
Monitor Configuration |  |  | 0.29867 | 58.73 |  |  |  | 1382 | 26.1578 | 0.013154 | 0.29867 | 67.555 |  |  |  | 1329 | 26.1578 | 0.030417
Serial No: |  | N/A | 0.37334 | 58.73 |  |  |  |  | 29.4275 | 0.010196 | 0.37334 | 71.415 |  |  |  |  | 29.4275 | 0.022606
ROM Ver: |  | 4.00 | 0.448 | 75.45 |  |  |  | 0 | 32.6973 | 0.003162 | 0.448 | 77.795 |  |  |  | 0 | 32.6973 | 0.039821
Detector: |  | N/A | 0.52267 | 76.74 |  |  |  | 0 | 35.967 | 0.0046012 | 0.52267 | 79.12 |  |  |  | 0 | 35.967 | 0.062826
Chamb Kq: |  | N/A | 0.59734 | 78.12 |  |  |  |  | 39.2367 | 0.0099346 | 0.59734 | 79.685 |  |  |  |  | 39.2367 | 0.07824
Atm Corr: |  | Automatic | 0.672 | 79.04 |  |  |  |  | 42.5064 | 0.011253 | 0.672 | 79.895 |  |  |  |  | 42.5064 | 0.07831
Temp (°C): |  | 22 | 0.74667 | 79.36 |  |  |  |  | 45.7762 | 0.0092732 | 0.74667 | 80.325 |  |  |  |  | 45.7762 | 0.059502
Press (kPa): |  | 101.3 | 0.82134 | 79.635 |  |  |  |  | 49.0459 | 0.0070274 | 0.82134 | 80.73 |  |  |  |  | 49.0459 | 0.02483
|  |  | 0.896 | 80.03 |  |  |  |  | 52.3156 | 0.0071393 | 0.896 | 81.225 |  |  |  |  | 52.3156 | 0.024519
|  |  | 0.97067 | 80.245 |  |  |  |  | 55.5853 | 0.0073726 | 0.97067 | 81.455 |  |  |  |  | 55.5853 | 0.065619
|  |  | 1.0453 | 80.32 |  |  |  |  | 58.8551 | 0.0063995 | 1.0453 | 81.52 |  |  |  |  | 58.8551 | 0.094169
|  |  | 1.12 | 80.535 |  |  |  |  | 62.1248 | 0.0065666 | 1.12 | 81.39 |  |  |  |  | 62.1248 | 0.10212
|  |  | 1.1947 | 80.605 |  |  |  |  | 65.3945 | 0.0082067 | 1.1947 | 81.535 |  |  |  |  | 65.3945 | 0.088924
|  |  | 1.2693 | 80.63 |  |  |  |  | 68.6642 | 0.0078706 | 1.2693 | 81.935 |  |  |  |  | 68.6642 | 0.060786
|  |  | 1.344 | 80.645 |  |  |  |  | 71.934 | 0.0036753 | 1.344 | 82.165 |  |  |  |  | 71.934 | 0.02826
|  |  | 1.4187 | 80.85 |  |  |  |  | 75.2037 | 0.00399 | 1.4187 | 82.02 |  |  |  |  | 75.2037 | 0.004476
|  |  | 1.4933 | 80.895 |  |  |  |  | 78.4734 | 0.012569 | 1.4933 | 81.395 |  |  |  |  | 78.4734 | 0.0116
|  |  | 1.568 | 80.86 |  |  |  |  | 81.7431 | 0.019005 | 1.568 | 81.345 |  |  |  |  | 81.7431 | 0.0084301
|  |  | 1.6427 | 80.85 |  |  |  |  | 85.0129 | 0.020746 | 1.6427 | 80.765 |  |  |  |  | 85.0129 | 0.0067177
|  |  | 1.7173 | 80.905 |  |  |  |  | 88.2826 | 0.017427 | 1.7173 | 81.465 |  |  |  |  | 88.2826 | 0.025221
|  |  | 1.792 | 80.945 |  |  |  |  | 91.5523 | 0.013104 | 1.792 | 81.845 |  |  |  |  | 91.5523 | 0.039865
|  |  | 1.8667 | 81.03 |  |  |  |  | 94.8221 | 0.016133 | 1.8667 | 82.49 |  |  |  |  | 94.8221 | 0.04506
|  |  | 1.9413 | 81.095 |  |  |  |  | 98.0918 | 0.022899 | 1.9413 | 81.905 |  |  |  |  | 98.0918 | 0.039523
|  |  | 2.016 | 81.085 |  |  |  |  | 101.3615 | 0.025821 | 2.016 | 81.37 |  |  |  |  | 101.3615 | 0.025894
|  |  | 2.0907 | 81.055 |  |  |  |  | 104.6312 | 0.022417 | 2.0907 | 81.635 |  |  |  |  | 104.6312 | 0.0089315
|  |  | 2.1653 | 81.105 |  |  |  |  | 107.901 | 0.015089 | 2.1653 | 81.31 |  |  |  |  | 107.901 | 0.0071519
|  |  | 2.24 | 81.155 |  |  |  |  | 111.1707 | 0.015047 | 2.24 | 82.07 |  |  |  |  | 111.1707 | 0.019647
|  |  | 2.3147 | 81.185 |  |  |  |  | 114.4404 | 0.024791 | 2.3147 | 82.28 |  |  |  |  | 114.4404 | 0.028603
|  |  | 2.3893 | 81.215 |  |  |  |  | 117.7101 | 0.032579 | 2.3893 | 82.485 |  |  |  |  | 117.7101 | 0.034482
|  |  | 2.464 | 81.265 |  |  |  |  | 120.9799 | 0.033506 | 2.464 | 81.73 |  |  |  |  | 120.9799 | 0.03712
|  |  | 2.5387 | 81.235 |  |  |  |  | 124.2496 | 0.026794 | 2.5387 | 81.6 |  |  |  |  | 124.2496 | 0.036068
|  |  | 2.6133 | 81.17 |  |  |  |  | 127.5193 | 0.014609 | 2.6133 | 81.975 |  |  |  |  | 127.5193 | 0.031806
|  |  | 2.688 | 81.07 |  |  |  |  | 130.789 | 0.0052534 | 2.688 | 81.275 |  |  |  |  | 130.789 | 0.026336
|  |  | 2.7627 | 81.035 |  |  |  |  | 134.0588 | 0.014579 | 2.7627 | 81.395 |  |  |  |  | 134.0588 | 0.023155
|  |  | 2.8373 | 81.045 |  |  |  |  | 137.3285 | 0.021455 | 2.8373 | 81.62 |  |  |  |  | 137.3285 | 0.026139
|  |  | 2.912 | 81.115 |  |  |  |  | 140.5982 | 0.02189 | 2.912 | 82.325 |  |  |  |  | 140.5982 | 0.03537
|  |  | 2.9867 | 81.135 |  |  |  |  | 143.8679 | 0.016931 | 2.9867 | 82.21 |  |  |  |  | 143.8679 | 0.047011
|  |  | 3.0613 | 81.035 |  |  |  |  | 147.1377 | 0.010752 | 3.0613 | 81.24 |  |  |  |  | 147.1377 | 0.056253
|  |  | 3.136 | 81.035 |  |  |  |  | 150.4074 | 0.010297 | 3.136 | 81.345 |  |  |  |  | 150.4074 | 0.058515
|  |  | 3.2107 | 81.075 |  |  |  |  | 153.6771 | 0.013133 | 3.2107 | 82.39 |  |  |  |  | 153.6771 | 0.05079
|  |  | 3.2853 | 81.21 |  |  |  |  | 156.9468 | 0.013022 | 3.2853 | 83.01 |  |  |  |  | 156.9468 | 0.033591
|  |  | 3.36 | 81.235 |  |  |  |  | 160.2166 | 0.0098145 | 3.36 | 82.52 |  |  |  |  | 160.2166 | 0.019719
|  |  | 3.4347 | 81.255 |  |  |  |  | 163.4863 | 0.0082674 | 3.4347 | 81.81 |  |  |  |  | 163.4863 | 0.03918
|  |  | 3.5093 | 81.16 |  |  |  |  | 166.756 | 0.011241 | 3.5093 | 81.83 |  |  |  |  | 166.756 | 0.067121
|  |  | 3.584 | 81.15 |  |  |  |  | 170.0257 | 0.013003 | 3.584 | 81.295 |  |  |  |  | 170.0257 | 0.088327
|  |  | 3.6587 | 81.17 |  |  |  |  | 173.2955 | 0.010392 | 3.6587 | 81.405 |  |  |  |  | 173.2955 | 0.097313
|  |  | 3.7334 | 81.235 |  |  |  |  | 176.5652 | 0.0040731 | 3.7334 | 81.55 |  |  |  |  | 176.5652 | 0.092214
|  |  | 3.808 | 81.285 |  |  |  |  | 179.8349 | 0.0068949 | 3.808 | 82.26 |  |  |  |  | 179.8349 | 0.074457
|  |  | 3.8827 | 81.24 |  |  |  |  | 183.1047 | 0.014872 | 3.8827 | 82.38 |  |  |  |  | 183.1047 | 0.048281
|  |  | 3.9574 | 81.145 |  |  |  |  | 186.3744 | 0.019471 | 3.9574 | 82.42 |  |  |  |  | 186.3744 | 0.019712
|  |  | 4.032 | 81.11 |  |  |  |  | 189.6441 | 0.019325 | 4.032 | 81.795 |  |  |  |  | 189.6441 | 0.0050813
|  |  | 4.1067 | 81.155 |  |  |  |  | 192.9138 | 0.01539 | 4.1067 | 81.47 |  |  |  |  | 192.9138 | 0.020546
|  |  | 4.1814 | 81.15 |  |  |  |  | 196.1836 | 0.01074 | 4.1814 | 81.25 |  |  |  |  | 196.1836 | 0.024503
|  |  | 4.256 | 81.055 |  |  |  |  | 199.4533 | 0.0089242 | 4.256 | 81.485 |  |  |  |  | 199.4533 | 0.018064
|  |  | 4.3307 | 80.965 |  |  |  |  | 202.723 | 0.0088203 | 4.3307 | 81.665 |  |  |  |  | 202.723 | 0.010268
|  |  | 4.4054 | 81.07 |  |  |  |  | 205.9927 | 0.0075268 | 4.4054 | 81.595 |  |  |  |  | 205.9927 | 0.019901
|  |  | 4.48 | 81.215 |  |  |  |  | 209.2625 | 0.0063182 | 4.48 | 81.375 |  |  |  |  | 209.2625 | 0.031235
|  |  | 4.5547 | 81.305 |  |  |  |  | 212.5322 | 0.0075756 | 4.5547 | 81.87 |  |  |  |  | 212.5322 | 0.034536
|  |  | 4.6294 | 81.21 |  |  |  |  | 215.8019 | 0.0085381 | 4.6294 | 81.78 |  |  |  |  | 215.8019 | 0.027939
|  |  | 4.704 | 80.995 |  |  |  |  | 219.0716 | 0.0066072 | 4.704 | 81.675 |  |  |  |  | 219.0716 | 0.014534
|  |  | 4.7787 | 81.02 |  |  |  |  | 222.3414 | 0.002019 | 4.7787 | 82.365 |  |  |  |  | 222.3414 | 0.010816
|  |  | 4.8534 | 81.065 |  |  |  |  | 225.6111 | 0.0042148 | 4.8534 | 82.605 |  |  |  |  | 225.6111 | 0.02221
|  |  | 4.928 | 81.23 |  |  |  |  | 228.8808 | 0.0083645 | 4.928 | 82.555 |  |  |  |  | 228.8808 | 0.026891
|  |  | 5.0027 | 81.31 |  |  |  |  | 232.1505 | 0.0093885 | 5.0027 | 82.76 |  |  |  |  | 232.1505 | 0.020647
|  |  | 5.0774 | 81.405 |  |  |  |  | 235.4203 | 0.0081943 | 5.0774 | 81.955 |  |  |  |  | 235.4203 | 0.0050117
|  |  | 5.152 | 81.405 |  |  |  |  | 238.69 | 0.0092227 | 5.152 | 81.58 |  |  |  |  | 238.69 | 0.015217
|  |  | 5.2267 | 81.4 |  |  |  |  | 241.9597 | 0.012876 | 5.2267 | 81.405 |  |  |  |  | 241.9597 | 0.032746
|  |  | 5.3014 | 81.365 |  |  |  |  | 245.2294 | 0.014577 | 5.3014 | 81.365 |  |  |  |  | 245.2294 | 0.041637
|  |  | 5.376 | 81.335 |  |  |  |  | 248.4992 | 0.011961 | 5.376 | 81.36 |  |  |  |  | 248.4992 | 0.038677
|  |  | 5.4507 | 81.32 |  |  |  |  | 251.7689 | 0.0053655 | 5.4507 | 81.91 |  |  |  |  | 251.7689 | 0.024924
|  |  | 5.5254 | 81.275 |  |  |  |  | 255.0386 | 0.0029388 | 5.5254 | 82.875 |  |  |  |  | 255.0386 | 0.010208
|  |  | 5.6 | 81.325 |  |  |  |  | 258.3083 | 0.0099612 | 5.6 | 82.155 |  |  |  |  | 258.3083 | 0.024618
|  |  | 5.6747 | 81.395 |  |  |  |  | 261.5781 | 0.013698 | 5.6747 | 82.215 |  |  |  |  | 261.5781 | 0.042843
|  |  | 5.7494 | 81.42 |  |  |  |  | 264.8478 | 0.01434 | 5.7494 | 81.555 |  |  |  |  | 264.8478 | 0.054636
|  |  | 5.824 | 81.41 |  |  |  |  | 268.1175 | 0.01425 | 5.824 | 81.195 |  |  |  |  | 268.1175 | 0.059681
|  |  | 5.8987 | 81.38 |  |  |  |  | 271.3873 | 0.014886 | 5.8987 | 81.825 |  |  |  |  | 271.3873 | 0.06067
|  |  | 5.9734 | 81.315 |  |  |  |  | 274.657 | 0.014339 | 5.9734 | 82.34 |  |  |  |  | 274.657 | 0.060956
|  |  | 6.048 | 81.215 |  |  |  |  | 277.9267 | 0.010529 | 6.048 | 82.065 |  |  |  |  | 277.9267 | 0.062408
|  |  | 6.1227 | 81.29 |  |  |  |  | 281.1964 | 0.0040041 | 6.1227 | 82.35 |  |  |  |  | 281.1964 | 0.06503
|  |  | 6.1974 | 81.335 |  |  |  |  | 284.4662 | 0.0047474 | 6.1974 | 82.29 |  |  |  |  | 284.4662 | 0.067931
|  |  | 6.272 | 81.25 |  |  |  |  | 287.7359 | 0.0095584 | 6.272 | 81.505 |  |  |  |  | 287.7359 | 0.069792
|  |  | 6.3467 | 81.175 |  |  |  |  | 291.0056 | 0.0098059 | 6.3467 | 81.12 |  |  |  |  | 291.0056 | 0.068926
|  |  | 6.4214 | 81.16 |  |  |  |  | 294.2753 | 0.0050444 | 6.4214 | 81.435 |  |  |  |  | 294.2753 | 0.063852
|  |  | 6.496 | 81.16 |  |  |  |  | 297.5451 | 0.0028491 | 6.496 | 82.56 |  |  |  |  | 297.5451 | 0.054334
|  |  | 6.5707 | 81.22 |  |  |  |  | 300.8148 | 0.010479 | 6.5707 | 82.875 |  |  |  |  | 300.8148 | 0.042788
|  |  | 6.6454 | 81.27 |  |  |  |  | 304.0845 | 0.014801 | 6.6454 | 82.195 |  |  |  |  | 304.0845 | 0.036286
|  |  | 6.72 | 81.25 |  |  |  |  | 307.3542 | 0.01426 | 6.72 | 81.265 |  |  |  |  | 307.3542 | 0.041819
|  |  | 6.7947 | 81.295 |  |  |  |  | 310.624 | 0.0095218 | 6.7947 | 81.295 |  |  |  |  | 310.624 | 0.053371
|  |  | 6.8694 | 81.33 |  |  |  |  | 313.8937 | 0.0029996 | 6.8694 | 81.305 |  |  |  |  | 313.8937 | 0.061478
|  |  | 6.944 | 81.33 |  |  |  |  | 317.1634 | 0.002539 | 6.944 | 81.24 |  |  |  |  | 317.1634 | 0.061125
|  |  | 7.0187 | 81.29 |  |  |  |  | 320.4331 | 0.0056928 | 7.0187 | 81.6 |  |  |  |  | 320.4331 | 0.05129
|  |  | 7.0934 | 81.26 |  |  |  |  | 323.7029 | 0.0077162 | 7.0934 | 82.09 |  |  |  |  | 323.7029 | 0.034346
|  |  | 7.168 | 81.275 |  |  |  |  | 326.9726 | 0.010437 | 7.168 | 82.15 |  |  |  |  | 326.9726 | 0.015013
|  |  | 7.2427 | 81.355 |  |  |  |  | 330.2423 | 0.012524 | 7.2427 | 82.39 |  |  |  |  | 330.2423 | 0.0017026
|  |  | 7.3174 | 81.41 |  |  |  |  | 333.512 | 0.011541 | 7.3174 | 83.085 |  |  |  |  | 333.512 | 0.012568
|  |  | 7.392 | 81.31 |  |  |  |  | 336.7818 | 0.006842 | 7.392 | 82.55 |  |  |  |  | 336.7818 | 0.019444
|  |  | 7.4667 | 81.235 |  |  |  |  | 340.0515 | 0.0031972 | 7.4667 | 82.315 |  |  |  |  | 340.0515 | 0.026979
|  |  | 7.5414 | 81.28 |  |  |  |  | 343.3212 | 0.0089419 | 7.5414 | 82.57 |  |  |  |  | 343.3212 | 0.035675
|  |  | 7.616 | 81.415 |  |  |  |  | 346.5909 | 0.012526 | 7.616 | 82.67 |  |  |  |  | 346.5909 | 0.04069
|  |  | 7.6907 | 81.42 |  |  |  |  | 349.8607 | 0.011312 | 7.6907 | 82.195 |  |  |  |  | 349.8607 | 0.037404
|  |  | 7.7654 | 81.3 |  |  |  |  | 353.1304 | 0.0066537 | 7.7654 | 81.895 |  |  |  |  | 353.1304 | 0.024937
|  |  | 7.84 | 81.245 |  |  |  |  | 356.4001 | 0.007375 | 7.84 | 81.65 |  |  |  |  | 356.4001 | 0.012353
|  |  | 7.9147 | 81.24 |  |  |  |  | 359.6699 | 0.013662 | 7.9147 | 81.535 |  |  |  |  | 359.6699 | 0.027162
|  |  | 7.9894 | 81.24 |  |  |  |  | 362.9396 | 0.017471 | 7.9894 | 81.925 |  |  |  |  | 362.9396 | 0.047378
|  |  | 8.064 | 81.245 |  |  |  |  | 366.2093 | 0.017104 | 8.064 | 82.165 |  |  |  |  | 366.2093 | 0.061629
|  |  | 8.1387 | 81.29 |  |  |  |  | 369.479 | 0.013462 | 8.1387 | 81.66 |  |  |  |  | 369.479 | 0.068179
|  |  | 8.2134 | 81.295 |  |  |  |  | 372.7488 | 0.0086559 | 8.2134 | 81.855 |  |  |  |  | 372.7488 | 0.068981
|  |  | 8.288 | 81.235 |  |  |  |  | 376.0185 | 0.0043391 | 8.288 | 81.975 |  |  |  |  | 376.0185 | 0.067507
|  |  | 8.3627 | 81.21 |  |  |  |  | 379.2882 | 0.00079436 | 8.3627 | 81.615 |  |  |  |  | 379.2882 | 0.065534
|  |  | 8.4374 | 81.33 |  |  |  |  | 382.5579 | 0.0039322 | 8.4374 | 81.535 |  |  |  |  | 382.5579 | 0.061392
|  |  | 8.512 | 81.42 |  |  |  |  | 385.8277 | 0.0077602 | 8.512 | 81.65 |  |  |  |  | 385.8277 | 0.051846
|  |  | 8.5867 | 81.455 |  |  |  |  | 389.0974 | 0.0098366 | 8.5867 | 81.6 |  |  |  |  | 389.0974 | 0.035169
|  |  | 8.6614 | 81.35 |  |  |  |  | 392.3671 | 0.0093779 | 8.6614 | 81.49 |  |  |  |  | 392.3671 | 0.013276
|  |  | 8.736 | 81.35 |  |  |  |  | 395.6368 | 0.0076658 | 8.736 | 81.88 |  |  |  |  | 395.6368 | 0.014184
|  |  | 8.8107 | 81.395 |  |  |  |  | 398.9066 | 0.0081321 | 8.8107 | 82.215 |  |  |  |  | 398.9066 | 0.033421
|  |  | 8.8854 | 81.32 |  |  |  |  | 402.1763 | 0.010593 | 8.8854 | 82.615 |  |  |  |  | 402.1763 | 0.043317
|  |  | 8.96 | 81.19 |  |  |  |  | 405.446 | 0.011772 | 8.96 | 82.295 |  |  |  |  | 405.446 | 0.042318
|  |  | 9.0347 | 81.12 |  |  |  |  | 408.7157 | 0.010464 | 9.0347 | 81.88 |  |  |  |  | 408.7157 | 0.040975
|  |  | 9.1094 | 81.285 |  |  |  |  | 411.9855 | 0.0073776 | 9.1094 | 81.89 |  |  |  |  | 411.9855 | 0.057156
|  |  | 9.184 | 81.305 |  |  |  |  | 415.2552 | 0.0041371 | 9.184 | 82.13 |  |  |  |  | 415.2552 | 0.083658
|  |  | 9.2587 | 81.27 |  |  |  |  | 418.5249 | 0.0020041 | 9.2587 | 82.165 |  |  |  |  | 418.5249 | 0.10357
|  |  | 9.3334 | 81.285 |  |  |  |  | 421.7946 | 0.00051078 | 9.3334 | 82.055 |  |  |  |  | 421.7946 | 0.10638
|  |  | 9.408 | 81.305 |  |  |  |  | 425.0644 | 0.0015176 | 9.408 | 82.045 |  |  |  |  | 425.0644 | 0.088855
|  |  | 9.4827 | 81.285 |  |  |  |  | 428.3341 | 0.003578 | 9.4827 | 81.86 |  |  |  |  | 428.3341 | 0.055249
|  |  | 9.5574 | 81.195 |  |  |  |  | 431.6038 | 0.005094 | 9.5574 | 82.08 |  |  |  |  | 431.6038 | 0.020186
|  |  | 9.632 | 81.15 |  |  |  |  | 434.8735 | 0.0058962 | 9.632 | 82.345 |  |  |  |  | 434.8735 | 0.032959
|  |  | 9.7067 | 81.225 |  |  |  |  | 438.1433 | 0.0060132 | 9.7067 | 82.375 |  |  |  |  | 438.1433 | 0.055071
|  |  | 9.7814 | 81.39 |  |  |  |  | 441.413 | 0.0049536 | 9.7814 | 81.44 |  |  |  |  | 441.413 | 0.060668
|  |  | 9.856 | 81.39 |  |  |  |  | 444.6827 | 0.0028333 | 9.856 | 81.59 |  |  |  |  | 444.6827 | 0.051912
|  |  | 9.9307 | 81.385 |  |  |  |  | 447.9525 | 0.0056697 | 9.9307 | 82.145 |  |  |  |  | 447.9525 | 0.042658
|  |  | 10.0054 | 81.315 |  |  |  |  | 451.2222 | 0.011792 | 10.0054 | 82.675 |  |  |  |  | 451.2222 | 0.050417
|  |  | 10.08 | 81.29 |  |  |  |  | 454.4919 | 0.016789 | 10.08 | 82.375 |  |  |  |  | 454.4919 | 0.067307
|  |  | 10.1547 | 81.325 |  |  |  |  | 457.7616 | 0.018274 | 10.1547 | 81.855 |  |  |  |  | 457.7616 | 0.079268
|  |  | 10.2294 | 81.325 |  |  |  |  | 461.0314 | 0.015246 | 10.2294 | 81.64 |  |  |  |  | 461.0314 | 0.081883
|  |  | 10.304 | 81.225 |  |  |  |  | 464.3011 | 0.0086061 | 10.304 | 81.635 |  |  |  |  | 464.3011 | 0.076129
|  |  | 10.3787 | 81.18 |  |  |  |  | 467.5708 | 0.0025022 | 10.3787 | 81.225 |  |  |  |  | 467.5708 | 0.064972
|  |  | 10.4534 | 81.28 |  |  |  |  | 470.8405 | 0.0070376 | 10.4534 | 82.265 |  |  |  |  | 470.8405 | 0.051597
|  |  | 10.528 | 81.35 |  |  |  |  | 474.1103 | 0.01014 | 10.528 | 82.15 |  |  |  |  | 474.1103 | 0.038955
|  |  | 10.6027 | 81.36 |  |  |  |  | 477.38 | 0.01074 | 10.6027 | 82.16 |  |  |  |  | 477.38 | 0.029586
|  |  | 10.6774 | 81.405 |  |  |  |  | 480.6497 | 0.012492 | 10.6774 | 81.845 |  |  |  |  | 480.6497 | 0.024964
|  |  | 10.752 | 81.39 |  |  |  |  | 483.9194 | 0.016685 | 10.752 | 81.28 |  |  |  |  | 483.9194 | 0.026051
|  |  | 10.8267 | 81.285 |  |  |  |  | 487.1892 | 0.020056 | 10.8267 | 81.68 |  |  |  |  | 487.1892 | 0.033007
|  |  | 10.9014 | 81.2 |  |  |  |  | 490.4589 | 0.020173 | 10.9014 | 82.455 |  |  |  |  | 490.4589 | 0.042027
|  |  | 10.976 | 81.285 |  |  |  |  | 493.7286 | 0.017253 | 10.976 | 83.215 |  |  |  |  | 493.7286 | 0.046987
|  |  | 11.0507 | 81.385 |  |  |  |  | 496.9983 | 0.014049 | 11.0507 | 83.49 |  |  |  |  | 496.9983 | 0.042932
|  |  | 11.1254 | 81.36 |  |  |  |  | 500.2681 | 0.013664 | 11.1254 | 82.775 |  |  |  |  | 500.2681 | 0.028279
|  |  | 11.2001 | 81.165 |  |  |  |  | 503.5378 | 0.015099 | 11.2001 | 82.375 |  |  |  |  | 503.5378 | 0.0056692
|  |  | 11.2747 | 81.175 |  |  |  |  | 506.8075 | 0.016042 | 11.2747 | 82.065 |  |  |  |  | 506.8075 | 0.018838
|  |  | 11.3494 | 81.37 |  |  |  |  | 510.0772 | 0.016687 | 11.3494 | 81.76 |  |  |  |  | 510.0772 | 0.03806
|  |  | 11.4241 | 81.39 |  |  |  |  | 513.347 | 0.018016 | 11.4241 | 81.63 |  |  |  |  | 513.347 | 0.046685
|  |  | 11.4987 | 81.375 |  |  |  |  | 516.6167 | 0.019372 | 11.4987 | 81.68 |  |  |  |  | 516.6167 | 0.043917
|  |  | 11.5734 | 81.38 |  |  |  |  | 519.8864 | 0.019173 | 11.5734 | 81.895 |  |  |  |  | 519.8864 | 0.035318
|  |  | 11.6481 | 81.405 |  |  |  |  | 523.1561 | 0.016694 | 11.6481 | 82.405 |  |  |  |  | 523.1561 | 0.032115
|  |  | 11.7227 | 81.475 |  |  |  |  | 526.4259 | 0.012504 | 11.7227 | 82.355 |  |  |  |  | 526.4259 | 0.037243
|  |  | 11.7974 | 81.45 |  |  |  |  | 529.6956 | 0.0079744 | 11.7974 | 81.795 |  |  |  |  | 529.6956 | 0.041077
|  |  | 11.8721 | 81.27 |  |  |  |  | 532.9653 | 0.0045039 | 11.8721 | 81.965 |  |  |  |  | 532.9653 | 0.039773
|  |  | 11.9467 | 81.18 |  |  |  |  | 536.2351 | 0.0025662 | 11.9467 | 82.455 |  |  |  |  | 536.2351 | 0.037882
|  |  | 12.0214 | 81.26 |  |  |  |  | 539.5048 | 0.0017499 | 12.0214 | 82.59 |  |  |  |  | 539.5048 | 0.040267
|  |  | 12.0961 | 81.325 |  |  |  |  | 542.7745 | 0.0032421 | 12.0961 | 82.47 |  |  |  |  | 542.7745 | 0.042769
|  |  | 12.1707 | 81.3 |  |  |  |  | 546.0442 | 0.0058965 | 12.1707 | 82.145 |  |  |  |  | 546.0442 | 0.038433
|  |  | 12.2454 | 81.34 |  |  |  |  | 549.314 | 0.0081366 | 12.2454 | 81.765 |  |  |  |  | 549.314 | 0.025786
|  |  | 12.3201 | 81.39 |  |  |  |  | 552.5837 | 0.0091706 | 12.3201 | 82.445 |  |  |  |  | 552.5837 | 0.015162
|  |  | 12.3947 | 81.42 |  |  |  |  | 555.8534 | 0.0090015 | 12.3947 | 82.615 |  |  |  |  | 555.8534 | 0.026788
|  |  | 12.4694 | 81.46 |  |  |  |  | 559.1231 | 0.0085515 | 12.4694 | 82.67 |  |  |  |  | 559.1231 | 0.040405
|  |  | 12.5441 | 81.565 |  |  |  |  | 562.3929 | 0.0089103 | 12.5441 | 82.245 |  |  |  |  | 562.3929 | 0.044864
|  |  | 12.6187 | 81.455 |  |  |  |  | 565.6626 | 0.0099674 | 12.6187 | 81.8 |  |  |  |  | 565.6626 | 0.040297
|  |  | 12.6934 | 81.285 |  |  |  |  | 568.9323 | 0.011059 | 12.6934 | 81.165 |  |  |  |  | 568.9323 | 0.035444
|  |  | 12.7681 | 81.215 |  |  |  |  | 572.202 | 0.012159 | 12.7681 | 81.41 |  |  |  |  | 572.202 | 0.041893
|  |  | 12.8427 | 81.28 |  |  |  |  | 575.4718 | 0.013295 | 12.8427 | 82.475 |  |  |  |  | 575.4718 | 0.054509
|  |  | 12.9174 | 81.26 |  |  |  |  | 578.7415 | 0.013748 | 12.9174 | 83.285 |  |  |  |  | 578.7415 | 0.062083
|  |  | 12.9921 | 81.125 |  |  |  |  | 582.0112 | 0.012774 | 12.9921 | 83.45 |  |  |  |  | 582.0112 | 0.059366
|  |  | 13.0667 | 81.075 |  |  |  |  | 585.2809 | 0.011046 | 13.0667 | 82.375 |  |  |  |  | 585.2809 | 0.046108
|  |  | 13.1414 | 81.09 |  |  |  |  | 588.5507 | 0.011022 | 13.1414 | 81.885 |  |  |  |  | 588.5507 | 0.025461
|  |  | 13.2161 | 81.155 |  |  |  |  | 591.8204 | 0.013247 | 13.2161 | 81.73 |  |  |  |  | 591.8204 | 0.005032
|  |  | 13.2907 | 81.21 |  |  |  |  | 595.0901 | 0.014958 | 13.2907 | 81.605 |  |  |  |  | 595.0901 | 0.020614
|  |  | 13.3654 | 81.19 |  |  |  |  | 598.3598 | 0.014276 | 13.3654 | 81.925 |  |  |  |  | 598.3598 | 0.035485
|  |  | 13.4401 | 81.26 |  |  |  |  | 601.6296 | 0.011407 | 13.4401 | 81.8 |  |  |  |  | 601.6296 | 0.041735
|  |  | 13.5147 | 81.275 |  |  |  |  | 604.8993 | 0.0082647 | 13.5147 | 81.165 |  |  |  |  | 604.8993 | 0.038423
|  |  | 13.5894 | 81.27 |  |  |  |  | 608.169 | 0.0075149 | 13.5894 | 81.245 |  |  |  |  | 608.169 | 0.026374
|  |  | 13.6641 | 81.275 |  |  |  |  | 611.4387 | 0.0096368 | 13.6641 | 81.4 |  |  |  |  | 611.4387 | 0.0083027
|  |  | 13.7387 | 81.205 |  |  |  |  | 614.7085 | 0.013245 | 13.7387 | 81.62 |  |  |  |  | 614.7085 | 0.014596
|  |  | 13.8134 | 81.185 |  |  |  |  | 617.9782 | 0.017008 | 13.8134 | 81.765 |  |  |  |  | 617.9782 | 0.034419
|  |  | 13.8881 | 81.195 |  |  |  |  | 621.2479 | 0.019045 | 13.8881 | 82.485 |  |  |  |  | 621.2479 | 0.048264
|  |  | 13.9627 | 81.205 |  |  |  |  | 624.5177 | 0.017748 | 13.9627 | 81.665 |  |  |  |  | 624.5177 | 0.052218
|  |  | 14.0374 | 81.265 |  |  |  |  | 627.7874 | 0.013118 | 14.0374 | 81.15 |  |  |  |  | 627.7874 | 0.044653
|  |  | 14.1121 | 81.165 |  |  |  |  | 631.0571 | 0.0080827 | 14.1121 | 80.985 |  |  |  |  | 631.0571 | 0.027046
|  |  | 14.1867 | 81.13 |  |  |  |  | 634.3268 | 0.008464 | 14.1867 | 81.355 |  |  |  |  | 634.3268 | 0.006895
|  |  | 14.2614 | 81.205 |  |  |  |  | 637.5966 | 0.010809 | 14.2614 | 81.665 |  |  |  |  | 637.5966 | 0.022459
|  |  | 14.3361 | 81.265 |  |  |  |  | 640.8663 | 0.009786 | 14.3361 | 82.2 |  |  |  |  | 640.8663 | 0.039733
|  |  | 14.4107 | 81.165 |  |  |  |  | 644.136 | 0.0055729 | 14.4107 | 81.595 |  |  |  |  | 644.136 | 0.047752
|  |  | 14.4854 | 81.075 |  |  |  |  | 647.4057 | 0.0051855 | 14.4854 | 81.86 |  |  |  |  | 647.4057 | 0.047373
|  |  | 14.5601 | 81.08 |  |  |  |  | 650.6755 | 0.0099329 | 14.5601 | 81.435 |  |  |  |  | 650.6755 | 0.044629
|  |  | 14.6347 | 81.2 |  |  |  |  | 653.9452 | 0.012423 | 14.6347 | 81.36 |  |  |  |  | 653.9452 | 0.045522
|  |  | 14.7094 | 81.235 |  |  |  |  | 657.2149 | 0.012142 | 14.7094 | 81.435 |  |  |  |  | 657.2149 | 0.047498
|  |  | 14.7841 | 81.225 |  |  |  |  | 660.4846 | 0.011882 | 14.7841 | 82.025 |  |  |  |  | 660.4846 | 0.04367
|  |  | 14.8587 | 81.225 |  |  |  |  | 663.7544 | 0.013485 | 14.8587 | 82.18 |  |  |  |  | 663.7544 | 0.032946
|  |  | 14.9334 | 81.175 |  |  |  |  | 667.0241 | 0.014314 | 14.9334 | 82.175 |  |  |  |  | 667.0241 | 0.030367
|  |  | 15.0081 | 81.205 |  |  |  |  | 670.2938 | 0.011789 | 15.0081 | 82.13 |  |  |  |  | 670.2938 | 0.052314
|  |  | 15.0827 | 81.27 |  |  |  |  | 673.5635 | 0.0058967 | 15.0827 | 81.81 |  |  |  |  | 673.5635 | 0.083259
|  |  | 15.1574 | 81.235 |  |  |  |  | 676.8333 | 0.0016354 | 15.1574 | 81.695 |  |  |  |  | 676.8333 | 0.11291
|  |  | 15.2321 | 81.285 |  |  |  |  | 680.103 | 0.007815 | 15.2321 | 81.605 |  |  |  |  | 680.103 | 0.13656
|  |  | 15.3067 | 81.39 |  |  |  |  | 683.3727 | 0.011688 | 15.3067 | 81.845 |  |  |  |  | 683.3727 | 0.15037
|  |  | 15.3814 | 81.36 |  |  |  |  | 686.6424 | 0.013882 | 15.3814 | 81.96 |  |  |  |  | 686.6424 | 0.15011
|  |  | 15.4561 | 81.195 |  |  |  |  | 689.9122 | 0.016059 | 15.4561 | 82 |  |  |  |  | 689.9122 | 0.1324
|  |  | 15.5307 | 81.1 |  |  |  |  | 693.1819 | 0.017938 | 15.5307 | 82.47 |  |  |  |  | 693.1819 | 0.097137
|  |  | 15.6054 | 81.14 |  |  |  |  | 696.4516 | 0.017578 | 15.6054 | 81.815 |  |  |  |  | 696.4516 | 0.049442
|  |  | 15.6801 | 81.28 |  |  |  |  | 699.7213 | 0.014195 | 15.6801 | 81.73 |  |  |  |  | 699.7213 | 0.012041
|  |  | 15.7547 | 81.4 |  |  |  |  | 702.9911 | 0.010114 | 15.7547 | 81.95 |  |  |  |  | 702.9911 | 0.049964
|  |  | 15.8294 | 81.335 |  |  |  |  | 706.2608 | 0.010363 | 15.8294 | 81.905 |  |  |  |  | 706.2608 | 0.077867
|  |  | 15.9041 | 81.34 |  |  |  |  | 709.5305 | 0.013019 | 15.9041 | 82.045 |  |  |  |  | 709.5305 | 0.086325
|  |  | 15.9787 | 81.4 |  |  |  |  | 712.8002 | 0.012927 | 15.9787 | 81.58 |  |  |  |  | 712.8002 | 0.078764
|  |  | 16.0534 | 81.39 |  |  |  |  | 716.07 | 0.0088637 | 16.0534 | 81.345 |  |  |  |  | 716.07 | 0.064644
|  |  | 16.1281 | 81.345 |  |  |  |  | 719.3397 | 0.0021835 | 16.1281 | 81.395 |  |  |  |  | 719.3397 | 0.053091
|  |  | 16.2027 | 81.29 |  |  |  |  | 722.6094 | 0.0047663 | 16.2027 | 81.475 |  |  |  |  | 722.6094 | 0.043626
|  |  | 16.2774 | 81.325 |  |  |  |  | 725.8792 | 0.0096454 | 16.2774 | 82.98 |  |  |  |  | 725.8792 | 0.02916
|  |  | 16.3521 | 81.31 |  |  |  |  | 729.1489 | 0.011508 | 16.3521 | 82.82 |  |  |  |  | 729.1489 | 0.0081347
|  |  | 16.4267 | 81.26 |  |  |  |  | 732.4186 | 0.010909 | 16.4267 | 82.44 |  |  |  |  | 732.4186 | 0.01614
|  |  | 16.5014 | 81.315 |  |  |  |  | 735.6883 | 0.0097711 | 16.5014 | 81.715 |  |  |  |  | 735.6883 | 0.031934
|  |  | 16.5761 | 81.3 |  |  |  |  | 738.9581 | 0.010548 | 16.5761 | 81.46 |  |  |  |  | 738.9581 | 0.033895
|  |  | 16.6507 | 81.245 |  |  |  |  | 742.2278 | 0.013823 | 16.6507 | 81.58 |  |  |  |  | 742.2278 | 0.02336
|  |  | 16.7254 | 81.315 |  |  |  |  | 745.4975 | 0.017975 | 16.7254 | 80.975 |  |  |  |  | 745.4975 | 0.028497
|  |  | 16.8001 | 81.395 |  |  |  |  | 748.7672 | 0.021008 | 16.8001 | 81.63 |  |  |  |  | 748.7672 | 0.057816
|  |  | 16.8747 | 81.395 |  |  |  |  | 752.037 | 0.021228 | 16.8747 | 82.135 |  |  |  |  | 752.037 | 0.085041
|  |  | 16.9494 | 81.39 |  |  |  |  | 755.3067 | 0.017805 | 16.9494 | 82.575 |  |  |  |  | 755.3067 | 0.099323
|  |  | 17.0241 | 81.44 |  |  |  |  | 758.5764 | 0.011211 | 17.0241 | 82.33 |  |  |  |  | 758.5764 | 0.098189
|  |  | 17.0987 | 81.43 |  |  |  |  | 761.8461 | 0.0032039 | 17.0987 | 82.165 |  |  |  |  | 761.8461 | 0.086649
|  |  | 17.1734 | 81.37 |  |  |  |  | 765.1159 | 0.0048152 | 17.1734 | 81.815 |  |  |  |  | 765.1159 | 0.074547
|  |  | 17.2481 | 81.28 |  |  |  |  | 768.3856 | 0.010079 | 17.2481 | 81.07 |  |  |  |  | 768.3856 | 0.068008
|  |  | 17.3227 | 81.29 |  |  |  |  | 771.6553 | 0.012407 | 17.3227 | 81.21 |  |  |  |  | 771.6553 | 0.062057
|  |  | 17.3974 | 81.4 |  |  |  |  | 774.925 | 0.012076 | 17.3974 | 81.2 |  |  |  |  | 774.925 | 0.048896
|  |  | 17.4721 | 81.5 |  |  |  |  | 778.1948 | 0.0099875 | 17.4721 | 81.365 |  |  |  |  | 778.1948 | 0.028578
|  |  | 17.5467 | 81.46 |  |  |  |  | 781.4645 | 0.0071509 | 17.5467 | 81.62 |  |  |  |  | 781.4645 | 0.020964
|  |  | 17.6214 | 81.295 |  |  |  |  | 784.7342 | 0.0050374 | 17.6214 | 81.975 |  |  |  |  | 784.7342 | 0.039891
|  |  | 17.6961 | 81.28 |  |  |  |  | 788.0039 | 0.0056718 | 17.6961 | 82.065 |  |  |  |  | 788.0039 | 0.05578
|  |  | 17.7707 | 81.435 |  |  |  |  | 791.2737 | 0.0075575 | 17.7707 | 81.795 |  |  |  |  | 791.2737 | 0.059938
|  |  | 17.8454 | 81.555 |  |  |  |  | 794.5434 | 0.0083393 | 17.8454 | 81.62 |  |  |  |  | 794.5434 | 0.052711
|  |  | 17.9201 | 81.515 |  |  |  |  | 797.8131 | 0.0074646 | 17.9201 | 81.91 |  |  |  |  | 797.8131 | 0.039104
|  |  | 17.9947 | 81.43 |  |  |  |  | 801.0828 | 0.0059171 | 17.9947 | 82.285 |  |  |  |  | 801.0828 | 0.025875
|  |  | 18.0694 | 81.33 |  |  |  |  | 804.3526 | 0.0054543 | 18.0694 | 82.475 |  |  |  |  | 804.3526 | 0.016446
|  |  | 18.1441 | 81.325 |  |  |  |  | 807.6223 | 0.0056732 | 18.1441 | 82.12 |  |  |  |  | 807.6223 | 0.007491
|  |  | 18.2187 | 81.37 |  |  |  |  | 810.892 | 0.0045952 | 18.2187 | 81.5 |  |  |  |  | 810.892 | 0.0076324
|  |  | 18.2934 | 81.49 |  |  |  |  | 814.1618 | 0.0015645 | 18.2934 | 81.425 |  |  |  |  | 814.1618 | 0.022129
|  |  | 18.3681 | 81.49 |  |  |  |  | 817.4315 | 0.0027849 | 18.3681 | 81.925 |  |  |  |  | 817.4315 | 0.033859
|  |  | 18.4427 | 81.37 |  |  |  |  | 820.7012 | 0.0070142 | 18.4427 | 82.475 |  |  |  |  | 820.7012 | 0.037689
|  |  | 18.5174 | 81.145 |  |  |  |  | 823.9709 | 0.0098092 | 18.5174 | 82.38 |  |  |  |  | 823.9709 | 0.032232
|  |  | 18.5921 | 81.17 |  |  |  |  | 827.2407 | 0.010369 | 18.5921 | 81.96 |  |  |  |  | 827.2407 | 0.020641
|  |  | 18.6667 | 81.28 |  |  |  |  | 830.5104 | 0.0086786 | 18.6667 | 82.02 |  |  |  |  | 830.5104 | 0.010148
|  |  | 18.7414 | 81.355 |  |  |  |  | 833.7801 | 0.0053801 | 18.7414 | 82.035 |  |  |  |  | 833.7801 | 0.010371
|  |  | 18.8161 | 81.375 |  |  |  |  | 837.0498 | 0.0014411 | 18.8161 | 81.825 |  |  |  |  | 837.0498 | 0.021396
|  |  | 18.8908 | 81.345 |  |  |  |  | 840.3196 | 0.0022096 | 18.8908 | 81.96 |  |  |  |  | 840.3196 | 0.042697
|  |  | 18.9654 | 81.29 |  |  |  |  | 843.5893 | 0.0049298 | 18.9654 | 82.105 |  |  |  |  | 843.5893 | 0.069149
|  |  | 19.0401 | 81.255 |  |  |  |  | 846.859 | 0.0064941 | 19.0401 | 82.275 |  |  |  |  | 846.859 | 0.090477
|  |  | 19.1148 | 81.28 |  |  |  |  | 850.1287 | 0.0070816 | 19.1148 | 82.505 |  |  |  |  | 850.1287 | 0.096596
|  |  | 19.1894 | 81.35 |  |  |  |  | 853.3985 | 0.0070674 | 19.1894 | 82.295 |  |  |  |  | 853.3985 | 0.082702
|  |  | 19.2641 | 81.31 |  |  |  |  | 856.6682 | 0.0066188 | 19.2641 | 81.97 |  |  |  |  | 856.6682 | 0.054885
|  |  | 19.3388 | 81.31 |  |  |  |  | 859.9379 | 0.0055227 | 19.3388 | 82.65 |  |  |  |  | 859.9379 | 0.044024
|  |  | 19.4134 | 81.44 |  |  |  |  | 863.2076 | 0.0035624 | 19.4134 | 82.395 |  |  |  |  | 863.2076 | 0.069527
|  |  | 19.4881 | 81.425 |  |  |  |  | 866.4774 | 0.0010899 | 19.4881 | 82.275 |  |  |  |  | 866.4774 | 0.091362
|  |  | 19.5628 | 81.445 |  |  |  |  | 869.7471 | 0.0020457 | 19.5628 | 82.245 |  |  |  |  | 869.7471 | 0.091599
|  |  | 19.6374 | 81.435 |  |  |  |  | 873.0168 | 0.004192 | 19.6374 | 81.96 |  |  |  |  | 873.0168 | 0.070047
|  |  | 19.7121 | 81.27 |  |  |  |  | 876.2865 | 0.006108 | 19.7121 | 81.765 |  |  |  |  | 876.2865 | 0.038734
|  |  | 19.7868 | 81.17 |  |  |  |  | 879.5563 | 0.0087491 | 19.7868 | 82 |  |  |  |  | 879.5563 | 0.025494
|  |  | 19.8614 | 81.26 |  |  |  |  | 882.826 | 0.012238 | 19.8614 | 82.12 |  |  |  |  | 882.826 | 0.032844
|  |  | 19.9361 | 81.195 |  |  |  |  | 886.0957 | 0.015274 | 19.9361 | 82.295 |  |  |  |  | 886.0957 | 0.024787
|  |  | 20.0108 | 81.135 |  |  |  |  | 889.3654 | 0.016258 | 20.0108 | 82.485 |  |  |  |  | 889.3654 | 0.00553
|  |  | 20.0854 | 81.345 |  |  |  |  | 892.6352 | 0.014246 | 20.0854 | 82.305 |  |  |  |  | 892.6352 | 0.045708
|  |  | 20.1601 | 81.45 |  |  |  |  | 895.9049 | 0.0094211 | 20.1601 | 81.855 |  |  |  |  | 895.9049 | 0.084099
|  |  | 20.2348 | 81.475 |  |  |  |  | 899.1746 | 0.0034375 | 20.2348 | 82.005 |  |  |  |  | 899.1746 | 0.1063
|  |  | 20.3094 | 81.305 |  |  |  |  | 902.4444 | 0.0046131 | 20.3094 | 82.26 |  |  |  |  | 902.4444 | 0.10502
|  |  | 20.3841 | 81.185 |  |  |  |  | 905.7141 | 0.0088935 | 20.3841 | 81.795 |  |  |  |  | 905.7141 | 0.081866
|  |  | 20.4588 | 81.15 |  |  |  |  | 908.9838 | 0.010544 | 20.4588 | 81.505 |  |  |  |  | 908.9838 | 0.046816
|  |  | 20.5334 | 81.13 |  |  |  |  | 912.2535 | 0.0093945 | 20.5334 | 81.61 |  |  |  |  | 912.2535 | 0.021631
|  |  | 20.6081 | 81.21 |  |  |  |  | 915.5233 | 0.0064738 | 20.6081 | 81.395 |  |  |  |  | 915.5233 | 0.035381
|  |  | 20.6828 | 81.235 |  |  |  |  | 918.793 | 0.0036232 | 20.6828 | 81.38 |  |  |  |  | 918.793 | 0.051059
|  |  | 20.7574 | 81.11 |  |  |  |  | 922.0627 | 0.0030206 | 20.7574 | 82.47 |  |  |  |  | 922.0627 | 0.057308
|  |  | 20.8321 | 80.965 |  |  |  |  | 925.3324 | 0.0038002 | 20.8321 | 82.715 |  |  |  |  | 925.3324 | 0.055104
|  |  | 20.9068 | 81.085 |  |  |  |  | 928.6022 | 0.0048318 | 20.9068 | 82.795 |  |  |  |  | 928.6022 | 0.044233
|  |  | 20.9814 | 81.25 |  |  |  |  | 931.8719 | 0.0065654 | 20.9814 | 82.28 |  |  |  |  | 931.8719 | 0.024933
|  |  | 21.0561 | 81.18 |  |  |  |  | 935.1416 | 0.0084534 | 21.0561 | 81.97 |  |  |  |  | 935.1416 | 0.014531
|  |  | 21.1308 | 81.105 |  |  |  |  | 938.4113 | 0.0094984 | 21.1308 | 82.13 |  |  |  |  | 938.4113 | 0.035468
|  |  | 21.2054 | 81.18 |  |  |  |  | 941.6811 | 0.009391 | 21.2054 | 82.57 |  |  |  |  | 941.6811 | 0.050033
|  |  | 21.2801 | 81.345 |  |  |  |  | 944.9508 | 0.0088878 | 21.2801 | 81.995 |  |  |  |  | 944.9508 | 0.046917
|  |  | 21.3548 | 81.27 |  |  |  |  | 948.2205 | 0.0091842 | 21.3548 | 81.625 |  |  |  |  | 948.2205 | 0.025671
|  |  | 21.4294 | 81.23 |  |  |  |  | 951.4902 | 0.0102 | 21.4294 | 81 |  |  |  |  | 951.4902 | 0.013482
|  |  | 21.5041 | 81.305 |  |  |  |  | 954.76 | 0.010712 | 21.5041 | 80.91 |  |  |  |  | 954.76 | 0.043252
|  |  | 21.5788 | 81.31 |  |  |  |  | 958.0297 | 0.0099808 | 21.5788 | 80.91 |  |  |  |  | 958.0297 | 0.060739
|  |  | 21.6534 | 81.265 |  |  |  |  | 961.2994 | 0.0081888 | 21.6534 | 81.535 |  |  |  |  | 961.2994 | 0.055178
|  |  | 21.7281 | 81.245 |  |  |  |  | 964.5691 | 0.0060898 | 21.7281 | 82.28 |  |  |  |  | 964.5691 | 0.026409
|  |  | 21.8028 | 81.31 |  |  |  |  | 967.8389 | 0.0047976 | 21.8028 | 82.825 |  |  |  |  | 967.8389 | 0.018573
|  |  | 21.8774 | 81.39 |  |  |  |  | 971.1086 | 0.0059532 | 21.8774 | 82.535 |  |  |  |  | 971.1086 | 0.062298
|  |  | 21.9521 | 81.41 |  |  |  |  | 974.3783 | 0.0093168 | 21.9521 | 82.21 |  |  |  |  | 974.3783 | 0.09213
|  |  | 22.0268 | 81.345 |  |  |  |  | 977.648 | 0.012951 | 22.0268 | 81.945 |  |  |  |  | 977.648 | 0.098012
|  |  | 22.1014 | 81.255 |  |  |  |  | 980.9178 | 0.015024 | 22.1014 | 81.97 |  |  |  |  | 980.9178 | 0.07863
|  |  | 22.1761 | 81.315 |  |  |  |  | 984.1875 | 0.014208 | 22.1761 | 82.575 |  |  |  |  | 984.1875 | 0.04174
|  |  | 22.2508 | 81.38 |  |  |  |  | 987.4572 | 0.010178 | 22.2508 | 82.585 |  |  |  |  | 987.4572 | 0.017953
|  |  | 22.3254 | 81.385 |  |  |  |  | 990.727 | 0.0038719 | 22.3254 | 82.33 |  |  |  |  | 990.727 | 0.049356
|  |  | 22.4001 | 81.42 |  |  |  |  | 993.9967 | 0.002805 | 22.4001 | 81.795 |  |  |  |  | 993.9967 | 0.073359
|  |  | 22.4748 | 81.315 |  |  |  |  | 997.2664 | 0.0075679 | 22.4748 | 81.425 |  |  |  |  | 997.2664 | 0.080909
|  |  | 22.5494 | 81.16 |  |  |  |  | 1000.5361 | 0.0088365 | 22.5494 | 81.875 |  |  |  |  | 1000.5361 | 0.074532
|  |  | 22.6241 | 81.165 |  |  |  |  | 1003.8059 | 0.0064 | 22.6241 | 82.165 |  |  |  |  | 1003.8059 | 0.060357
|  |  | 22.6988 | 81.255 |  |  |  |  | 1007.0756 | 0.0034104 | 22.6988 | 82.49 |  |  |  |  | 1007.0756 | 0.044693
|  |  | 22.7734 | 81.27 |  |  |  |  | 1010.3453 | 0.0076274 | 22.7734 | 81.745 |  |  |  |  | 1010.3453 | 0.031708
|  |  | 22.8481 | 81.185 |  |  |  |  | 1013.615 | 0.012358 | 22.8481 | 81.6 |  |  |  |  | 1013.615 | 0.022588
|  |  | 22.9228 | 81.25 |  |  |  |  | 1016.8848 | 0.014161 | 22.9228 | 81.775 |  |  |  |  | 1016.8848 | 0.016431
|  |  | 22.9974 | 81.285 |  |  |  |  | 1020.1545 | 0.01243 | 22.9974 | 81.85 |  |  |  |  | 1020.1545 | 0.015141
|  |  | 23.0721 | 81.35 |  |  |  |  | 1023.4242 | 0.0080206 | 23.0721 | 81.63 |  |  |  |  | 1023.4242 | 0.023947
|  |  | 23.1468 | 81.33 |  |  |  |  | 1026.6939 | 0.0030287 | 23.1468 | 81.72 |  |  |  |  | 1026.6939 | 0.038756
|  |  | 23.2214 | 81.225 |  |  |  |  | 1029.9637 | 0.0036532 | 23.2214 | 81.75 |  |  |  |  | 1029.9637 | 0.053443
|  |  | 23.2961 | 81.15 |  |  |  |  | 1033.2334 | 0.0065566 | 23.2961 | 82.83 |  |  |  |  | 1033.2334 | 0.063535
|  |  | 23.3708 | 81.19 |  |  |  |  | 1036.5031 | 0.0078341 | 23.3708 | 83.145 |  |  |  |  | 1036.5031 | 0.066281
|  |  | 23.4454 | 81.24 |  |  |  |  | 1039.7728 | 0.0078871 | 23.4454 | 83.24 |  |  |  |  | 1039.7728 | 0.060956
|  |  | 23.5201 | 81.175 |  |  |  |  | 1043.0426 | 0.0076242 | 23.5201 | 81.71 |  |  |  |  | 1043.0426 | 0.04897
|  |  | 23.5948 | 81.145 |  |  |  |  | 1046.3123 | 0.007809 | 23.5948 | 82.175 |  |  |  |  | 1046.3123 | 0.034703
|  |  | 23.6694 | 81.16 |  |  |  |  | 1049.582 | 0.0087107 | 23.6694 | 82.15 |  |  |  |  | 1049.582 | 0.02731
|  |  | 23.7441 | 81.305 |  |  |  |  | 1052.8517 | 0.0099612 | 23.7441 | 80.78 |  |  |  |  | 1052.8517 | 0.030331
|  |  | 23.8188 | 81.335 |  |  |  |  | 1056.1215 | 0.010881 | 23.8188 | 80.81 |  |  |  |  | 1056.1215 | 0.030488
|  |  | 23.8934 | 81.25 |  |  |  |  | 1059.3912 | 0.011052 | 23.8934 | 82.085 |  |  |  |  | 1059.3912 | 0.018019
|  |  | 23.9681 | 81.12 |  |  |  |  | 1062.6609 | 0.010556 | 23.9681 | 82.65 |  |  |  |  | 1062.6609 | 0.011246
|  |  | 24.0428 | 81.025 |  |  |  |  | 1065.9306 | 0.009597 | 24.0428 | 82.21 |  |  |  |  | 1065.9306 | 0.046485
|  |  | 24.1174 | 81.115 |  |  |  |  | 1069.2004 | 0.0080439 | 24.1174 | 81.695 |  |  |  |  | 1069.2004 | 0.08081
|  |  | 24.1921 | 81.225 |  |  |  |  | 1072.4701 | 0.0064263 | 24.1921 | 82.06 |  |  |  |  | 1072.4701 | 0.10112
|  |  | 24.2668 | 81.245 |  |  |  |  | 1075.7398 | 0.0081942 | 24.2668 | 82.26 |  |  |  |  | 1075.7398 | 0.098042
|  |  | 24.3414 | 81.265 |  |  |  |  | 1079.0096 | 0.013718 | 24.3414 | 82.39 |  |  |  |  | 1079.0096 | 0.0692
|  |  | 24.4161 | 81.27 |  |  |  |  | 1082.2793 | 0.019337 | 24.4161 | 82.355 |  |  |  |  | 1082.2793 | 0.02111
|  |  | 24.4908 | 81.295 |  |  |  |  | 1085.549 | 0.022372 | 24.4908 | 82.635 |  |  |  |  | 1085.549 | 0.035434
|  |  | 24.5654 | 81.33 |  |  |  |  | 1088.8187 | 0.021344 | 24.5654 | 82.825 |  |  |  |  | 1088.8187 | 0.081978
|  |  | 24.6401 | 81.33 |  |  |  |  | 1092.0885 | 0.016218 | 24.6401 | 82.57 |  |  |  |  | 1092.0885 | 0.1079
|  |  | 24.7148 | 81.295 |  |  |  |  | 1095.3582 | 0.0083932 | 24.7148 | 81.665 |  |  |  |  | 1095.3582 | 0.10822
|  |  | 24.7894 | 81.25 |  |  |  |  | 1098.6279 | 0.00067556 | 24.7894 | 81.385 |  |  |  |  | 1098.6279 | 0.086872
|  |  | 24.8641 | 81.24 |  |  |  |  | 1101.8976 | 0.0063153 | 24.8641 | 81.44 |  |  |  |  | 1101.8976 | 0.05446
|  |  | 24.9388 | 81.17 |  |  |  |  | 1105.1674 | 0.0094148 | 24.9388 | 81.19 |  |  |  |  | 1105.1674 | 0.023847
|  |  | 25.0134 | 81.175 |  |  |  |  | 1108.4371 | 0.0092252 | 25.0134 | 81.83 |  |  |  |  | 1108.4371 | 0.014876
|  |  | 25.0881 | 81.2 |  |  |  |  | 1111.7068 | 0.0070215 | 25.0881 | 82.08 |  |  |  |  | 1111.7068 | 0.032058
|  |  | 25.1628 | 81.295 |  |  |  |  | 1114.9765 | 0.004716 | 25.1628 | 82.215 |  |  |  |  | 1114.9765 | 0.05348
|  |  | 25.2374 | 81.405 |  |  |  |  | 1118.2463 | 0.003488 | 25.2374 | 82.49 |  |  |  |  | 1118.2463 | 0.073773
|  |  | 25.3121 | 81.42 |  |  |  |  | 1121.516 | 0.0031555 | 25.3121 | 82.7 |  |  |  |  | 1121.516 | 0.085887
|  |  | 25.3868 | 81.29 |  |  |  |  | 1124.7857 | 0.004244 | 25.3868 | 82.2 |  |  |  |  | 1124.7857 | 0.085704
|  |  | 25.4614 | 81.325 |  |  |  |  | 1128.0554 | 0.0061616 | 25.4614 | 81.645 |  |  |  |  | 1128.0554 | 0.07607
|  |  | 25.5361 | 81.34 |  |  |  |  | 1131.3252 | 0.0071944 | 25.5361 | 81.33 |  |  |  |  | 1131.3252 | 0.065493
|  |  | 25.6108 | 81.255 |  |  |  |  | 1134.5949 | 0.006487 | 25.6108 | 81.615 |  |  |  |  | 1134.5949 | 0.05895
|  |  | 25.6854 | 81.29 |  |  |  |  | 1137.8646 | 0.004542 | 25.6854 | 82.05 |  |  |  |  | 1137.8646 | 0.050473
|  |  | 25.7601 | 81.32 |  |  |  |  | 1141.1343 | 0.0034738 | 25.7601 | 82.1 |  |  |  |  | 1141.1343 | 0.032668
|  |  | 25.8348 | 81.3 |  |  |  |  | 1144.4041 | 0.0042088 | 25.8348 | 81.74 |  |  |  |  | 1144.4041 | 0.0063417
|  |  | 25.9094 | 81.36 |  |  |  |  | 1147.6738 | 0.0045467 | 25.9094 | 81.41 |  |  |  |  | 1147.6738 | 0.022818
|  |  | 25.9841 | 81.42 |  |  |  |  | 1150.9435 | 0.0047648 | 25.9841 | 81.455 |  |  |  |  | 1150.9435 | 0.04333
|  |  | 26.0588 | 81.405 |  |  |  |  | 1154.2132 | 0.0066955 | 26.0588 | 81.31 |  |  |  |  | 1154.2132 | 0.049714
|  |  | 26.1334 | 81.37 |  |  |  |  | 1157.483 | 0.0091844 | 26.1334 | 81.845 |  |  |  |  | 1157.483 | 0.042847
|  |  | 26.2081 | 81.255 |  |  |  |  | 1160.7527 | 0.010113 | 26.2081 | 82.17 |  |  |  |  | 1160.7527 | 0.034419
|  |  | 26.2828 | 81.275 |  |  |  |  | 1164.0224 | 0.008625 | 26.2828 | 82.565 |  |  |  |  | 1164.0224 | 0.039811
|  |  | 26.3575 | 81.405 |  |  |  |  | 1167.2922 | 0.0058237 | 26.3575 | 82.665 |  |  |  |  | 1167.2922 | 0.049981
|  |  | 26.4321 | 81.465 |  |  |  |  | 1170.5619 | 0.0059695 | 26.4321 | 82.535 |  |  |  |  | 1170.5619 | 0.051479
|  |  | 26.5068 | 81.42 |  |  |  |  | 1173.8316 | 0.0092754 | 26.5068 | 82.15 |  |  |  |  | 1173.8316 | 0.041899
|  |  | 26.5815 | 81.325 |  |  |  |  | 1177.1013 | 0.011641 | 26.5815 | 81.88 |  |  |  |  | 1177.1013 | 0.026254
|  |  | 26.6561 | 81.255 |  |  |  |  | 1180.3711 | 0.011808 | 26.6561 | 81.83 |  |  |  |  | 1180.3711 | 0.015422
|  |  | 26.7308 | 81.23 |  |  |  |  | 1183.6408 | 0.010111 | 26.7308 | 81.44 |  |  |  |  | 1183.6408 | 0.019837
|  |  | 26.8055 | 81.32 |  |  |  |  | 1186.9105 | 0.0075287 | 26.8055 | 81.135 |  |  |  |  | 1186.9105 | 0.028674
|  |  | 26.8801 | 81.355 |  |  |  |  | 1190.1802 | 0.0048295 | 26.8801 | 82.02 |  |  |  |  | 1190.1802 | 0.036105
|  |  | 26.9548 | 81.37 |  |  |  |  | 1193.45 | 0.0021643 | 26.9548 | 81.905 |  |  |  |  | 1193.45 | 0.039023
|  |  | 27.0295 | 81.21 |  |  |  |  | 1196.7197 | 0.00068063 | 27.0295 | 81.055 |  |  |  |  | 1196.7197 | 0.033305
|  |  | 27.1041 | 81.18 |  |  |  |  | 1199.9894 | 0.0026803 | 27.1041 | 81.745 |  |  |  |  | 1199.9894 | 0.01851
|  |  | 27.1788 | 81.265 |  |  |  |  | 1203.2591 | 0.0034259 | 27.1788 | 81.84 |  |  |  |  | 1203.2591 | 0.017809
|  |  | 27.2535 | 81.395 |  |  |  |  | 1206.5289 | 0.0027694 | 27.2535 | 81.485 |  |  |  |  | 1206.5289 | 0.041433
|  |  | 27.3281 | 81.39 |  |  |  |  | 1209.7986 | 0.0041515 | 27.3281 | 81.745 |  |  |  |  | 1209.7986 | 0.06019
|  |  | 27.4028 | 81.315 |  |  |  |  | 1213.0683 | 0.0086671 | 27.4028 | 82.17 |  |  |  |  | 1213.0683 | 0.06535
|  |  | 27.4775 | 81.225 |  |  |  |  | 1216.338 | 0.013415 | 27.4775 | 82.28 |  |  |  |  | 1216.338 | 0.055509
|  |  | 27.5521 | 81.215 |  |  |  |  | 1219.6078 | 0.016609 | 27.5521 | 82.405 |  |  |  |  | 1219.6078 | 0.03587
|  |  | 27.6268 | 81.27 |  |  |  |  | 1222.8775 | 0.01724 | 27.6268 | 82.36 |  |  |  |  | 1222.8775 | 0.016944
|  |  | 27.7015 | 81.44 |  |  |  |  | 1226.1472 | 0.015206 | 27.7015 | 82.105 |  |  |  |  | 1226.1472 | 0.010773
|  |  | 27.7761 | 81.425 |  |  |  |  | 1229.4169 | 0.011326 | 27.7761 | 81.94 |  |  |  |  | 1229.4169 | 0.014054
|  |  | 27.8508 | 81.345 |  |  |  |  | 1232.6867 | 0.0073501 | 27.8508 | 81.325 |  |  |  |  | 1232.6867 | 0.0325
|  |  | 27.9255 | 81.285 |  |  |  |  | 1235.9564 | 0.0060861 | 27.9255 | 81.315 |  |  |  |  | 1235.9564 | 0.063227
|  |  | 28.0001 | 81.31 |  |  |  |  | 1239.2261 | 0.0077524 | 28.0001 | 82.375 |  |  |  |  | 1239.2261 | 0.09338
|  |  | 28.0748 | 81.26 |  |  |  |  | 1242.4958 | 0.0094131 | 28.0748 | 82.38 |  |  |  |  | 1242.4958 | 0.11082
|  |  | 28.1495 | 81.35 |  |  |  |  | 1245.7656 | 0.0099653 | 28.1495 | 82.285 |  |  |  |  | 1245.7656 | 0.10764
|  |  | 28.2241 | 81.49 |  |  |  |  | 1249.0353 | 0.0093537 | 28.2241 | 81.66 |  |  |  |  | 1249.0353 | 0.082805
|  |  | 28.2988 | 81.585 |  |  |  |  | 1252.305 | 0.0077412 | 28.2988 | 81.425 |  |  |  |  | 1252.305 | 0.042534
|  |  | 28.3735 | 81.455 |  |  |  |  | 1255.5748 | 0.0054146 | 28.3735 | 81.83 |  |  |  |  | 1255.5748 | 0.0054077
|  |  | 28.4481 | 81.33 |  |  |  |  | 1258.8445 | 0.0036494 | 28.4481 | 81.895 |  |  |  |  | 1258.8445 | 0.040888
|  |  | 28.5228 | 81.295 |  |  |  |  | 1262.1142 | 0.0053847 | 28.5228 | 81.87 |  |  |  |  | 1262.1142 | 0.063662
|  |  | 28.5975 | 81.34 |  |  |  |  | 1265.3839 | 0.0088728 | 28.5975 | 82.42 |  |  |  |  | 1265.3839 | 0.070033
|  |  | 28.6721 | 81.34 |  |  |  |  | 1268.6537 | 0.011815 | 28.6721 | 82.75 |  |  |  |  | 1268.6537 | 0.06509
|  |  | 28.7468 | 81.28 |  |  |  |  | 1271.9234 | 0.013165 | 28.7468 | 82.13 |  |  |  |  | 1271.9234 | 0.056355
|  |  | 28.8215 | 81.265 |  |  |  |  | 1275.1931 | 0.012471 | 28.8215 | 81.605 |  |  |  |  | 1275.1931 | 0.046339
|  |  | 28.8961 | 81.265 |  |  |  |  | 1278.4628 | 0.0099085 | 28.8961 | 81.195 |  |  |  |  | 1278.4628 | 0.031419
|  |  | 28.9708 | 81.24 |  |  |  |  | 1281.7326 | 0.006335 | 28.9708 | 82.155 |  |  |  |  | 1281.7326 | 0.016427
|  |  | 29.0455 | 81.24 |  |  |  |  | 1285.0023 | 0.0034643 | 29.0455 | 82.195 |  |  |  |  | 1285.0023 | 0.035432
|  |  | 29.1201 | 81.205 |  |  |  |  | 1288.272 | 0.0033991 | 29.1201 | 82.5 |  |  |  |  | 1288.272 | 0.063786
|  |  | 29.1948 | 81.225 |  |  |  |  | 1291.5417 | 0.0038606 | 29.1948 | 81.99 |  |  |  |  | 1291.5417 | 0.081856
|  |  | 29.2695 | 81.18 |  |  |  |  | 1294.8115 | 0.0028118 | 29.2695 | 81.985 |  |  |  |  | 1294.8115 | 0.082036
|  |  | 29.3441 | 81.125 |  |  |  |  | 1298.0812 | 0.00034144 | 29.3441 | 81.94 |  |  |  |  | 1298.0812 | 0.063357
|  |  | 29.4188 | 81.145 |  |  |  |  | 1301.3509 | 0.0028791 | 29.4188 | 82.33 |  |  |  |  | 1301.3509 | 0.031365
|  |  | 29.4935 | 81.195 |  |  |  |  | 1304.6206 | 0.0058408 | 29.4935 | 81.935 |  |  |  |  | 1304.6206 | 0.0049908
|  |  | 29.5681 | 81.2 |  |  |  |  | 1307.8904 | 0.007901 | 29.5681 | 81.785 |  |  |  |  | 1307.8904 | 0.032587
|  |  | 29.6428 | 81.17 |  |  |  |  | 1311.1601 | 0.0086804 | 29.6428 | 81.99 |  |  |  |  | 1311.1601 | 0.046052
|  |  | 29.7175 | 81.17 |  |  |  |  | 1314.4298 | 0.0079713 | 29.7175 | 81.795 |  |  |  |  | 1314.4298 | 0.043104
|  |  | 29.7921 | 81.185 |  |  |  |  | 1317.6995 | 0.005665 | 29.7921 | 82.165 |  |  |  |  | 1317.6995 | 0.028564
|  |  | 29.8668 | 81.195 |  |  |  |  | 1320.9693 | 0.0024487 | 29.8668 | 82.07 |  |  |  |  | 1320.9693 | 0.014638
|  |  | 29.9415 | 81.225 |  |  |  |  | 1324.239 | 0.0045636 | 29.9415 | 81.91 |  |  |  |  | 1324.239 | 0.017608
|  |  | 30.0161 | 81.265 |  |  |  |  | 1327.5087 | 0.00925 | 30.0161 | 81.895 |  |  |  |  | 1327.5087 | 0.021366
|  |  | 30.0908 | 81.23 |  |  |  |  | 1330.7784 | 0.012346 | 30.0908 | 81.425 |  |  |  |  | 1330.7784 | 0.01754
|  |  | 30.1655 | 81.175 |  |  |  |  | 1334.0482 | 0.012301 | 30.1655 | 81.99 |  |  |  |  | 1334.0482 | 0.013989
|  |  | 30.2401 | 81.195 |  |  |  |  | 1337.3179 | 0.0086114 | 30.2401 | 81.92 |  |  |  |  | 1337.3179 | 0.018231
|  |  | 30.3148 | 81.26 |  |  |  |  | 1340.5876 | 0.0022627 | 30.3148 | 81.725 |  |  |  |  | 1340.5876 | 0.019225
|  |  | 30.3895 | 81.25 |  |  |  |  | 1343.8574 | 0.0059338 | 30.3895 | 81.58 |  |  |  |  | 1343.8574 | 0.011555
|  |  | 30.4641 | 81.235 |  |  |  |  | 1347.1271 | 0.012353 | 30.4641 | 81.385 |  |  |  |  | 1347.1271 | 0.017004
|  |  | 30.5388 | 81.165 |  |  |  |  | 1350.3968 | 0.015882 | 30.5388 | 81.135 |  |  |  |  | 1350.3968 | 0.040147
|  |  | 30.6135 | 81.155 |  |  |  |  | 1353.6665 | 0.016118 | 30.6135 | 81.145 |  |  |  |  | 1353.6665 | 0.060563
|  |  | 30.6881 | 81.21 |  |  |  |  | 1356.9363 | 0.014441 | 30.6881 | 81.68 |  |  |  |  | 1356.9363 | 0.069635
|  |  | 30.7628 | 81.195 |  |  |  |  | 1360.206 | 0.013214 | 30.7628 | 81.485 |  |  |  |  | 1360.206 | 0.063478
|  |  | 30.8375 | 81.135 |  |  |  |  | 1363.4757 | 0.013136 | 30.8375 | 81.08 |  |  |  |  | 1363.4757 | 0.043399
|  |  | 30.9121 | 81.155 |  |  |  |  | 1366.7454 | 0.012551 | 30.9121 | 81.11 |  |  |  |  | 1366.7454 | 0.015327
|  |  | 30.9868 | 81.195 |  |  |  |  | 1370.0152 | 0.01054 | 30.9868 | 81.315 |  |  |  |  | 1370.0152 | 0.015169
|  |  | 31.0615 | 81.2 |  |  |  |  | 1373.2849 | 0.0083781 | 31.0615 | 81.55 |  |  |  |  | 1373.2849 | 0.037795
|  |  | 31.1361 | 81.22 |  |  |  |  | 1376.5546 | 0.0079572 | 31.1361 | 81.29 |  |  |  |  | 1376.5546 | 0.050773
|  |  | 31.2108 | 81.255 |  |  |  |  | 1379.8243 | 0.0081554 | 31.2108 | 81.595 |  |  |  |  | 1379.8243 | 0.053765
|  |  | 31.2855 | 81.34 |  |  |  |  | 1383.0941 | 0.0070336 | 31.2855 | 81.615 |  |  |  |  | 1383.0941 | 0.049557
|  |  | 31.3601 | 81.35 |  |  |  |  | 1386.3638 | 0.0055591 | 31.3601 | 81.485 |  |  |  |  | 1386.3638 | 0.041905
|  |  | 31.4348 | 81.29 |  |  |  |  | 1389.6335 | 0.0071079 | 31.4348 | 81.765 |  |  |  |  | 1389.6335 | 0.033591
|  |  | 31.5095 | 81.22 |  |  |  |  | 1392.9032 | 0.010174 | 31.5095 | 81.915 |  |  |  |  | 1392.9032 | 0.025517
|  |  | 31.5841 | 81.085 |  |  |  |  | 1396.173 | 0.011873 | 31.5841 | 81.625 |  |  |  |  | 1396.173 | 0.017086
|  |  | 31.6588 | 81.03 |  |  |  |  | 1399.4427 | 0.011474 | 31.6588 | 81.885 |  |  |  |  | 1399.4427 | 0.0079735
|  |  | 31.7335 | 81.01 |  |  |  |  | 1402.7124 | 0.0095838 | 31.7335 | 81.96 |  |  |  |  | 1402.7124 | 0.0099406
|  |  | 31.8081 | 81.045 |  |  |  |  | 1405.9821 | 0.0074337 | 31.8081 | 81.795 |  |  |  |  | 1405.9821 | 0.023563
|  |  | 31.8828 | 81.095 |  |  |  |  | 1409.2519 | 0.0067455 | 31.8828 | 82.01 |  |  |  |  | 1409.2519 | 0.036539
|  |  | 31.9575 | 81.265 |  |  |  |  | 1412.5216 | 0.0090249 | 31.9575 | 81.66 |  |  |  |  | 1412.5216 | 0.044023
|  |  | 32.0321 | 81.42 |  |  |  |  | 1415.7913 | 0.013037 | 32.0321 | 81.98 |  |  |  |  | 1415.7913 | 0.042666
|  |  | 32.1068 | 81.335 |  |  |  |  | 1419.061 | 0.016511 | 32.1068 | 82.29 |  |  |  |  | 1419.061 | 0.032101
|  |  | 32.1815 | 81.275 |  |  |  |  | 1422.3308 | 0.017667 | 32.1815 | 81.5 |  |  |  |  | 1422.3308 | 0.017989
|  |  | 32.2561 | 81.3 |  |  |  |  | 1425.6005 | 0.015691 | 32.2561 | 81.81 |  |  |  |  | 1425.6005 | 0.020505
|  |  | 32.3308 | 81.265 |  |  |  |  | 1428.8702 | 0.011062 | 32.3308 | 82.305 |  |  |  |  | 1428.8702 | 0.034293
|  |  | 32.4055 | 81.275 |  |  |  |  | 1432.14 | 0.0054099 | 32.4055 | 81.29 |  |  |  |  | 1432.14 | 0.041312
|  |  | 32.4801 | 81.33 |  |  |  |  | 1435.4097 | 0.0019056 | 32.4801 | 80.925 |  |  |  |  | 1435.4097 | 0.037646
|  |  | 32.5548 | 81.255 |  |  |  |  | 1438.6794 | 0.0040917 | 32.5548 | 81.465 |  |  |  |  | 1438.6794 | 0.026776
|  |  | 32.6295 | 81.2 |  |  |  |  | 1441.9491 | 0.0067374 | 32.6295 | 81.565 |  |  |  |  | 1441.9491 | 0.023291
|  |  | 32.7041 | 81.26 |  |  |  |  | 1445.2189 | 0.010482 | 32.7041 | 81.39 |  |  |  |  | 1445.2189 | 0.035396
|  |  | 32.7788 | 81.385 |  |  |  |  | 1448.4886 | 0.014916 | 32.7788 | 81.165 |  |  |  |  | 1448.4886 | 0.046985
|  |  | 32.8535 | 81.37 |  |  |  |  | 1451.7583 | 0.018296 | 32.8535 | 81.215 |  |  |  |  | 1451.7583 | 0.049723
|  |  | 32.9281 | 81.305 |  |  |  |  | 1455.028 | 0.019338 | 32.9281 | 81.81 |  |  |  |  | 1455.028 | 0.042175
|  |  | 33.0028 | 81.235 |  |  |  |  | 1458.2978 | 0.017997 | 33.0028 | 81.825 |  |  |  |  | 1458.2978 | 0.02655
|  |  | 33.0775 | 81.175 |  |  |  |  | 1461.5675 | 0.015474 | 33.0775 | 81.825 |  |  |  |  | 1461.5675 | 0.0088257
|  |  | 33.1521 | 81.155 |  |  |  |  | 1464.8372 | 0.013459 | 33.1521 | 81.605 |  |  |  |  | 1464.8372 | 0.019766
|  |  | 33.2268 | 81.225 |  |  |  |  | 1468.1069 | 0.012646 | 33.2268 | 81.985 |  |  |  |  | 1468.1069 | 0.041174
|  |  | 33.3015 | 81.28 |  |  |  |  | 1471.3767 | 0.012291 | 33.3015 | 82.26 |  |  |  |  | 1471.3767 | 0.060683
|  |  | 33.3761 | 81.345 |  |  |  |  | 1474.6464 | 0.011728 | 33.3761 | 81.83 |  |  |  |  | 1474.6464 | 0.075049
|  |  | 33.4508 | 81.345 |  |  |  |  | 1477.9161 | 0.011041 | 33.4508 | 82.065 |  |  |  |  | 1477.9161 | 0.080734
|  |  | 33.5255 | 81.26 |  |  |  |  | 1481.1858 | 0.010372 | 33.5255 | 81.925 |  |  |  |  | 1481.1858 | 0.075125
|  |  | 33.6002 | 81.235 |  |  |  |  | 1484.4556 | 0.0094505 | 33.6002 | 81.59 |  |  |  |  | 1484.4556 | 0.058365
|  |  | 33.6748 | 81.31 |  |  |  |  | 1487.7253 | 0.0079806 | 33.6748 | 81.695 |  |  |  |  | 1487.7253 | 0.034705
|  |  | 33.7495 | 81.295 |  |  |  |  | 1490.995 | 0.0060686 | 33.7495 | 81.94 |  |  |  |  | 1490.995 | 0.01631
|  |  | 33.8242 | 81.26 |  |  |  |  | 1494.2647 | 0.0041351 | 33.8242 | 81.595 |  |  |  |  | 1494.2647 | 0.023354
|  |  | 33.8988 | 81.26 |  |  |  |  | 1497.5345 | 0.0025311 | 33.8988 | 81.73 |  |  |  |  | 1497.5345 | 0.031244
|  |  | 33.9735 | 81.225 |  |  |  |  | 1500.8042 | 0.0013061 | 33.9735 | 81.84 |  |  |  |  | 1500.8042 | 0.028734
|  |  | 34.0482 | 81.195 |  |  |  |  | 1504.0739 | 0.00070215 | 34.0482 | 81.78 |  |  |  |  | 1504.0739 | 0.022471
|  |  | 34.1228 | 81.33 |  |  |  |  | 1507.3436 | 0.0011349 | 34.1228 | 81.92 |  |  |  |  | 1507.3436 | 0.029134
|  |  | 34.1975 | 81.44 |  |  |  |  | 1510.6134 | 0.0013942 | 34.1975 | 81.87 |  |  |  |  | 1510.6134 | 0.043043
|  |  | 34.2722 | 81.405 |  |  |  |  | 1513.8831 | 0.0014914 | 34.2722 | 81.88 |  |  |  |  | 1513.8831 | 0.05053
|  |  | 34.3468 | 81.32 |  |  |  |  | 1517.1528 | 0.0028149 | 34.3468 | 81.975 |  |  |  |  | 1517.1528 | 0.046905
|  |  | 34.4215 | 81.28 |  |  |  |  | 1520.4226 | 0.0053267 | 34.4215 | 81.695 |  |  |  |  | 1520.4226 | 0.033871
|  |  | 34.4962 | 81.275 |  |  |  |  | 1523.6923 | 0.0081513 | 34.4962 | 81.905 |  |  |  |  | 1523.6923 | 0.01822
|  |  | 34.5708 | 81.31 |  |  |  |  | 1526.962 | 0.010505 | 34.5708 | 82.59 |  |  |  |  | 1526.962 | 0.01173
|  |  | 34.6455 | 81.355 |  |  |  |  | 1530.2317 | 0.011716 | 34.6455 | 82.96 |  |  |  |  | 1530.2317 | 0.013028
|  |  | 34.7202 | 81.37 |  |  |  |  | 1533.5015 | 0.011344 | 34.7202 | 82.5 |  |  |  |  | 1533.5015 | 0.0080157
|  |  | 34.7948 | 81.39 |  |  |  |  | 1536.7712 | 0.0093039 | 34.7948 | 82.28 |  |  |  |  | 1536.7712 | 0.0040019
|  |  | 34.8695 | 81.325 |  |  |  |  | 1540.0409 | 0.0059291 | 34.8695 | 82.015 |  |  |  |  | 1540.0409 | 0.01739
|  |  | 34.9442 | 81.34 |  |  |  |  | 1543.3106 | 0.0019647 | 34.9442 | 81.62 |  |  |  |  | 1543.3106 | 0.026038
|  |  | 35.0188 | 81.34 |  |  |  |  | 1546.5804 | 0.0018737 | 35.0188 | 81.94 |  |  |  |  | 1546.5804 | 0.02596
|  |  | 35.0935 | 81.275 |  |  |  |  | 1549.8501 | 0.0043238 | 35.0935 | 82.285 |  |  |  |  | 1549.8501 | 0.0187
|  |  | 35.1682 | 81.25 |  |  |  |  | 1553.1198 | 0.0050873 | 35.1682 | 81.665 |  |  |  |  | 1553.1198 | 0.016113
|  |  | 35.2428 | 81.215 |  |  |  |  | 1556.3895 | 0.0043504 | 35.2428 | 81.195 |  |  |  |  | 1556.3895 | 0.024394
|  |  | 35.3175 | 81.255 |  |  |  |  | 1559.6593 | 0.0035093 | 35.3175 | 81.245 |  |  |  |  | 1559.6593 | 0.028726
|  |  | 35.3922 | 81.345 |  |  |  |  | 1562.929 | 0.0042765 | 35.3922 | 81.13 |  |  |  |  | 1562.929 | 0.022418
|  |  | 35.4668 | 81.385 |  |  |  |  | 1566.1987 | 0.0053679 | 35.4668 | 81.41 |  |  |  |  | 1566.1987 | 0.0091981
|  |  | 35.5415 | 81.375 |  |  |  |  | 1569.4684 | 0.005267 | 35.5415 | 81.255 |  |  |  |  | 1569.4684 | 0.020662
|  |  | 35.6162 | 81.3 |  |  |  |  | 1572.7382 | 0.0038349 | 35.6162 | 81.645 |  |  |  |  | 1572.7382 | 0.038979
|  |  | 35.6908 | 81.225 |  |  |  |  | 1576.0079 | 0.002017 | 35.6908 | 82.68 |  |  |  |  | 1576.0079 | 0.048093
|  |  | 35.7655 | 81.155 |  |  |  |  | 1579.2776 | 0.0014721 | 35.7655 | 82.96 |  |  |  |  | 1579.2776 | 0.045691
|  |  | 35.8402 | 81.205 |  |  |  |  | 1582.5473 | 0.0022225 | 35.8402 | 82.28 |  |  |  |  | 1582.5473 | 0.039453
|  |  | 35.9148 | 81.32 |  |  |  |  | 1585.8171 | 0.0051128 | 35.9148 | 81.565 |  |  |  |  | 1585.8171 | 0.045295
|  |  | 35.9895 | 81.375 |  |  |  |  | 1589.0868 | 0.010279 | 35.9895 | 81.805 |  |  |  |  | 1589.0868 | 0.060802
|  |  | 36.0642 | 81.31 |  |  |  |  | 1592.3565 | 0.016183 | 36.0642 | 82.07 |  |  |  |  | 1592.3565 | 0.070589
|  |  | 36.1388 | 81.215 |  |  |  |  | 1595.6262 | 0.020965 | 36.1388 | 81.83 |  |  |  |  | 1595.6262 | 0.066664
|  |  | 36.2135 | 81.285 |  |  |  |  | 1598.896 | 0.023131 | 36.2135 | 81.4 |  |  |  |  | 1598.896 | 0.049143
|  |  | 36.2882 | 81.4 |  |  |  |  | 1602.1657 | 0.022122 | 36.2882 | 81.665 |  |  |  |  | 1602.1657 | 0.026413
|  |  | 36.3628 | 81.41 |  |  |  |  | 1605.4354 | 0.018512 | 36.3628 | 82.345 |  |  |  |  | 1605.4354 | 0.02441
|  |  | 36.4375 | 81.275 |  |  |  |  | 1608.7052 | 0.013695 | 36.4375 | 82.525 |  |  |  |  | 1608.7052 | 0.041139
|  |  | 36.5122 | 81.215 |  |  |  |  | 1611.9749 | 0.0091655 | 36.5122 | 82.485 |  |  |  |  | 1611.9749 | 0.051305
|  |  | 36.5868 | 81.235 |  |  |  |  | 1615.2446 | 0.0058678 | 36.5868 | 82.305 |  |  |  |  | 1615.2446 | 0.051815
|  |  | 36.6615 | 81.21 |  |  |  |  | 1618.5143 | 0.004823 | 36.6615 | 81.93 |  |  |  |  | 1618.5143 | 0.046986
|  |  | 36.7362 | 81.23 |  |  |  |  | 1621.7841 | 0.0068041 | 36.7362 | 81.495 |  |  |  |  | 1621.7841 | 0.042572
|  |  | 36.8108 | 81.36 |  |  |  |  | 1625.0538 | 0.0097426 | 36.8108 | 81.095 |  |  |  |  | 1625.0538 | 0.039818
|  |  | 36.8855 | 81.435 |  |  |  |  | 1628.3235 | 0.011761 | 36.8855 | 81.21 |  |  |  |  | 1628.3235 | 0.035452
|  |  | 36.9602 | 81.36 |  |  |  |  | 1631.5932 | 0.011954 | 36.9602 | 81.2 |  |  |  |  | 1631.5932 | 0.02772
|  |  | 37.0348 | 81.27 |  |  |  |  | 1634.863 | 0.010216 | 37.0348 | 81.11 |  |  |  |  | 1634.863 | 0.018613
|  |  | 37.1095 | 81.16 |  |  |  |  | 1638.1327 | 0.0071118 | 37.1095 | 81.38 |  |  |  |  | 1638.1327 | 0.011426
|  |  | 37.1842 | 81.195 |  |  |  |  | 1641.4024 | 0.0039804 | 37.1842 | 81.405 |  |  |  |  | 1641.4024 | 0.0073943
|  |  | 37.2588 | 81.33 |  |  |  |  | 1644.6721 | 0.0041417 | 37.2588 | 81.88 |  |  |  |  | 1644.6721 | 0.011255
|  |  | 37.3335 | 81.315 |  |  |  |  | 1647.9419 | 0.0068671 | 37.3335 | 81.89 |  |  |  |  | 1647.9419 | 0.02288
|  |  | 37.4082 | 81.215 |  |  |  |  | 1651.2116 | 0.0089235 | 37.4082 | 81.685 |  |  |  |  | 1651.2116 | 0.034916
|  |  | 37.4828 | 81.18 |  |  |  |  | 1654.4813 | 0.0092712 | 37.4828 | 81.335 |  |  |  |  | 1654.4813 | 0.041509
|  |  | 37.5575 | 81.265 |  |  |  |  | 1657.751 | 0.0076337 | 37.5575 | 81.965 |  |  |  |  | 1657.751 | 0.038916
|  |  | 37.6322 | 81.35 |  |  |  |  | 1661.0208 | 0.0043633 | 37.6322 | 82.47 |  |  |  |  | 1661.0208 | 0.027353
|  |  | 37.7068 | 81.345 |  |  |  |  | 1664.2905 | 0.0012703 | 37.7068 | 82.575 |  |  |  |  | 1664.2905 | 0.01508
|  |  | 37.7815 | 81.24 |  |  |  |  | 1667.5602 | 0.0043095 | 37.7815 | 81.64 |  |  |  |  | 1667.5602 | 0.022336
|  |  | 37.8562 | 81.195 |  |  |  |  | 1670.8299 | 0.0071303 | 37.8562 | 81.32 |  |  |  |  | 1670.8299 | 0.034318
|  |  | 37.9308 | 81.23 |  |  |  |  | 1674.0997 | 0.0082089 | 37.9308 | 81.495 |  |  |  |  | 1674.0997 | 0.038296
|  |  | 38.0055 | 81.23 |  |  |  |  | 1677.3694 | 0.0073461 | 38.0055 | 81.66 |  |  |  |  | 1677.3694 | 0.035783
|  |  | 38.0802 | 81.215 |  |  |  |  | 1680.6391 | 0.0050253 | 38.0802 | 81.975 |  |  |  |  | 1680.6391 | 0.037992
|  |  | 38.1548 | 81.2 |  |  |  |  | 1683.9088 | 0.0033493 | 38.1548 | 82.09 |  |  |  |  | 1683.9088 | 0.050203
|  |  | 38.2295 | 81.31 |  |  |  |  | 1687.1786 | 0.005313 | 38.2295 | 81.505 |  |  |  |  | 1687.1786 | 0.061158
|  |  | 38.3042 | 81.29 |  |  |  |  | 1690.4483 | 0.0077981 | 38.3042 | 81.34 |  |  |  |  | 1690.4483 | 0.060197
|  |  | 38.3788 | 81.22 |  |  |  |  | 1693.718 | 0.0087214 | 38.3788 | 81.635 |  |  |  |  | 1693.718 | 0.043104
|  |  | 38.4535 | 81.345 |  |  |  |  | 1696.9878 | 0.0076733 | 38.4535 | 82.055 |  |  |  |  | 1696.9878 | 0.012592
|  |  | 38.5282 | 81.33 |  |  |  |  | 1700.2575 | 0.0052053 | 38.5282 | 82.71 |  |  |  |  | 1700.2575 | 0.027579
|  |  | 38.6028 | 81.27 |  |  |  |  | 1703.5272 | 0.0030632 | 38.6028 | 82.705 |  |  |  |  | 1703.5272 | 0.063034
|  |  | 38.6775 | 81.29 |  |  |  |  | 1706.7969 | 0.0043561 | 38.6775 | 82.675 |  |  |  |  | 1706.7969 | 0.086907
|  |  | 38.7522 | 81.28 |  |  |  |  | 1710.0667 | 0.007444 | 38.7522 | 82.23 |  |  |  |  | 1710.0667 | 0.093507
|  |  | 38.8268 | 81.26 |  |  |  |  | 1713.3364 | 0.010737 | 38.8268 | 82.265 |  |  |  |  | 1713.3364 | 0.082827
|  |  | 38.9015 | 81.235 |  |  |  |  | 1716.6061 | 0.013492 | 38.9015 | 81.315 |  |  |  |  | 1716.6061 | 0.0611
|  |  | 38.9762 | 81.24 |  |  |  |  | 1719.8758 | 0.014748 | 38.9762 | 80.535 |  |  |  |  | 1719.8758 | 0.042246
|  |  | 39.0508 | 81.34 |  |  |  |  | 1723.1456 | 0.013895 | 39.0508 | 80.985 |  |  |  |  | 1723.1456 | 0.044098
|  |  | 39.1255 | 81.36 |  |  |  |  | 1726.4153 | 0.01139 | 39.1255 | 81.37 |  |  |  |  | 1726.4153 | 0.057799
|  |  | 39.2002 | 81.2 |  |  |  |  | 1729.685 | 0.0088855 | 39.2002 | 81.47 |  |  |  |  | 1729.685 | 0.067131
|  |  | 39.2748 | 81.175 |  |  |  |  | 1732.9547 | 0.0076621 | 39.2748 | 82.175 |  |  |  |  | 1732.9547 | 0.06777
|  |  | 39.3495 | 81.3 |  |  |  |  | 1736.2245 | 0.0063422 | 39.3495 | 83.25 |  |  |  |  | 1736.2245 | 0.061522
|  |  | 39.4242 | 81.31 |  |  |  |  | 1739.4942 | 0.0031532 | 39.4242 | 82.62 |  |  |  |  | 1739.4942 | 0.056235
|  |  | 39.4988 | 81.26 |  |  |  |  | 1742.7639 | 0.0021482 | 39.4988 | 81.91 |  |  |  |  | 1742.7639 | 0.064332
|  |  | 39.5735 | 81.19 |  |  |  |  | 1746.0336 | 0.0071698 | 39.5735 | 81.365 |  |  |  |  | 1746.0336 | 0.08548
|  |  | 39.6482 | 81.17 |  |  |  |  | 1749.3034 | 0.010261 | 39.6482 | 81.62 |  |  |  |  | 1749.3034 | 0.10629
|  |  | 39.7228 | 81.18 |  |  |  |  | 1752.5731 | 0.0099226 | 39.7228 | 81.645 |  |  |  |  | 1752.5731 | 0.11482
|  |  | 39.7975 | 81.12 |  |  |  |  | 1755.8428 | 0.0064322 | 39.7975 | 81.995 |  |  |  |  | 1755.8428 | 0.1045
|  |  | 39.8722 | 81.115 |  |  |  |  | 1759.1125 | 0.0042728 | 39.8722 | 81.865 |  |  |  |  | 1759.1125 | 0.075285
|  |  | 39.9468 | 81.155 |  |  |  |  | 1762.3823 | 0.0084864 | 39.9468 | 81.3 |  |  |  |  | 1762.3823 | 0.034267
|  |  | 40.0215 | 81.22 |  |  |  |  | 1765.652 | 0.012061 | 40.0215 | 81.135 |  |  |  |  | 1765.652 | 0.018383
|  |  | 40.0962 | 81.22 |  |  |  |  | 1768.9217 | 0.012695 | 40.0962 | 81.81 |  |  |  |  | 1768.9217 | 0.050134
|  |  | 40.1708 | 81.22 |  |  |  |  | 1772.1914 | 0.010746 | 40.1708 | 82.105 |  |  |  |  | 1772.1914 | 0.067602
|  |  | 40.2455 | 81.21 |  |  |  |  | 1775.4612 | 0.0076675 | 40.2455 | 82.86 |  |  |  |  | 1775.4612 | 0.06607
|  |  | 40.3202 | 81.325 |  |  |  |  | 1778.7309 | 0.0046209 | 40.3202 | 82.485 |  |  |  |  | 1778.7309 | 0.050701
|  |  | 40.3948 | 81.325 |  |  |  |  | 1782.0006 | 0.0030258 | 40.3948 | 82.09 |  |  |  |  | 1782.0006 | 0.033312
|  |  | 40.4695 | 81.205 |  |  |  |  | 1785.2704 | 0.006802 | 40.4695 | 82.065 |  |  |  |  | 1785.2704 | 0.026654
|  |  | 40.5442 | 81.055 |  |  |  |  | 1788.5401 | 0.01244 | 40.5442 | 82.185 |  |  |  |  | 1788.5401 | 0.025882
|  |  | 40.6188 | 81.08 |  |  |  |  | 1791.8098 | 0.016983 | 40.6188 | 82.205 |  |  |  |  | 1791.8098 | 0.022426
|  |  | 40.6935 | 81.21 |  |  |  |  | 1795.0795 | 0.018594 | 40.6935 | 82.315 |  |  |  |  | 1795.0795 | 0.02655
|  |  | 40.7682 | 81.305 |  |  |  |  | 1798.3493 | 0.016626 | 40.7682 | 81.99 |  |  |  |  | 1798.3493 | 0.040351
|  |  | 40.8428 | 81.255 |  |  |  |  | 1801.619 | 0.011825 | 40.8428 | 82.055 |  |  |  |  | 1801.619 | 0.049856
|  |  | 40.9175 | 81.16 |  |  |  |  | 1804.8887 | 0.0059068 | 40.9175 | 81.95 |  |  |  |  | 1804.8887 | 0.047135
|  |  | 40.9922 | 81.17 |  |  |  |  | 1808.1584 | 0.0013978 | 40.9922 | 82.18 |  |  |  |  | 1808.1584 | 0.032164
|  |  | 41.0668 | 81.225 |  |  |  |  | 1811.4282 | 0.0039393 | 41.0668 | 82.07 |  |  |  |  | 1811.4282 | 0.015702
|  |  | 41.1415 | 81.275 |  |  |  |  | 1814.6979 | 0.0063204 | 41.1415 | 82.245 |  |  |  |  | 1814.6979 | 0.025385
|  |  | 41.2162 | 81.31 |  |  |  |  | 1817.9676 | 0.0074713 | 41.2162 | 82.44 |  |  |  |  | 1817.9676 | 0.040393
|  |  | 41.2909 | 81.28 |  |  |  |  | 1821.2373 | 0.0073285 | 41.2909 | 82.03 |  |  |  |  | 1821.2373 | 0.047643
|  |  | 41.3655 | 81.2 |  |  |  |  | 1824.5071 | 0.0065691 | 41.3655 | 81.92 |  |  |  |  | 1824.5071 | 0.050235
|  |  | 41.4402 | 81.12 |  |  |  |  | 1827.7768 | 0.0076353 | 41.4402 | 82.085 |  |  |  |  | 1827.7768 | 0.054095
|  |  | 41.5149 | 81.09 |  |  |  |  | 1831.0465 | 0.010808 | 41.5149 | 82.225 |  |  |  |  | 1831.0465 | 0.058598
|  |  | 41.5895 | 81.165 |  |  |  |  | 1834.3162 | 0.013338 | 41.5895 | 82.135 |  |  |  |  | 1834.3162 | 0.058423
|  |  | 41.6642 | 81.275 |  |  |  |  | 1837.586 | 0.01338 | 41.6642 | 82.34 |  |  |  |  | 1837.586 | 0.052896
|  |  | 41.7389 | 81.27 |  |  |  |  | 1840.8557 | 0.01072 | 41.7389 | 81.83 |  |  |  |  | 1840.8557 | 0.050043
|  |  | 41.8135 | 81.24 |  |  |  |  | 1844.1254 | 0.0065868 | 41.8135 | 81.815 |  |  |  |  | 1844.1254 | 0.057507
|  |  | 41.8882 | 81.13 |  |  |  |  | 1847.3951 | 0.0030043 | 41.8882 | 82.44 |  |  |  |  | 1847.3951 | 0.069007
|  |  | 41.9629 | 81.175 |  |  |  |  | 1850.6649 | 0.0020966 | 41.9629 | 82.52 |  |  |  |  | 1850.6649 | 0.075589
|  |  | 42.0375 | 81.23 |  |  |  |  | 1853.9346 | 0.0042247 | 42.0375 | 82.005 |  |  |  |  | 1853.9346 | 0.074721
|  |  | 42.1122 | 81.21 |  |  |  |  | 1857.2043 | 0.0082581 | 42.1122 | 81.85 |  |  |  |  | 1857.2043 | 0.068068
|  |  | 42.1869 | 81.21 |  |  |  |  | 1860.474 | 0.012486 | 42.1869 | 81.81 |  |  |  |  | 1860.474 | 0.0567
|  |  | 42.2615 | 81.235 |  |  |  |  | 1863.7438 | 0.014842 | 42.2615 | 81.395 |  |  |  |  | 1863.7438 | 0.039189
|  |  | 42.3362 | 81.325 |  |  |  |  | 1867.0135 | 0.014117 | 42.3362 | 81.255 |  |  |  |  | 1867.0135 | 0.016704
|  |  | 42.4109 | 81.41 |  |  |  |  | 1870.2832 | 0.01082 | 42.4109 | 81.355 |  |  |  |  | 1870.2832 | 0.023276
|  |  | 42.4855 | 81.39 |  |  |  |  | 1873.553 | 0.0078571 | 42.4855 | 81.485 |  |  |  |  | 1873.553 | 0.049362
|  |  | 42.5602 | 81.325 |  |  |  |  | 1876.8227 | 0.0088713 | 42.5602 | 82 |  |  |  |  | 1876.8227 | 0.065056
|  |  | 42.6349 | 81.29 |  |  |  |  | 1880.0924 | 0.011467 | 42.6349 | 82.195 |  |  |  |  | 1880.0924 | 0.062411
|  |  | 42.7095 | 81.265 |  |  |  |  | 1883.3621 | 0.013027 | 42.7095 | 81.75 |  |  |  |  | 1883.3621 | 0.040364
|  |  | 42.7842 | 81.21 |  |  |  |  | 1886.6319 | 0.013573 | 42.7842 | 81.49 |  |  |  |  | 1886.6319 | 0.0081336
|  |  | 42.8589 | 81.175 |  |  |  |  | 1889.9016 | 0.013356 | 42.8589 | 81.045 |  |  |  |  | 1889.9016 | 0.037432
|  |  | 42.9335 | 81.26 |  |  |  |  | 1893.1713 | 0.011719 | 42.9335 | 81.3 |  |  |  |  | 1893.1713 | 0.06928
|  |  | 43.0082 | 81.415 |  |  |  |  | 1896.441 | 0.0081224 | 43.0082 | 82 |  |  |  |  | 1896.441 | 0.085206
|  |  | 43.0829 | 81.4 |  |  |  |  | 1899.7108 | 0.0047195 | 43.0829 | 82.475 |  |  |  |  | 1899.7108 | 0.082655
|  |  | 43.1575 | 81.31 |  |  |  |  | 1902.9805 | 0.00779 | 43.1575 | 82.09 |  |  |  |  | 1902.9805 | 0.065765
|  |  | 43.2322 | 81.28 |  |  |  |  | 1906.2502 | 0.012359 | 43.2322 | 81.94 |  |  |  |  | 1906.2502 | 0.043307
|  |  | 43.3069 | 81.33 |  |  |  |  | 1909.5199 | 0.01447 | 43.3069 | 81.78 |  |  |  |  | 1909.5199 | 0.025092
|  |  | 43.3815 | 81.26 |  |  |  |  | 1912.7897 | 0.013546 | 43.3815 | 82.25 |  |  |  |  | 1912.7897 | 0.016578
|  |  | 43.4562 | 81.155 |  |  |  |  | 1916.0594 | 0.010623 | 43.4562 | 82.37 |  |  |  |  | 1916.0594 | 0.015296
|  |  | 43.5309 | 81.11 |  |  |  |  | 1919.3291 | 0.0074845 | 43.5309 | 82.49 |  |  |  |  | 1919.3291 | 0.02122
|  |  | 43.6055 | 81.105 |  |  |  |  | 1922.5988 | 0.0051912 | 43.6055 | 81.695 |  |  |  |  | 1922.5988 | 0.031387
|  |  | 43.6802 | 81.19 |  |  |  |  | 1925.8686 | 0.0032343 | 43.6802 | 81.36 |  |  |  |  | 1925.8686 | 0.039506
|  |  | 43.7549 | 81.205 |  |  |  |  | 1929.1383 | 0.0029275 | 43.7549 | 81.725 |  |  |  |  | 1929.1383 | 0.043471
|  |  | 43.8295 | 81.155 |  |  |  |  | 1932.408 | 0.00619 | 43.8295 | 81.42 |  |  |  |  | 1932.408 | 0.045848
|  |  | 43.9042 | 81.21 |  |  |  |  | 1935.6777 | 0.0098751 | 43.9042 | 81.53 |  |  |  |  | 1935.6777 | 0.049382
|  |  | 43.9789 | 81.41 |  |  |  |  | 1938.9475 | 0.012637 | 43.9789 | 81.34 |  |  |  |  | 1938.9475 | 0.051905
|  |  | 44.0535 | 81.42 |  |  |  |  | 1942.2172 | 0.014088 | 44.0535 | 82.035 |  |  |  |  | 1942.2172 | 0.048789
|  |  | 44.1282 | 81.385 |  |  |  |  | 1945.4869 | 0.01414 | 44.1282 | 82.24 |  |  |  |  | 1945.4869 | 0.038033
|  |  | 44.2029 | 81.27 |  |  |  |  | 1948.7566 | 0.012696 | 44.2029 | 82.2 |  |  |  |  | 1948.7566 | 0.021955
|  |  | 44.2775 | 81.19 |  |  |  |  | 1952.0264 | 0.0099299 | 44.2775 | 81.36 |  |  |  |  | 1952.0264 | 0.0076748
|  |  | 44.3522 | 81.205 |  |  |  |  | 1955.2961 | 0.00715 | 44.3522 | 81.535 |  |  |  |  | 1955.2961 | 0.010857
|  |  | 44.4269 | 81.225 |  |  |  |  | 1958.5658 | 0.0069822 | 44.4269 | 81.585 |  |  |  |  | 1958.5658 | 0.014066
|  |  | 44.5015 | 81.25 |  |  |  |  | 1961.8355 | 0.0086763 | 44.5015 | 82.04 |  |  |  |  | 1961.8355 | 0.010265
|  |  | 44.5762 | 81.26 |  |  |  |  | 1965.1053 | 0.009191 | 44.5762 | 82.325 |  |  |  |  | 1965.1053 | 0.010097
|  |  | 44.6509 | 81.255 |  |  |  |  | 1968.375 | 0.0076371 | 44.6509 | 82.355 |  |  |  |  | 1968.375 | 0.02153
|  |  | 44.7255 | 81.32 |  |  |  |  | 1971.6447 | 0.0052354 | 44.7255 | 82.035 |  |  |  |  | 1971.6447 | 0.032985
|  |  | 44.8002 | 81.285 |  |  |  |  | 1974.9145 | 0.0050779 | 44.8002 | 81.365 |  |  |  |  | 1974.9145 | 0.040253
|  |  | 44.8749 | 81.19 |  |  |  |  | 1978.1842 | 0.0070681 | 44.8749 | 81.735 |  |  |  |  | 1978.1842 | 0.0428
|  |  | 44.9495 | 81.13 |  |  |  |  | 1981.4539 | 0.0088983 | 44.9495 | 82.15 |  |  |  |  | 1981.4539 | 0.042216
|  |  | 45.0242 | 81.17 |  |  |  |  | 1984.7236 | 0.010664 | 45.0242 | 83.05 |  |  |  |  | 1984.7236 | 0.041113
|  |  | 45.0989 | 81.205 |  |  |  |  | 1987.9934 | 0.01278 | 45.0989 | 82.65 |  |  |  |  | 1987.9934 | 0.0421
|  |  | 45.1735 | 81.2 |  |  |  |  | 1991.2631 | 0.014395 | 45.1735 | 81.975 |  |  |  |  | 1991.2631 | 0.047036
|  |  | 45.2482 | 81.175 |  |  |  |  | 1994.5328 | 0.014169 | 45.2482 | 81.39 |  |  |  |  | 1994.5328 | 0.05636
|  |  | 45.3229 | 81.23 |  |  |  |  | 1997.8025 | 0.011484 | 45.3229 | 81.275 |  |  |  |  | 1997.8025 | 0.068568
|  |  | 45.3975 | 81.31 |  |  |  |  | 2001.0723 | 0.0068668 | 45.3975 | 81.15 |  |  |  |  | 2001.0723 | 0.080275
|  |  | 45.4722 | 81.24 |  |  |  |  | 2004.342 | 0.0022607 | 45.4722 | 81.575 |  |  |  |  | 2004.342 | 0.087025
|  |  | 45.5469 | 81.19 |  |  |  |  | 2007.6117 | 0.0039793 | 45.5469 | 82.28 |  |  |  |  | 2007.6117 | 0.084767
|  |  | 45.6215 | 81.14 |  |  |  |  | 2010.8814 | 0.0064224 | 45.6215 | 82.345 |  |  |  |  | 2010.8814 | 0.071622
|  |  | 45.6962 | 81.09 |  |  |  |  | 2014.1512 | 0.006771 | 45.6962 | 82.22 |  |  |  |  | 2014.1512 | 0.049289
|  |  | 45.7709 | 81.185 |  |  |  |  | 2017.4209 | 0.00544 | 45.7709 | 82.54 |  |  |  |  | 2017.4209 | 0.023821
|  |  | 45.8455 | 81.47 |  |  |  |  | 2020.6906 | 0.0036577 | 45.8455 | 81.65 |  |  |  |  | 2020.6906 | 0.012209
|  |  | 45.9202 | 81.53 |  |  |  |  | 2023.9603 | 0.003265 | 45.9202 | 81.94 |  |  |  |  | 2023.9603 | 0.020884
|  |  | 45.9949 | 81.365 |  |  |  |  | 2027.2301 | 0.0046024 | 45.9949 | 82.375 |  |  |  |  | 2027.2301 | 0.020163
|  |  | 46.0695 | 80.985 |  |  |  |  | 2030.4998 | 0.0064498 | 46.0695 | 82.43 |  |  |  |  | 2030.4998 | 0.0072165
|  |  | 46.1442 | 80.985 |  |  |  |  | 2033.7695 | 0.0083758 | 46.1442 | 81.71 |  |  |  |  | 2033.7695 | 0.01448
|  |  | 46.2189 | 81.075 |  |  |  |  | 2037.0392 | 0.010067 | 46.2189 | 82.115 |  |  |  |  | 2037.0392 | 0.034179
|  |  | 46.2935 | 81.135 |  |  |  |  | 2040.309 | 0.011109 | 46.2935 | 82.46 |  |  |  |  | 2040.309 | 0.045298
|  |  | 46.3682 | 81.18 |  |  |  |  | 2043.5787 | 0.011191 | 46.3682 | 81.605 |  |  |  |  | 2043.5787 | 0.044101
|  |  | 46.4429 | 81.35 |  |  |  |  | 2046.8484 | 0.010252 | 46.4429 | 81.985 |  |  |  |  | 2046.8484 | 0.034503
|  |  | 46.5175 | 81.425 |  |  |  |  | 2050.1181 | 0.0084798 | 46.5175 | 81.665 |  |  |  |  | 2050.1181 | 0.030572
|  |  | 46.5922 | 81.315 |  |  |  |  | 2053.3879 | 0.0061952 | 46.5922 | 81.565 |  |  |  |  | 2053.3879 | 0.038487
|  |  | 46.6669 | 81.285 |  |  |  |  | 2056.6576 | 0.0037174 | 46.6669 | 81.955 |  |  |  |  | 2056.6576 | 0.042801
|  |  | 46.7415 | 81.345 |  |  |  |  | 2059.9273 | 0.0016858 | 46.7415 | 82.995 |  |  |  |  | 2059.9273 | 0.034095
|  |  | 46.8162 | 81.365 |  |  |  |  | 2063.1971 | 0.002502 | 46.8162 | 82.475 |  |  |  |  | 2063.1971 | 0.012765
|  |  | 46.8909 | 81.45 |  |  |  |  | 2066.4668 | 0.0042058 | 46.8909 | 80.985 |  |  |  |  | 2066.4668 | 0.016281
|  |  | 46.9655 | 81.47 |  |  |  |  | 2069.7365 | 0.0047112 | 46.9655 | 80.09 |  |  |  |  | 2069.7365 | 0.041646
|  |  | 47.0402 | 81.38 |  |  |  |  | 2073.0062 | 0.0032324 | 47.0402 | 80.24 |  |  |  |  | 2073.0062 | 0.056729
|  |  | 47.1149 | 81.26 |  |  |  |  | 2076.276 | 0.00044599 | 47.1149 | 80.985 |  |  |  |  | 2076.276 | 0.057731
|  |  | 47.1895 | 81.16 |  |  |  |  | 2079.5457 | 0.0055644 | 47.1895 | 81.605 |  |  |  |  | 2079.5457 | 0.046485
|  |  | 47.2642 | 81.155 |  |  |  |  | 2082.8154 | 0.010925 | 47.2642 | 82.1 |  |  |  |  | 2082.8154 | 0.029224
|  |  | 47.3389 | 81.14 |  |  |  |  | 2086.0851 | 0.014886 | 47.3389 | 82.745 |  |  |  |  | 2086.0851 | 0.013799
|  |  | 47.4135 | 81.19 |  |  |  |  | 2089.3549 | 0.01617 | 47.4135 | 82.815 |  |  |  |  | 2089.3549 | 0.0062756
|  |  | 47.4882 | 81.31 |  |  |  |  | 2092.6246 | 0.014376 | 47.4882 | 82.7 |  |  |  |  | 2092.6246 | 0.0068949
|  |  | 47.5629 | 81.3 |  |  |  |  | 2095.8943 | 0.010249 | 47.5629 | 82.34 |  |  |  |  | 2095.8943 | 0.014148
|  |  | 47.6375 | 81.24 |  |  |  |  | 2099.164 | 0.0057315 | 47.6375 | 82.485 |  |  |  |  | 2099.164 | 0.02303
|  |  | 47.7122 | 81.21 |  |  |  |  | 2102.4338 | 0.0041621 | 47.7122 | 82.475 |  |  |  |  | 2102.4338 | 0.027531
|  |  | 47.7869 | 81.23 |  |  |  |  | 2105.7035 | 0.0050898 | 47.7869 | 82.005 |  |  |  |  | 2105.7035 | 0.024603
|  |  | 47.8615 | 81.305 |  |  |  |  | 2108.9732 | 0.0047439 | 47.8615 | 81.94 |  |  |  |  | 2108.9732 | 0.016487
|  |  | 47.9362 | 81.305 |  |  |  |  | 2112.2429 | 0.0028072 | 47.9362 | 81.8 |  |  |  |  | 2112.2429 | 0.015375
|  |  | 48.0109 | 81.28 |  |  |  |  | 2115.5127 | 0.00047491 | 48.0109 | 81.325 |  |  |  |  | 2115.5127 | 0.025205
|  |  | 48.0855 | 81.28 |  |  |  |  | 2118.7824 | 0.0012531 | 48.0855 | 81.05 |  |  |  |  | 2118.7824 | 0.033386
|  |  | 48.1602 | 81.31 |  |  |  |  | 2122.0521 | 0.0012459 | 48.1602 | 81.33 |  |  |  |  | 2122.0521 | 0.036182
|  |  | 48.2349 | 81.28 |  |  |  |  | 2125.3218 | 0.0003064 | 48.2349 | 81.955 |  |  |  |  | 2125.3218 | 0.035133
|  |  | 48.3095 | 81.27 |  |  |  |  | 2128.5916 | 0.0023572 | 48.3095 | 82.2 |  |  |  |  | 2128.5916 | 0.034217
|  |  | 48.3842 | 81.34 |  |  |  |  | 2131.8613 | 0.0040319 | 48.3842 | 81.745 |  |  |  |  | 2131.8613 | 0.037215
|  |  | 48.4589 | 81.37 |  |  |  |  | 2135.131 | 0.0048653 | 48.4589 | 81.195 |  |  |  |  | 2135.131 | 0.04505
|  |  | 48.5336 | 81.315 |  |  |  |  | 2138.4007 | 0.005809 | 48.5336 | 81.31 |  |  |  |  | 2138.4007 | 0.055477
|  |  | 48.6082 | 81.305 |  |  |  |  | 2141.6705 | 0.0081373 | 48.6082 | 81.58 |  |  |  |  | 2141.6705 | 0.06472
|  |  | 48.6829 | 81.3 |  |  |  |  | 2144.9402 | 0.011242 | 48.6829 | 81.965 |  |  |  |  | 2144.9402 | 0.069354
|  |  | 48.7575 | 81.315 |  |  |  |  | 2148.2099 | 0.013766 | 48.7575 | 81.485 |  |  |  |  | 2148.2099 | 0.068037
|  |  | 48.8322 | 81.25 |  |  |  |  | 2151.4797 | 0.014796 | 48.8322 | 81.5 |  |  |  |  | 2151.4797 | 0.062113
|  |  | 48.9069 | 81.295 |  |  |  |  | 2154.7494 | 0.013896 | 48.9069 | 81.485 |  |  |  |  | 2154.7494 | 0.054117
|  |  | 48.9816 | 81.455 |  |  |  |  | 2158.0191 | 0.010903 | 48.9816 | 81.435 |  |  |  |  | 2158.0191 | 0.045079
|  |  | 49.0562 | 81.44 |  |  |  |  | 2161.2888 | 0.0059326 | 49.0562 | 82.02 |  |  |  |  | 2161.2888 | 0.034005
|  |  | 49.1309 | 81.395 |  |  |  |  | 2164.5586 | 0.0016848 | 49.1309 | 82.065 |  |  |  |  | 2164.5586 | 0.020786
|  |  | 49.2056 | 81.38 |  |  |  |  | 2167.8283 | 0.0077355 | 49.2056 | 82.34 |  |  |  |  | 2167.8283 | 0.009705
|  |  | 49.2802 | 81.385 |  |  |  |  | 2171.098 | 0.013059 | 49.2802 | 81.84 |  |  |  |  | 2171.098 | 0.01101
|  |  | 49.3549 | 81.33 |  |  |  |  | 2174.3677 | 0.015646 | 49.3549 | 81.31 |  |  |  |  | 2174.3677 | 0.017537
|  |  | 49.4296 | 81.26 |  |  |  |  | 2177.6375 | 0.01538 | 49.4296 | 81.16 |  |  |  |  | 2177.6375 | 0.027861
|  |  | 49.5042 | 81.32 |  |  |  |  | 2180.9072 | 0.014289 | 49.5042 | 82.05 |  |  |  |  | 2180.9072 | 0.043827
|  |  | 49.5789 | 81.42 |  |  |  |  | 2184.1769 | 0.015291 | 49.5789 | 81.97 |  |  |  |  | 2184.1769 | 0.059144
|  |  | 49.6536 | 81.425 |  |  |  |  | 2187.4466 | 0.017666 | 49.6536 | 81.86 |  |  |  |  | 2187.4466 | 0.065273
|  |  | 49.7282 | 81.315 |  |  |  |  | 2190.7164 | 0.018128 | 49.7282 | 81.36 |  |  |  |  | 2190.7164 | 0.056799
|  |  | 49.8029 | 81.255 |  |  |  |  | 2193.9861 | 0.01485 | 49.8029 | 81.37 |  |  |  |  | 2193.9861 | 0.035064
|  |  | 49.8776 | 81.275 |  |  |  |  | 2197.2558 | 0.0081635 | 49.8776 | 81.345 |  |  |  |  | 2197.2558 | 0.019878
|  |  | 49.9522 | 81.295 |  |  |  |  | 2200.5255 | 0.0003969 | 49.9522 | 80.67 |  |  |  |  | 2200.5255 | 0.041353
|  |  | 50.0269 | 81.21 |  |  |  |  | 2203.7953 | 0.0076705 | 50.0269 | 81.195 |  |  |  |  | 2203.7953 | 0.061003
|  |  | 50.1016 | 81.145 |  |  |  |  | 2207.065 | 0.012852 | 50.1016 | 82.245 |  |  |  |  | 2207.065 | 0.064932
|  |  | 50.1762 | 81.235 |  |  |  |  | 2210.3347 | 0.015255 | 50.1762 | 82.085 |  |  |  |  | 2210.3347 | 0.052908
|  |  | 50.2509 | 81.31 |  |  |  |  | 2213.6044 | 0.015409 | 50.2509 | 81.94 |  |  |  |  | 2213.6044 | 0.032773
|  |  | 50.3256 | 81.31 |  |  |  |  | 2216.8742 | 0.01383 | 50.3256 | 81.59 |  |  |  |  | 2216.8742 | 0.02116
|  |  | 50.4002 | 81.345 |  |  |  |  | 2220.1439 | 0.010585 | 50.4002 | 81.24 |  |  |  |  | 2220.1439 | 0.026167
|  |  | 50.4749 | 81.41 |  |  |  |  | 2223.4136 | 0.0065765 | 50.4749 | 81.215 |  |  |  |  | 2223.4136 | 0.030017
|  |  | 50.5496 | 81.335 |  |  |  |  | 2226.6833 | 0.0068489 | 50.5496 | 81.535 |  |  |  |  | 2226.6833 | 0.032986
|  |  | 50.6242 | 81.32 |  |  |  |  | 2229.9531 | 0.011364 | 50.6242 | 81.24 |  |  |  |  | 2229.9531 | 0.042822
|  |  | 50.6989 | 81.27 |  |  |  |  | 2233.2228 | 0.014311 | 50.6989 | 81.665 |  |  |  |  | 2233.2228 | 0.054338
|  |  | 50.7736 | 81.165 |  |  |  |  | 2236.4925 | 0.013667 | 50.7736 | 81.485 |  |  |  |  | 2236.4925 | 0.058122
|  |  | 50.8482 | 81.16 |  |  |  |  | 2239.7623 | 0.0097105 | 50.8482 | 80.97 |  |  |  |  | 2239.7623 | 0.05045
|  |  | 50.9229 | 81.235 |  |  |  |  | 2243.032 | 0.0048617 | 50.9229 | 81.21 |  |  |  |  | 2243.032 | 0.033987
|  |  | 50.9976 | 81.33 |  |  |  |  | 2246.3017 | 0.0045567 | 50.9976 | 81.37 |  |  |  |  | 2246.3017 | 0.015989
|  |  | 51.0722 | 81.295 |  |  |  |  | 2249.5714 | 0.0063283 | 51.0722 | 81.76 |  |  |  |  | 2249.5714 | 0.010808
|  |  | 51.1469 | 81.275 |  |  |  |  | 2252.8412 | 0.0058753 | 51.1469 | 81.735 |  |  |  |  | 2252.8412 | 0.019934
|  |  | 51.2216 | 81.255 |  |  |  |  | 2256.1109 | 0.0048296 | 51.2216 | 81.405 |  |  |  |  | 2256.1109 | 0.029809
|  |  | 51.2962 | 81.2 |  |  |  |  | 2259.3806 | 0.0061225 | 51.2962 | 82.065 |  |  |  |  | 2259.3806 | 0.038232
|  |  | 51.3709 | 81.24 |  |  |  |  | 2262.6503 | 0.0072651 | 51.3709 | 82.07 |  |  |  |  | 2262.6503 | 0.040298
|  |  | 51.4456 | 81.305 |  |  |  |  | 2265.9201 | 0.0057133 | 51.4456 | 81.55 |  |  |  |  | 2265.9201 | 0.031208
|  |  | 51.5202 | 81.13 |  |  |  |  | 2269.1898 | 0.0027708 | 51.5202 | 81.135 |  |  |  |  | 2269.1898 | 0.011802
|  |  | 51.5949 | 81.095 |  |  |  |  | 2272.4595 | 0.0067806 | 51.5949 | 81.87 |  |  |  |  | 2272.4595 | 0.020964
|  |  | 51.6696 | 81.075 |  |  |  |  | 2275.7292 | 0.012063 | 51.6696 | 82.31 |  |  |  |  | 2275.7292 | 0.047222
|  |  | 51.7442 | 81.035 |  |  |  |  | 2278.999 | 0.01475 | 51.7442 | 82.65 |  |  |  |  | 2278.999 | 0.064724
|  |  | 51.8189 | 81.12 |  |  |  |  | 2282.2687 | 0.014075 | 51.8189 | 82.165 |  |  |  |  | 2282.2687 | 0.069243
|  |  | 51.8936 | 81.245 |  |  |  |  | 2285.5384 | 0.01136 | 51.8936 | 81.99 |  |  |  |  | 2285.5384 | 0.063696
|  |  | 51.9682 | 81.31 |  |  |  |  | 2288.8081 | 0.0097182 | 51.9682 | 81.97 |  |  |  |  | 2288.8081 | 0.055912
|  |  | 52.0429 | 81.16 |  |  |  |  | 2292.0779 | 0.010464 | 52.0429 | 81.91 |  |  |  |  | 2292.0779 | 0.05021
|  |  | 52.1176 | 81.11 |  |  |  |  | 2295.3476 | 0.010802 | 52.1176 | 82.22 |  |  |  |  | 2295.3476 | 0.041679
|  |  | 52.1922 | 81.12 |  |  |  |  | 2298.6173 | 0.0089323 | 52.1922 | 82.535 |  |  |  |  | 2298.6173 | 0.026197
|  |  | 52.2669 | 81.225 |  |  |  |  | 2301.887 | 0.0060946 | 52.2669 | 81.86 |  |  |  |  | 2301.887 | 0.021243
|  |  | 52.3416 | 81.235 |  |  |  |  | 2305.1568 | 0.0071233 | 52.3416 | 81.645 |  |  |  |  | 2305.1568 | 0.0434
|  |  | 52.4162 | 81.17 |  |  |  |  | 2308.4265 | 0.011309 | 52.4162 | 81.76 |  |  |  |  | 2308.4265 | 0.063923
|  |  | 52.4909 | 81.255 |  |  |  |  | 2311.6962 | 0.014494 | 52.4909 | 82.155 |  |  |  |  | 2311.6962 | 0.072366
|  |  | 52.5656 | 81.435 |  |  |  |  | 2314.9659 | 0.015217 | 52.5656 | 81.68 |  |  |  |  | 2314.9659 | 0.068802
|  |  | 52.6402 | 81.495 |  |  |  |  | 2318.2357 | 0.013497 | 52.6402 | 81.13 |  |  |  |  | 2318.2357 | 0.060165
|  |  | 52.7149 | 81.365 |  |  |  |  | 2321.5054 | 0.010709 | 52.7149 | 81.895 |  |  |  |  | 2321.5054 | 0.053232
|  |  | 52.7896 | 81.285 |  |  |  |  | 2324.7751 | 0.0094383 | 52.7896 | 81.91 |  |  |  |  | 2324.7751 | 0.045373
|  |  | 52.8642 | 81.21 |  |  |  |  | 2328.0449 | 0.01032 | 52.8642 | 82.11 |  |  |  |  | 2328.0449 | 0.029089
|  |  | 52.9389 | 81.105 |  |  |  |  | 2331.3146 | 0.010511 | 52.9389 | 82.225 |  |  |  |  | 2331.3146 | 0.0046454
|  |  | 53.0136 | 81.09 |  |  |  |  | 2334.5843 | 0.0080646 | 53.0136 | 82.005 |  |  |  |  | 2334.5843 | 0.029986
|  |  | 53.0882 | 81.135 |  |  |  |  | 2337.854 | 0.0036974 | 53.0882 | 81.915 |  |  |  |  | 2337.854 | 0.057201
|  |  | 53.1629 | 81.26 |  |  |  |  | 2341.1238 | 0.0054165 | 53.1629 | 81.52 |  |  |  |  | 2341.1238 | 0.07214
|  |  | 53.2376 | 81.355 |  |  |  |  | 2344.3935 | 0.010495 | 53.2376 | 81.855 |  |  |  |  | 2344.3935 | 0.071661
|  |  | 53.3122 | 81.365 |  |  |  |  | 2347.6632 | 0.012924 | 53.3122 | 81.745 |  |  |  |  | 2347.6632 | 0.059216
|  |  | 53.3869 | 81.31 |  |  |  |  | 2350.9329 | 0.011548 | 53.3869 | 81.91 |  |  |  |  | 2350.9329 | 0.043596
|  |  | 53.4616 | 81.305 |  |  |  |  | 2354.2027 | 0.0071796 | 53.4616 | 81.555 |  |  |  |  | 2354.2027 | 0.034102
|  |  | 53.5362 | 81.25 |  |  |  |  | 2357.4724 | 0.003981 | 53.5362 | 81.26 |  |  |  |  | 2357.4724 | 0.031478
|  |  | 53.6109 | 81.08 |  |  |  |  | 2360.7421 | 0.0071581 | 53.6109 | 81.9 |  |  |  |  | 2360.7421 | 0.030998
|  |  | 53.6856 | 81.05 |  |  |  |  | 2364.0118 | 0.0097093 | 53.6856 | 82.435 |  |  |  |  | 2364.0118 | 0.032762
|  |  | 53.7602 | 81.115 |  |  |  |  | 2367.2816 | 0.0096027 | 53.7602 | 82.67 |  |  |  |  | 2367.2816 | 0.035204
|  |  | 53.8349 | 81.2 |  |  |  |  | 2370.5513 | 0.0086102 | 53.8349 | 81.99 |  |  |  |  | 2370.5513 | 0.032531
|  |  | 53.9096 | 81.255 |  |  |  |  | 2373.821 | 0.010042 | 53.9096 | 81.975 |  |  |  |  | 2373.821 | 0.023116
|  |  | 53.9842 | 81.26 |  |  |  |  | 2377.0907 | 0.013188 | 53.9842 | 81.67 |  |  |  |  | 2377.0907 | 0.021428
|  |  | 54.0589 | 81.255 |  |  |  |  | 2380.3605 | 0.015105 | 54.0589 | 81.615 |  |  |  |  | 2380.3605 | 0.038246
|  |  | 54.1336 | 81.235 |  |  |  |  | 2383.6302 | 0.014475 | 54.1336 | 81.84 |  |  |  |  | 2383.6302 | 0.054071
|  |  | 54.2082 | 81.24 |  |  |  |  | 2386.8999 | 0.011703 | 54.2082 | 81.81 |  |  |  |  | 2386.8999 | 0.058902
|  |  | 54.2829 | 81.25 |  |  |  |  | 2390.1696 | 0.0090225 | 54.2829 | 81.665 |  |  |  |  | 2390.1696 | 0.051178
|  |  | 54.3576 | 81.255 |  |  |  |  | 2393.4394 | 0.0095809 | 54.3576 | 81.86 |  |  |  |  | 2393.4394 | 0.037645
|  |  | 54.4322 | 81.25 |  |  |  |  | 2396.7091 | 0.012045 | 54.4322 | 81.73 |  |  |  |  | 2396.7091 | 0.034955
|  |  | 54.5069 | 81.26 |  |  |  |  | 2399.9788 | 0.013023 | 54.5069 | 81.57 |  |  |  |  | 2399.9788 | 0.047286
|  |  | 54.5816 | 81.375 |  |  |  |  | 2403.2485 | 0.011071 | 54.5816 | 81.21 |  |  |  |  | 2403.2485 | 0.059086
|  |  | 54.6562 | 81.505 |  |  |  |  | 2406.5183 | 0.0064031 | 54.6562 | 81.79 |  |  |  |  | 2406.5183 | 0.063167
|  |  | 54.7309 | 81.575 |  |  |  |  | 2409.788 | 0.00058704 | 54.7309 | 81.955 |  |  |  |  | 2409.788 | 0.059387
|  |  | 54.8056 | 81.48 |  |  |  |  | 2413.0577 | 0.0045737 | 54.8056 | 81.955 |  |  |  |  | 2413.0577 | 0.051016
|  |  | 54.8802 | 81.305 |  |  |  |  | 2416.3275 | 0.0066683 | 54.8802 | 82.04 |  |  |  |  | 2416.3275 | 0.044573
|  |  | 54.9549 | 81.23 |  |  |  |  | 2419.5972 | 0.0049198 | 54.9549 | 82.09 |  |  |  |  | 2419.5972 | 0.048064
|  |  | 55.0296 | 81.23 |  |  |  |  | 2422.8669 | 0.0030658 | 55.0296 | 81.875 |  |  |  |  | 2422.8669 | 0.059841
|  |  | 55.1042 | 81.255 |  |  |  |  | 2426.1366 | 0.00961 | 55.1042 | 81.955 |  |  |  |  | 2426.1366 | 0.069858
|  |  | 55.1789 | 81.32 |  |  |  |  | 2429.4064 | 0.016506 | 55.1789 | 82.34 |  |  |  |  | 2429.4064 | 0.070137
|  |  | 55.2536 | 81.215 |  |  |  |  | 2432.6761 | 0.020599 | 55.2536 | 82.105 |  |  |  |  | 2432.6761 | 0.05749
|  |  | 55.3282 | 81.18 |  |  |  |  | 2435.9458 | 0.020699 | 55.3282 | 81.81 |  |  |  |  | 2435.9458 | 0.033948
|  |  | 55.4029 | 81.145 |  |  |  |  | 2439.2155 | 0.017343 | 55.4029 | 81.895 |  |  |  |  | 2439.2155 | 0.010138
|  |  | 55.4776 | 81.11 |  |  |  |  | 2442.4853 | 0.012588 | 55.4776 | 81.79 |  |  |  |  | 2442.4853 | 0.02653
|  |  | 55.5522 | 81.11 |  |  |  |  | 2445.755 | 0.0090283 | 55.5522 | 82.24 |  |  |  |  | 2445.755 | 0.044521
|  |  | 55.6269 | 81.165 |  |  |  |  | 2449.0247 | 0.0074011 | 55.6269 | 81.97 |  |  |  |  | 2449.0247 | 0.051095
|  |  | 55.7016 | 81.22 |  |  |  |  | 2452.2944 | 0.0061893 | 55.7016 | 81.545 |  |  |  |  | 2452.2944 | 0.046933
|  |  | 55.7762 | 81.155 |  |  |  |  | 2455.5642 | 0.0060419 | 55.7762 | 81.16 |  |  |  |  | 2455.5642 | 0.036634
|  |  | 55.8509 | 81.135 |  |  |  |  | 2458.8339 | 0.00833 | 55.8509 | 81.455 |  |  |  |  | 2458.8339 | 0.025856
|  |  | 55.9256 | 81.285 |  |  |  |  | 2462.1036 | 0.010684 | 55.9256 | 81.505 |  |  |  |  | 2462.1036 | 0.017399
|  |  | 56.0002 | 81.29 |  |  |  |  | 2465.3733 | 0.011037 | 56.0002 | 81.635 |  |  |  |  | 2465.3733 | 0.010194
|  |  | 56.0749 | 81.24 |  |  |  |  | 2468.6431 | 0.0092239 | 56.0749 | 81.63 |  |  |  |  | 2468.6431 | 0.0099494
|  |  | 56.1496 | 81.095 |  |  |  |  | 2471.9128 | 0.0071783 | 56.1496 | 81.815 |  |  |  |  | 2471.9128 | 0.02094
|  |  | 56.2243 | 81.075 |  |  |  |  | 2475.1825 | 0.00776 | 56.2243 | 81.795 |  |  |  |  | 2475.1825 | 0.032703
|  |  | 56.2989 | 81.09 |  |  |  |  | 2478.4522 | 0.0095873 | 56.2989 | 81.905 |  |  |  |  | 2478.4522 | 0.040836
|  |  | 56.3736 | 81.205 |  |  |  |  | 2481.722 | 0.0099264 | 56.3736 | 81.78 |  |  |  |  | 2481.722 | 0.043712
|  |  | 56.4482 | 81.28 |  |  |  |  | 2484.9917 | 0.0080627 | 56.4482 | 82.44 |  |  |  |  | 2484.9917 | 0.041595
|  |  | 56.5229 | 81.285 |  |  |  |  | 2488.2614 | 0.004644 | 56.5229 | 82.765 |  |  |  |  | 2488.2614 | 0.035887
|  |  | 56.5976 | 81.24 |  |  |  |  | 2491.5311 | 0.00094554 | 56.5976 | 82.665 |  |  |  |  | 2491.5311 | 0.028214
|  |  | 56.6723 | 81.3 |  |  |  |  | 2494.8009 | 0.0018938 | 56.6723 | 82.26 |  |  |  |  | 2494.8009 | 0.019793
|  |  | 56.7469 | 81.32 |  |  |  |  | 2498.0706 | 0.0031056 | 56.7469 | 81.475 |  |  |  |  | 2498.0706 | 0.011437
|  |  | 56.8216 | 81.295 |  |  |  |  | 2501.3403 | 0.0028266 | 56.8216 | 81.26 |  |  |  |  | 2501.3403 | 0.0063052
|  |  | 56.8963 | 81.265 |  |  |  |  | 2504.6101 | 0.0017853 | 56.8963 | 81.23 |  |  |  |  | 2504.6101 | 0.012803
|  |  | 56.9709 | 81.23 |  |  |  |  | 2507.8798 | 0.00099511 | 56.9709 | 81.37 |  |  |  |  | 2507.8798 | 0.023796
|  |  | 57.0456 | 81.245 |  |  |  |  | 2511.1495 | 0.0013043 | 57.0456 | 82.1 |  |  |  |  | 2511.1495 | 0.035391
|  |  | 57.1203 | 81.225 |  |  |  |  | 2514.4192 | 0.0030431 | 57.1203 | 82.62 |  |  |  |  | 2514.4192 | 0.045494
|  |  | 57.1949 | 81.22 |  |  |  |  | 2517.689 | 0.005916 | 57.1949 | 82.33 |  |  |  |  | 2517.689 | 0.051959
|  |  | 57.2696 | 81.23 |  |  |  |  | 2520.9587 | 0.009121 | 57.2696 | 81.825 |  |  |  |  | 2520.9587 | 0.053273
|  |  | 57.3443 | 81.17 |  |  |  |  | 2524.2284 | 0.01167 | 57.3443 | 81.25 |  |  |  |  | 2524.2284 | 0.049696
|  |  | 57.4189 | 81.205 |  |  |  |  | 2527.4981 | 0.012756 | 57.4189 | 81.18 |  |  |  |  | 2527.4981 | 0.044594
|  |  | 57.4936 | 81.26 |  |  |  |  | 2530.7679 | 0.012019 | 57.4936 | 81.185 |  |  |  |  | 2530.7679 | 0.04409
|  |  | 57.5683 | 81.305 |  |  |  |  | 2534.0376 | 0.0096291 | 57.5683 | 81.83 |  |  |  |  | 2534.0376 | 0.050319
|  |  | 57.6429 | 81.33 |  |  |  |  | 2537.3073 | 0.0062054 | 57.6429 | 82.405 |  |  |  |  | 2537.3073 | 0.05743
|  |  | 57.7176 | 81.255 |  |  |  |  | 2540.577 | 0.0027413 | 57.7176 | 81.815 |  |  |  |  | 2540.577 | 0.058741
|  |  | 57.7923 | 81.17 |  |  |  |  | 2543.8468 | 0.0017407 | 57.7923 | 81.725 |  |  |  |  | 2543.8468 | 0.051016
|  |  | 57.8669 | 81.125 |  |  |  |  | 2547.1165 | 0.0029721 | 57.8669 | 81.535 |  |  |  |  | 2547.1165 | 0.035163
|  |  | 57.9416 | 81.11 |  |  |  |  | 2550.3862 | 0.0028639 | 57.9416 | 81.765 |  |  |  |  | 2550.3862 | 0.019115
|  |  | 58.0163 | 81.18 |  |  |  |  | 2553.6559 | 0.0015142 | 58.0163 | 82.065 |  |  |  |  | 2553.6559 | 0.026117
|  |  | 58.0909 | 81.225 |  |  |  |  | 2556.9257 | 0.0027372 | 58.0909 | 81.825 |  |  |  |  | 2556.9257 | 0.04455
|  |  | 58.1656 | 81.335 |  |  |  |  | 2560.1954 | 0.005474 | 58.1656 | 81.68 |  |  |  |  | 2560.1954 | 0.05793
|  |  | 58.2403 | 81.355 |  |  |  |  | 2563.4651 | 0.0069593 | 58.2403 | 81.98 |  |  |  |  | 2563.4651 | 0.062108
|  |  | 58.3149 | 81.31 |  |  |  |  | 2566.7348 | 0.0060973 | 58.3149 | 81.745 |  |  |  |  | 2566.7348 | 0.05637
|  |  | 58.3896 | 81.345 |  |  |  |  | 2570.0046 | 0.0029757 | 58.3896 | 82.435 |  |  |  |  | 2570.0046 | 0.042782
|  |  | 58.4643 | 81.375 |  |  |  |  | 2573.2743 | 0.0038016 | 58.4643 | 82.195 |  |  |  |  | 2573.2743 | 0.028027
|  |  | 58.5389 | 81.305 |  |  |  |  | 2576.544 | 0.0090819 | 58.5389 | 81.565 |  |  |  |  | 2576.544 | 0.026425
|  |  | 58.6136 | 81.17 |  |  |  |  | 2579.8137 | 0.013014 | 58.6136 | 81.425 |  |  |  |  | 2579.8137 | 0.036214
|  |  | 58.6883 | 81.295 |  |  |  |  | 2583.0835 | 0.014032 | 58.6883 | 82.155 |  |  |  |  | 2583.0835 | 0.041285
|  |  | 58.7629 | 81.485 |  |  |  |  | 2586.3532 | 0.011932 | 58.7629 | 82.155 |  |  |  |  | 2586.3532 | 0.036331
|  |  | 58.8376 | 81.535 |  |  |  |  | 2589.6229 | 0.0082 | 58.8376 | 82.305 |  |  |  |  | 2589.6229 | 0.026031
|  |  | 58.9123 | 81.5 |  |  |  |  | 2592.8927 | 0.0065689 | 58.9123 | 82.18 |  |  |  |  | 2592.8927 | 0.028681
|  |  | 58.9869 | 81.46 |  |  |  |  | 2596.1624 | 0.0083401 | 58.9869 | 81.82 |  |  |  |  | 2596.1624 | 0.045531
|  |  | 59.0616 | 81.4 |  |  |  |  | 2599.4321 | 0.0093177 | 59.0616 | 81.55 |  |  |  |  | 2599.4321 | 0.060379
|  |  | 59.1363 | 81.2 |  |  |  |  | 2602.7018 | 0.0077412 | 59.1363 | 81.57 |  |  |  |  | 2602.7018 | 0.068316
|  |  | 59.2109 | 81.185 |  |  |  |  | 2605.9716 | 0.0044183 | 59.2109 | 81.67 |  |  |  |  | 2605.9716 | 0.072035
|  |  | 59.2856 | 81.2 |  |  |  |  | 2609.2413 | 0.0037456 | 59.2856 | 81.315 |  |  |  |  | 2609.2413 | 0.076141
|  |  | 59.3603 | 81.165 |  |  |  |  | 2612.511 | 0.0067425 | 59.3603 | 81.04 |  |  |  |  | 2612.511 | 0.080204
|  |  | 59.4349 | 81.085 |  |  |  |  | 2615.7807 | 0.0084623 | 59.4349 | 81.225 |  |  |  |  | 2615.7807 | 0.078912
|  |  | 59.5096 | 81.005 |  |  |  |  | 2619.0505 | 0.0078942 | 59.5096 | 81.76 |  |  |  |  | 2619.0505 | 0.068511
|  |  | 59.5843 | 80.955 |  |  |  |  | 2622.3202 | 0.0057758 | 59.5843 | 82.01 |  |  |  |  | 2622.3202 | 0.050689
|  |  | 59.6589 | 81.115 |  |  |  |  | 2625.5899 | 0.0045595 | 59.6589 | 82.195 |  |  |  |  | 2625.5899 | 0.032389
|  |  | 59.7336 | 81.285 |  |  |  |  | 2628.8596 | 0.0059186 | 59.7336 | 81.945 |  |  |  |  | 2628.8596 | 0.021851
|  |  | 59.8083 | 81.29 |  |  |  |  | 2632.1294 | 0.0072164 | 59.8083 | 82.82 |  |  |  |  | 2632.1294 | 0.017515
|  |  | 59.8829 | 81.195 |  |  |  |  | 2635.3991 | 0.0069248 | 59.8829 | 82.78 |  |  |  |  | 2635.3991 | 0.014019
|  |  | 59.9576 | 81.13 |  |  |  |  | 2638.6688 | 0.005399 | 59.9576 | 81.765 |  |  |  |  | 2638.6688 | 0.021792
|  |  | 60.0323 | 81.2 |  |  |  |  | 2641.9385 | 0.0044516 | 60.0323 | 81.025 |  |  |  |  | 2641.9385 | 0.036229
|  |  | 60.1069 | 81.35 |  |  |  |  | 2645.2083 | 0.0051072 | 60.1069 | 81.39 |  |  |  |  | 2645.2083 | 0.045467
|  |  | 60.1816 | 81.415 |  |  |  |  | 2648.478 | 0.0055808 | 60.1816 | 82.145 |  |  |  |  | 2648.478 | 0.044875
|  |  | 60.2563 | 81.37 |  |  |  |  | 2651.7477 | 0.0053951 | 60.2563 | 81.91 |  |  |  |  | 2651.7477 | 0.037857
|  |  | 60.3309 | 81.305 |  |  |  |  | 2655.0174 | 0.0063648 | 60.3309 | 81.745 |  |  |  |  | 2655.0174 | 0.036406
|  |  | 60.4056 | 81.3 |  |  |  |  | 2658.2872 | 0.0086215 | 60.4056 | 81.455 |  |  |  |  | 2658.2872 | 0.04512
|  |  | 60.4803 | 81.32 |  |  |  |  | 2661.5569 | 0.0096901 | 60.4803 | 81.895 |  |  |  |  | 2661.5569 | 0.052673
|  |  | 60.5549 | 81.285 |  |  |  |  | 2664.8266 | 0.0077284 | 60.5549 | 81.845 |  |  |  |  | 2664.8266 | 0.052254
|  |  | 60.6296 | 81.25 |  |  |  |  | 2668.0963 | 0.0024533 | 60.6296 | 81.86 |  |  |  |  | 2668.0963 | 0.046273
|  |  | 60.7043 | 81.255 |  |  |  |  | 2671.3661 | 0.005108 | 60.7043 | 81.92 |  |  |  |  | 2671.3661 | 0.042411
|  |  | 60.7789 | 81.33 |  |  |  |  | 2674.6358 | 0.01233 | 60.7789 | 81.895 |  |  |  |  | 2674.6358 | 0.04281
|  |  | 60.8536 | 81.385 |  |  |  |  | 2677.9055 | 0.016858 | 60.8536 | 82.28 |  |  |  |  | 2677.9055 | 0.040255
|  |  | 60.9283 | 81.355 |  |  |  |  | 2681.1753 | 0.017046 | 60.9283 | 82.07 |  |  |  |  | 2681.1753 | 0.029777
|  |  | 61.0029 | 81.3 |  |  |  |  | 2684.445 | 0.012882 | 61.0029 | 82.08 |  |  |  |  | 2684.445 | 0.016426
|  |  | 61.0776 | 81.145 |  |  |  |  | 2687.7147 | 0.00637 | 61.0776 | 81.795 |  |  |  |  | 2687.7147 | 0.022166
|  |  | 61.1523 | 80.95 |  |  |  |  | 2690.9844 | 0.0046612 | 61.1523 | 81.75 |  |  |  |  | 2690.9844 | 0.036729
|  |  | 61.2269 | 80.99 |  |  |  |  | 2694.2542 | 0.0087433 | 61.2269 | 81.33 |  |  |  |  | 2694.2542 | 0.04401
|  |  | 61.3016 | 81.17 |  |  |  |  | 2697.5239 | 0.0099937 | 61.3016 | 81.175 |  |  |  |  | 2697.5239 | 0.041252
|  |  | 61.3763 | 81.325 |  |  |  |  | 2700.7936 | 0.0076294 | 61.3763 | 80.885 |  |  |  |  | 2700.7936 | 0.03068
|  |  | 61.4509 | 81.325 |  |  |  |  | 2704.0633 | 0.003614 | 61.4509 | 80.63 |  |  |  |  | 2704.0633 | 0.019204
|  |  | 61.5256 | 81.325 |  |  |  |  | 2707.3331 | 0.0044243 | 61.5256 | 81 |  |  |  |  | 2707.3331 | 0.023505
|  |  | 61.6003 | 81.36 |  |  |  |  | 2710.6028 | 0.0074926 | 61.6003 | 81.405 |  |  |  |  | 2710.6028 | 0.041472
|  |  | 61.6749 | 81.375 |  |  |  |  | 2713.8725 | 0.0083354 | 61.6749 | 81.29 |  |  |  |  | 2713.8725 | 0.059967
|  |  | 61.7496 | 81.34 |  |  |  |  | 2717.1422 | 0.0070462 | 61.7496 | 80.985 |  |  |  |  | 2717.1422 | 0.072191
|  |  | 61.8243 | 81.28 |  |  |  |  | 2720.412 | 0.0055847 | 61.8243 | 81.245 |  |  |  |  | 2720.412 | 0.07296
|  |  | 61.8989 | 81.315 |  |  |  |  | 2723.6817 | 0.0061397 | 61.8989 | 81.515 |  |  |  |  | 2723.6817 | 0.059738
|  |  | 61.9736 | 81.275 |  |  |  |  | 2726.9514 | 0.0073273 | 61.9736 | 82.105 |  |  |  |  | 2726.9514 | 0.034333
|  |  | 62.0483 | 81.3 |  |  |  |  | 2730.2211 | 0.0071511 | 62.0483 | 82.065 |  |  |  |  | 2730.2211 | 0.010535
|  |  | 62.1229 | 81.435 |  |  |  |  | 2733.4909 | 0.0051389 | 62.1229 | 82.81 |  |  |  |  | 2733.4909 | 0.035599
|  |  | 62.1976 | 81.365 |  |  |  |  | 2736.7606 | 0.0019283 | 62.1976 | 82.575 |  |  |  |  | 2736.7606 | 0.060741
|  |  | 62.2723 | 81.235 |  |  |  |  | 2740.0303 | 0.0023395 | 62.2723 | 82.47 |  |  |  |  | 2740.0303 | 0.074547
|  |  | 62.3469 | 81.035 |  |  |  |  | 2743.3 | 0.0048609 | 62.3469 | 82.68 |  |  |  |  | 2743.3 | 0.076189
|  |  | 62.4216 | 80.98 |  |  |  |  | 2746.5698 | 0.0052852 | 62.4216 | 82.435 |  |  |  |  | 2746.5698 | 0.068781
|  |  | 62.4963 | 81.015 |  |  |  |  | 2749.8395 | 0.0030419 | 62.4963 | 82.12 |  |  |  |  | 2749.8395 | 0.057803
|  |  | 62.5709 | 81.18 |  |  |  |  | 2753.1092 | 0.0020264 | 62.5709 | 82.145 |  |  |  |  | 2753.1092 | 0.04893
|  |  | 62.6456 | 81.21 |  |  |  |  | 2756.3789 | 0.0073386 | 62.6456 | 82.08 |  |  |  |  | 2756.3789 | 0.044655
|  |  | 62.7203 | 81.15 |  |  |  |  | 2759.6487 | 0.011797 | 62.7203 | 82.16 |  |  |  |  | 2759.6487 | 0.042413
|  |  | 62.7949 | 81.175 |  |  |  |  | 2762.9184 | 0.013581 | 62.7949 | 82.005 |  |  |  |  | 2762.9184 | 0.038507
|  |  | 62.8696 | 81.175 |  |  |  |  | 2766.1881 | 0.011998 | 62.8696 | 82.055 |  |  |  |  | 2766.1881 | 0.032547
|  |  | 62.9443 | 81.195 |  |  |  |  | 2769.4579 | 0.0077425 | 62.9443 | 82.18 |  |  |  |  | 2769.4579 | 0.027772
|  |  | 63.0189 | 81.255 |  |  |  |  | 2772.7276 | 0.0030824 | 63.0189 | 82.17 |  |  |  |  | 2772.7276 | 0.025788
|  |  | 63.0936 | 81.245 |  |  |  |  | 2775.9973 | 0.0034935 | 63.0936 | 81.97 |  |  |  |  | 2775.9973 | 0.02151
|  |  | 63.1683 | 81.225 |  |  |  |  | 2779.267 | 0.0049581 | 63.1683 | 81.405 |  |  |  |  | 2779.267 | 0.0095837
|  |  | 63.2429 | 81.245 |  |  |  |  | 2782.5368 | 0.0039632 | 63.2429 | 81.83 |  |  |  |  | 2782.5368 | 0.013875
|  |  | 63.3176 | 81.265 |  |  |  |  | 2785.8065 | 0.0019039 | 63.3176 | 82.155 |  |  |  |  | 2785.8065 | 0.039816
|  |  | 63.3923 | 81.235 |  |  |  |  | 2789.0762 | 0.0029658 | 63.3923 | 82.085 |  |  |  |  | 2789.0762 | 0.063296
|  |  | 63.4669 | 81.115 |  |  |  |  | 2792.3459 | 0.0039274 | 63.4669 | 81.53 |  |  |  |  | 2792.3459 | 0.07618
|  |  | 63.5416 | 81.19 |  |  |  |  | 2795.6157 | 0.0024894 | 63.5416 | 81.58 |  |  |  |  | 2795.6157 | 0.073757
|  |  | 63.6163 | 81.325 |  |  |  |  | 2798.8854 | 0.0019347 | 63.6163 | 82.09 |  |  |  |  | 2798.8854 | 0.056917
|  |  | 63.6909 | 81.325 |  |  |  |  | 2802.1551 | 0.0063221 | 63.6909 | 82.215 |  |  |  |  | 2802.1551 | 0.033509
|  |  | 63.7656 | 81.19 |  |  |  |  | 2805.4248 | 0.0099968 | 63.7656 | 81.78 |  |  |  |  | 2805.4248 | 0.022561
|  |  | 63.8403 | 81.185 |  |  |  |  | 2808.6946 | 0.011602 | 63.8403 | 81.395 |  |  |  |  | 2808.6946 | 0.030909
|  |  | 63.915 | 81.295 |  |  |  |  | 2811.9643 | 0.011211 | 63.915 | 81.1 |  |  |  |  | 2811.9643 | 0.034005
|  |  | 63.9896 | 81.255 |  |  |  |  | 2815.234 | 0.010094 | 63.9896 | 81.23 |  |  |  |  | 2815.234 | 0.028042
|  |  | 64.0643 | 81.22 |  |  |  |  | 2818.5037 | 0.0093559 | 64.0643 | 81.54 |  |  |  |  | 2818.5037 | 0.026304
|  |  | 64.139 | 81.225 |  |  |  |  | 2821.7735 | 0.0084109 | 64.139 | 81.82 |  |  |  |  | 2821.7735 | 0.037476
|  |  | 64.2136 | 81.22 |  |  |  |  | 2825.0432 | 0.0061534 | 64.2136 | 81.31 |  |  |  |  | 2825.0432 | 0.047648
|  |  | 64.2883 | 81.2 |  |  |  |  | 2828.3129 | 0.0028137 | 64.2883 | 80.51 |  |  |  |  | 2828.3129 | 0.049159
|  |  | 64.363 | 81.125 |  |  |  |  | 2831.5826 | 0.0024735 | 64.363 | 81.88 |  |  |  |  | 2831.5826 | 0.045022
|  |  | 64.4376 | 81.135 |  |  |  |  | 2834.8524 | 0.0047172 | 64.4376 | 82.7 |  |  |  |  | 2834.8524 | 0.044966
|  |  | 64.5123 | 81.1 |  |  |  |  | 2838.1221 | 0.0051339 | 64.5123 | 82.345 |  |  |  |  | 2838.1221 | 0.051469
|  |  | 64.587 | 81.06 |  |  |  |  | 2841.3918 | 0.00441 | 64.587 | 82.15 |  |  |  |  | 2841.3918 | 0.055329
|  |  | 64.6616 | 81.07 |  |  |  |  | 2844.6615 | 0.0061921 | 64.6616 | 82.005 |  |  |  |  | 2844.6615 | 0.049695
|  |  | 64.7363 | 81.08 |  |  |  |  | 2847.9313 | 0.010147 | 64.7363 | 81.715 |  |  |  |  | 2847.9313 | 0.034723
|  |  | 64.811 | 81.165 |  |  |  |  | 2851.201 | 0.013402 | 64.811 | 81.57 |  |  |  |  | 2851.201 | 0.016282
|  |  | 64.8856 | 81.275 |  |  |  |  | 2854.4707 | 0.014486 | 64.8856 | 81.86 |  |  |  |  | 2854.4707 | 0.0091785
|  |  | 64.9603 | 81.29 |  |  |  |  | 2857.7405 | 0.012967 | 64.9603 | 81.1 |  |  |  |  | 2857.7405 | 0.015227
|  |  | 65.035 | 81.36 |  |  |  |  | 2861.0102 | 0.0091848 | 65.035 | 81.23 |  |  |  |  | 2861.0102 | 0.015363
|  |  | 65.1096 | 81.32 |  |  |  |  | 2864.2799 | 0.0039929 | 65.1096 | 81.63 |  |  |  |  | 2864.2799 | 0.015406
|  |  | 65.1843 | 81.245 |  |  |  |  | 2867.5496 | 0.0015881 | 65.1843 | 81.855 |  |  |  |  | 2867.5496 | 0.025232
|  |  | 65.259 | 81.25 |  |  |  |  | 2870.8194 | 0.0063381 | 65.259 | 82.2 |  |  |  |  | 2870.8194 | 0.037366
|  |  | 65.3336 | 81.26 |  |  |  |  | 2874.0891 | 0.0095339 | 65.3336 | 81.89 |  |  |  |  | 2874.0891 | 0.044432
|  |  | 65.4083 | 81.205 |  |  |  |  | 2877.3588 | 0.010793 | 65.4083 | 81.63 |  |  |  |  | 2877.3588 | 0.043338
|  |  | 65.483 | 81.14 |  |  |  |  | 2880.6285 | 0.010301 | 65.483 | 81.68 |  |  |  |  | 2880.6285 | 0.033713
|  |  | 65.5576 | 81.095 |  |  |  |  | 2883.8983 | 0.0086297 | 65.5576 | 81.665 |  |  |  |  | 2883.8983 | 0.018185
|  |  | 65.6323 | 81.19 |  |  |  |  | 2887.168 | 0.0064554 | 65.6323 | 82.365 |  |  |  |  | 2887.168 | 0.017327
|  |  | 65.707 | 81.295 |  |  |  |  | 2890.4377 | 0.0047538 | 65.707 | 82.1 |  |  |  |  | 2890.4377 | 0.040066
|  |  | 65.7816 | 81.215 |  |  |  |  | 2893.7074 | 0.0054056 | 65.7816 | 81.08 |  |  |  |  | 2893.7074 | 0.062417
|  |  | 65.8563 | 81.23 |  |  |  |  | 2896.9772 | 0.0080568 | 65.8563 | 81.2 |  |  |  |  | 2896.9772 | 0.076143
|  |  | 65.931 | 81.3 |  |  |  |  | 2900.2469 | 0.010508 | 65.931 | 80.975 |  |  |  |  | 2900.2469 | 0.076534
|  |  | 66.0056 | 81.25 |  |  |  |  | 2903.5166 | 0.011381 | 66.0056 | 81.455 |  |  |  |  | 2903.5166 | 0.062717
|  |  | 66.0803 | 81.19 |  |  |  |  | 2906.7863 | 0.010088 | 66.0803 | 81.555 |  |  |  |  | 2906.7863 | 0.038928
|  |  | 66.155 | 81.28 |  |  |  |  | 2910.0561 | 0.0068028 | 66.155 | 81.825 |  |  |  |  | 2910.0561 | 0.019217
|  |  | 66.2296 | 81.265 |  |  |  |  | 2913.3258 | 0.0024188 | 66.2296 | 81.94 |  |  |  |  | 2913.3258 | 0.029737
|  |  | 66.3043 | 81.14 |  |  |  |  | 2916.5955 | 0.0034027 | 66.3043 | 81.7 |  |  |  |  | 2916.5955 | 0.044117
|  |  | 66.379 | 81.12 |  |  |  |  | 2919.8652 | 0.0081736 | 66.379 | 81.7 |  |  |  |  | 2919.8652 | 0.04831
|  |  | 66.4536 | 81.305 |  |  |  |  | 2923.135 | 0.012411 | 66.4536 | 81.735 |  |  |  |  | 2923.135 | 0.043169
|  |  | 66.5283 | 81.35 |  |  |  |  | 2926.4047 | 0.015581 | 66.5283 | 81.93 |  |  |  |  | 2926.4047 | 0.033441
|  |  | 66.603 | 81.25 |  |  |  |  | 2929.6744 | 0.017203 | 66.603 | 82.27 |  |  |  |  | 2929.6744 | 0.023698
|  |  | 66.6776 | 81.19 |  |  |  |  | 2932.9441 | 0.016973 | 66.6776 | 82.42 |  |  |  |  | 2932.9441 | 0.018782
|  |  | 66.7523 | 81.13 |  |  |  |  | 2936.2139 | 0.015042 | 66.7523 | 82.33 |  |  |  |  | 2936.2139 | 0.027051
|  |  | 66.827 | 81.2 |  |  |  |  | 2939.4836 | 0.01208 | 66.827 | 82.205 |  |  |  |  | 2939.4836 | 0.043325
|  |  | 66.9016 | 81.255 |  |  |  |  | 2942.7533 | 0.0090214 | 66.9016 | 81.945 |  |  |  |  | 2942.7533 | 0.056749
|  |  | 66.9763 | 81.415 |  |  |  |  | 2946.0231 | 0.0065869 | 66.9763 | 82.225 |  |  |  |  | 2946.0231 | 0.059437
|  |  | 67.051 | 81.41 |  |  |  |  | 2949.2928 | 0.0048712 | 67.051 | 82.375 |  |  |  |  | 2949.2928 | 0.047519
|  |  | 67.1256 | 81.305 |  |  |  |  | 2952.5625 | 0.0034148 | 67.1256 | 82.715 |  |  |  |  | 2952.5625 | 0.022679
|  |  | 67.2003 | 81.265 |  |  |  |  | 2955.8322 | 0.001762 | 67.2003 | 81.62 |  |  |  |  | 2955.8322 | 0.0095744
|  |  | 67.275 | 81.165 |  |  |  |  | 2959.102 | 0.0015947 | 67.275 | 81.59 |  |  |  |  | 2959.102 | 0.036555
|  |  | 67.3496 | 81.16 |  |  |  |  | 2962.3717 | 0.0043443 | 67.3496 | 82.135 |  |  |  |  | 2962.3717 | 0.051165
|  |  | 67.4243 | 81.24 |  |  |  |  | 2965.6414 | 0.007133 | 67.4243 | 82.765 |  |  |  |  | 2965.6414 | 0.048093
|  |  | 67.499 | 81.305 |  |  |  |  | 2968.9111 | 0.0086944 | 67.499 | 82.42 |  |  |  |  | 2968.9111 | 0.030282
|  |  | 67.5736 | 81.3 |  |  |  |  | 2972.1809 | 0.0081028 | 67.5736 | 82.085 |  |  |  |  | 2972.1809 | 0.022659
|  |  | 67.6483 | 81.17 |  |  |  |  | 2975.4506 | 0.0051262 | 67.6483 | 81.29 |  |  |  |  | 2975.4506 | 0.046373
|  |  | 67.723 | 81.085 |  |  |  |  | 2978.7203 | 0.00046374 | 67.723 | 81.61 |  |  |  |  | 2978.7203 | 0.066589
|  |  | 67.7976 | 81.115 |  |  |  |  | 2981.99 | 0.0044997 | 67.7976 | 82.1 |  |  |  |  | 2981.99 | 0.071743
|  |  | 67.8723 | 81.33 |  |  |  |  | 2985.2598 | 0.0081275 | 67.8723 | 82.305 |  |  |  |  | 2985.2598 | 0.062089
|  |  | 67.947 | 81.4 |  |  |  |  | 2988.5295 | 0.0093521 | 67.947 | 81.78 |  |  |  |  | 2988.5295 | 0.047693
|  |  | 68.0216 | 81.405 |  |  |  |  | 2991.7992 | 0.008161 | 68.0216 | 81.57 |  |  |  |  | 2991.7992 | 0.045612
|  |  | 68.0963 | 81.355 |  |  |  |  | 2995.0689 | 0.005968 | 68.0963 | 81.03 |  |  |  |  | 2995.0689 | 0.054627
|  |  | 68.171 | 81.32 |  |  |  |  | 2998.3387 | 0.0054581 | 68.171 | 81.145 |  |  |  |  | 2998.3387 | 0.057803
|  |  | 68.2456 | 81.325 |  |  |  |  | 3001.6084 | 0.0066977 | 68.2456 | 81.345 |  |  |  |  | 3001.6084 | 0.047797
|  |  | 68.3203 | 81.33 |  |  |  |  | 3004.8781 | 0.0072231 | 68.3203 | 82.265 |  |  |  |  | 3004.8781 | 0.026242
|  |  | 68.395 | 81.39 |  |  |  |  | 3008.1478 | 0.0062909 | 68.395 | 82.475 |  |  |  |  | 3008.1478 | 0.0010946
|  |  | 68.4696 | 81.39 |  |  |  |  | 3011.4176 | 0.0045593 | 68.4696 | 82.58 |  |  |  |  | 3011.4176 | 0.024323
|  |  | 68.5443 | 81.415 |  |  |  |  | 3014.6873 | 0.003637 | 68.5443 | 82.07 |  |  |  |  | 3014.6873 | 0.039165
|  |  | 68.619 | 81.445 |  |  |  |  | 3017.957 | 0.0047304 | 68.619 | 81.645 |  |  |  |  | 3017.957 | 0.042352
|  |  | 68.6936 | 81.385 |  |  |  |  | 3021.2267 | 0.0065704 | 68.6936 | 81.87 |  |  |  |  | 3021.2267 | 0.035638
|  |  | 68.7683 | 81.27 |  |  |  |  | 3024.4965 | 0.0078835 | 68.7683 | 81.97 |  |  |  |  | 3024.4965 | 0.025583
|  |  | 68.843 | 81.315 |  |  |  |  | 3027.7662 | 0.007849 | 68.843 | 81.665 |  |  |  |  | 3027.7662 | 0.023318
|  |  | 68.9176 | 81.43 |  |  |  |  | 3031.0359 | 0.0061589 | 68.9176 | 81.34 |  |  |  |  | 3031.0359 | 0.029641
|  |  | 68.9923 | 81.43 |  |  |  |  | 3034.3057 | 0.0040748 | 68.9923 | 81.525 |  |  |  |  | 3034.3057 | 0.034208
|  |  | 69.067 | 81.3 |  |  |  |  | 3037.5754 | 0.0057348 | 69.067 | 82.44 |  |  |  |  | 3037.5754 | 0.033104
|  |  | 69.1416 | 81.17 |  |  |  |  | 3040.8451 | 0.0096122 | 69.1416 | 82.665 |  |  |  |  | 3040.8451 | 0.028106
|  |  | 69.2163 | 81.09 |  |  |  |  | 3044.1148 | 0.012784 | 69.2163 | 83.075 |  |  |  |  | 3044.1148 | 0.02415
|  |  | 69.291 | 81.155 |  |  |  |  | 3047.3846 | 0.014606 | 69.291 | 81.995 |  |  |  |  | 3047.3846 | 0.023597
|  |  | 69.3656 | 81.24 |  |  |  |  | 3050.6543 | 0.015428 | 69.3656 | 81.355 |  |  |  |  | 3050.6543 | 0.021738
|  |  | 69.4403 | 81.285 |  |  |  |  | 3053.924 | 0.015703 | 69.4403 | 81.99 |  |  |  |  | 3053.924 | 0.013847
|  |  | 69.515 | 81.335 |  |  |  |  | 3057.1937 | 0.015269 | 69.515 | 82.475 |  |  |  |  | 3057.1937 | 0.0054533
|  |  | 69.5896 | 81.43 |  |  |  |  | 3060.4635 | 0.013575 | 69.5896 | 82.92 |  |  |  |  | 3060.4635 | 0.021537
|  |  | 69.6643 | 81.385 |  |  |  |  | 3063.7332 | 0.010448 | 69.6643 | 83.18 |  |  |  |  | 3063.7332 | 0.037863
|  |  | 69.739 | 81.265 |  |  |  |  | 3067.0029 | 0.0064659 | 69.739 | 82.845 |  |  |  |  | 3067.0029 | 0.046524
|  |  | 69.8136 | 81.2 |  |  |  |  | 3070.2726 | 0.0028518 | 69.8136 | 82.655 |  |  |  |  | 3070.2726 | 0.044073
|  |  | 69.8883 | 81.165 |  |  |  |  | 3073.5424 | 0.0018865 | 69.8883 | 82.135 |  |  |  |  | 3073.5424 | 0.031758
|  |  | 69.963 | 81.195 |  |  |  |  | 3076.8121 | 0.00281 | 69.963 | 82.07 |  |  |  |  | 3076.8121 | 0.01992
|  |  | 70.0376 | 81.215 |  |  |  |  | 3080.0818 | 0.0031719 | 70.0376 | 82.31 |  |  |  |  | 3080.0818 | 0.027814
|  |  | 70.1123 | 81.255 |  |  |  |  | 3083.3515 | 0.0032476 | 70.1123 | 82.245 |  |  |  |  | 3083.3515 | 0.040747
|  |  | 70.187 | 81.335 |  |  |  |  | 3086.6213 | 0.0031448 | 70.187 | 82.33 |  |  |  |  | 3086.6213 | 0.044897
|  |  | 70.2617 | 81.37 |  |  |  |  | 3089.891 | 0.0025609 | 70.2617 | 82.42 |  |  |  |  | 3089.891 | 0.038064
|  |  | 70.3363 | 81.25 |  |  |  |  | 3093.1607 | 0.0016378 | 70.3363 | 82.855 |  |  |  |  | 3093.1607 | 0.023534
|  |  | 70.411 | 81.205 |  |  |  |  | 3096.4304 | 0.0017109 | 70.411 | 82.525 |  |  |  |  | 3096.4304 | 0.0077354
|  |  | 70.4856 | 81.215 |  |  |  |  | 3099.7002 | 0.0031257 | 70.4856 | 82.05 |  |  |  |  | 3099.7002 | 0.0041292
|  |  | 70.5603 | 81.22 |  |  |  |  | 3102.9699 | 0.0051285 | 70.5603 | 81.66 |  |  |  |  | 3102.9699 | 0.0061262
|  |  | 70.635 | 81.195 |  |  |  |  | 3106.2396 | 0.0076301 | 70.635 | 81.09 |  |  |  |  | 3106.2396 | 0.0072808
|  |  | 70.7096 | 81.23 |  |  |  |  | 3109.5093 | 0.010067 | 70.7096 | 81.44 |  |  |  |  | 3109.5093 | 0.020015
|  |  | 70.7843 | 81.265 |  |  |  |  | 3112.7791 | 0.011401 | 70.7843 | 81.675 |  |  |  |  | 3112.7791 | 0.034271
|  |  | 70.859 | 81.19 |  |  |  |  | 3116.0488 | 0.010738 | 70.859 | 82.115 |  |  |  |  | 3116.0488 | 0.044155
|  |  | 70.9336 | 81.135 |  |  |  |  | 3119.3185 | 0.0080052 | 70.9336 | 81.815 |  |  |  |  | 3119.3185 | 0.049023
|  |  | 71.0083 | 81.135 |  |  |  |  | 3122.5883 | 0.0050938 | 71.0083 | 81.6 |  |  |  |  | 3122.5883 | 0.053627
|  |  | 71.083 | 81.19 |  |  |  |  | 3125.858 | 0.0067438 | 71.083 | 81.875 |  |  |  |  | 3125.858 | 0.061956
|  |  | 71.1577 | 81.105 |  |  |  |  | 3129.1277 | 0.010438 | 71.1577 | 82.495 |  |  |  |  | 3129.1277 | 0.069799
|  |  | 71.2323 | 81.055 |  |  |  |  | 3132.3974 | 0.012382 | 71.2323 | 82.405 |  |  |  |  | 3132.3974 | 0.069041
|  |  | 71.307 | 81.18 |  |  |  |  | 3135.6672 | 0.011546 | 71.307 | 82.545 |  |  |  |  | 3135.6672 | 0.054796
|  |  | 71.3817 | 81.265 |  |  |  |  | 3138.9369 | 0.0083456 | 71.3817 | 82.49 |  |  |  |  | 3138.9369 | 0.028055
|  |  | 71.4563 | 81.285 |  |  |  |  | 3142.2066 | 0.0047591 | 71.4563 | 81.865 |  |  |  |  | 3142.2066 | 0.005546
|  |  | 71.531 | 81.26 |  |  |  |  | 3145.4763 | 0.0051957 | 71.531 | 81.395 |  |  |  |  | 3145.4763 | 0.034885
|  |  | 71.6057 | 81.27 |  |  |  |  | 3148.7461 | 0.0080847 | 71.6057 | 81.165 |  |  |  |  | 3148.7461 | 0.052668
|  |  | 71.6803 | 81.24 |  |  |  |  | 3152.0158 | 0.010188 | 71.6803 | 81.07 |  |  |  |  | 3152.0158 | 0.053695
|  |  | 71.755 | 81.165 |  |  |  |  | 3155.2855 | 0.011543 | 71.755 | 80.765 |  |  |  |  | 3155.2855 | 0.040047
|  |  | 71.8297 | 81.195 |  |  |  |  | 3158.5552 | 0.012766 | 71.8297 | 81.065 |  |  |  |  | 3158.5552 | 0.026852
|  |  | 71.9043 | 81.28 |  |  |  |  | 3161.825 | 0.013601 | 71.9043 | 81.435 |  |  |  |  | 3161.825 | 0.039393
|  |  | 71.979 | 81.36 |  |  |  |  | 3165.0947 | 0.013067 | 71.979 | 81.775 |  |  |  |  | 3165.0947 | 0.058585
|  |  | 72.0537 | 81.34 |  |  |  |  | 3168.3644 | 0.010534 | 72.0537 | 80.97 |  |  |  |  | 3168.3644 | 0.067852
|  |  | 72.1283 | 81.315 |  |  |  |  | 3171.6341 | 0.0064194 | 72.1283 | 81.21 |  |  |  |  | 3171.6341 | 0.064587
|  |  | 72.203 | 81.195 |  |  |  |  | 3174.9039 | 0.002865 | 72.203 | 82.005 |  |  |  |  | 3174.9039 | 0.052371
|  |  | 72.2777 | 81.13 |  |  |  |  | 3178.1736 | 0.004068 | 72.2777 | 82.03 |  |  |  |  | 3178.1736 | 0.037095
|  |  | 72.3523 | 81.15 |  |  |  |  | 3181.4433 | 0.0054293 | 72.3523 | 82.07 |  |  |  |  | 3181.4433 | 0.02303
|  |  | 72.427 | 81.23 |  |  |  |  | 3184.713 | 0.0050038 | 72.427 | 82.535 |  |  |  |  | 3184.713 | 0.012094
|  |  | 72.5017 | 81.255 |  |  |  |  | 3187.9828 | 0.0053616 | 72.5017 | 82.855 |  |  |  |  | 3187.9828 | 0.011697
|  |  | 72.5763 | 81.185 |  |  |  |  | 3191.2525 | 0.0086028 | 72.5763 | 82.895 |  |  |  |  | 3191.2525 | 0.018393
|  |  | 72.651 | 81.13 |  |  |  |  | 3194.5222 | 0.012015 | 72.651 | 82.345 |  |  |  |  | 3194.5222 | 0.018796
|  |  | 72.7257 | 81.155 |  |  |  |  | 3197.7919 | 0.013437 | 72.7257 | 81.955 |  |  |  |  | 3197.7919 | 0.0081425
|  |  | 72.8003 | 81.095 |  |  |  |  | 3201.0617 | 0.012135 | 72.8003 | 82.315 |  |  |  |  | 3201.0617 | 0.013069
|  |  | 72.875 | 81 |  |  |  |  | 3204.3314 | 0.0085399 | 72.875 | 82.015 |  |  |  |  | 3204.3314 | 0.037737
|  |  | 72.9497 | 80.99 |  |  |  |  | 3207.6011 | 0.0039337 | 72.9497 | 82.26 |  |  |  |  | 3207.6011 | 0.056592
|  |  | 73.0243 | 81.02 |  |  |  |  | 3210.8708 | 0.001616 | 73.0243 | 81.78 |  |  |  |  | 3210.8708 | 0.060463
|  |  | 73.099 | 81.09 |  |  |  |  | 3214.1406 | 0.0041419 | 73.099 | 81.63 |  |  |  |  | 3214.1406 | 0.04527
|  |  | 73.1737 | 81.07 |  |  |  |  | 3217.4103 | 0.0054037 | 73.1737 | 82.135 |  |  |  |  | 3217.4103 | 0.020835
|  |  | 73.2483 | 81.045 |  |  |  |  | 3220.68 | 0.0051344 | 73.2483 | 82.375 |  |  |  |  | 3220.68 | 0.041888
|  |  | 73.323 | 81.065 |  |  |  |  | 3223.9498 | 0.003928 | 73.323 | 81.585 |  |  |  |  | 3223.9498 | 0.077669
|  |  | 73.3977 | 81.215 |  |  |  |  | 3227.2195 | 0.0029246 | 73.3977 | 81.76 |  |  |  |  | 3227.2195 | 0.097548
|  |  | 73.4723 | 81.34 |  |  |  |  | 3230.4892 | 0.003307 | 73.4723 | 82.03 |  |  |  |  | 3230.4892 | 0.092628
|  |  | 73.547 | 81.34 |  |  |  |  | 3233.7589 | 0.0040734 | 73.547 | 81.975 |  |  |  |  | 3233.7589 | 0.063401
|  |  | 73.6217 | 81.225 |  |  |  |  | 3237.0287 | 0.0043878 | 73.6217 | 81.895 |  |  |  |  | 3237.0287 | 0.018996
|  |  | 73.6963 | 81.215 |  |  |  |  | 3240.2984 | 0.0057194 | 73.6963 | 82.085 |  |  |  |  | 3240.2984 | 0.028568
|  |  | 73.771 | 81.335 |  |  |  |  | 3243.5681 | 0.0093638 | 73.771 | 81.715 |  |  |  |  | 3243.5681 | 0.062678
|  |  | 73.8457 | 81.36 |  |  |  |  | 3246.8378 | 0.01371 | 73.8457 | 82.235 |  |  |  |  | 3246.8378 | 0.075661
|  |  | 73.9203 | 81.24 |  |  |  |  | 3250.1076 | 0.016608 | 73.9203 | 82.795 |  |  |  |  | 3250.1076 | 0.065866
|  |  | 73.995 | 81.21 |  |  |  |  | 3253.3773 | 0.016541 | 73.995 | 82.89 |  |  |  |  | 3253.3773 | 0.039058
|  |  | 74.0697 | 81.195 |  |  |  |  | 3256.647 | 0.013155 | 74.0697 | 82.865 |  |  |  |  | 3256.647 | 0.0051234
|  |  | 74.1443 | 81.27 |  |  |  |  | 3259.9167 | 0.0075972 | 74.1443 | 82.66 |  |  |  |  | 3259.9167 | 0.027086
|  |  | 74.219 | 81.245 |  |  |  |  | 3263.1865 | 0.0035367 | 74.219 | 82.455 |  |  |  |  | 3263.1865 | 0.052498
|  |  | 74.2937 | 81.2 |  |  |  |  | 3266.4562 | 0.0055617 | 74.2937 | 82.07 |  |  |  |  | 3266.4562 | 0.071055
|  |  | 74.3683 | 81.305 |  |  |  |  | 3269.7259 | 0.0065765 | 74.3683 | 82.105 |  |  |  |  | 3269.7259 | 0.083254
|  |  | 74.443 | 81.41 |  |  |  |  | 3272.9956 | 0.004772 | 74.443 | 81.645 |  |  |  |  | 3272.9956 | 0.087063
|  |  | 74.5177 | 81.34 |  |  |  |  | 3276.2654 | 0.0041602 | 74.5177 | 80.71 |  |  |  |  | 3276.2654 | 0.079918
|  |  | 74.5923 | 81.295 |  |  |  |  | 3279.5351 | 0.0080879 | 74.5923 | 81.02 |  |  |  |  | 3279.5351 | 0.064038
|  |  | 74.667 | 81.22 |  |  |  |  | 3282.8048 | 0.011271 | 74.667 | 81.455 |  |  |  |  | 3282.8048 | 0.050952
|  |  | 74.7417 | 81.245 |  |  |  |  | 3286.0745 | 0.011508 | 74.7417 | 81.88 |  |  |  |  | 3286.0745 | 0.05317
|  |  | 74.8163 | 81.28 |  |  |  |  | 3289.3443 | 0.0086944 | 74.8163 | 82.12 |  |  |  |  | 3289.3443 | 0.060002
|  |  | 74.891 | 81.21 |  |  |  |  | 3292.614 | 0.0041479 | 74.891 | 81.865 |  |  |  |  | 3292.614 | 0.056186
|  |  | 74.9657 | 81.175 |  |  |  |  | 3295.8837 | 0.0010567 | 74.9657 | 81.455 |  |  |  |  | 3295.8837 | 0.040315
|  |  | 75.0403 | 81.175 |  |  |  |  | 3299.1534 | 0.002907 | 75.0403 | 81.83 |  |  |  |  | 3299.1534 | 0.030401
|  |  | 75.115 | 81.22 |  |  |  |  | 3302.4232 | 0.0022849 | 75.115 | 82.345 |  |  |  |  | 3302.4232 | 0.046343
|  |  | 75.1897 | 81.3 |  |  |  |  | 3305.6929 | 0.0011255 | 75.1897 | 82.06 |  |  |  |  | 3305.6929 | 0.063501
|  |  | 75.2643 | 81.285 |  |  |  |  | 3308.9626 | 0.0059067 | 75.2643 | 82.105 |  |  |  |  | 3308.9626 | 0.067861
|  |  | 75.339 | 81.265 |  |  |  |  | 3312.2324 | 0.010587 | 75.339 | 82.345 |  |  |  |  | 3312.2324 | 0.060522
|  |  | 75.4137 | 81.255 |  |  |  |  | 3315.5021 | 0.013642 | 75.4137 | 82.695 |  |  |  |  | 3315.5021 | 0.052084
|  |  | 75.4883 | 81.235 |  |  |  |  | 3318.7718 | 0.014161 | 75.4883 | 82.56 |  |  |  |  | 3318.7718 | 0.052111
|  |  | 75.563 | 81.23 |  |  |  |  | 3322.0415 | 0.011982 | 75.563 | 81.975 |  |  |  |  | 3322.0415 | 0.0543
|  |  | 75.6377 | 81.165 |  |  |  |  | 3325.3113 | 0.0077341 | 75.6377 | 81.95 |  |  |  |  | 3325.3113 | 0.048799
|  |  | 75.7123 | 81.18 |  |  |  |  | 3328.581 | 0.0035065 | 75.7123 | 82.005 |  |  |  |  | 3328.581 | 0.035489
|  |  | 75.787 | 81.115 |  |  |  |  | 3331.8507 | 0.0053956 | 75.787 | 81.8 |  |  |  |  | 3331.8507 | 0.023938
|  |  | 75.8617 | 80.945 |  |  |  |  | 3335.1204 | 0.0090954 | 75.8617 | 81.685 |  |  |  |  | 3335.1204 | 0.02199
|  |  | 75.9363 | 81.025 |  |  |  |  | 3338.3902 | 0.011193 | 75.9363 | 81.795 |  |  |  |  | 3338.3902 | 0.017623
|  |  | 76.011 | 81.17 |  |  |  |  | 3341.6599 | 0.011923 | 76.011 | 81.86 |  |  |  |  | 3341.6599 | 0.010056
|  |  | 76.0857 | 81.445 |  |  |  |  | 3344.9296 | 0.012513 | 76.0857 | 82.23 |  |  |  |  | 3344.9296 | 0.033939
|  |  | 76.1603 | 81.5 |  |  |  |  | 3348.1993 | 0.013384 | 76.1603 | 81.92 |  |  |  |  | 3348.1993 | 0.065206
|  |  | 76.235 | 81.475 |  |  |  |  | 3351.4691 | 0.013387 | 76.235 | 81.89 |  |  |  |  | 3351.4691 | 0.088991
|  |  | 76.3097 | 81.35 |  |  |  |  | 3354.7388 | 0.011301 | 76.3097 | 82.235 |  |  |  |  | 3354.7388 | 0.096715
|  |  | 76.3843 | 81.295 |  |  |  |  | 3358.0085 | 0.0070939 | 76.3843 | 82.28 |  |  |  |  | 3358.0085 | 0.085795
|  |  | 76.459 | 81.305 |  |  |  |  | 3361.2782 | 0.0020783 | 76.459 | 82.47 |  |  |  |  | 3361.2782 | 0.061528
|  |  | 76.5337 | 81.355 |  |  |  |  | 3364.548 | 0.0025122 | 76.5337 | 82.67 |  |  |  |  | 3364.548 | 0.040923
|  |  | 76.6083 | 81.27 |  |  |  |  | 3367.8177 | 0.0042406 | 76.6083 | 82.97 |  |  |  |  | 3367.8177 | 0.047517
|  |  | 76.683 | 81.2 |  |  |  |  | 3371.0874 | 0.0031537 | 76.683 | 82.99 |  |  |  |  | 3371.0874 | 0.063631
|  |  | 76.7577 | 81.205 |  |  |  |  | 3374.3571 | 0.0015967 | 76.7577 | 82.425 |  |  |  |  | 3374.3571 | 0.068759
|  |  | 76.8323 | 81.16 |  |  |  |  | 3377.6269 | 0.0047007 | 76.8323 | 81.04 |  |  |  |  | 3377.6269 | 0.059035
|  |  | 76.907 | 81.155 |  |  |  |  | 3380.8966 | 0.0069834 | 76.907 | 80.925 |  |  |  |  | 3380.8966 | 0.038292
|  |  | 76.9817 | 81.225 |  |  |  |  | 3384.1663 | 0.006675 | 76.9817 | 82.065 |  |  |  |  | 3384.1663 | 0.01745
|  |  | 77.0563 | 81.27 |  |  |  |  | 3387.436 | 0.0040387 | 77.0563 | 82.32 |  |  |  |  | 3387.436 | 0.02321
|  |  | 77.131 | 81.375 |  |  |  |  | 3390.7058 | 0.0030597 | 77.131 | 81.86 |  |  |  |  | 3390.7058 | 0.037114
|  |  | 77.2057 | 81.41 |  |  |  |  | 3393.9755 | 0.0062097 | 77.2057 | 82.065 |  |  |  |  | 3393.9755 | 0.041364
|  |  | 77.2803 | 81.33 |  |  |  |  | 3397.2452 | 0.008015 | 77.2803 | 82.15 |  |  |  |  | 3397.2452 | 0.034597
|  |  | 77.355 | 81.22 |  |  |  |  | 3400.515 | 0.0070592 | 77.355 | 81.63 |  |  |  |  | 3400.515 | 0.020578
|  |  | 77.4297 | 81.11 |  |  |  |  | 3403.7847 | 0.0036591 | 77.4297 | 81.79 |  |  |  |  | 3403.7847 | 0.012123
|  |  | 77.5043 | 81.155 |  |  |  |  | 3407.0544 | 0.0012126 | 77.5043 | 82.055 |  |  |  |  | 3407.0544 | 0.021791
|  |  | 77.579 | 81.235 |  |  |  |  | 3410.3241 | 0.0056346 | 77.579 | 82.005 |  |  |  |  | 3410.3241 | 0.029117
|  |  | 77.6537 | 81.265 |  |  |  |  | 3413.5939 | 0.0090813 | 77.6537 | 81.81 |  |  |  |  | 3413.5939 | 0.028302
|  |  | 77.7283 | 81.255 |  |  |  |  | 3416.8636 | 0.011155 | 77.7283 | 81.855 |  |  |  |  | 3416.8636 | 0.024706
|  |  | 77.803 | 81.235 |  |  |  |  | 3420.1333 | 0.011882 | 77.803 | 81.49 |  |  |  |  | 3420.1333 | 0.031993
|  |  | 77.8777 | 81.26 |  |  |  |  | 3423.403 | 0.011249 | 77.8777 | 81.28 |  |  |  |  | 3423.403 | 0.047969
|  |  | 77.9523 | 81.225 |  |  |  |  | 3426.6728 | 0.0092997 | 77.9523 | 81.555 |  |  |  |  | 3426.6728 | 0.060497
|  |  | 78.027 | 81.27 |  |  |  |  | 3429.9425 | 0.0065361 | 78.027 | 81.425 |  |  |  |  | 3429.9425 | 0.062346
|  |  | 78.1017 | 81.28 |  |  |  |  | 3433.2122 | 0.0040596 | 78.1017 | 81.42 |  |  |  |  | 3433.2122 | 0.051076
|  |  | 78.1763 | 81.12 |  |  |  |  | 3436.4819 | 0.0029241 | 78.1763 | 81.685 |  |  |  |  | 3436.4819 | 0.028874
|  |  | 78.251 | 81.06 |  |  |  |  | 3439.7517 | 0.0030091 | 78.251 | 81.88 |  |  |  |  | 3439.7517 | 0.0039942
|  |  | 78.3257 | 81.16 |  |  |  |  | 3443.0214 | 0.0046416 | 78.3257 | 82.215 |  |  |  |  | 3443.0214 | 0.023101
|  |  | 78.4004 | 81.255 |  |  |  |  | 3446.2911 | 0.007436 | 78.4004 | 82.085 |  |  |  |  | 3446.2911 | 0.036382
|  |  | 78.475 | 81.255 |  |  |  |  | 3449.5608 | 0.0097123 | 78.475 | 82.5 |  |  |  |  | 3449.5608 | 0.035027
|  |  | 78.5497 | 81.23 |  |  |  |  | 3452.8306 | 0.01011 | 78.5497 | 81.745 |  |  |  |  | 3452.8306 | 0.022214
|  |  | 78.6244 | 81.115 |  |  |  |  | 3456.1003 | 0.0083454 | 78.6244 | 82.13 |  |  |  |  | 3456.1003 | 0.021743
|  |  | 78.699 | 81.085 |  |  |  |  | 3459.37 | 0.0058607 | 78.699 | 81.715 |  |  |  |  | 3459.37 | 0.041265
|  |  | 78.7737 | 81.215 |  |  |  |  | 3462.6397 | 0.0058663 | 78.7737 | 81.49 |  |  |  |  | 3462.6397 | 0.053755
|  |  | 78.8484 | 81.23 |  |  |  |  | 3465.9095 | 0.0076629 | 78.8484 | 81.815 |  |  |  |  | 3465.9095 | 0.049277
|  |  | 78.923 | 81.265 |  |  |  |  | 3469.1792 | 0.0081204 | 78.923 | 81.34 |  |  |  |  | 3469.1792 | 0.027501
|  |  | 78.9977 | 81.22 |  |  |  |  | 3472.4489 | 0.0064698 | 78.9977 | 80.91 |  |  |  |  | 3472.4489 | 0.017447
|  |  | 79.0723 | 81.265 |  |  |  |  | 3475.7186 | 0.0034818 | 79.0723 | 81.715 |  |  |  |  | 3475.7186 | 0.0513
|  |  | 79.147 | 81.275 |  |  |  |  | 3478.9884 | 0.0015796 | 79.147 | 82.13 |  |  |  |  | 3478.9884 | 0.078651
|  |  | 79.2217 | 81.19 |  |  |  |  | 3482.2581 | 0.0038121 | 79.2217 | 81.82 |  |  |  |  | 3482.2581 | 0.087929
|  |  | 79.2964 | 81.225 |  |  |  |  | 3485.5278 | 0.0067347 | 79.2964 | 81.51 |  |  |  |  | 3485.5278 | 0.076947
|  |  | 79.371 | 81.36 |  |  |  |  | 3488.7976 | 0.010146 | 79.371 | 81.66 |  |  |  |  | 3488.7976 | 0.051554
|  |  | 79.4457 | 81.37 |  |  |  |  | 3492.0673 | 0.013541 | 79.4457 | 82.135 |  |  |  |  | 3492.0673 | 0.026364
|  |  | 79.5204 | 81.21 |  |  |  |  | 3495.337 | 0.015845 | 79.5204 | 81.99 |  |  |  |  | 3495.337 | 0.025842
|  |  | 79.595 | 81.09 |  |  |  |  | 3498.6067 | 0.016335 | 79.595 | 82.13 |  |  |  |  | 3498.6067 | 0.034076
|  |  | 79.6697 | 81.035 |  |  |  |  | 3501.8765 | 0.015286 | 79.6697 | 82.68 |  |  |  |  | 3501.8765 | 0.033903
|  |  | 79.7444 | 81.17 |  |  |  |  | 3505.1462 | 0.014026 | 79.7444 | 82.79 |  |  |  |  | 3505.1462 | 0.035231
|  |  | 79.819 | 81.325 |  |  |  |  | 3508.4159 | 0.013897 | 79.819 | 81.87 |  |  |  |  | 3508.4159 | 0.048298
|  |  | 79.8937 | 81.24 |  |  |  |  | 3511.6856 | 0.014561 | 79.8937 | 81.73 |  |  |  |  | 3511.6856 | 0.063116
|  |  | 79.9684 | 81.135 |  |  |  |  | 3514.9554 | 0.014451 | 79.9684 | 81.67 |  |  |  |  | 3514.9554 | 0.069621
|  |  | 80.043 | 81.085 |  |  |  |  | 3518.2251 | 0.012434 | 80.043 | 81.12 |  |  |  |  | 3518.2251 | 0.065651
|  |  | 80.1177 | 81.095 |  |  |  |  | 3521.4948 | 0.008401 | 80.1177 | 80.825 |  |  |  |  | 3521.4948 | 0.054977
|  |  | 80.1924 | 81.135 |  |  |  |  | 3524.7645 | 0.0041135 | 80.1924 | 81.44 |  |  |  |  | 3524.7645 | 0.043969
|  |  | 80.267 | 81.395 |  |  |  |  | 3528.0343 | 0.0067062 | 80.267 | 81.745 |  |  |  |  | 3528.0343 | 0.037796
|  |  | 80.3417 | 81.435 |  |  |  |  | 3531.304 | 0.012301 | 80.3417 | 81.315 |  |  |  |  | 3531.304 | 0.037689
|  |  | 80.4164 | 81.29 |  |  |  |  | 3534.5737 | 0.016575 | 80.4164 | 81.355 |  |  |  |  | 3534.5737 | 0.040437
|  |  | 80.491 | 81.045 |  |  |  |  | 3537.8434 | 0.018473 | 80.491 | 81.83 |  |  |  |  | 3537.8434 | 0.039984
|  |  | 80.5657 | 81.05 |  |  |  |  | 3541.1132 | 0.017809 | 80.5657 | 82.41 |  |  |  |  | 3541.1132 | 0.030681
|  |  | 80.6404 | 81.205 |  |  |  |  | 3544.3829 | 0.014973 | 80.6404 | 82.535 |  |  |  |  | 3544.3829 | 0.010693
|  |  | 80.715 | 81.25 |  |  |  |  | 3547.6526 | 0.010684 | 80.715 | 81.67 |  |  |  |  | 3547.6526 | 0.01646
|  |  | 80.7897 | 81.12 |  |  |  |  | 3550.9223 | 0.0058343 | 80.7897 | 81.225 |  |  |  |  | 3550.9223 | 0.042561
|  |  | 80.8644 | 81.115 |  |  |  |  | 3554.1921 | 0.0020083 | 80.8644 | 81.21 |  |  |  |  | 3554.1921 | 0.058587
|  |  | 80.939 | 81.2 |  |  |  |  | 3557.4618 | 0.0034561 | 80.939 | 81.47 |  |  |  |  | 3557.4618 | 0.059082
|  |  | 81.0137 | 81.175 |  |  |  |  | 3560.7315 | 0.0051295 | 81.0137 | 81.96 |  |  |  |  | 3560.7315 | 0.047332
|  |  | 81.0884 | 81.08 |  |  |  |  | 3564.0012 | 0.0052146 | 81.0884 | 82.085 |  |  |  |  | 3564.0012 | 0.041094
|  |  | 81.163 | 80.965 |  |  |  |  | 3567.271 | 0.005253 | 81.163 | 81.68 |  |  |  |  | 3567.271 | 0.053364
|  |  | 81.2377 | 81 |  |  |  |  | 3570.5407 | 0.0073445 | 81.2377 | 82.3 |  |  |  |  | 3570.5407 | 0.064299
|  |  | 81.3124 | 81.05 |  |  |  |  | 3573.8104 | 0.010149 | 81.3124 | 83.085 |  |  |  |  | 3573.8104 | 0.058927
|  |  | 81.387 | 81.14 |  |  |  |  | 3577.0802 | 0.01158 | 81.387 | 83.3 |  |  |  |  | 3577.0802 | 0.037394
|  |  | 81.4617 | 81.15 |  |  |  |  | 3580.3499 | 0.010704 | 81.4617 | 82.405 |  |  |  |  | 3580.3499 | 0.026602
|  |  | 81.5364 | 81.2 |  |  |  |  | 3583.6196 | 0.0077677 | 81.5364 | 81.95 |  |  |  |  | 3583.6196 | 0.053427
|  |  | 81.611 | 81.27 |  |  |  |  | 3586.8893 | 0.0039319 | 81.611 | 82.13 |  |  |  |  | 3586.8893 | 0.077119
|  |  | 81.6857 | 81.21 |  |  |  |  | 3590.1591 | 0.00071225 | 81.6857 | 82.31 |  |  |  |  | 3590.1591 | 0.083978
|  |  | 81.7604 | 81.22 |  |  |  |  | 3593.4288 | 0.00077923 | 81.7604 | 81.615 |  |  |  |  | 3593.4288 | 0.074871
|  |  | 81.835 | 81.315 |  |  |  |  | 3596.6985 | 0.0011754 | 81.835 | 81.155 |  |  |  |  | 3596.6985 | 0.058957
|  |  | 81.9097 | 81.34 |  |  |  |  | 3599.9682 | 0.0034744 | 81.9097 | 81.765 |  |  |  |  | 3599.9682 | 0.046539
|  |  | 81.9844 | 81.27 |  |  |  |  | 3603.238 | 0.0066085 | 81.9844 | 81.99 |  |  |  |  | 3603.238 | 0.03737
|  |  | 82.059 | 81.065 |  |  |  |  | 3606.5077 | 0.009088 | 82.059 | 81.815 |  |  |  |  | 3606.5077 | 0.025756
|  |  | 82.1337 | 81.09 |  |  |  |  | 3609.7774 | 0.010119 | 82.1337 | 81.865 |  |  |  |  | 3609.7774 | 0.026862
|  |  | 82.2084 | 81.075 |  |  |  |  | 3613.0471 | 0.0096886 | 82.2084 | 81.45 |  |  |  |  | 3613.0471 | 0.048897
|  |  | 82.283 | 81.05 |  |  |  |  | 3616.3169 | 0.0083797 | 82.283 | 81.255 |  |  |  |  | 3616.3169 | 0.067801
|  |  | 82.3577 | 81.035 |  |  |  |  | 3619.5866 | 0.0069318 | 82.3577 | 81.18 |  |  |  |  | 3619.5866 | 0.071048
|  |  | 82.4324 | 81.04 |  |  |  |  | 3622.8563 | 0.0059408 | 82.4324 | 81.01 |  |  |  |  | 3622.8563 | 0.05538
|  |  | 82.507 | 81.04 |  |  |  |  | 3626.126 | 0.0057827 | 82.507 | 81.26 |  |  |  |  | 3626.126 | 0.026471
|  |  | 82.5817 | 81.085 |  |  |  |  | 3629.3958 | 0.006112 | 82.5817 | 81.13 |  |  |  |  | 3629.3958 | 0.01996
|  |  | 82.6564 | 81.15 |  |  |  |  | 3632.6655 | 0.0058277 | 82.6564 | 81.375 |  |  |  |  | 3632.6655 | 0.046757
|  |  | 82.731 | 81.215 |  |  |  |  | 3635.9352 | 0.0040026 | 82.731 | 81.47 |  |  |  |  | 3635.9352 | 0.061946
|  |  | 82.8057 | 81.17 |  |  |  |  | 3639.2049 | 0.00047076 | 82.8057 | 81.74 |  |  |  |  | 3639.2049 | 0.059621
|  |  | 82.8804 | 81.19 |  |  |  |  | 3642.4747 | 0.0041107 | 82.8804 | 82.05 |  |  |  |  | 3642.4747 | 0.042005
|  |  | 82.955 | 81.26 |  |  |  |  | 3645.7444 | 0.0085521 | 82.955 | 82.585 |  |  |  |  | 3645.7444 | 0.015848
|  |  | 83.0297 | 81.34 |  |  |  |  | 3649.0141 | 0.011659 | 83.0297 | 82.34 |  |  |  |  | 3649.0141 | 0.013102
|  |  | 83.1044 | 81.365 |  |  |  |  | 3652.2838 | 0.012731 | 83.1044 | 82.005 |  |  |  |  | 3652.2838 | 0.034588
|  |  | 83.179 | 81.345 |  |  |  |  | 3655.5536 | 0.011808 | 83.179 | 81.765 |  |  |  |  | 3655.5536 | 0.046497
|  |  | 83.2537 | 81.265 |  |  |  |  | 3658.8233 | 0.0095665 | 83.2537 | 81.485 |  |  |  |  | 3658.8233 | 0.046733
|  |  | 83.3284 | 81.28 |  |  |  |  | 3662.093 | 0.0069539 | 83.3284 | 81.34 |  |  |  |  | 3662.093 | 0.036857
|  |  | 83.403 | 81.31 |  |  |  |  | 3665.3628 | 0.0048428 | 83.403 | 81.86 |  |  |  |  | 3665.3628 | 0.027162
|  |  | 83.4777 | 81.33 |  |  |  |  | 3668.6325 | 0.0037922 | 83.4777 | 82.2 |  |  |  |  | 3668.6325 | 0.038374
|  |  | 83.5524 | 81.35 |  |  |  |  | 3671.9022 | 0.0034018 | 83.5524 | 81.375 |  |  |  |  | 3671.9022 | 0.059756
|  |  | 83.627 | 81.265 |  |  |  |  | 3675.1719 | 0.0024306 | 83.627 | 81.29 |  |  |  |  | 3675.1719 | 0.074269
|  |  | 83.7017 | 81.09 |  |  |  |  | 3678.4417 | 0.00035519 | 83.7017 | 81.055 |  |  |  |  | 3678.4417 | 0.074444
|  |  | 83.7764 | 81.1 |  |  |  |  | 3681.7114 | 0.003801 | 83.7764 | 81.29 |  |  |  |  | 3681.7114 | 0.05902
|  |  | 83.851 | 81.215 |  |  |  |  | 3684.9811 | 0.0081067 | 83.851 | 81.35 |  |  |  |  | 3684.9811 | 0.0326
|  |  | 83.9257 | 81.36 |  |  |  |  | 3688.2508 | 0.011548 | 83.9257 | 81.29 |  |  |  |  | 3688.2508 | 0.0084527
|  |  | 84.0004 | 81.34 |  |  |  |  | 3691.5206 | 0.012733 | 84.0004 | 81.935 |  |  |  |  | 3691.5206 | 0.022382
|  |  | 84.075 | 81.27 |  |  |  |  | 3694.7903 | 0.010878 | 84.075 | 82.385 |  |  |  |  | 3694.7903 | 0.030601
|  |  | 84.1497 | 81.24 |  |  |  |  | 3698.06 | 0.0062681 | 84.1497 | 82.46 |  |  |  |  | 3698.06 | 0.028667
|  |  | 84.2244 | 81.25 |  |  |  |  | 3701.3297 | 0.0023802 | 84.2244 | 82.09 |  |  |  |  | 3701.3297 | 0.041797
|  |  | 84.299 | 81.32 |  |  |  |  | 3704.5995 | 0.0074823 | 84.299 | 82.435 |  |  |  |  | 3704.5995 | 0.076302
|  |  | 84.3737 | 81.32 |  |  |  |  | 3707.8692 | 0.011898 | 84.3737 | 82.455 |  |  |  |  | 3707.8692 | 0.11247
|  |  | 84.4484 | 81.255 |  |  |  |  | 3711.1389 | 0.013335 | 84.4484 | 82.245 |  |  |  |  | 3711.1389 | 0.1358
|  |  | 84.523 | 81.28 |  |  |  |  | 3714.4086 | 0.011551 | 84.523 | 82.265 |  |  |  |  | 3714.4086 | 0.13742
|  |  | 84.5977 | 81.335 |  |  |  |  | 3717.6784 | 0.0074472 | 84.5977 | 82.81 |  |  |  |  | 3717.6784 | 0.1151
|  |  | 84.6724 | 81.29 |  |  |  |  | 3720.9481 | 0.0032224 | 84.6724 | 82.25 |  |  |  |  | 3720.9481 | 0.074145
|  |  | 84.747 | 81.275 |  |  |  |  | 3724.2178 | 0.004065 | 84.747 | 81.155 |  |  |  |  | 3724.2178 | 0.03038
|  |  | 84.8217 | 81.195 |  |  |  |  | 3727.4875 | 0.0064826 | 84.8217 | 80.875 |  |  |  |  | 3727.4875 | 0.039059
|  |  | 84.8964 | 81.14 |  |  |  |  | 3730.7573 | 0.0073667 | 84.8964 | 81.395 |  |  |  |  | 3730.7573 | 0.069153
|  |  | 84.971 | 81.145 |  |  |  |  | 3734.027 | 0.0072502 | 84.971 | 81.71 |  |  |  |  | 3734.027 | 0.082122
|  |  | 85.0457 | 81.23 |  |  |  |  | 3737.2967 | 0.0072483 | 85.0457 | 81.635 |  |  |  |  | 3737.2967 | 0.076444
|  |  | 85.1204 | 81.285 |  |  |  |  | 3740.5664 | 0.0075299 | 85.1204 | 81.305 |  |  |  |  | 3740.5664 | 0.059138
|  |  | 85.195 | 81.29 |  |  |  |  | 3743.8362 | 0.0073196 | 85.195 | 80.92 |  |  |  |  | 3743.8362 | 0.041037
|  |  | 85.2697 | 81.26 |  |  |  |  | 3747.1059 | 0.0062448 | 85.2697 | 81.515 |  |  |  |  | 3747.1059 | 0.029928
|  |  | 85.3444 | 81.285 |  |  |  |  | 3750.3756 | 0.0048269 | 85.3444 | 81.975 |  |  |  |  | 3750.3756 | 0.022639
|  |  | 85.419 | 81.33 |  |  |  |  | 3753.6454 | 0.0039021 | 85.419 | 81.83 |  |  |  |  | 3753.6454 | 0.014345
|  |  | 85.4937 | 81.275 |  |  |  |  | 3756.9151 | 0.0032357 | 85.4937 | 81.79 |  |  |  |  | 3756.9151 | 0.01346
|  |  | 85.5684 | 81.235 |  |  |  |  | 3760.1848 | 0.0018118 | 85.5684 | 81.86 |  |  |  |  | 3760.1848 | 0.020135
|  |  | 85.6431 | 81.275 |  |  |  |  | 3763.4545 | 0.00095459 | 85.6431 | 81.845 |  |  |  |  | 3763.4545 | 0.021455
|  |  | 85.7177 | 81.35 |  |  |  |  | 3766.7243 | 0.0038508 | 85.7177 | 81.825 |  |  |  |  | 3766.7243 | 0.016603
|  |  | 85.7924 | 81.29 |  |  |  |  | 3769.994 | 0.0063168 | 85.7924 | 81.985 |  |  |  |  | 3769.994 | 0.018342
|  |  | 85.8671 | 81.23 |  |  |  |  | 3773.2637 | 0.0073901 | 85.8671 | 82.7 |  |  |  |  | 3773.2637 | 0.029123
|  |  | 85.9417 | 81.185 |  |  |  |  | 3776.5334 | 0.006766 | 85.9417 | 82.425 |  |  |  |  | 3776.5334 | 0.035014
|  |  | 86.0164 | 81.23 |  |  |  |  | 3779.8032 | 0.0048777 | 86.0164 | 82.505 |  |  |  |  | 3779.8032 | 0.030018
|  |  | 86.091 | 81.295 |  |  |  |  | 3783.0729 | 0.002724 | 86.091 | 82.88 |  |  |  |  | 3783.0729 | 0.016508
|  |  | 86.1657 | 81.33 |  |  |  |  | 3786.3426 | 0.0014419 | 86.1657 | 82.955 |  |  |  |  | 3786.3426 | 0.014901
|  |  | 86.2404 | 81.24 |  |  |  |  | 3789.6123 | 0.0017462 | 86.2404 | 82.295 |  |  |  |  | 3789.6123 | 0.028295
|  |  | 86.315 | 81.225 |  |  |  |  | 3792.8821 | 0.0036678 | 86.315 | 81.445 |  |  |  |  | 3792.8821 | 0.033191
|  |  | 86.3897 | 81.3 |  |  |  |  | 3796.1518 | 0.0064872 | 86.3897 | 81.6 |  |  |  |  | 3796.1518 | 0.024735
|  |  | 86.4644 | 81.365 |  |  |  |  | 3799.4215 | 0.0089578 | 86.4644 | 81.095 |  |  |  |  | 3799.4215 | 0.0059171
|  |  | 86.5391 | 81.33 |  |  |  |  | 3802.6912 | 0.0098612 | 86.5391 | 82.135 |  |  |  |  | 3802.6912 | 0.017271
|  |  | 86.6137 | 81.26 |  |  |  |  | 3805.961 | 0.008531 | 86.6137 | 83.09 |  |  |  |  | 3805.961 | 0.034172
|  |  | 86.6884 | 81.195 |  |  |  |  | 3809.2307 | 0.0051619 | 86.6884 | 82.6 |  |  |  |  | 3809.2307 | 0.038974
|  |  | 86.7631 | 81.155 |  |  |  |  | 3812.5004 | 0.00095336 | 86.7631 | 82.925 |  |  |  |  | 3812.5004 | 0.029815
|  |  | 86.8377 | 81.08 |  |  |  |  | 3815.7701 | 0.0035921 | 86.8377 | 82.695 |  |  |  |  | 3815.7701 | 0.010425
|  |  | 86.9124 | 81.16 |  |  |  |  | 3819.0399 | 0.0063338 | 86.9124 | 82.225 |  |  |  |  | 3819.0399 | 0.01433
|  |  | 86.9871 | 81.22 |  |  |  |  | 3822.3096 | 0.0074003 | 86.9871 | 81.635 |  |  |  |  | 3822.3096 | 0.034223
|  |  | 87.0617 | 81.215 |  |  |  |  | 3825.5793 | 0.007548 | 87.0617 | 81.51 |  |  |  |  | 3825.5793 | 0.046476
|  |  | 87.1364 | 81.27 |  |  |  |  | 3828.849 | 0.0078642 | 87.1364 | 81.55 |  |  |  |  | 3828.849 | 0.050611
|  |  | 87.2111 | 81.375 |  |  |  |  | 3832.1188 | 0.0081854 | 87.2111 | 81.52 |  |  |  |  | 3832.1188 | 0.049362
|  |  | 87.2857 | 81.4 |  |  |  |  | 3835.3885 | 0.0073932 | 87.2857 | 81.695 |  |  |  |  | 3835.3885 | 0.045242
|  |  | 87.3604 | 81.24 |  |  |  |  | 3838.6582 | 0.0049617 | 87.3604 | 81.53 |  |  |  |  | 3838.6582 | 0.03817
|  |  | 87.4351 | 81.18 |  |  |  |  | 3841.928 | 0.0015768 | 87.4351 | 81.765 |  |  |  |  | 3841.928 | 0.026933
|  |  | 87.5097 | 81.245 |  |  |  |  | 3845.1977 | 0.0023058 | 87.5097 | 81.725 |  |  |  |  | 3845.1977 | 0.011999
|  |  | 87.5844 | 81.415 |  |  |  |  | 3848.4674 | 0.0040849 | 87.5844 | 81.565 |  |  |  |  | 3848.4674 | 0.006667
|  |  | 87.6591 | 81.45 |  |  |  |  | 3851.7371 | 0.0035901 | 87.6591 | 81.4 |  |  |  |  | 3851.7371 | 0.021681
|  |  | 87.7337 | 81.38 |  |  |  |  | 3855.0069 | 0.0016282 | 87.7337 | 81.695 |  |  |  |  | 3855.0069 | 0.034038
|  |  | 87.8084 | 81.185 |  |  |  |  | 3858.2766 | 0.0042372 | 87.8084 | 81.89 |  |  |  |  | 3858.2766 | 0.04206
|  |  | 87.8831 | 81.11 |  |  |  |  | 3861.5463 | 0.0081223 | 87.8831 | 81.765 |  |  |  |  | 3861.5463 | 0.045553
|  |  | 87.9577 | 81.085 |  |  |  |  | 3864.816 | 0.010596 | 87.9577 | 81.665 |  |  |  |  | 3864.816 | 0.045474
|  |  | 88.0324 | 81.085 |  |  |  |  | 3868.0858 | 0.010981 | 88.0324 | 82.28 |  |  |  |  | 3868.0858 | 0.044144
|  |  | 88.1071 | 81.15 |  |  |  |  | 3871.3555 | 0.0096723 | 88.1071 | 81.92 |  |  |  |  | 3871.3555 | 0.043568
|  |  | 88.1817 | 81.26 |  |  |  |  | 3874.6252 | 0.0079687 | 88.1817 | 80.82 |  |  |  |  | 3874.6252 | 0.042324
|  |  | 88.2564 | 81.3 |  |  |  |  | 3877.8949 | 0.0074918 | 88.2564 | 81.24 |  |  |  |  | 3877.8949 | 0.036465
|  |  | 88.3311 | 81.245 |  |  |  |  | 3881.1647 | 0.0087091 | 88.3311 | 82.28 |  |  |  |  | 3881.1647 | 0.023936
|  |  | 88.4057 | 81.135 |  |  |  |  | 3884.4344 | 0.010663 | 88.4057 | 82.43 |  |  |  |  | 3884.4344 | 0.0094763
|  |  | 88.4804 | 81.06 |  |  |  |  | 3887.7041 | 0.01226 | 88.4804 | 82.035 |  |  |  |  | 3887.7041 | 0.017456
|  |  | 88.5551 | 81.01 |  |  |  |  | 3890.9738 | 0.012606 | 88.5551 | 82 |  |  |  |  | 3890.9738 | 0.029497
|  |  | 88.6297 | 81.01 |  |  |  |  | 3894.2436 | 0.011362 | 88.6297 | 81.895 |  |  |  |  | 3894.2436 | 0.032604
|  |  | 88.7044 | 81.02 |  |  |  |  | 3897.5133 | 0.009346 | 88.7044 | 82.55 |  |  |  |  | 3897.5133 | 0.027029
|  |  | 88.7791 | 81.075 |  |  |  |  | 3900.783 | 0.008456 | 88.7791 | 82.945 |  |  |  |  | 3900.783 | 0.020108
|  |  | 88.8537 | 81.165 |  |  |  |  | 3904.0527 | 0.0089942 | 88.8537 | 83.045 |  |  |  |  | 3904.0527 | 0.021994
|  |  | 88.9284 | 81.18 |  |  |  |  | 3907.3225 | 0.008749 | 88.9284 | 82.305 |  |  |  |  | 3907.3225 | 0.030129
|  |  | 89.0031 | 81.125 |  |  |  |  | 3910.5922 | 0.0063373 | 89.0031 | 81.78 |  |  |  |  | 3910.5922 | 0.043125
|  |  | 89.0777 | 81.205 |  |  |  |  | 3913.8619 | 0.0022558 | 89.0777 | 81.855 |  |  |  |  | 3913.8619 | 0.063531
|  |  | 89.1524 | 81.245 |  |  |  |  | 3917.1316 | 0.0027643 | 89.1524 | 81.815 |  |  |  |  | 3917.1316 | 0.086122
|  |  | 89.2271 | 81.25 |  |  |  |  | 3920.4014 | 0.0057404 | 89.2271 | 82.085 |  |  |  |  | 3920.4014 | 0.10098
|  |  | 89.3017 | 81.23 |  |  |  |  | 3923.6711 | 0.0065716 | 89.3017 | 82.315 |  |  |  |  | 3923.6711 | 0.10076
|  |  | 89.3764 | 81.085 |  |  |  |  | 3926.9408 | 0.0066001 | 89.3764 | 82.025 |  |  |  |  | 3926.9408 | 0.085949
|  |  | 89.4511 | 81.11 |  |  |  |  | 3930.2106 | 0.0086367 | 89.4511 | 81.68 |  |  |  |  | 3930.2106 | 0.068679
|  |  | 89.5257 | 81.155 |  |  |  |  | 3933.4803 | 0.01169 | 89.5257 | 82.215 |  |  |  |  | 3933.4803 | 0.067338
|  |  | 89.6004 | 81.105 |  |  |  |  | 3936.75 | 0.013077 | 89.6004 | 82.425 |  |  |  |  | 3936.75 | 0.076899
|  |  | 89.6751 | 80.94 |  |  |  |  | 3940.0197 | 0.011828 | 89.6751 | 81.785 |  |  |  |  | 3940.0197 | 0.076762
|  |  | 89.7497 | 80.93 |  |  |  |  | 3943.2895 | 0.0093096 | 89.7497 | 81.21 |  |  |  |  | 3943.2895 | 0.058373
|  |  | 89.8244 | 81.04 |  |  |  |  | 3946.5592 | 0.0088729 | 89.8244 | 80.925 |  |  |  |  | 3946.5592 | 0.026132
|  |  | 89.8991 | 81.145 |  |  |  |  | 3949.8289 | 0.010685 | 89.8991 | 80.75 |  |  |  |  | 3949.8289 | 0.016914
|  |  | 89.9737 | 81.135 |  |  |  |  | 3953.0986 | 0.011414 | 89.9737 | 81.07 |  |  |  |  | 3953.0986 | 0.043546
|  |  | 90.0484 | 81.085 |  |  |  |  | 3956.3684 | 0.0098253 | 90.0484 | 81.39 |  |  |  |  | 3956.3684 | 0.055029
|  |  | 90.1231 | 80.97 |  |  |  |  | 3959.6381 | 0.0072299 | 90.1231 | 81.65 |  |  |  |  | 3959.6381 | 0.049344
|  |  | 90.1977 | 80.975 |  |  |  |  | 3962.9078 | 0.0065828 | 90.1977 | 81.64 |  |  |  |  | 3962.9078 | 0.038688
|  |  | 90.2724 | 81.085 |  |  |  |  | 3966.1775 | 0.0079157 | 90.2724 | 81.75 |  |  |  |  | 3966.1775 | 0.045037
|  |  | 90.3471 | 81.25 |  |  |  |  | 3969.4473 | 0.0086301 | 90.3471 | 81.62 |  |  |  |  | 3969.4473 | 0.06151
|  |  | 90.4217 | 81.315 |  |  |  |  | 3972.717 | 0.00811 | 90.4217 | 81.145 |  |  |  |  | 3972.717 | 0.069985
|  |  | 90.4964 | 81.335 |  |  |  |  | 3975.9867 | 0.00685 | 90.4964 | 81.44 |  |  |  |  | 3975.9867 | 0.065321
|  |  | 90.5711 | 81.235 |  |  |  |  | 3979.2564 | 0.005014 | 90.5711 | 81.46 |  |  |  |  | 3979.2564 | 0.051081
|  |  | 90.6457 | 81.105 |  |  |  |  | 3982.5262 | 0.0029223 | 90.6457 | 81.55 |  |  |  |  | 3982.5262 | 0.035561
|  |  | 90.7204 | 81.14 |  |  |  |  | 3985.7959 | 0.0044979 | 90.7204 | 81.43 |  |  |  |  | 3985.7959 | 0.027715
|  |  | 90.7951 | 81.25 |  |  |  |  | 3989.0656 | 0.0085353 | 90.7951 | 81.64 |  |  |  |  | 3989.0656 | 0.030536
|  |  | 90.8697 | 81.32 |  |  |  |  | 3992.3353 | 0.011645 | 90.8697 | 81.635 |  |  |  |  | 3992.3353 | 0.041069
|  |  | 90.9444 | 81.355 |  |  |  |  | 3995.6051 | 0.012759 | 90.9444 | 82.05 |  |  |  |  | 3995.6051 | 0.055568
|  |  | 91.0191 | 81.315 |  |  |  |  | 3998.8748 | 0.012529 | 91.0191 | 81.845 |  |  |  |  | 3998.8748 | 0.067598
|  |  | 91.0937 | 81.265 | [6 empty] | 91.0937 | 82.07
|  |  | 91.1684 | 81.165 | [6 empty] | 91.1684 | 81.47
|  |  | 91.2431 | 81.15 | [6 empty] | 91.2431 | 81.555
|  |  | 91.3177 | 81.22 | [6 empty] | 91.3177 | 81.71
|  |  | 91.3924 | 81.27 | [6 empty] | 91.3924 | 81.925
|  |  | 91.4671 | 81.295 | [6 empty] | 91.4671 | 81.17
|  |  | 91.5417 | 81.275 | [6 empty] | 91.5417 | 81.33
|  |  | 91.6164 | 81.26 | [6 empty] | 91.6164 | 81.35
|  |  | 91.6911 | 81.245 | [6 empty] | 91.6911 | 81.595
|  |  | 91.7657 | 81.275 | [6 empty] | 91.7657 | 81.58
|  |  | 91.8404 | 81.355 | [6 empty] | 91.8404 | 81.91
|  |  | 91.9151 | 81.4 | [6 empty] | 91.9151 | 82.81
|  |  | 91.9897 | 81.385 | [6 empty] | 91.9897 | 82.755
|  |  | 92.0644 | 81.365 | [6 empty] | 92.0644 | 82.245
|  |  | 92.1391 | 81.345 | [6 empty] | 92.1391 | 81.51
|  |  | 92.2137 | 81.305 | [6 empty] | 92.2137 | 81.58
|  |  | 92.2884 | 81.23 | [6 empty] | 92.2884 | 81.55
|  |  | 92.3631 | 81.135 | [6 empty] | 92.3631 | 81.91
|  |  | 92.4377 | 81.115 | [6 empty] | 92.4377 | 82.28
|  |  | 92.5124 | 81.21 | [6 empty] | 92.5124 | 82.135
|  |  | 92.5871 | 81.295 | [6 empty] | 92.5871 | 81.285
|  |  | 92.6617 | 81.195 | [6 empty] | 92.6617 | 81.68
|  |  | 92.7364 | 81.135 | [6 empty] | 92.7364 | 82.145
|  |  | 92.8111 | 81.14 | [6 empty] | 92.8111 | 82.255
|  |  | 92.8857 | 81.115 | [6 empty] | 92.8857 | 82.55
|  |  | 92.9604 | 81.12 | [6 empty] | 92.9604 | 82.555
|  |  | 93.0351 | 81.175 | [6 empty] | 93.0351 | 82.615
|  |  | 93.1097 | 81.145 | [6 empty] | 93.1097 | 81.83
|  |  | 93.1844 | 81.275 | [6 empty] | 93.1844 | 81.195
|  |  | 93.2591 | 81.385 | [6 empty] | 93.2591 | 80.445
|  |  | 93.3337 | 81.34 | [6 empty] | 93.3337 | 81.59
|  |  | 93.4084 | 81.335 | [6 empty] | 93.4084 | 82.425
|  |  | 93.4831 | 81.32 | [6 empty] | 93.4831 | 82.245
|  |  | 93.5577 | 81.28 | [6 empty] | 93.5577 | 81.97
|  |  | 93.6324 | 81.155 | [6 empty] | 93.6324 | 82.115
|  |  | 93.7071 | 81.09 | [6 empty] | 93.7071 | 82.75
|  |  | 93.7818 | 81.145 | [6 empty] | 93.7818 | 82.685
|  |  | 93.8564 | 81.22 | [6 empty] | 93.8564 | 82.215
|  |  | 93.9311 | 81.245 | [6 empty] | 93.9311 | 81.6
|  |  | 94.0058 | 81.18 | [6 empty] | 94.0058 | 81.49
|  |  | 94.0804 | 81.29 | [6 empty] | 94.0804 | 81.47
|  |  | 94.1551 | 81.31 | [6 empty] | 94.1551 | 81.615
|  |  | 94.2298 | 81.26 | [6 empty] | 94.2298 | 82.025
|  |  | 94.3044 | 81.21 | [6 empty] | 94.3044 | 81.92
|  |  | 94.3791 | 81.145 | [6 empty] | 94.3791 | 81.85
|  |  | 94.4538 | 81.125 | [6 empty] | 94.4538 | 82.16
|  |  | 94.5284 | 81.135 | [6 empty] | 94.5284 | 82.41
|  |  | 94.6031 | 81.22 | [6 empty] | 94.6031 | 82.015
|  |  | 94.6778 | 81.305 | [6 empty] | 94.6778 | 81.84
|  |  | 94.7524 | 81.215 | [6 empty] | 94.7524 | 81.455
|  |  | 94.8271 | 81.205 | [6 empty] | 94.8271 | 81.36
|  |  | 94.9018 | 81.2 | [6 empty] | 94.9018 | 81.345
|  |  | 94.9764 | 81.245 | [6 empty] | 94.9764 | 81.475
|  |  | 95.0511 | 81.29 | [6 empty] | 95.0511 | 81.765
|  |  | 95.1258 | 81.395 | [6 empty] | 95.1258 | 82.03
|  |  | 95.2004 | 81.34 | [6 empty] | 95.2004 | 81.795
|  |  | 95.2751 | 81.16 | [6 empty] | 95.2751 | 82.59
|  |  | 95.3498 | 81.165 | [6 empty] | 95.3498 | 82.535
|  |  | 95.4244 | 81.335 | [6 empty] | 95.4244 | 82.225
|  |  | 95.4991 | 81.445 | [6 empty] | 95.4991 | 81.7
|  |  | 95.5738 | 81.43 | [6 empty] | 95.5738 | 81.875
|  |  | 95.6484 | 81.19 | [6 empty] | 95.6484 | 82.005
|  |  | 95.7231 | 81.1 | [6 empty] | 95.7231 | 82.23
|  |  | 95.7978 | 81.105 | [6 empty] | 95.7978 | 82.445
|  |  | 95.8724 | 81.12 | [6 empty] | 95.8724 | 82.15
|  |  | 95.9471 | 81.125 | [6 empty] | 95.9471 | 80.945
|  |  | 96.0218 | 81.22 | [6 empty] | 96.0218 | 80.875
|  |  | 96.0964 | 81.24 | [6 empty] | 96.0964 | 85.205
|  |  | 96.1711 | 81.235 | [6 empty] | 96.1711 | 85.205
|  |  | 96.2458 | 81.215 | [6 empty] | 96.2458 | 0
|  |  | 96.3204 | 81.195 | [6 empty] | 96.3204 | 0
|  |  | 96.3951 | 81.17 | [6 empty] | 96.3951 | 0
|  |  | 96.4698 | 81.115 | [6 empty] | 96.4698 | 0
|  |  | 96.5444 | 81.35
|  |  | 96.6191 | 81.515
|  |  | 96.6938 | 81.495
|  |  | 96.7684 | 81.325
|  |  | 96.8431 | 81.28
|  |  | 96.9178 | 81.21
|  |  | 96.9924 | 81.195
|  |  | 97.0671 | 81.22
|  |  | 97.1418 | 81.19
|  |  | 97.2164 | 81.13
|  |  | 97.2911 | 81.045
|  |  | 97.3658 | 81.005
|  |  | 97.4404 | 81.045
|  |  | 97.5151 | 81.165
|  |  | 97.5898 | 81.265
|  |  | 97.6644 | 81.135
|  |  | 97.7391 | 81.045
|  |  | 97.8138 | 81.07
|  |  | 97.8884 | 81.16
|  |  | 97.9631 | 81.275
|  |  | 98.0378 | 81.195
|  |  | 98.1124 | 80.985
|  |  | 98.1871 | 81.015
|  |  | 98.2618 | 81.14
|  |  | 98.3364 | 81.115
|  |  | 98.4111 | 81.07
|  |  | 98.4858 | 81.095
|  |  | 98.5604 | 81.19
|  |  | 98.6351 | 81.25
|  |  | 98.7098 | 81.195
|  |  | 98.7844 | 81.215
|  |  | 98.8591 | 81.1
|  |  | 98.9338 | 81.125
|  |  | 99.0084 | 81.195
|  |  | 99.0831 | 81.04
|  |  | 99.1578 | 80.935
|  |  | 99.2324 | 80.97
|  |  | 99.3071 | 81.255
|  |  | 99.3818 | 81.3
|  |  | 99.4564 | 81.2
|  |  | 99.5311 | 81.19
|  |  | 99.6058 | 81.24
|  |  | 99.6804 | 81.255
|  |  | 99.7551 | 81.235
|  |  | 99.8298 | 81.13
|  |  | 99.9044 | 81.125
|  |  | 99.9791 | 80.405
|  |  | 100.0538 | 79.83
|  |  | 100.1284 | 79.83
|  |  | 100.2031 | 0
|  |  | 100.2778 | 0
|  |  | 100.3524 | 0
|  |  | 100.4271 | 0
Note on E9: Start Measurement Here: 1. Place cursor in green cell 2. Ctrl-Shift-K 3. Allow cells to be erased 3. Expose
Note on F7: Enter all machine setings: mA, kV, ms, etc.

## Sample X-ray Waveform Data
XLPRO ™ Version 4.06
Operation Settings Area |  |  |  | Accu-kV Waveform Capture Sample Data
[5 empty] | Date: | 2003-07-12
Units |  |  |  |  | By: | JRL
Radiation: |  | R
/Time: |  | /min |  |  | Manufacturer |  | Model |  | Other
Scale: |  |  |  | GENERATOR: | General Electric |  | DXR-750 |  | 3Ph-12P
Pulse Rate: |  | 15 |  | TUBE: | Machlett |  | Dynamax 69D |  | 12 Degree Anode
Mode
Units Display: |  | Rescale |  | Generator Setting |  |  |  | FS | Filtration mm Al |  | Generator Setting |  |  |  | FS | Filtration mm Al |  |  | Generator Setting |  |  |  | FS | Filtration mm Al
Cell Update: |  | Incremental |  | kVp | ms | mA | FS | Dist (cm) | Inherent | Total | kVp | ms | mA | FS | Dist (cm) | Inherent | Total |  | kVp | ms | mA | FS | Dist (cm) | Inherent | Total
Auto Zero: |  | Disabled |  | 44 | 200 | 25 | Auto | 45 | 1.7 | 2.5 | 54 | 200 | 600 | Auto | 45 | 1.7 | 2.5 |  | 140 | 200 | 25 | Auto | 45 | 1.7 | 2.5
Capture Mode: |  | Auto |  | NOTES: | [6 empty] | NOTES: | [7 empty] | NOTES:
Screen Tracking: |  | Yes
kVp |  | FT
Config: | kVS | Time+kV | 199.7 | ms | 43.36 | kVp |  |  |  |  | 200.6 | ms | 48.61 | kVp |  |  |  |  | 205.1 | ms | 137.9 | kVp
Scale: |  | 4083 | Peak= | 43.355 | kVp | PPV= | 42.6 | kV | Control= | 32832 | Peak= | 48.605 | kVp | PPV= | 46 | kV | Control= | 32832 | Peak= | 137.945 | kVp | PPV= | 133.6 | kV | Control= | 32832
Start delay (ms): |  | 0 | Rel Dose= | N/A | /s | Practical= | 42.52 | kV |  |  | Rel Dose= | N/A | /s | Practical= | 47.235 | kV |  |  | Rel Dose= | N/A | /s | Practical= | 135.375 | kV
Width (ms): |  | 4778 | Pk Freq= | 117.6197 | Hz | Rate= | 1.502 | Hz | Avg. mA= | N/A | Pk Freq= | 149.4089 | Hz | Rate= | 1.4953 | Hz | Avg. mA= | 0 | Pk Freq= | 114.4408 | Hz | Rate= | 6.3381 | Hz | Avg. mA= | N/A
Filtration (mm-Al): |  | 2.5 | Width= | 199.7342 | ms | Pulses= | 1 |  | mAs= | N/A | Width= | 200.6302 | ms | Pulses= | 1 |  | mAs= | 0 | Width= | 205.1102 | ms | Pulses= | 2 |  | mAs= | N/A
Dental threshold: |  |  | ms | kV |  |  |  | LINES | Hz | kV | ms | kV |  |  |  | LINES | Hz | kV | ms | kV |  |  |  | LINES | Hz | kV
Device Index |  | 0 | -0.29867 | 0 |  |  |  | -0.29867 | 0 | -3.4386e-15 | -0.29867 | 0 |  |  |  | -0.29867 | 0 | 1.782e-15 | -0.29867 | 0 |  |  |  | -0.29867 | 0 | 1.3508e-14
Monitor Configuration |  |  | -0.224 | 0 |  |  |  | 200.0329 | 3.2697 | 0.03635 | -0.224 | 0 |  |  |  | 200.9289 | 2.4361 | 0.028129 | -0.224 | 0 |  |  |  | 218.3263 | 3.2697 | 0.30122
Serial No: |  | N/A | -0.14933 | 0 |  |  |  | 43.355 | 6.5395 | 0.052843 | -0.14933 | 0 |  |  |  | 48.605 | 4.8722 | 0.079395 | -0.14933 | 0 |  |  |  | 137.945 | 6.5395 | 0.37245
ROM Ver: |  | 4.00 | -0.074667 | 0 |  |  |  | 43.355 | 9.8092 | 0.046356 | -0.074667 | 0 |  |  |  | 48.605 | 7.3083 | 0.12473 | -0.074667 | 0 |  |  |  | 137.945 | 9.8092 | 0.23614
Detector: |  | N/A | 0 | 37.76 |  |  |  | 42.52 | 13.0789 | 0.031165 | 0 | 37.76 |  |  |  | 47.235 | 9.7445 | 0.12865 | 0 | 37.76 |  |  |  | 135.375 | 13.0789 | 0.32574
Chamb Kq: |  | N/A | 0.074667 | 37.76 |  |  |  | 42.52 | 16.3486 | 0.019021 | 0.074667 | 37.76 |  |  |  | 47.235 | 12.1806 | 0.074944 | 0.074667 | 37.76 |  |  |  | 135.375 | 16.3486 | 0.42997
Atm Corr: |  | Automatic | 0.14933 | 37.76 |  |  |  | 0 | 19.6184 | 0.01427 | 0.14933 | 37.76 |  |  |  | 0 | 14.6167 | 0.032134 | 0.14933 | 37.76 |  |  |  | 0 | 19.6184 | 0.2774
Temp (°C): |  | 22 | 0.224 | 39.31 |  |  |  | 0 | 22.8881 | 0.026993 | 0.224 | 37.76 |  |  |  | 0 | 17.0528 | 0.13282 | 0.224 | 37.76 |  |  |  | 0 | 22.8881 | 0.21855
Press (kPa): |  | 101.3 | 0.29867 | 39.83 |  |  |  | 2693 | 26.1578 | 0.034838 | 0.29867 | 37.76 |  |  |  | 2707 | 19.4889 | 0.19502 | 0.29867 | 0 |  |  |  | 2939 | 26.1578 | 0.42322
|  |  | 0.37334 | 40.775 |  |  |  |  | 29.4275 | 0.030524 | 0.37334 | 37.76 |  |  |  |  | 21.925 | 0.18373 | 0.37334 | 0 |  |  |  |  | 29.4275 | 0.37042
|  |  | 0.448 | 42.19 |  |  |  | 0 | 32.6973 | 0.020555 | 0.448 | 37.76 |  |  |  | 0 | 24.3611 | 0.10206 | 0.448 | 0 |  |  |  | 0 | 32.6973 | 0.15413
|  |  | 0.52267 | 43.71 |  |  |  | 0 | 35.967 | 0.01532 | 0.52267 | 37.76 |  |  |  | 0 | 26.7973 | 0.04926 | 0.52267 | 0 |  |  |  | 0 | 35.967 | 0.36359
|  |  | 0.59734 | 43.675 |  |  |  |  | 39.2367 | 0.026655 | 0.59734 | 37.76 |  |  |  |  | 29.2334 | 0.13957 | 0.59734 | 0 |  |  |  |  | 39.2367 | 0.43593
|  |  | 0.672 | 42.925 |  |  |  |  | 42.5064 | 0.035141 | 0.672 | 37.76 |  |  |  |  | 31.6695 | 0.17374 | 0.672 | 0 |  |  |  |  | 42.5064 | 0.20066
|  |  | 0.74667 | 44.285 |  |  |  |  | 45.7762 | 0.041411 | 0.74667 | 37.845 |  |  |  |  | 34.1056 | 0.11915 | 0.74667 | 0 |  |  |  |  | 45.7762 | 0.2496
|  |  | 0.82134 | 42.685 |  |  |  |  | 49.0459 | 0.076085 | 0.82134 | 38.24 |  |  |  |  | 36.5417 | 0.025844 | 0.82134 | 0 |  |  |  |  | 49.0459 | 0.48027
|  |  | 0.896 | 42.685 |  |  |  |  | 52.3156 | 0.11681 | 0.896 | 38.5 |  |  |  |  | 38.9778 | 0.16131 | 0.896 | 0 |  |  |  |  | 52.3156 | 0.38599
|  |  | 0.97067 | 44.52 |  |  |  |  | 55.5853 | 0.1402 | 0.97067 | 38.92 |  |  |  |  | 41.4139 | 0.26934 | 0.97067 | 0 |  |  |  |  | 55.5853 | 0.060561
|  |  | 1.0453 | 43.56 |  |  |  |  | 58.8551 | 0.16332 | 1.0453 | 39.33 |  |  |  |  | 43.8501 | 0.29402 | 1.0453 | 0 |  |  |  |  | 58.8551 | 0.37666
|  |  | 1.12 | 43.12 |  |  |  |  | 62.1248 | 0.19416 | 1.12 | 39.745 |  |  |  |  | 46.2862 | 0.23354 | 1.12 | 0 |  |  |  |  | 62.1248 | 0.58547
|  |  | 1.1947 | 43.99 |  |  |  |  | 65.3945 | 0.19097 | 1.1947 | 40.07 |  |  |  |  | 48.7223 | 0.1709 | 1.1947 | 0 |  |  |  |  | 65.3945 | 0.61318
|  |  | 1.2693 | 44.51 |  |  |  |  | 68.6642 | 0.12846 | 1.2693 | 40.355 |  |  |  |  | 51.1584 | 0.26187 | 1.2693 | 0 |  |  |  |  | 68.6642 | 0.56882
|  |  | 1.344 | 44.54 |  |  |  |  | 71.934 | 0.037563 | 1.344 | 40.68 |  |  |  |  | 53.5945 | 0.42037 | 1.344 | 0 |  |  |  |  | 71.934 | 0.44948
|  |  | 1.4187 | 43.99 |  |  |  |  | 75.2037 | 0.021705 | 1.4187 | 41.005 |  |  |  |  | 56.0306 | 0.55612 | 1.4187 | 0 |  |  |  |  | 75.2037 | 0.32932
|  |  | 1.4933 | 44.965 |  |  |  |  | 78.4734 | 0.023066 | 1.4933 | 41.01 |  |  |  |  | 58.4667 | 0.63984 | 1.4933 | 0 |  |  |  |  | 78.4734 | 0.44237
|  |  | 1.568 | 44.89 |  |  |  |  | 81.7431 | 0.075145 | 1.568 | 40.99 |  |  |  |  | 60.9029 | 0.65235 | 1.568 | 0 |  |  |  |  | 81.7431 | 0.47425
|  |  | 1.6427 | 44.045 |  |  |  |  | 85.0129 | 0.12319 | 1.6427 | 40.955 |  |  |  |  | 63.339 | 0.57563 | 1.6427 | 0 |  |  |  |  | 85.0129 | 0.30184
|  |  | 1.7173 | 44.03 |  |  |  |  | 88.2826 | 0.098461 | 1.7173 | 40.93 |  |  |  |  | 65.7751 | 0.40606 | 1.7173 | 0 |  |  |  |  | 88.2826 | 0.45123
|  |  | 1.792 | 44.355 |  |  |  |  | 91.5523 | 0.038298 | 1.792 | 40.68 |  |  |  |  | 68.2112 | 0.17271 | 1.792 | 0 |  |  |  |  | 91.5523 | 0.67032
|  |  | 1.8667 | 43.475 |  |  |  |  | 94.8221 | 0.089126 | 1.8667 | 40.405 |  |  |  |  | 70.6473 | 0.10546 | 1.8667 | 0 |  |  |  |  | 94.8221 | 0.51676
|  |  | 1.9413 | 43.075 |  |  |  |  | 98.0918 | 0.075661 | 1.9413 | 40.27 |  |  |  |  | 73.0834 | 0.29022 | 1.9413 | 0 |  |  |  |  | 98.0918 | 0.2706
|  |  | 2.016 | 43.83 |  |  |  |  | 101.3615 | 0.043138 | 2.016 | 40.21 |  |  |  |  | 75.5195 | 0.37571 | 2.016 | 39.385 |  |  |  |  | 101.3615 | 0.66348
|  |  | 2.0907 | 44.41 |  |  |  |  | 104.6312 | 0.14523 | 2.0907 | 39.905 |  |  |  |  | 77.9557 | 0.33885 | 2.0907 | 42.495 |  |  |  |  | 104.6312 | 0.74484
|  |  | 2.1653 | 42.8 |  |  |  |  | 107.901 | 0.12118 | 2.1653 | 39.39 |  |  |  |  | 80.3918 | 0.21563 | 2.1653 | 41.955 |  |  |  |  | 107.901 | 0.26874
|  |  | 2.24 | 42.875 |  |  |  |  | 111.1707 | 0.09463 | 2.24 | 38.67 |  |  |  |  | 82.8279 | 0.12989 | 2.24 | 44.01 |  |  |  |  | 111.1707 | 0.83087
|  |  | 2.3147 | 43.59 |  |  |  |  | 114.4404 | 0.4267 | 2.3147 | 38.36 |  |  |  |  | 85.264 | 0.20966 | 2.3147 | 47.5 |  |  |  |  | 114.4404 | 1.7809
|  |  | 2.3893 | 44.175 |  |  |  |  | 117.7101 | 0.70111 | 2.3893 | 38.56 |  |  |  |  | 87.7001 | 0.28084 | 2.3893 | 50.36 |  |  |  |  | 117.7101 | 2.2922
|  |  | 2.464 | 43.55 |  |  |  |  | 120.9799 | 0.75699 | 2.464 | 38.625 |  |  |  |  | 90.1362 | 0.28776 | 2.464 | 54.265 |  |  |  |  | 120.9799 | 2.2014
|  |  | 2.5387 | 42.98 |  |  |  |  | 124.2496 | 0.56048 | 2.5387 | 38.785 |  |  |  |  | 92.5723 | 0.2762 | 2.5387 | 55.98 |  |  |  |  | 124.2496 | 1.6644
|  |  | 2.6133 | 43.4 |  |  |  |  | 127.5193 | 0.23032 | 2.6133 | 38.9 |  |  |  |  | 95.0084 | 0.34805 | 2.6133 | 58.3 |  |  |  |  | 127.5193 | 0.96803
|  |  | 2.688 | 42 |  |  |  |  | 130.789 | 0.072153 | 2.688 | 38.835 |  |  |  |  | 97.4446 | 0.52876 | 2.688 | 63.045 |  |  |  |  | 130.789 | 0.3245
|  |  | 2.7627 | 41.745 |  |  |  |  | 134.0588 | 0.16437 | 2.7627 | 38.89 |  |  |  |  | 99.8807 | 0.75934 | 2.7627 | 65.62 |  |  |  |  | 134.0588 | 0.18943
|  |  | 2.8373 | 42.965 |  |  |  |  | 137.3285 | 0.089907 | 2.8373 | 39.115 |  |  |  |  | 102.3168 | 0.98137 | 2.8373 | 68.02 |  |  |  |  | 137.3285 | 0.4264
|  |  | 2.912 | 44.06 |  |  |  |  | 140.5982 | 0.04841 | 2.912 | 38.955 |  |  |  |  | 104.7529 | 1.1489 | 2.912 | 68.02 |  |  |  |  | 140.5982 | 0.37007
|  |  | 2.9867 | 42.58 |  |  |  |  | 143.8679 | 0.12811 | 2.9867 | 38.865 |  |  |  |  | 107.189 | 1.2425 | 2.9867 | 76.285 |  |  |  |  | 143.8679 | 0.13379
|  |  | 3.0613 | 42.8 |  |  |  |  | 147.1377 | 0.10349 | 3.0613 | 38.97 |  |  |  |  | 109.6251 | 1.2804 | 3.0613 | 80.9 |  |  |  |  | 147.1377 | 0.24387
|  |  | 3.136 | 42.905 |  |  |  |  | 150.4074 | 0.023955 | 3.136 | 38.885 |  |  |  |  | 112.0612 | 1.2975 | 3.136 | 84.37 |  |  |  |  | 150.4074 | 0.33198
|  |  | 3.2107 | 42.24 |  |  |  |  | 153.6771 | 0.046992 | 3.2107 | 38.49 |  |  |  |  | 114.4974 | 1.2968 | 3.2107 | 86.55 |  |  |  |  | 153.6771 | 0.21985
|  |  | 3.2853 | 42.465 |  |  |  |  | 156.9468 | 0.048144 | 3.2853 | 38.065 |  |  |  |  | 116.9335 | 1.2363 | 3.2853 | 90.205 |  |  |  |  | 156.9468 | 0.18583
|  |  | 3.36 | 42.135 |  |  |  |  | 160.2166 | 0.01278 | 3.36 | 37.95 |  |  |  |  | 119.3696 | 1.0814 | 3.36 | 91.44 |  |  |  |  | 160.2166 | 0.2967
|  |  | 3.4347 | 41.97 |  |  |  |  | 163.4863 | 0.016192 | 3.4347 | 37.895 |  |  |  |  | 121.8057 | 0.86991 | 3.4347 | 93.625 |  |  |  |  | 163.4863 | 0.30779
|  |  | 3.5093 | 42.65 |  |  |  |  | 166.756 | 0.046031 | 3.5093 | 37.875 |  |  |  |  | 124.2418 | 0.73526 | 3.5093 | 95.695 |  |  |  |  | 166.756 | 0.28464
|  |  | 3.584 | 42.66 |  |  |  |  | 170.0257 | 0.095412 | 3.584 | 37.76 |  |  |  |  | 126.6779 | 0.76878 | 3.584 | 95.695 |  |  |  |  | 170.0257 | 0.24291
|  |  | 3.6587 | 41.88 |  |  |  |  | 173.2955 | 0.11822 | 3.6587 | 37.76 |  |  |  |  | 129.114 | 0.84466 | 3.6587 | 103.905 |  |  |  |  | 173.2955 | 0.22388
|  |  | 3.7334 | 41.355 |  |  |  |  | 176.5652 | 0.092914 | 3.7334 | 37.76 |  |  |  |  | 131.5502 | 0.82857 | 3.7334 | 105.89 |  |  |  |  | 176.5652 | 0.46312
|  |  | 3.808 | 40.45 |  |  |  |  | 179.8349 | 0.070386 | 3.808 | 37.76 |  |  |  |  | 133.9863 | 0.71681 | 3.808 | 110.005 |  |  |  |  | 179.8349 | 0.66986
|  |  | 3.8827 | 40.74 |  |  |  |  | 183.1047 | 0.090922 | 3.8827 | 37.76 |  |  |  |  | 136.4224 | 0.60511 | 3.8827 | 112.41 |  |  |  |  | 183.1047 | 0.62219
|  |  | 3.9574 | 41.73 |  |  |  |  | 186.3744 | 0.081084 | 3.9574 | 37.76 |  |  |  |  | 138.8585 | 0.57117 | 3.9574 | 114.42 |  |  |  |  | 186.3744 | 0.33788
|  |  | 4.032 | 40.695 |  |  |  |  | 189.6441 | 0.028586 | 4.032 | 37.76 |  |  |  |  | 141.2946 | 0.5507 | 4.032 | 116.8 |  |  |  |  | 189.6441 | 0.18215
|  |  | 4.1067 | 40.74 |  |  |  |  | 192.9138 | 0.041004 | 4.1067 | 37.76 |  |  |  |  | 143.7307 | 0.45533 | 4.1067 | 119.61 |  |  |  |  | 192.9138 | 0.31008
|  |  | 4.1814 | 41.125 |  |  |  |  | 196.1836 | 0.081605 | 4.1814 | 37.76 |  |  |  |  | 146.1668 | 0.30649 | 4.1814 | 120.675 |  |  |  |  | 196.1836 | 0.23899
|  |  | 4.256 | 40.905 |  |  |  |  | 199.4533 | 0.090096 | 4.256 | 37.76 |  |  |  |  | 148.603 | 0.25917 | 4.256 | 121.785 |  |  |  |  | 199.4533 | 0.156
|  |  | 4.3307 | 40.95 |  |  |  |  | 202.723 | 0.071608 | 4.3307 | 37.76 |  |  |  |  | 151.0391 | 0.33332 | 4.3307 | 123.46 |  |  |  |  | 202.723 | 0.28852
|  |  | 4.4054 | 41.46 |  |  |  |  | 205.9927 | 0.034194 | 4.4054 | 37.76 |  |  |  |  | 153.4752 | 0.34008 | 4.4054 | 124.885 |  |  |  |  | 205.9927 | 0.27872
|  |  | 4.48 | 41.3 |  |  |  |  | 209.2625 | 0.0087221 | 4.48 | 37.76 |  |  |  |  | 155.9113 | 0.22056 | 4.48 | 127.41 |  |  |  |  | 209.2625 | 0.16721
|  |  | 4.5547 | 40.48 |  |  |  |  | 212.5322 | 0.036188 | 4.5547 | 37.76 |  |  |  |  | 158.3474 | 0.060467 | 4.5547 | 127.915 |  |  |  |  | 212.5322 | 0.22801
|  |  | 4.6294 | 41.035 |  |  |  |  | 215.8019 | 0.057451 | 4.6294 | 37.76 |  |  |  |  | 160.7835 | 0.23214 | 4.6294 | 128.62 |  |  |  |  | 215.8019 | 0.25077
|  |  | 4.704 | 40.695 |  |  |  |  | 219.0716 | 0.084358 | 4.704 | 37.76 |  |  |  |  | 163.2196 | 0.37232 | 4.704 | 130.085 |  |  |  |  | 219.0716 | 0.19476
|  |  | 4.7787 | 41.3 |  |  |  |  | 222.3414 | 0.093101 | 4.7787 | 37.76 |  |  |  |  | 165.6558 | 0.38622 | 4.7787 | 131 |  |  |  |  | 222.3414 | 0.28431
|  |  | 4.8534 | 41.64 |  |  |  |  | 225.6111 | 0.06611 | 4.8534 | 37.76 |  |  |  |  | 168.0919 | 0.27936 | 4.8534 | 131.49 |  |  |  |  | 225.6111 | 0.32714
|  |  | 4.928 | 41.505 |  |  |  |  | 228.8808 | 0.03695 | 4.928 | 37.76 |  |  |  |  | 170.528 | 0.12178 | 4.928 | 131.81 |  |  |  |  | 228.8808 | 0.22537
|  |  | 5.0027 | 42.135 |  |  |  |  | 232.1505 | 0.044243 | 5.0027 | 38.025 |  |  |  |  | 172.9641 | 0.11554 | 5.0027 | 132.16 |  |  |  |  | 232.1505 | 0.28054
|  |  | 5.0774 | 41.525 |  |  |  |  | 235.4203 | 0.020652 | 5.0774 | 37.825 |  |  |  |  | 175.4002 | 0.18304 | 5.0774 | 132.475 |  |  |  |  | 235.4203 | 0.42185
|  |  | 5.152 | 42.075 |  |  |  |  | 238.69 | 0.052694 | 5.152 | 37.84 |  |  |  |  | 177.8363 | 0.1575 | 5.152 | 133.145 |  |  |  |  | 238.69 | 0.37585
|  |  | 5.2267 | 40.96 |  |  |  |  | 241.9597 | 0.11797 | 5.2267 | 37.76 |  |  |  |  | 180.2724 | 0.073525 | 5.2267 | 133.73 |  |  |  |  | 241.9597 | 0.22744
|  |  | 5.3014 | 40.515 |  |  |  |  | 245.2294 | 0.13807 | 5.3014 | 37.83 |  |  |  |  | 182.7086 | 0.1445 | 5.3014 | 133.805 |  |  |  |  | 245.2294 | 0.2941
|  |  | 5.376 | 41.73 |  |  |  |  | 248.4992 | 0.10214 | 5.376 | 37.925 |  |  |  |  | 185.1447 | 0.25158 | 5.376 | 133.535 |  |  |  |  | 248.4992 | 0.33802
|  |  | 5.4507 | 39.775 |  |  |  |  | 251.7689 | 0.03972 | 5.4507 | 37.76 |  |  |  |  | 187.5808 | 0.28347 | 5.4507 | 133.1 |  |  |  |  | 251.7689 | 0.21041
|  |  | 5.5254 | 40.27 |  |  |  |  | 255.0386 | 0.0095293 | 5.5254 | 38.685 |  |  |  |  | 190.0169 | 0.2312 | 5.5254 | 133.185 |  |  |  |  | 255.0386 | 0.17328
|  |  | 5.6 | 40.96 |  |  |  |  | 258.3083 | 0.021421 | 5.6 | 39.65 |  |  |  |  | 192.453 | 0.14818 | 5.6 | 133.63 |  |  |  |  | 258.3083 | 0.26772
|  |  | 5.6747 | 41.89 |  |  |  |  | 261.5781 | 0.025061 | 5.6747 | 40.02 |  |  |  |  | 194.8891 | 0.15269 | 5.6747 | 134.205 |  |  |  |  | 261.5781 | 0.20495
|  |  | 5.7494 | 42.29 |  |  |  |  | 264.8478 | 0.042679 | 5.7494 | 40.375 |  |  |  |  | 197.3252 | 0.20922 | 5.7494 | 134.4 |  |  |  |  | 264.8478 | 0.11369
|  |  | 5.824 | 40.98 |  |  |  |  | 268.1175 | 0.055573 | 5.824 | 40.525 |  |  |  |  | 199.7614 | 0.21486 | 5.824 | 134.255 |  |  |  |  | 268.1175 | 0.22625
|  |  | 5.8987 | 40.185 |  |  |  |  | 271.3873 | 0.05968 | 5.8987 | 40.69 |  |  |  |  | 202.1975 | 0.15384 | 5.8987 | 134.19 |  |  |  |  | 271.3873 | 0.23168
|  |  | 5.9734 | 40.585 |  |  |  |  | 274.657 | 0.054842 | 5.9734 | 40.85 |  |  |  |  | 204.6336 | 0.052028 | 5.9734 | 134.485 |  |  |  |  | 274.657 | 0.15685
|  |  | 6.048 | 41.065 |  |  |  |  | 277.9267 | 0.052486 | 6.048 | 40.99 |  |  |  |  | 207.0697 | 0.052164 | 6.048 | 134.435 |  |  |  |  | 277.9267 | 0.22307
|  |  | 6.1227 | 40.66 |  |  |  |  | 281.1964 | 0.06542 | 6.1227 | 41.055 |  |  |  |  | 209.5058 | 0.12327 | 6.1227 | 133.875 |  |  |  |  | 281.1964 | 0.25355
|  |  | 6.1974 | 40.895 |  |  |  |  | 284.4662 | 0.073019 | 6.1974 | 41 |  |  |  |  | 211.9419 | 0.14933 | 6.1974 | 133.88 |  |  |  |  | 284.4662 | 0.17698
|  |  | 6.272 | 40.165 |  |  |  |  | 287.7359 | 0.064356 | 6.272 | 40.885 |  |  |  |  | 214.378 | 0.13817 | 6.272 | 134.52 |  |  |  |  | 287.7359 | 0.20864
|  |  | 6.3467 | 40.205 |  |  |  |  | 291.0056 | 0.049701 | 6.3467 | 40.875 |  |  |  |  | 216.8142 | 0.11021 | 6.3467 | 134.7 |  |  |  |  | 291.0056 | 0.27695
|  |  | 6.4214 | 39.91 |  |  |  |  | 294.2753 | 0.038121 | 6.4214 | 40.93 |  |  |  |  | 219.2503 | 0.086868 | 6.4214 | 134.565 |  |  |  |  | 294.2753 | 0.2118
|  |  | 6.496 | 40.52 |  |  |  |  | 297.5451 | 0.040986 | 6.496 | 40.985 |  |  |  |  | 221.6864 | 0.085093 | 6.496 | 134 |  |  |  |  | 297.5451 | 0.13061
|  |  | 6.5707 | 40.57 |  |  |  |  | 300.8148 | 0.049385 | 6.5707 | 40.885 |  |  |  |  | 224.1225 | 0.10501 | 6.5707 | 133.53 |  |  |  |  | 300.8148 | 0.20812
|  |  | 6.6454 | 40.74 |  |  |  |  | 304.0845 | 0.032209 | 6.6454 | 40.59 |  |  |  |  | 226.5586 | 0.12619 | 6.6454 | 133.725 |  |  |  |  | 304.0845 | 0.20449
|  |  | 6.72 | 42.31 |  |  |  |  | 307.3542 | 0.02778 | 6.72 | 40.365 |  |  |  |  | 228.9947 | 0.1292 | 6.72 | 133.825 |  |  |  |  | 307.3542 | 0.11829
|  |  | 6.7947 | 41.8 |  |  |  |  | 310.624 | 0.07648 | 6.7947 | 40.135 |  |  |  |  | 231.4308 | 0.1067 | 6.7947 | 133.47 |  |  |  |  | 310.624 | 0.16429
|  |  | 6.8694 | 41.8 |  |  |  |  | 313.8937 | 0.091772 | 6.8694 | 40.045 |  |  |  |  | 233.867 | 0.065664 | 6.8694 | 132.91 |  |  |  |  | 313.8937 | 0.1981
|  |  | 6.944 | 41.58 |  |  |  |  | 317.1634 | 0.058422 | 6.944 | 40.445 |  |  |  |  | 236.3031 | 0.023533 | 6.944 | 133.105 |  |  |  |  | 317.1634 | 0.13328
|  |  | 7.0187 | 40.555 |  |  |  |  | 320.4331 | 0.0066429 | 7.0187 | 41.085 |  |  |  |  | 238.7392 | 0.005222 | 7.0187 | 133.53 |  |  |  |  | 320.4331 | 0.13156
|  |  | 7.0934 | 40.36 |  |  |  |  | 323.7029 | 0.033833 | 7.0934 | 41.565 |  |  |  |  | 241.1753 | 0.0019973 | 7.0934 | 132.92 |  |  |  |  | 323.7029 | 0.18277
|  |  | 7.168 | 41.735 |  |  |  |  | 326.9726 | 0.041624 | 7.168 | 42.025 |  |  |  |  | 243.6114 | 0.027637 | 7.168 | 132.25 |  |  |  |  | 326.9726 | 0.16999
|  |  | 7.2427 | 43.075 |  |  |  |  | 330.2423 | 0.075451 | 7.2427 | 42.135 |  |  |  |  | 246.0475 | 0.066747 | 7.2427 | 131.87 |  |  |  |  | 330.2423 | 0.18281
|  |  | 7.3174 | 41.755 |  |  |  |  | 333.512 | 0.10618 | 7.3174 | 42.195 |  |  |  |  | 248.4836 | 0.098585 | 7.3174 | 131.145 |  |  |  |  | 333.512 | 0.20952
|  |  | 7.392 | 41.365 |  |  |  |  | 336.7818 | 0.086036 | 7.392 | 42.355 |  |  |  |  | 250.9198 | 0.10959 | 7.392 | 130.52 |  |  |  |  | 336.7818 | 0.14998
|  |  | 7.4667 | 40.49 |  |  |  |  | 340.0515 | 0.034802 | 7.4667 | 42.4 |  |  |  |  | 253.3559 | 0.097864 | 7.4667 | 129.82 |  |  |  |  | 340.0515 | 0.098892
|  |  | 7.5414 | 42.14 |  |  |  |  | 343.3212 | 0.069173 | 7.5414 | 42.32 |  |  |  |  | 255.792 | 0.074646 | 7.5414 | 128.955 |  |  |  |  | 343.3212 | 0.17409
|  |  | 7.616 | 41.86 |  |  |  |  | 346.5909 | 0.078707 | 7.616 | 42.225 |  |  |  |  | 258.2281 | 0.058812 | 7.616 | 128.7 |  |  |  |  | 346.5909 | 0.17921
|  |  | 7.6907 | 42.27 |  |  |  |  | 349.8607 | 0.015682 | 7.6907 | 42.125 |  |  |  |  | 260.6642 | 0.057307 | 7.6907 | 128.375 |  |  |  |  | 349.8607 | 0.17764
|  |  | 7.7654 | 42.24 |  |  |  |  | 353.1304 | 0.092598 | 7.7654 | 42.03 |  |  |  |  | 263.1003 | 0.059333 | 7.7654 | 128.08 |  |  |  |  | 353.1304 | 0.26384
|  |  | 7.84 | 42.61 |  |  |  |  | 356.4001 | 0.18333 | 7.84 | 42.035 |  |  |  |  | 265.5364 | 0.06203 | 7.84 | 127.445 |  |  |  |  | 356.4001 | 0.27759
|  |  | 7.9147 | 43.535 |  |  |  |  | 359.6699 | 0.21263 | 7.9147 | 41.825 |  |  |  |  | 267.9725 | 0.066582 | 7.9147 | 127.395 |  |  |  |  | 359.6699 | 0.17342
|  |  | 7.9894 | 43.415 |  |  |  |  | 362.9396 | 0.17539 | 7.9894 | 41.285 |  |  |  |  | 270.4087 | 0.066909 | 7.9894 | 127.305 |  |  |  |  | 362.9396 | 0.15818
|  |  | 8.064 | 43.445 |  |  |  |  | 366.2093 | 0.10163 | 8.064 | 40.945 |  |  |  |  | 272.8448 | 0.057888 | 8.064 | 126.515 |  |  |  |  | 366.2093 | 0.23805
|  |  | 8.1387 | 44.015 |  |  |  |  | 369.479 | 0.030412 | 8.1387 | 40.585 |  |  |  |  | 275.2809 | 0.044027 | 8.1387 | 125.75 |  |  |  |  | 369.479 | 0.18868
|  |  | 8.2134 | 43.035 |  |  |  |  | 372.7488 | 0.023031 | 8.2134 | 40.2 |  |  |  |  | 277.717 | 0.036327 | 8.2134 | 124.975 |  |  |  |  | 372.7488 | 0.059821
|  |  | 8.288 | 42.325 |  |  |  |  | 376.0185 | 0.041866 | 8.288 | 40.23 |  |  |  |  | 280.1531 | 0.034373 | 8.288 | 124.72 |  |  |  |  | 376.0185 | 0.14788
|  |  | 8.3627 | 43.315 |  |  |  |  | 379.2882 | 0.038702 | 8.3627 | 40.475 |  |  |  |  | 282.5892 | 0.02866 | 8.3627 | 124.72 |  |  |  |  | 379.2882 | 0.16386
|  |  | 8.4374 | 43.625 |  |  |  |  | 382.5579 | 0.022001 | 8.4374 | 40.785 |  |  |  |  | 285.0253 | 0.024105 | 8.4374 | 124.75 |  |  |  |  | 382.5579 | 0.043284
|  |  | 8.512 | 42.7 |  |  |  |  | 385.8277 | 0.021214 | 8.512 | 41.07 |  |  |  |  | 287.4615 | 0.029396 | 8.512 | 124.595 |  |  |  |  | 385.8277 | 0.11507
|  |  | 8.5867 | 42.76 |  |  |  |  | 389.0974 | 0.034717 | 8.5867 | 41.11 |  |  |  |  | 289.8976 | 0.030192 | 8.5867 | 124.485 |  |  |  |  | 389.0974 | 0.18257
|  |  | 8.6614 | 44.17 |  |  |  |  | 392.3671 | 0.027275 | 8.6614 | 41.14 |  |  |  |  | 292.3337 | 0.015535 | 8.6614 | 124.375 |  |  |  |  | 392.3671 | 0.11015
|  |  | 8.736 | 44.04 |  |  |  |  | 395.6368 | 0.012915 | 8.736 | 41.17 |  |  |  |  | 294.7698 | 0.013827 | 8.736 | 123.8 |  |  |  |  | 395.6368 | 0.072668
|  |  | 8.8107 | 42.995 |  |  |  |  | 398.9066 | 0.061461 | 8.8107 | 41.08 |  |  |  |  | 297.2059 | 0.044134 | 8.8107 | 123.59 |  |  |  |  | 398.9066 | 0.17758
|  |  | 8.8854 | 42.245 |  |  |  |  | 402.1763 | 0.096119 | 8.8854 | 40.885 |  |  |  |  | 299.642 | 0.062462 | 8.8854 | 123.865 |  |  |  |  | 402.1763 | 0.16581
|  |  | 8.96 | 43.55 |  |  |  |  | 405.446 | 0.089989 | 8.96 | 40.645 |  |  |  |  | 302.0781 | 0.058823 | 8.96 | 124.23 |  |  |  |  | 405.446 | 0.077522
|  |  | 9.0347 | 43.53 |  |  |  |  | 408.7157 | 0.045805 | 9.0347 | 40.75 |  |  |  |  | 304.5143 | 0.033568 | 9.0347 | 124.3 |  |  |  |  | 408.7157 | 0.12998
|  |  | 9.1094 | 43.195 |  |  |  |  | 411.9855 | 0.029767 | 9.1094 | 41.02 |  |  |  |  | 306.9504 | 0.0074357 | 9.1094 | 124.325 |  |  |  |  | 411.9855 | 0.16103
|  |  | 9.184 | 43.265 |  |  |  |  | 415.2552 | 0.050243 | 9.184 | 41.375 |  |  |  |  | 309.3865 | 0.038753 | 9.184 | 124.095 |  |  |  |  | 415.2552 | 0.10427
|  |  | 9.2587 | 42.705 |  |  |  |  | 418.5249 | 0.026973 | 9.2587 | 41.51 |  |  |  |  | 311.8226 | 0.061481 | 9.2587 | 123.595 |  |  |  |  | 418.5249 | 0.10661
|  |  | 9.3334 | 42.705 |  |  |  |  | 421.7946 | 0.043555 | 9.3334 | 41.585 |  |  |  |  | 314.2587 | 0.079683 | 9.3334 | 124.02 |  |  |  |  | 421.7946 | 0.14391
|  |  | 9.408 | 42.33 |  |  |  |  | 425.0644 | 0.09774 | 9.408 | 41.57 |  |  |  |  | 316.6948 | 0.10589 | 9.408 | 124.705 |  |  |  |  | 425.0644 | 0.11044
|  |  | 9.4827 | 43.02 |  |  |  |  | 428.3341 | 0.11133 | 9.4827 | 41.475 |  |  |  |  | 319.1309 | 0.13438 | 9.4827 | 125.435 |  |  |  |  | 428.3341 | 0.12622
|  |  | 9.5574 | 43.77 |  |  |  |  | 431.6038 | 0.080455 | 9.5574 | 41.245 |  |  |  |  | 321.5671 | 0.14643 | 9.5574 | 125.805 |  |  |  |  | 431.6038 | 0.18764
|  |  | 9.632 | 42.97 |  |  |  |  | 434.8735 | 0.050794 | 9.632 | 41.11 |  |  |  |  | 324.0032 | 0.12948 | 9.632 | 126.21 |  |  |  |  | 434.8735 | 0.16193
|  |  | 9.7067 | 43.51 |  |  |  |  | 438.1433 | 0.070204 | 9.7067 | 41.54 |  |  |  |  | 326.4393 | 0.085346 | 9.7067 | 126.99 |  |  |  |  | 438.1433 | 0.075574
|  |  | 9.7814 | 42.645 |  |  |  |  | 441.413 | 0.08919 | 9.7814 | 42.375 |  |  |  |  | 328.8754 | 0.029848 | 9.7814 | 127.085 |  |  |  |  | 441.413 | 0.12738
|  |  | 9.856 | 42.615 |  |  |  |  | 444.6827 | 0.098519 | 9.856 | 42.62 |  |  |  |  | 331.3115 | 0.022525 | 9.856 | 126.295 |  |  |  |  | 444.6827 | 0.17359
|  |  | 9.9307 | 43.765 |  |  |  |  | 447.9525 | 0.10203 | 9.9307 | 42.865 |  |  |  |  | 333.7476 | 0.045825 | 9.9307 | 125.945 |  |  |  |  | 447.9525 | 0.13436
|  |  | 10.0054 | 44.245 |  |  |  |  | 451.2222 | 0.089257 | 10.0054 | 43.345 |  |  |  |  | 336.1837 | 0.050447 | 10.0054 | 125.96 |  |  |  |  | 451.2222 | 0.088512
|  |  | 10.08 | 42.675 |  |  |  |  | 454.4919 | 0.056931 | 10.08 | 43.68 |  |  |  |  | 338.6199 | 0.064933 | 10.08 | 127.055 |  |  |  |  | 454.4919 | 0.10481
|  |  | 10.1547 | 43.635 |  |  |  |  | 457.7616 | 0.016545 | 10.1547 | 43.94 |  |  |  |  | 341.056 | 0.095774 | 10.1547 | 127.725 |  |  |  |  | 457.7616 | 0.08813
|  |  | 10.2294 | 43.325 |  |  |  |  | 461.0314 | 0.02235 | 10.2294 | 44.13 |  |  |  |  | 343.4921 | 0.11555 | 10.2294 | 127.81 |  |  |  |  | 461.0314 | 0.059365
|  |  | 10.304 | 43.54 |  |  |  |  | 464.3011 | 0.05019 | 10.304 | 44.145 |  |  |  |  | 345.9282 | 0.1066 | 10.304 | 127.7 |  |  |  |  | 464.3011 | 0.097639
|  |  | 10.3787 | 42.75 |  |  |  |  | 467.5708 | 0.070743 | 10.3787 | 44.09 |  |  |  |  | 348.3643 | 0.069365 | 10.3787 | 127.86 |  |  |  |  | 467.5708 | 0.10838
|  |  | 10.4534 | 42.445 |  |  |  |  | 470.8405 | 0.082809 | 10.4534 | 44.055 |  |  |  |  | 350.8004 | 0.025099 | 10.4534 | 127.95 |  |  |  |  | 470.8405 | 0.066463
|  |  | 10.528 | 42.48 |  |  |  |  | 474.1103 | 0.081149 | 10.528 | 44.055 |  |  |  |  | 353.2365 | 0.039146 | 10.528 | 128.51 |  |  |  |  | 474.1103 | 0.063736
|  |  | 10.6027 | 41.45 |  |  |  |  | 477.38 | 0.062843 | 10.6027 | 43.85 |  |  |  |  | 355.6727 | 0.05633 | 10.6027 | 129.04 |  |  |  |  | 477.38 | 0.13964
|  |  | 10.6774 | 41.385 |  |  |  |  | 480.6497 | 0.037304 | 10.6774 | 43.625 |  |  |  |  | 358.1088 | 0.043047 | 10.6774 | 129.535 |  |  |  |  | 480.6497 | 0.19052
|  |  | 10.752 | 41.91 |  |  |  |  | 483.9194 | 0.029855 | 10.752 | 43.495 |  |  |  |  | 360.5449 | 0.011057 | 10.752 | 130.05 |  |  |  |  | 483.9194 | 0.17797
|  |  | 10.8267 | 43.21 |  |  |  |  | 487.1892 | 0.034862 | 10.8267 | 43.41 |  |  |  |  | 362.981 | 0.041988 | 10.8267 | 130.69 |  |  |  |  | 487.1892 | 0.097353
|  |  | 10.9014 | 43.835 |  |  |  |  | 490.4589 | 0.036496 | 10.9014 | 42.82 |  |  |  |  | 365.4171 | 0.070635 | 10.9014 | 131.045 |  |  |  |  | 490.4589 | 0.052559
|  |  | 10.976 | 41.97 |  |  |  |  | 493.7286 | 0.04356 | 10.976 | 42.515 |  |  |  |  | 367.8532 | 0.071208 | 10.976 | 130.725 |  |  |  |  | 493.7286 | 0.12727
|  |  | 11.0507 | 42.545 |  |  |  |  | 496.9983 | 0.045433 | 11.0507 | 42.38 |  |  |  |  | 370.2893 | 0.043785 | 11.0507 | 130.88 |  |  |  |  | 496.9983 | 0.12992
|  |  | 11.1254 | 43.255 |  |  |  |  | 500.2681 | 0.040277 | 11.1254 | 42.625 |  |  |  |  | 372.7255 | 0.016903 | 11.1254 | 131.415 |  |  |  |  | 500.2681 | 0.060982
|  |  | 11.2001 | 43.545 |  |  |  |  | 503.5378 | 0.059318 | 11.2001 | 43.33 |  |  |  |  | 375.1616 | 0.049008 | 11.2001 | 131.675 |  |  |  |  | 503.5378 | 0.067953
|  |  | 11.2747 | 42.795 |  |  |  |  | 506.8075 | 0.086911 | 11.2747 | 43.605 |  |  |  |  | 377.5977 | 0.069039 | 11.2747 | 132.26 |  |  |  |  | 506.8075 | 0.10444
|  |  | 11.3494 | 42.47 |  |  |  |  | 510.0772 | 0.093597 | 11.3494 | 43.765 |  |  |  |  | 380.0338 | 0.061761 | 11.3494 | 132.89 |  |  |  |  | 510.0772 | 0.067554
|  |  | 11.4241 | 43.09 |  |  |  |  | 513.347 | 0.077564 | 11.4241 | 43.825 |  |  |  |  | 382.4699 | 0.034085 | 11.4241 | 134.3 |  |  |  |  | 513.347 | 0.066459
|  |  | 11.4987 | 41.885 |  |  |  |  | 516.6167 | 0.055831 | 11.4987 | 43.775 |  |  |  |  | 384.906 | 0.024091 | 11.4987 | 134.6 |  |  |  |  | 516.6167 | 0.12465
|  |  | 11.5734 | 41.235 |  |  |  |  | 519.8864 | 0.043302 | 11.5734 | 43.64 |  |  |  |  | 387.3421 | 0.04725 | 11.5734 | 134.13 |  |  |  |  | 519.8864 | 0.11456
|  |  | 11.6481 | 42.8 |  |  |  |  | 523.1561 | 0.04506 | 11.6481 | 43.555 |  |  |  |  | 389.7783 | 0.057288 | 11.6481 | 134.345 |  |  |  |  | 523.1561 | 0.059025
|  |  | 11.7227 | 43.555 |  |  |  |  | 526.4259 | 0.058548 | 11.7227 | 43.49 |  |  |  |  | 392.2144 | 0.055267 | 11.7227 | 134.72 |  |  |  |  | 526.4259 | 0.089269
|  |  | 11.7974 | 42.305 |  |  |  |  | 529.6956 | 0.064255 | 11.7974 | 43.165 |  |  |  |  | 394.6505 | 0.062195 | 11.7974 | 135.04 |  |  |  |  | 529.6956 | 0.10387
|  |  | 11.8721 | 42.495 |  |  |  |  | 532.9653 | 0.049562 | 11.8721 | 43.075 |  |  |  |  | 397.0866 | 0.079813 | 11.8721 | 135.56 |  |  |  |  | 532.9653 | 0.039745
|  |  | 11.9467 | 42.1 |  |  |  |  | 536.2351 | 0.021483 | 11.9467 | 42.975 |  |  |  |  | 399.5227 | 0.086495 | 11.9467 | 136.38 |  |  |  |  | 536.2351 | 0.063693
|  |  | 12.0214 | 42.21 |  |  |  |  | 539.5048 | 0.020865 | 12.0214 | 42.715 |  |  |  |  | 401.9588 | 0.069639 | 12.0214 | 136.945 |  |  |  |  | 539.5048 | 0.13601
|  |  | 12.0961 | 41.505 |  |  |  |  | 542.7745 | 0.040633 | 12.0961 | 42.38 |  |  |  |  | 404.3949 | 0.031663 | 12.0961 | 136.82 |  |  |  |  | 542.7745 | 0.13719
|  |  | 12.1707 | 41.96 |  |  |  |  | 546.0442 | 0.050591 | 12.1707 | 41.775 |  |  |  |  | 406.8311 | 0.017952 | 12.1707 | 136.615 |  |  |  |  | 546.0442 | 0.094763
|  |  | 12.2454 | 41.6 |  |  |  |  | 549.314 | 0.047141 | 12.2454 | 41.34 |  |  |  |  | 409.2672 | 0.053453 | 12.2454 | 136.83 |  |  |  |  | 549.314 | 0.091716
|  |  | 12.3201 | 41.415 |  |  |  |  | 552.5837 | 0.028452 | 12.3201 | 41.07 |  |  |  |  | 411.7033 | 0.065173 | 12.3201 | 137.18 |  |  |  |  | 552.5837 | 0.10791
|  |  | 12.3947 | 40.825 |  |  |  |  | 555.8534 | 0.01085 | 12.3947 | 40.765 |  |  |  |  | 414.1394 | 0.049569 | 12.3947 | 137.555 |  |  |  |  | 555.8534 | 0.093836
|  |  | 12.4694 | 41.77 |  |  |  |  | 559.1231 | 0.021927 | 12.4694 | 40.77 |  |  |  |  | 416.5755 | 0.043085 | 12.4694 | 137.815 |  |  |  |  | 559.1231 | 0.069179
|  |  | 12.5441 | 42.865 |  |  |  |  | 562.3929 | 0.013498 | 12.5441 | 40.955 |  |  |  |  | 419.0116 | 0.09051 | 12.5441 | 138.04 |  |  |  |  | 562.3929 | 0.062768
|  |  | 12.6187 | 43 |  |  |  |  | 565.6626 | 0.023674 | 12.6187 | 41.365 |  |  |  |  | 421.4477 | 0.14283 | 12.6187 | 138.31 |  |  |  |  | 565.6626 | 0.077698
|  |  | 12.6934 | 42.405 |  |  |  |  | 568.9323 | 0.064586 | 12.6934 | 42.13 |  |  |  |  | 423.8839 | 0.17359 | 12.6934 | 138.645 |  |  |  |  | 568.9323 | 0.089357
|  |  | 12.7681 | 42.91 |  |  |  |  | 572.202 | 0.089592 | 12.7681 | 42.305 |  |  |  |  | 426.32 | 0.1741 | 12.7681 | 138.925 |  |  |  |  | 572.202 | 0.064176
|  |  | 12.8427 | 42.29 |  |  |  |  | 575.4718 | 0.08994 | 12.8427 | 42.31 |  |  |  |  | 428.7561 | 0.14955 | 12.8427 | 138.985 |  |  |  |  | 575.4718 | 0.010568
|  |  | 12.9174 | 42.575 |  |  |  |  | 578.7415 | 0.0762 | 12.9174 | 42.18 |  |  |  |  | 431.1922 | 0.11655 | 12.9174 | 139.09 |  |  |  |  | 578.7415 | 0.078914
|  |  | 12.9921 | 42.97 |  |  |  |  | 582.0112 | 0.056595 | 12.9921 | 42.205 |  |  |  |  | 433.6283 | 0.093514 | 12.9921 | 138.985 |  |  |  |  | 582.0112 | 0.1126
|  |  | 13.0667 | 42.315 |  |  |  |  | 585.2809 | 0.02883 | 13.0667 | 42.37 |  |  |  |  | 436.0644 | 0.08349 | 13.0667 | 138.77 |  |  |  |  | 585.2809 | 0.086091
|  |  | 13.1414 | 41.01 |  |  |  |  | 588.5507 | 0.031668 | 13.1414 | 42.63 |  |  |  |  | 438.5005 | 0.079444 | 13.1414 | 138.3 |  |  |  |  | 588.5507 | 0.049772
|  |  | 13.2161 | 42.69 |  |  |  |  | 591.8204 | 0.04884 | 13.2161 | 42.58 |  |  |  |  | 440.9366 | 0.085844 | 13.2161 | 138.39 |  |  |  |  | 591.8204 | 0.086644
|  |  | 13.2907 | 43.19 |  |  |  |  | 595.0901 | 0.027282 | 13.2907 | 42.355 |  |  |  |  | 443.3728 | 0.1005 | 13.2907 | 138.405 |  |  |  |  | 595.0901 | 0.10084
|  |  | 13.3654 | 44.08 |  |  |  |  | 598.3598 | 0.033064 | 13.3654 | 42.035 |  |  |  |  | 445.8089 | 0.10294 | 13.3654 | 138.4 |  |  |  |  | 598.3598 | 0.067395
|  |  | 13.4401 | 42.455 |  |  |  |  | 601.6296 | 0.093798 | 13.4401 | 41.935 |  |  |  |  | 448.245 | 0.080609 | 13.4401 | 138.53 |  |  |  |  | 601.6296 | 0.023596
|  |  | 13.5147 | 41.855 |  |  |  |  | 604.8993 | 0.11211 | 13.5147 | 41.81 |  |  |  |  | 450.6811 | 0.07019 | 13.5147 | 138.81 |  |  |  |  | 604.8993 | 0.046971
|  |  | 13.5894 | 41.11 |  |  |  |  | 608.169 | 0.067574 | 13.5894 | 41.62 |  |  |  |  | 453.1172 | 0.13323 | 13.5894 | 138.87 |  |  |  |  | 608.169 | 0.073548
|  |  | 13.6641 | 41.28 |  |  |  |  | 611.4387 | 0.023586 | 13.6641 | 41.55 |  |  |  |  | 455.5533 | 0.21053 | 13.6641 | 138.815 |  |  |  |  | 611.4387 | 0.082569
|  |  | 13.7387 | 42.765 |  |  |  |  | 614.7085 | 0.10141 | 13.7387 | 41.54 |  |  |  |  | 457.9894 | 0.25371 | 13.7387 | 138.61 |  |  |  |  | 614.7085 | 0.064692
|  |  | 13.8134 | 42.115 |  |  |  |  | 617.9782 | 0.12957 | 13.8134 | 41.82 |  |  |  |  | 460.4256 | 0.24009 | 13.8134 | 138.45 |  |  |  |  | 617.9782 | 0.02801
|  |  | 13.8881 | 42.545 |  |  |  |  | 621.2479 | 0.096686 | 13.8881 | 42.125 |  |  |  |  | 462.8617 | 0.17874 | 13.8881 | 138.67 |  |  |  |  | 621.2479 | 0.043152
|  |  | 13.9627 | 42.465 |  |  |  |  | 624.5177 | 0.032517 | 13.9627 | 42.49 |  |  |  |  | 465.2978 | 0.14902 | 13.9627 | 138.835 |  |  |  |  | 624.5177 | 0.062564
|  |  | 14.0374 | 43.935 |  |  |  |  | 627.7874 | 0.018006 | 14.0374 | 42.825 |  |  |  |  | 467.7339 | 0.23078 | 14.0374 | 138.37 |  |  |  |  | 627.7874 | 0.06255
|  |  | 14.1121 | 43.015 |  |  |  |  | 631.0571 | 0.030641 | 14.1121 | 43.03 |  |  |  |  | 470.17 | 0.32653 | 14.1121 | 138.015 |  |  |  |  | 631.0571 | 0.077098
|  |  | 14.1867 | 42.725 |  |  |  |  | 634.3268 | 0.046482 | 14.1867 | 43.245 |  |  |  |  | 472.6061 | 0.36603 | 14.1867 | 137.79 |  |  |  |  | 634.3268 | 0.084581
|  |  | 14.2614 | 42.035 |  |  |  |  | 637.5966 | 0.092518 | 14.2614 | 43.395 |  |  |  |  | 475.0422 | 0.32784 | 14.2614 | 137.86 |  |  |  |  | 637.5966 | 0.054154
|  |  | 14.3361 | 42 |  |  |  |  | 640.8663 | 0.12541 | 14.3361 | 43.585 |  |  |  |  | 477.4784 | 0.22432 | 14.3361 | 137.655 |  |  |  |  | 640.8663 | 0.03429
|  |  | 14.4107 | 41.65 |  |  |  |  | 644.136 | 0.12654 | 14.4107 | 43.66 |  |  |  |  | 479.9145 | 0.10832 | 14.4107 | 136.915 |  |  |  |  | 644.136 | 0.067325
|  |  | 14.4854 | 41.57 |  |  |  |  | 647.4057 | 0.10074 | 14.4854 | 43.655 |  |  |  |  | 482.3506 | 0.1218 | 14.4854 | 136.87 |  |  |  |  | 647.4057 | 0.062624
|  |  | 14.5601 | 42.385 |  |  |  |  | 650.6755 | 0.0623 | 14.5601 | 43.395 |  |  |  |  | 484.7867 | 0.19294 | 14.5601 | 136.69 |  |  |  |  | 650.6755 | 0.057485
|  |  | 14.6347 | 42.465 |  |  |  |  | 653.9452 | 0.023978 | 14.6347 | 43.33 |  |  |  |  | 487.2228 | 0.20858 | 14.6347 | 136.23 |  |  |  |  | 653.9452 | 0.10169
|  |  | 14.7094 | 42.385 |  |  |  |  | 657.2149 | 0.016751 | 14.7094 | 43.35 |  |  |  |  | 489.6589 | 0.16628 | 14.7094 | 135.795 |  |  |  |  | 657.2149 | 0.11025
|  |  | 14.7841 | 44.065 |  |  |  |  | 660.4846 | 0.03884 | 14.7841 | 43.315 |  |  |  |  | 492.095 | 0.12412 | 14.7841 | 135.785 |  |  |  |  | 660.4846 | 0.056934
|  |  | 14.8587 | 41.87 |  |  |  |  | 663.7544 | 0.04978 | 14.8587 | 43.225 |  |  |  |  | 494.5312 | 0.16927 | 14.8587 | 136.465 |  |  |  |  | 663.7544 | 0.052156
|  |  | 14.9334 | 42.34 |  |  |  |  | 667.0241 | 0.044279 | 14.9334 | 43.1 |  |  |  |  | 496.9673 | 0.24116 | 14.9334 | 137.575 |  |  |  |  | 667.0241 | 0.11442
|  |  | 15.0081 | 42.025 |  |  |  |  | 670.2938 | 0.022029 | 15.0081 | 42.895 |  |  |  |  | 499.4034 | 0.27565 | 15.0081 | 137.67 |  |  |  |  | 670.2938 | 0.11651
|  |  | 15.0827 | 42.835 |  |  |  |  | 673.5635 | 0.018075 | 15.0827 | 42.75 |  |  |  |  | 501.8395 | 0.26107 | 15.0827 | 137.065 |  |  |  |  | 673.5635 | 0.061744
|  |  | 15.1574 | 41.935 |  |  |  |  | 676.8333 | 0.047957 | 15.1574 | 42.735 |  |  |  |  | 504.2756 | 0.2173 | 15.1574 | 136.455 |  |  |  |  | 676.8333 | 0.011428
|  |  | 15.2321 | 41.685 |  |  |  |  | 680.103 | 0.064336 | 15.2321 | 43.165 |  |  |  |  | 506.7117 | 0.18475 | 15.2321 | 136.48 |  |  |  |  | 680.103 | 0.038899
|  |  | 15.3067 | 42.285 |  |  |  |  | 683.3727 | 0.063237 | 15.3067 | 43.74 |  |  |  |  | 509.1478 | 0.18924 | 15.3067 | 136.43 |  |  |  |  | 683.3727 | 0.027166
|  |  | 15.3814 | 42.75 |  |  |  |  | 686.6424 | 0.052127 | 15.3814 | 44.28 |  |  |  |  | 511.584 | 0.20766 | 15.3814 | 135.73 |  |  |  |  | 686.6424 | 0.035384
|  |  | 15.4561 | 42.58 |  |  |  |  | 689.9122 | 0.03998 | 15.4561 | 44.785 |  |  |  |  | 514.0201 | 0.21121 | 15.4561 | 135.47 |  |  |  |  | 689.9122 | 0.055051
|  |  | 15.5307 | 42.355 |  |  |  |  | 693.1819 | 0.030545 | 15.5307 | 45.055 |  |  |  |  | 516.4562 | 0.19104 | 15.5307 | 135.735 |  |  |  |  | 693.1819 | 0.041467
|  |  | 15.6054 | 42.735 |  |  |  |  | 696.4516 | 0.026214 | 15.6054 | 45.265 |  |  |  |  | 518.8923 | 0.1504 | 15.6054 | 135.67 |  |  |  |  | 696.4516 | 0.0079649
|  |  | 15.6801 | 43.665 |  |  |  |  | 699.7213 | 0.028503 | 15.6801 | 45.455 |  |  |  |  | 521.3284 | 0.10777 | 15.6801 | 135.085 |  |  |  |  | 699.7213 | 0.020359
|  |  | 15.7547 | 43.22 |  |  |  |  | 702.9911 | 0.03405 | 15.7547 | 45.645 |  |  |  |  | 523.7645 | 0.11706 | 15.7547 | 134.905 |  |  |  |  | 702.9911 | 0.034334
|  |  | 15.8294 | 42.92 |  |  |  |  | 706.2608 | 0.038849 | 15.8294 | 45.79 |  |  |  |  | 526.2006 | 0.18067 | 15.8294 | 134.48 |  |  |  |  | 706.2608 | 0.064543
|  |  | 15.9041 | 43.18 |  |  |  |  | 709.5305 | 0.047835 | 15.9041 | 45.915 |  |  |  |  | 528.6368 | 0.23958 | 15.9041 | 134.515 |  |  |  |  | 709.5305 | 0.10001
|  |  | 15.9787 | 44.285 |  |  |  |  | 712.8002 | 0.073932 | 15.9787 | 45.96 |  |  |  |  | 531.0729 | 0.25991 | 15.9787 | 134.68 |  |  |  |  | 712.8002 | 0.13252
|  |  | 16.0534 | 43.635 |  |  |  |  | 716.07 | 0.11147 | 16.0534 | 45.84 |  |  |  |  | 533.509 | 0.23041 | 16.0534 | 134.475 |  |  |  |  | 716.07 | 0.17442
|  |  | 16.1281 | 42.55 |  |  |  |  | 719.3397 | 0.14335 | 16.1281 | 45.575 |  |  |  |  | 535.9451 | 0.16402 | 16.1281 | 134.65 |  |  |  |  | 719.3397 | 0.20848
|  |  | 16.2027 | 43.81 |  |  |  |  | 722.6094 | 0.15715 | 16.2027 | 45.335 |  |  |  |  | 538.3812 | 0.11398 | 16.2027 | 134.64 |  |  |  |  | 722.6094 | 0.19695
|  |  | 16.2774 | 43.07 |  |  |  |  | 725.8792 | 0.14789 | 16.2774 | 45.375 |  |  |  |  | 540.8173 | 0.1537 | 16.2774 | 134.04 |  |  |  |  | 725.8792 | 0.12899
|  |  | 16.3521 | 42.455 |  |  |  |  | 729.1489 | 0.11622 | 16.3521 | 45.18 |  |  |  |  | 543.2534 | 0.22899 | 16.3521 | 134.02 |  |  |  |  | 729.1489 | 0.038272
|  |  | 16.4267 | 42.79 |  |  |  |  | 732.4186 | 0.069445 | 16.4267 | 44.9 |  |  |  |  | 545.6896 | 0.2943 | 16.4267 | 134.095 |  |  |  |  | 732.4186 | 0.057174
|  |  | 16.5014 | 43.6 |  |  |  |  | 735.6883 | 0.028318 | 16.5014 | 44.68 |  |  |  |  | 548.1257 | 0.34081 | 16.5014 | 134.06 |  |  |  |  | 735.6883 | 0.082265
|  |  | 16.5761 | 43.13 |  |  |  |  | 738.9581 | 0.031573 | 16.5761 | 44.335 |  |  |  |  | 550.5618 | 0.3607 | 16.5761 | 134.435 |  |  |  |  | 738.9581 | 0.055887
|  |  | 16.6507 | 42.125 |  |  |  |  | 742.2278 | 0.038123 | 16.6507 | 44.505 |  |  |  |  | 552.9979 | 0.34117 | 16.6507 | 134.58 |  |  |  |  | 742.2278 | 0.038085
|  |  | 16.7254 | 42.315 |  |  |  |  | 745.4975 | 0.033919 | 16.7254 | 44.915 |  |  |  |  | 555.434 | 0.27604 | 16.7254 | 134.275 |  |  |  |  | 745.4975 | 0.089725
|  |  | 16.8001 | 42.865 |  |  |  |  | 748.7672 | 0.033068 | 16.8001 | 45.62 |  |  |  |  | 557.8701 | 0.18055 | 16.8001 | 134.185 |  |  |  |  | 748.7672 | 0.11998
|  |  | 16.8747 | 43.02 |  |  |  |  | 752.037 | 0.0231 | 16.8747 | 46.04 |  |  |  |  | 560.3062 | 0.10908 | 16.8747 | 133.855 |  |  |  |  | 752.037 | 0.10464
|  |  | 16.9494 | 43.76 |  |  |  |  | 755.3067 | 0.0098667 | 16.9494 | 46.295 |  |  |  |  | 562.7424 | 0.12364 | 16.9494 | 133.865 |  |  |  |  | 755.3067 | 0.069514
|  |  | 17.0241 | 43.55 |  |  |  |  | 758.5764 | 0.049555 | 17.0241 | 46.39 |  |  |  |  | 565.1785 | 0.14201 | 17.0241 | 134.01 |  |  |  |  | 758.5764 | 0.093918
|  |  | 17.0987 | 43.05 |  |  |  |  | 761.8461 | 0.075098 | 17.0987 | 46.32 |  |  |  |  | 567.6146 | 0.11053 | 17.0987 | 134 |  |  |  |  | 761.8461 | 0.12667
|  |  | 17.1734 | 43.57 |  |  |  |  | 765.1159 | 0.068128 | 17.1734 | 46.425 |  |  |  |  | 570.0507 | 0.057812 | 17.1734 | 133.97 |  |  |  |  | 765.1159 | 0.10959
|  |  | 17.2481 | 42.58 |  |  |  |  | 768.3856 | 0.036978 | 17.2481 | 46.425 |  |  |  |  | 572.4868 | 0.09559 | 17.2481 | 134.195 |  |  |  |  | 768.3856 | 0.06205
|  |  | 17.3227 | 43.325 |  |  |  |  | 771.6553 | 0.046735 | 17.3227 | 46.325 |  |  |  |  | 574.9229 | 0.14403 | 17.3227 | 133.94 |  |  |  |  | 771.6553 | 0.061903
|  |  | 17.3974 | 43.12 |  |  |  |  | 774.925 | 0.07773 | 17.3974 | 46.04 |  |  |  |  | 577.359 | 0.13165 | 17.3974 | 133.505 |  |  |  |  | 774.925 | 0.065006
|  |  | 17.4721 | 42.09 |  |  |  |  | 778.1948 | 0.079257 | 17.4721 | 45.855 |  |  |  |  | 579.7952 | 0.061479 | 17.4721 | 133.17 |  |  |  |  | 778.1948 | 0.017021
|  |  | 17.5467 | 42.46 |  |  |  |  | 781.4645 | 0.04963 | 17.5467 | 45.65 |  |  |  |  | 582.2313 | 0.11798 | 17.5467 | 133.54 |  |  |  |  | 781.4645 | 0.055098
|  |  | 17.6214 | 42.735 |  |  |  |  | 784.7342 | 0.008121 | 17.6214 | 45.205 |  |  |  |  | 584.6674 | 0.25465 | 17.6214 | 133.955 |  |  |  |  | 784.7342 | 0.098852
|  |  | 17.6961 | 42.795 |  |  |  |  | 788.0039 | 0.026182 | 17.6961 | 44.835 |  |  |  |  | 587.1035 | 0.35031 | 17.6961 | 134.085 |  |  |  |  | 788.0039 | 0.083913
|  |  | 17.7707 | 43.505 |  |  |  |  | 791.2737 | 0.034301 | 17.7707 | 44.5 |  |  |  |  | 589.5396 | 0.36519 | 17.7707 | 134.1 |  |  |  |  | 791.2737 | 0.04754
|  |  | 17.8454 | 42.67 |  |  |  |  | 794.5434 | 0.022309 | 17.8454 | 44.005 |  |  |  |  | 591.9757 | 0.29177 | 17.8454 | 134.23 |  |  |  |  | 794.5434 | 0.083928
|  |  | 17.9201 | 43.295 |  |  |  |  | 797.8131 | 0.022741 | 17.9201 | 43.6 |  |  |  |  | 594.4118 | 0.15626 | 17.9201 | 135.23 |  |  |  |  | 797.8131 | 0.10331
|  |  | 17.9947 | 42.88 |  |  |  |  | 801.0828 | 0.036859 | 17.9947 | 43.2 |  |  |  |  | 596.848 | 0.062452 | 17.9947 | 135.405 |  |  |  |  | 801.0828 | 0.072829
|  |  | 18.0694 | 42.605 |  |  |  |  | 804.3526 | 0.035243 | 18.0694 | 42.855 |  |  |  |  | 599.2841 | 0.16746 | 18.0694 | 134.96 |  |  |  |  | 804.3526 | 0.040091
|  |  | 18.1441 | 42.835 |  |  |  |  | 807.6223 | 0.016286 | 18.1441 | 42.915 |  |  |  |  | 601.7202 | 0.24202 | 18.1441 | 134.57 |  |  |  |  | 807.6223 | 0.052686
|  |  | 18.2187 | 43.355 |  |  |  |  | 810.892 | 0.012135 | 18.2187 | 43.315 |  |  |  |  | 604.1563 | 0.25174 | 18.2187 | 134.665 |  |  |  |  | 810.892 | 0.052368
|  |  | 18.2934 | 43.33 |  |  |  |  | 814.1618 | 0.027474 | 18.2934 | 43.77 |  |  |  |  | 606.5924 | 0.21269 | 18.2934 | 134.73 |  |  |  |  | 814.1618 | 0.074502
|  |  | 18.3681 | 43.005 |  |  |  |  | 817.4315 | 0.027054 | 18.3681 | 44.035 |  |  |  |  | 609.0285 | 0.15873 | 18.3681 | 134.38 |  |  |  |  | 817.4315 | 0.11297
|  |  | 18.4427 | 43.215 |  |  |  |  | 820.7012 | 0.026091 | 18.4427 | 44.085 |  |  |  |  | 611.4646 | 0.11861 | 18.4427 | 133.825 |  |  |  |  | 820.7012 | 0.11369
|  |  | 18.5174 | 43.135 |  |  |  |  | 823.9709 | 0.041571 | 18.5174 | 44.18 |  |  |  |  | 613.9007 | 0.090791 | 18.5174 | 133.745 |  |  |  |  | 823.9709 | 0.068948
|  |  | 18.5921 | 42.67 |  |  |  |  | 827.2407 | 0.050453 | 18.5921 | 44.195 |  |  |  |  | 616.3369 | 0.061341 | 18.5921 | 134.12 |  |  |  |  | 827.2407 | 0.041861
|  |  | 18.6667 | 43.065 |  |  |  |  | 830.5104 | 0.041111 | 18.6667 | 44.155 |  |  |  |  | 618.773 | 0.049897 | 18.6667 | 134.305 |  |  |  |  | 830.5104 | 0.08015
|  |  | 18.7414 | 42.545 |  |  |  |  | 833.7801 | 0.019388 | 18.7414 | 44.075 |  |  |  |  | 621.2091 | 0.071849 | 18.7414 | 134.44 |  |  |  |  | 833.7801 | 0.088713
|  |  | 18.8161 | 41.88 |  |  |  |  | 837.0498 | 0.0023778 | 18.8161 | 43.96 |  |  |  |  | 623.6452 | 0.085345 | 18.8161 | 134.665 |  |  |  |  | 837.0498 | 0.060498
|  |  | 18.8908 | 41.45 |  |  |  |  | 840.3196 | 0.019845 | 18.8908 | 43.89 |  |  |  |  | 626.0813 | 0.071625 | 18.8908 | 134.6 |  |  |  |  | 840.3196 | 0.054212
|  |  | 18.9654 | 41.955 |  |  |  |  | 843.5893 | 0.037283 | 18.9654 | 43.755 |  |  |  |  | 628.5174 | 0.035057 | 18.9654 | 134.515 |  |  |  |  | 843.5893 | 0.10036
|  |  | 19.0401 | 41.705 |  |  |  |  | 846.859 | 0.049241 | 19.0401 | 43.31 |  |  |  |  | 630.9535 | 0.023048 | 19.0401 | 134.075 |  |  |  |  | 846.859 | 0.13119
|  |  | 19.1148 | 41.075 |  |  |  |  | 850.1287 | 0.044627 | 19.1148 | 43.09 |  |  |  |  | 633.3897 | 0.059818 | 19.1148 | 133.985 |  |  |  |  | 850.1287 | 0.12427
|  |  | 19.1894 | 41.3 |  |  |  |  | 853.3985 | 0.025809 | 19.1894 | 42.835 |  |  |  |  | 635.8258 | 0.080142 | 19.1894 | 134.19 |  |  |  |  | 853.3985 | 0.087645
|  |  | 19.2641 | 42.325 |  |  |  |  | 856.6682 | 0.026281 | 19.2641 | 42.29 |  |  |  |  | 638.2619 | 0.078146 | 19.2641 | 134.775 |  |  |  |  | 856.6682 | 0.052081
|  |  | 19.3388 | 43.155 |  |  |  |  | 859.9379 | 0.038365 | 19.3388 | 41.98 |  |  |  |  | 640.698 | 0.059619 | 19.3388 | 135.34 |  |  |  |  | 859.9379 | 0.041044
|  |  | 19.4134 | 42.5 |  |  |  |  | 863.2076 | 0.031311 | 19.4134 | 41.825 |  |  |  |  | 643.1341 | 0.035493 | 19.4134 | 135.45 |  |  |  |  | 863.2076 | 0.038222
|  |  | 19.4881 | 43.245 |  |  |  |  | 866.4774 | 0.01719 | 19.4881 | 42.19 |  |  |  |  | 645.5702 | 0.016898 | 19.4881 | 135.7 |  |  |  |  | 866.4774 | 0.044909
|  |  | 19.5628 | 43.8 |  |  |  |  | 869.7471 | 0.038086 | 19.5628 | 42.465 |  |  |  |  | 648.0063 | 0.010815 | 19.5628 | 135.775 |  |  |  |  | 869.7471 | 0.04587
|  |  | 19.6374 | 42.995 |  |  |  |  | 873.0168 | 0.062403 | 19.6374 | 42.81 |  |  |  |  | 650.4425 | 0.010159 | 19.6374 | 135.875 |  |  |  |  | 873.0168 | 0.033969
|  |  | 19.7121 | 42.685 |  |  |  |  | 876.2865 | 0.077977 | 19.7121 | 43.07 |  |  |  |  | 652.8786 | 0.011464 | 19.7121 | 136.445 |  |  |  |  | 876.2865 | 0.065997
|  |  | 19.7868 | 42.305 |  |  |  |  | 879.5563 | 0.082167 | 19.7868 | 43.19 |  |  |  |  | 655.3147 | 0.017794 | 19.7868 | 136.625 |  |  |  |  | 879.5563 | 0.10138
|  |  | 19.8614 | 41.4 |  |  |  |  | 882.826 | 0.070871 | 19.8614 | 43.27 |  |  |  |  | 657.7508 | 0.024427 | 19.8614 | 136.28 |  |  |  |  | 882.826 | 0.089796
|  |  | 19.9361 | 41.695 |  |  |  |  | 886.0957 | 0.047171 | 19.9361 | 43.26 |  |  |  |  | 660.1869 | 0.02624 | 19.9361 | 135.93 |  |  |  |  | 886.0957 | 0.024908
|  |  | 20.0108 | 42.825 |  |  |  |  | 889.3654 | 0.026787 | 20.0108 | 43.185 |  |  |  |  | 662.623 | 0.021501 | 20.0108 | 136.255 |  |  |  |  | 889.3654 | 0.062539
|  |  | 20.0854 | 41.735 |  |  |  |  | 892.6352 | 0.025377 | 20.0854 | 43.07 |  |  |  |  | 665.0591 | 0.018423 | 20.0854 | 136.445 |  |  |  |  | 892.6352 | 0.1241
|  |  | 20.1601 | 40.18 |  |  |  |  | 895.9049 | 0.043246 | 20.1601 | 42.93 |  |  |  |  | 667.4953 | 0.02985 | 20.1601 | 136.24 |  |  |  |  | 895.9049 | 0.1284
|  |  | 20.2348 | 41.055 |  |  |  |  | 899.1746 | 0.073418 | 20.2348 | 42.61 |  |  |  |  | 669.9314 | 0.042337 | 20.2348 | 135.99 |  |  |  |  | 899.1746 | 0.080181
|  |  | 20.3094 | 41.385 |  |  |  |  | 902.4444 | 0.094466 | 20.3094 | 42.325 |  |  |  |  | 672.3675 | 0.043037 | 20.3094 | 136.18 |  |  |  |  | 902.4444 | 0.032566
|  |  | 20.3841 | 41.2 |  |  |  |  | 905.7141 | 0.088398 | 20.3841 | 42.145 |  |  |  |  | 674.8036 | 0.02795 | 20.3841 | 136.38 |  |  |  |  | 905.7141 | 0.060858
|  |  | 20.4588 | 40.48 |  |  |  |  | 908.9838 | 0.053865 | 20.4588 | 41.94 |  |  |  |  | 677.2397 | 0.01573 | 20.4588 | 137.095 |  |  |  |  | 908.9838 | 0.066714
|  |  | 20.5334 | 41.095 |  |  |  |  | 912.2535 | 0.011643 | 20.5334 | 41.615 |  |  |  |  | 679.6758 | 0.043442 | 20.5334 | 137.56 |  |  |  |  | 912.2535 | 0.031193
|  |  | 20.6081 | 39.69 |  |  |  |  | 915.5233 | 0.042423 | 20.6081 | 41.35 |  |  |  |  | 682.1119 | 0.067536 | 20.6081 | 137.705 |  |  |  |  | 915.5233 | 0.02618
|  |  | 20.6828 | 40.56 |  |  |  |  | 918.793 | 0.065958 | 20.6828 | 41.015 |  |  |  |  | 684.5481 | 0.072927 | 20.6828 | 137.47 |  |  |  |  | 918.793 | 0.074397
|  |  | 20.7574 | 42.91 |  |  |  |  | 922.0627 | 0.065105 | 20.7574 | 40.7 |  |  |  |  | 686.9842 | 0.059534 | 20.7574 | 137.3 |  |  |  |  | 922.0627 | 0.087102
|  |  | 20.8321 | 42.43 |  |  |  |  | 925.3324 | 0.043226 | 20.8321 | 40.615 |  |  |  |  | 689.4203 | 0.04857 | 20.8321 | 137.375 |  |  |  |  | 925.3324 | 0.057235
|  |  | 20.9068 | 42.575 |  |  |  |  | 928.6022 | 0.012932 | 20.9068 | 41.06 |  |  |  |  | 691.8564 | 0.067455 | 20.9068 | 137.635 |  |  |  |  | 928.6022 | 0.036772
|  |  | 20.9814 | 43.08 |  |  |  |  | 931.8719 | 0.023469 | 20.9814 | 41.465 |  |  |  |  | 694.2925 | 0.091798 | 20.9814 | 137.905 |  |  |  |  | 931.8719 | 0.084661
|  |  | 21.0561 | 43.08 |  |  |  |  | 935.1416 | 0.036424 | 21.0561 | 41.84 |  |  |  |  | 696.7286 | 0.098474 | 21.0561 | 138.495 |  |  |  |  | 935.1416 | 0.10991
|  |  | 21.1308 | 43.675 |  |  |  |  | 938.4113 | 0.029923 | 21.1308 | 42.155 |  |  |  |  | 699.1647 | 0.083454 | 21.1308 | 138.72 |  |  |  |  | 938.4113 | 0.099866
|  |  | 21.2054 | 41.395 |  |  |  |  | 941.6811 | 0.01261 | 21.2054 | 42.74 |  |  |  |  | 701.6009 | 0.056534 | 21.2054 | 138.425 |  |  |  |  | 941.6811 | 0.082708
|  |  | 21.2801 | 42.005 |  |  |  |  | 944.9508 | 0.014312 | 21.2801 | 43.045 |  |  |  |  | 704.037 | 0.038218 | 21.2801 | 138.245 |  |  |  |  | 944.9508 | 0.076787
|  |  | 21.3548 | 41.595 |  |  |  |  | 948.2205 | 0.023582 | 21.3548 | 43.175 |  |  |  |  | 706.4731 | 0.037291 | 21.3548 | 137.6 |  |  |  |  | 948.2205 | 0.060493
|  |  | 21.4294 | 41.62 |  |  |  |  | 951.4902 | 0.025884 | 21.4294 | 43.06 |  |  |  |  | 708.9092 | 0.032039 | 21.4294 | 137.25 |  |  |  |  | 951.4902 | 0.039147
|  |  | 21.5041 | 42.575 |  |  |  |  | 954.76 | 0.023799 | 21.5041 | 43.05 |  |  |  |  | 711.3453 | 0.018373 | 21.5041 | 137.26 |  |  |  |  | 954.76 | 0.048249
|  |  | 21.5788 | 42.2 |  |  |  |  | 958.0297 | 0.019806 | 21.5788 | 43.055 |  |  |  |  | 713.7814 | 0.028536 | 21.5788 | 137.56 |  |  |  |  | 958.0297 | 0.054634
|  |  | 21.6534 | 41.635 |  |  |  |  | 961.2994 | 0.030637 | 21.6534 | 43.04 |  |  |  |  | 716.2175 | 0.048831 | 21.6534 | 137.49 |  |  |  |  | 961.2994 | 0.050482
|  |  | 21.7281 | 40.91 |  |  |  |  | 964.5691 | 0.050392 | 21.7281 | 43.065 |  |  |  |  | 718.6537 | 0.05614 | 21.7281 | 137.13 |  |  |  |  | 964.5691 | 0.060609
|  |  | 21.8028 | 41.905 |  |  |  |  | 967.8389 | 0.055613 | 21.8028 | 42.99 |  |  |  |  | 721.0898 | 0.050531 | 21.8028 | 137.025 |  |  |  |  | 967.8389 | 0.065917
|  |  | 21.8774 | 40.705 |  |  |  |  | 971.1086 | 0.033228 | 21.8774 | 42.97 |  |  |  |  | 723.5259 | 0.052084 | 21.8774 | 136.95 |  |  |  |  | 971.1086 | 0.047298
|  |  | 21.9521 | 41.065 |  |  |  |  | 974.3783 | 0.018128 | 21.9521 | 42.75 |  |  |  |  | 725.962 | 0.07355 | 21.9521 | 137.14 |  |  |  |  | 974.3783 | 0.017075
|  |  | 22.0268 | 41.565 |  |  |  |  | 977.648 | 0.072977 | 22.0268 | 42.435 |  |  |  |  | 728.3981 | 0.094817 | 22.0268 | 137.705 |  |  |  |  | 977.648 | 0.0061901
|  |  | 22.1014 | 42.305 |  |  |  |  | 980.9178 | 0.11172 | 22.1014 | 42.395 |  |  |  |  | 730.8342 | 0.099657 | 22.1014 | 138.105 |  |  |  |  | 980.9178 | 0.031191
|  |  | 22.1761 | 43.235 |  |  |  |  | 984.1875 | 0.11482 | 22.1761 | 42.62 |  |  |  |  | 733.2703 | 0.084275 | 22.1761 | 138.04 |  |  |  |  | 984.1875 | 0.065389
|  |  | 22.2508 | 41.82 |  |  |  |  | 987.4572 | 0.080952 | 22.2508 | 43.02 |  |  |  |  | 735.7065 | 0.05504 | 22.2508 | 137.71 |  |  |  |  | 987.4572 | 0.083508
|  |  | 22.3254 | 41.615 |  |  |  |  | 990.727 | 0.026388 | 22.3254 | 43.73 |  |  |  |  | 738.1426 | 0.028292 | 22.3254 | 137.535 |  |  |  |  | 990.727 | 0.064942
|  |  | 22.4001 | 41.5 |  |  |  |  | 993.9967 | 0.02509 | 22.4001 | 44.34 |  |  |  |  | 740.5787 | 0.030184 | 22.4001 | 137.735 |  |  |  |  | 993.9967 | 0.017051
|  |  | 22.4748 | 41.23 |  |  |  |  | 997.2664 | 0.055418 | 22.4748 | 44.62 |  |  |  |  | 743.0148 | 0.040762 | 22.4748 | 138.175 |  |  |  |  | 997.2664 | 0.029647
|  |  | 22.5494 | 40.415 |  |  |  |  | 1000.5361 | 0.060261 | 22.5494 | 44.875 |  |  |  |  | 745.4509 | 0.04078 | 22.5494 | 137.945 |  |  |  |  | 1000.5361 | 0.04381
|  |  | 22.6241 | 42.29 |  |  |  |  | 1003.8059 | 0.04679 | 22.6241 | 45.15 |  |  |  |  | 747.887 | 0.031642 | 22.6241 | 137.08 |  |  |  |  | 1003.8059 | 0.028794
|  |  | 22.6988 | 42.33 |  |  |  |  | 1007.0756 | 0.025555 | 22.6988 | 45.21 |  |  |  |  | 750.3231 | 0.019484 | 22.6988 | 136.59 |  |  |  |  | 1007.0756 | 0.046941
|  |  | 22.7734 | 42.08 |  |  |  |  | 1010.3453 | 0.0044612 | 22.7734 | 45.3 |  |  |  |  | 752.7593 | 0.014723 | 22.7734 | 136.14 |  |  |  |  | 1010.3453 | 0.069924
|  |  | 22.8481 | 41.52 |  |  |  |  | 1013.615 | 0.01602 | 22.8481 | 45.43 |  |  |  |  | 755.1954 | 0.028406 | 22.8481 | 135.995 |  |  |  |  | 1013.615 | 0.059765
|  |  | 22.9228 | 40.595 |  |  |  |  | 1016.8848 | 0.030482 | 22.9228 | 45.645 |  |  |  |  | 757.6315 | 0.050561 | 22.9228 | 135.9 |  |  |  |  | 1016.8848 | 0.029962
|  |  | 22.9974 | 41.42 |  |  |  |  | 1020.1545 | 0.035871 | 22.9974 | 45.815 |  |  |  |  | 760.0676 | 0.073315 | 22.9974 | 135.34 |  |  |  |  | 1020.1545 | 0.034089
|  |  | 23.0721 | 41.015 |  |  |  |  | 1023.4242 | 0.0277 | 23.0721 | 45.695 |  |  |  |  | 762.5037 | 0.089598 | 23.0721 | 135.145 |  |  |  |  | 1023.4242 | 0.048031
|  |  | 23.1468 | 41.275 |  |  |  |  | 1026.6939 | 0.01755 | 23.1468 | 45.58 |  |  |  |  | 764.9398 | 0.093536 | 23.1468 | 135.425 |  |  |  |  | 1026.6939 | 0.050877
|  |  | 23.2214 | 41.76 |  |  |  |  | 1029.9637 | 0.03959 | 23.2214 | 45.5 |  |  |  |  | 767.3759 | 0.08276 | 23.2214 | 135.05 |  |  |  |  | 1029.9637 | 0.048722
|  |  | 23.2961 | 41.56 |  |  |  |  | 1033.2334 | 0.061049 | 23.2961 | 45.43 |  |  |  |  | 769.8121 | 0.059589 | 23.2961 | 134.965 |  |  |  |  | 1033.2334 | 0.032223
|  |  | 23.3708 | 41.465 |  |  |  |  | 1036.5031 | 0.063234 | 23.3708 | 45.255 |  |  |  |  | 772.2482 | 0.030657 | 23.3708 | 135.15 |  |  |  |  | 1036.5031 | 0.022548
|  |  | 23.4454 | 43.08 |  |  |  |  | 1039.7728 | 0.045957 | 23.4454 | 45.05 |  |  |  |  | 774.6843 | 0.0078131 | 23.4454 | 134.38 |  |  |  |  | 1039.7728 | 0.050167
|  |  | 23.5201 | 42.5 |  |  |  |  | 1043.0426 | 0.020992 | 23.5201 | 44.75 |  |  |  |  | 777.1204 | 0.015582 | 23.5201 | 134 |  |  |  |  | 1043.0426 | 0.057115
|  |  | 23.5948 | 42.225 |  |  |  |  | 1046.3123 | 0.0068703 | 23.5948 | 44.815 |  |  |  |  | 779.5565 | 0.017748 | 23.5948 | 134.435 |  |  |  |  | 1046.3123 | 0.046142
|  |  | 23.6694 | 42.345 |  |  |  |  | 1049.582 | 0.01841 | 23.6694 | 45.465 |  |  |  |  | 781.9926 | 0.010415 | 23.6694 | 135.505 |  |  |  |  | 1049.582 | 0.065198
|  |  | 23.7441 | 42.345 |  |  |  |  | 1052.8517 | 0.029613 | 23.7441 | 46.735 |  |  |  |  | 784.4287 | 0.019113 | 23.7441 | 135.545 |  |  |  |  | 1052.8517 | 0.086172
|  |  | 23.8188 | 42.085 |  |  |  |  | 1056.1215 | 0.032869 | 23.8188 | 47.19 |  |  |  |  | 786.8648 | 0.036828 | 23.8188 | 135.03 |  |  |  |  | 1056.1215 | 0.072764
|  |  | 23.8934 | 42.1 |  |  |  |  | 1059.3912 | 0.022624 | 23.8934 | 47.525 |  |  |  |  | 789.301 | 0.047277 | 23.8934 | 134.175 |  |  |  |  | 1059.3912 | 0.042121
|  |  | 23.9681 | 43.355 |  |  |  |  | 1062.6609 | 0.0021658 | 23.9681 | 47.97 |  |  |  |  | 791.7371 | 0.045414 | 23.9681 | 133.645 |  |  |  |  | 1062.6609 | 0.048941
|  |  | 24.0428 | 43.385 |  |  |  |  | 1065.9306 | 0.027796 | 24.0428 | 48.15 |  |  |  |  | 794.1732 | 0.032307 | 24.0428 | 133.165 |  |  |  |  | 1065.9306 | 0.053458
|  |  | 24.1174 | 42.355 |  |  |  |  | 1069.2004 | 0.05683 | 24.1174 | 48.255 |  |  |  |  | 796.6093 | 0.015669 | 24.1174 | 132.885 |  |  |  |  | 1069.2004 | 0.033936
|  |  | 24.1921 | 42.325 |  |  |  |  | 1072.4701 | 0.081388 | 24.1921 | 48.42 |  |  |  |  | 799.0454 | 0.015735 | 24.1921 | 132.915 |  |  |  |  | 1072.4701 | 0.049873
|  |  | 24.2668 | 43.015 |  |  |  |  | 1075.7398 | 0.093055 | 24.2668 | 48.525 |  |  |  |  | 801.4815 | 0.024716 | 24.2668 | 132.97 |  |  |  |  | 1075.7398 | 0.075556
|  |  | 24.3414 | 43.05 |  |  |  |  | 1079.0096 | 0.085392 | 24.3414 | 48.49 |  |  |  |  | 803.9176 | 0.026569 | 24.3414 | 132.315 |  |  |  |  | 1079.0096 | 0.068935
|  |  | 24.4161 | 43.4 |  |  |  |  | 1082.2793 | 0.059352 | 24.4161 | 48.365 |  |  |  |  | 806.3538 | 0.022802 | 24.4161 | 131.92 |  |  |  |  | 1082.2793 | 0.038177
|  |  | 24.4908 | 43.86 |  |  |  |  | 1085.549 | 0.02306 | 24.4908 | 48.195 |  |  |  |  | 808.7899 | 0.019543 | 24.4908 | 131.63 |  |  |  |  | 1085.549 | 0.033759
|  |  | 24.5654 | 42.705 |  |  |  |  | 1088.8187 | 0.026325 | 24.5654 | 48.08 |  |  |  |  | 811.226 | 0.018285 | 24.5654 | 131.245 |  |  |  |  | 1088.8187 | 0.056171
|  |  | 24.6401 | 42.485 |  |  |  |  | 1092.0885 | 0.060863 | 24.6401 | 47.865 |  |  |  |  | 813.6621 | 0.015474 | 24.6401 | 131.255 |  |  |  |  | 1092.0885 | 0.067681
|  |  | 24.7148 | 42.365 |  |  |  |  | 1095.3582 | 0.079242 | 24.7148 | 47.745 |  |  |  |  | 816.0982 | 0.016698 | 24.7148 | 131.665 |  |  |  |  | 1095.3582 | 0.070089
|  |  | 24.7894 | 43.62 |  |  |  |  | 1098.6279 | 0.075437 | 24.7894 | 47.59 |  |  |  |  | 818.5343 | 0.028322 | 24.7894 | 131.72 |  |  |  |  | 1098.6279 | 0.070326
|  |  | 24.8641 | 43.94 |  |  |  |  | 1101.8976 | 0.060144 | 24.8641 | 47.575 |  |  |  |  | 820.9704 | 0.04055 | 24.8641 | 131.12 |  |  |  |  | 1101.8976 | 0.07273
|  |  | 24.9388 | 43.7 |  |  |  |  | 1105.1674 | 0.049276 | 24.9388 | 47.965 |  |  |  |  | 823.4066 | 0.045427 | 24.9388 | 130.415 |  |  |  |  | 1105.1674 | 0.068846
|  |  | 25.0134 | 44.44 |  |  |  |  | 1108.4371 | 0.03338 | 25.0134 | 48.335 |  |  |  |  | 825.8427 | 0.03989 | 25.0134 | 129.895 |  |  |  |  | 1108.4371 | 0.04537
|  |  | 25.0881 | 44.115 |  |  |  |  | 1111.7068 | 0.0075403 | 25.0881 | 48.65 |  |  |  |  | 828.2788 | 0.025632 | 25.0881 | 129.975 |  |  |  |  | 1111.7068 | 0.011181
|  |  | 25.1628 | 43.27 |  |  |  |  | 1114.9765 | 0.051614 | 25.1628 | 48.8 |  |  |  |  | 830.7149 | 0.0099374 | 25.1628 | 129.815 |  |  |  |  | 1114.9765 | 0.038448
|  |  | 25.2374 | 44.69 |  |  |  |  | 1118.2463 | 0.087855 | 25.2374 | 48.93 |  |  |  |  | 833.151 | 0.014429 | 25.2374 | 129.37 |  |  |  |  | 1118.2463 | 0.051993
|  |  | 25.3121 | 44.005 |  |  |  |  | 1121.516 | 0.095787 | 25.3121 | 49.015 |  |  |  |  | 835.5871 | 0.023318 | 25.3121 | 129.54 |  |  |  |  | 1121.516 | 0.036286
|  |  | 25.3868 | 43.685 |  |  |  |  | 1124.7857 | 0.075419 | 25.3868 | 49 |  |  |  |  | 838.0232 | 0.025096 | 25.3868 | 129.27 |  |  |  |  | 1124.7857 | 0.01959
|  |  | 25.4614 | 43.06 |  |  |  |  | 1128.0554 | 0.043422 | 25.4614 | 48.79 |  |  |  |  | 840.4594 | 0.023149 | 25.4614 | 128.875 |  |  |  |  | 1128.0554 | 0.028862
|  |  | 25.5361 | 43.25 |  |  |  |  | 1131.3252 | 0.033159 | 25.5361 | 48.665 |  |  |  |  | 842.8955 | 0.025157 | 25.5361 | 128.73 |  |  |  |  | 1131.3252 | 0.019664
|  |  | 25.6108 | 43.245 |  |  |  |  | 1134.5949 | 0.052427 | 25.6108 | 48.545 |  |  |  |  | 845.3316 | 0.030832 | 25.6108 | 129.155 |  |  |  |  | 1134.5949 | 0.024654
|  |  | 25.6854 | 43.59 |  |  |  |  | 1137.8646 | 0.06623 | 25.6854 | 48.4 |  |  |  |  | 847.7677 | 0.032132 | 25.6854 | 129.495 |  |  |  |  | 1137.8646 | 0.058548
|  |  | 25.7601 | 43.27 |  |  |  |  | 1141.1343 | 0.062852 | 25.7601 | 48.18 |  |  |  |  | 850.2038 | 0.023614 | 25.7601 | 129.325 |  |  |  |  | 1141.1343 | 0.072233
|  |  | 25.8348 | 43.385 |  |  |  |  | 1144.4041 | 0.044236 | 25.8348 | 47.87 |  |  |  |  | 852.6399 | 0.0061757 | 25.8348 | 129.37 |  |  |  |  | 1144.4041 | 0.06168
|  |  | 25.9094 | 43.18 |  |  |  |  | 1147.6738 | 0.023739 | 25.9094 | 47.34 |  |  |  |  | 855.076 | 0.02511 | 25.9094 | 129.44 |  |  |  |  | 1147.6738 | 0.062177
|  |  | 25.9841 | 44.05 |  |  |  |  | 1150.9435 | 0.020111 | 25.9841 | 47.005 |  |  |  |  | 857.5122 | 0.054264 | 25.9841 | 129.63 |  |  |  |  | 1150.9435 | 0.090631
|  |  | 26.0588 | 43.6 |  |  |  |  | 1154.2132 | 0.029734 | 26.0588 | 46.675 |  |  |  |  | 859.9483 | 0.077491 | 26.0588 | 129.62 |  |  |  |  | 1154.2132 | 0.10561
|  |  | 26.1334 | 42.42 |  |  |  |  | 1157.483 | 0.040862 | 26.1334 | 46.25 |  |  |  |  | 862.3844 | 0.087929 | 26.1334 | 129.645 |  |  |  |  | 1157.483 | 0.08523
|  |  | 26.2081 | 42.195 |  |  |  |  | 1160.7527 | 0.043495 | 26.2081 | 45.45 |  |  |  |  | 864.8205 | 0.082404 | 26.2081 | 129.74 |  |  |  |  | 1160.7527 | 0.040088
|  |  | 26.2828 | 42.42 |  |  |  |  | 1164.0224 | 0.030951 | 26.2828 | 45.19 |  |  |  |  | 867.2566 | 0.062678 | 26.2828 | 129.635 |  |  |  |  | 1164.0224 | 0.043849
|  |  | 26.3575 | 43.96 |  |  |  |  | 1167.2922 | 0.019206 | 26.3575 | 45.225 |  |  |  |  | 869.6927 | 0.035264 | 26.3575 | 129.52 |  |  |  |  | 1167.2922 | 0.066784
|  |  | 26.4321 | 44.325 |  |  |  |  | 1170.5619 | 0.040602 | 26.4321 | 45.455 |  |  |  |  | 872.1288 | 0.01248 | 26.4321 | 129.63 |  |  |  |  | 1170.5619 | 0.045392
|  |  | 26.5068 | 42.8 |  |  |  |  | 1173.8316 | 0.061039 | 26.5068 | 45.69 |  |  |  |  | 874.565 | 0.018528 | 26.5068 | 129.785 |  |  |  |  | 1173.8316 | 0.026039
|  |  | 26.5815 | 43.11 |  |  |  |  | 1177.1013 | 0.066542 | 26.5815 | 45.875 |  |  |  |  | 877.0011 | 0.025084 | 26.5815 | 129.815 |  |  |  |  | 1177.1013 | 0.08479
|  |  | 26.6561 | 42.255 |  |  |  |  | 1180.3711 | 0.055722 | 26.6561 | 45.925 |  |  |  |  | 879.4372 | 0.021321 | 26.6561 | 129.815 |  |  |  |  | 1180.3711 | 0.1086
|  |  | 26.7308 | 41.935 |  |  |  |  | 1183.6408 | 0.033456 | 26.7308 | 45.96 |  |  |  |  | 881.8733 | 0.013257 | 26.7308 | 129.81 |  |  |  |  | 1183.6408 | 0.078671
|  |  | 26.8055 | 42.61 |  |  |  |  | 1186.9105 | 0.019213 | 26.8055 | 45.825 |  |  |  |  | 884.3094 | 0.014708 | 26.8055 | 130.005 |  |  |  |  | 1186.9105 | 0.018805
|  |  | 26.8801 | 41.545 |  |  |  |  | 1190.1802 | 0.032948 | 26.8801 | 45.845 |  |  |  |  | 886.7455 | 0.021322 | 26.8801 | 130.315 |  |  |  |  | 1190.1802 | 0.034296
|  |  | 26.9548 | 41.955 |  |  |  |  | 1193.45 | 0.041177 | 26.9548 | 45.825 |  |  |  |  | 889.1816 | 0.022929 | 26.9548 | 130.4 |  |  |  |  | 1193.45 | 0.040478
|  |  | 27.0295 | 41.515 |  |  |  |  | 1196.7197 | 0.036347 | 27.0295 | 45.525 |  |  |  |  | 891.6178 | 0.01893 | 27.0295 | 130.17 |  |  |  |  | 1196.7197 | 0.037658
|  |  | 27.1041 | 42.41 |  |  |  |  | 1199.9894 | 0.030581 | 27.1041 | 45.255 |  |  |  |  | 894.0539 | 0.013325 | 27.1041 | 129.775 |  |  |  |  | 1199.9894 | 0.077812
|  |  | 27.1788 | 41.92 |  |  |  |  | 1203.2591 | 0.039247 | 27.1788 | 45.065 |  |  |  |  | 896.49 | 0.011859 | 27.1788 | 130.44 |  |  |  |  | 1203.2591 | 0.094853
|  |  | 27.2535 | 42.55 |  |  |  |  | 1206.5289 | 0.051029 | 27.2535 | 44.945 |  |  |  |  | 898.9261 | 0.013685 | 27.2535 | 131.395 |  |  |  |  | 1206.5289 | 0.066911
|  |  | 27.3281 | 41.815 |  |  |  |  | 1209.7986 | 0.053895 | 27.3281 | 44.615 |  |  |  |  | 901.3622 | 0.014149 | 27.3281 | 132.125 |  |  |  |  | 1209.7986 | 0.037273
|  |  | 27.4028 | 41.97 |  |  |  |  | 1213.0683 | 0.046531 | 27.4028 | 44.295 |  |  |  |  | 903.7983 | 0.012208 | 27.4028 | 132.395 |  |  |  |  | 1213.0683 | 0.063811
|  |  | 27.4775 | 42.075 |  |  |  |  | 1216.338 | 0.033442 | 27.4775 | 44.025 |  |  |  |  | 906.2344 | 0.0083454 | 27.4775 | 132.49 |  |  |  |  | 1216.338 | 0.070615
|  |  | 27.5521 | 42.94 |  |  |  |  | 1219.6078 | 0.022403 | 27.5521 | 43.46 |  |  |  |  | 908.6706 | 0.0038417 | 27.5521 | 131.885 |  |  |  |  | 1219.6078 | 0.051895
|  |  | 27.6268 | 41.32 |  |  |  |  | 1222.8775 | 0.027857 | 27.6268 | 42.96 |  |  |  |  | 911.1067 | 0.0073156 | 27.6268 | 131.905 |  |  |  |  | 1222.8775 | 0.036783
|  |  | 27.7015 | 42.215 |  |  |  |  | 1226.1472 | 0.041948 | 27.7015 | 42.625 |  |  |  |  | 913.5428 | 0.015624 | 27.7015 | 132.295 |  |  |  |  | 1226.1472 | 0.018322
|  |  | 27.7761 | 42.88 |  |  |  |  | 1229.4169 | 0.046016 | 27.7761 | 42.47 |  |  |  |  | 915.9789 | 0.023742 | 27.7761 | 133.23 |  |  |  |  | 1229.4169 | 0.023694
|  |  | 27.8508 | 42.15 |  |  |  |  | 1232.6867 | 0.033169 | 27.8508 | 42.83 |  |  |  |  | 918.415 | 0.030057 | 27.8508 | 133.9 |  |  |  |  | 1232.6867 | 0.060114
|  |  | 27.9255 | 42.725 |  |  |  |  | 1235.9564 | 0.0091606 | 27.9255 | 43.355 |  |  |  |  | 920.8511 | 0.034438 | 27.9255 | 134.155 |  |  |  |  | 1235.9564 | 0.063057
|  |  | 28.0001 | 43.78 |  |  |  |  | 1239.2261 | 0.013172 | 28.0001 | 43.875 |  |  |  |  | 923.2872 | 0.038642 | 28.0001 | 134.195 |  |  |  |  | 1239.2261 | 0.029403
|  |  | 28.0748 | 42.505 |  |  |  |  | 1242.4958 | 0.026531 | 28.0748 | 43.91 |  |  |  |  | 925.7234 | 0.044829 | 28.0748 | 134.22 |  |  |  |  | 1242.4958 | 0.044632
|  |  | 28.1495 | 43.36 |  |  |  |  | 1245.7656 | 0.041389 | 28.1495 | 43.85 |  |  |  |  | 928.1595 | 0.052266 | 28.1495 | 134.475 |  |  |  |  | 1245.7656 | 0.070672
|  |  | 28.2241 | 42.335 |  |  |  |  | 1249.0353 | 0.063268 | 28.2241 | 43.93 |  |  |  |  | 930.5956 | 0.056875 | 28.2241 | 134.755 |  |  |  |  | 1249.0353 | 0.047267
|  |  | 28.2988 | 41.47 |  |  |  |  | 1252.305 | 0.079948 | 28.2988 | 43.99 |  |  |  |  | 933.0317 | 0.054196 | 28.2988 | 134.985 |  |  |  |  | 1252.305 | 0.013362
|  |  | 28.3735 | 41.945 |  |  |  |  | 1255.5748 | 0.083125 | 28.3735 | 43.835 |  |  |  |  | 935.4678 | 0.042444 | 28.3735 | 135.06 |  |  |  |  | 1255.5748 | 0.068478
|  |  | 28.4481 | 42.225 |  |  |  |  | 1258.8445 | 0.075695 | 28.4481 | 43.62 |  |  |  |  | 937.9039 | 0.025136 | 28.4481 | 134.875 |  |  |  |  | 1258.8445 | 0.079589
|  |  | 28.5228 | 41.785 |  |  |  |  | 1262.1142 | 0.066493 | 28.5228 | 43.42 |  |  |  |  | 940.34 | 0.017428 | 28.5228 | 134.965 |  |  |  |  | 1262.1142 | 0.042724
|  |  | 28.5975 | 41.43 |  |  |  |  | 1265.3839 | 0.064633 | 28.5975 | 43.135 |  |  |  |  | 942.7762 | 0.025981 | 28.5975 | 135.82 |  |  |  |  | 1265.3839 | 0.016704
|  |  | 28.6721 | 42.405 |  |  |  |  | 1268.6537 | 0.072088 | 28.6721 | 42.66 |  |  |  |  | 945.2123 | 0.027386 | 28.6721 | 136.57 |  |  |  |  | 1268.6537 | 0.041012
|  |  | 28.7468 | 43.75 |  |  |  |  | 1271.9234 | 0.078492 | 28.7468 | 42.34 |  |  |  |  | 947.6484 | 0.014332 | 28.7468 | 137.105 |  |  |  |  | 1271.9234 | 0.022324
|  |  | 28.8215 | 43.205 |  |  |  |  | 1275.1931 | 0.074056 | 28.8215 | 41.96 |  |  |  |  | 950.0845 | 0.014697 | 28.8215 | 137.305 |  |  |  |  | 1275.1931 | 0.022809
|  |  | 28.8961 | 42.245 |  |  |  |  | 1278.4628 | 0.059605 | 28.8961 | 41.4 |  |  |  |  | 952.5206 | 0.044926 | 28.8961 | 137.285 |  |  |  |  | 1278.4628 | 0.053838
|  |  | 28.9708 | 41.77 |  |  |  |  | 1281.7326 | 0.044463 | 28.9708 | 41.015 |  |  |  |  | 954.9567 | 0.069455 | 28.9708 | 137.625 |  |  |  |  | 1281.7326 | 0.043537
|  |  | 29.0455 | 41.82 |  |  |  |  | 1285.0023 | 0.036547 | 29.0455 | 40.65 |  |  |  |  | 957.3928 | 0.078343 | 29.0455 | 137.955 |  |  |  |  | 1285.0023 | 0.023591
|  |  | 29.1201 | 42.335 |  |  |  |  | 1288.272 | 0.033641 | 29.1201 | 40.485 |  |  |  |  | 959.8289 | 0.067697 | 29.1201 | 138.045 |  |  |  |  | 1288.272 | 0.063392
|  |  | 29.1948 | 42.13 |  |  |  |  | 1291.5417 | 0.030078 | 29.1948 | 40.785 |  |  |  |  | 962.2651 | 0.041199 | 29.1948 | 138.57 |  |  |  |  | 1291.5417 | 0.076823
|  |  | 29.2695 | 41.95 |  |  |  |  | 1294.8115 | 0.02293 | 29.2695 | 41.395 |  |  |  |  | 964.7012 | 0.014305 | 29.2695 | 139.215 |  |  |  |  | 1294.8115 | 0.053528
|  |  | 29.3441 | 41.95 |  |  |  |  | 1298.0812 | 0.013362 | 29.3441 | 41.855 |  |  |  |  | 967.1373 | 0.030295 | 29.3441 | 139.28 |  |  |  |  | 1298.0812 | 0.042516
|  |  | 29.4188 | 42.08 |  |  |  |  | 1301.3509 | 0.021114 | 29.4188 | 42.17 |  |  |  |  | 969.5734 | 0.046965 | 29.4188 | 139.16 |  |  |  |  | 1301.3509 | 0.056182
|  |  | 29.4935 | 43.265 |  |  |  |  | 1304.6206 | 0.038399 | 29.4935 | 42.295 |  |  |  |  | 972.0095 | 0.048474 | 29.4935 | 139.125 |  |  |  |  | 1304.6206 | 0.044415
|  |  | 29.5681 | 42.535 |  |  |  |  | 1307.8904 | 0.044208 | 29.5681 | 42.5 |  |  |  |  | 974.4456 | 0.037425 | 29.5681 | 139.175 |  |  |  |  | 1307.8904 | 0.03267
|  |  | 29.6428 | 42.3 |  |  |  |  | 1311.1601 | 0.038198 | 29.6428 | 42.615 |  |  |  |  | 976.8817 | 0.021356 | 29.6428 | 139.41 |  |  |  |  | 1311.1601 | 0.051781
|  |  | 29.7175 | 42.265 |  |  |  |  | 1314.4298 | 0.055514 | 29.7175 | 42.515 |  |  |  |  | 979.3179 | 0.01034 | 29.7175 | 139.605 |  |  |  |  | 1314.4298 | 0.048778
|  |  | 29.7921 | 41.98 |  |  |  |  | 1317.6995 | 0.090419 | 29.7921 | 42.445 |  |  |  |  | 981.754 | 0.017544 | 29.7921 | 139.455 |  |  |  |  | 1317.6995 | 0.026254
|  |  | 29.8668 | 41.88 |  |  |  |  | 1320.9693 | 0.10603 | 29.8668 | 42.62 |  |  |  |  | 984.1901 | 0.032463 | 29.8668 | 138.88 |  |  |  |  | 1320.9693 | 0.041453
|  |  | 29.9415 | 41.79 |  |  |  |  | 1324.239 | 0.092794 | 29.9415 | 42.715 |  |  |  |  | 986.6262 | 0.046467 | 29.9415 | 139.01 |  |  |  |  | 1324.239 | 0.05361
|  |  | 30.0161 | 42.655 |  |  |  |  | 1327.5087 | 0.078454 | 30.0161 | 42.535 |  |  |  |  | 989.0623 | 0.051981 | 30.0161 | 139.455 |  |  |  |  | 1327.5087 | 0.038255
|  |  | 30.0908 | 42.3 |  |  |  |  | 1330.7784 | 0.094053 | 30.0908 | 42.135 |  |  |  |  | 991.4984 | 0.0433 | 30.0908 | 139.925 |  |  |  |  | 1330.7784 | 0.02086
|  |  | 30.1655 | 43.03 |  |  |  |  | 1334.0482 | 0.1091 | 30.1655 | 41.875 |  |  |  |  | 993.9345 | 0.022575 | 30.1655 | 139.395 |  |  |  |  | 1334.0482 | 0.02877
|  |  | 30.2401 | 42.915 |  |  |  |  | 1337.3179 | 0.099096 | 30.2401 | 41.58 |  |  |  |  | 996.3707 | 0.024278 | 30.2401 | 138.815 |  |  |  |  | 1337.3179 | 0.035467
|  |  | 30.3148 | 41.355 |  |  |  |  | 1340.5876 | 0.065951 | 30.3148 | 41.455 |  |  |  |  | 998.8068 | 0.051408 | 30.3148 | 138.87 |  |  |  |  | 1340.5876 | 0.044331
|  |  | 30.3895 | 41.84 |  |  |  |  | 1343.8574 | 0.021137 | 30.3895 | 41.78 |  |  |  |  | 1001.2429 | 0.067642 | 30.3895 | 138.92 |  |  |  |  | 1343.8574 | 0.041942
|  |  | 30.4641 | 41.91 |  |  |  |  | 1347.1271 | 0.019551 | 30.4641 | 42.07 |  |  |  |  | 1003.679 | 0.064089 | 30.4641 | 138.45 |  |  |  |  | 1347.1271 | 0.013376
|  |  | 30.5388 | 42.32 |  |  |  |  | 1350.3968 | 0.035759 | 30.5388 | 42.555 |  |  |  |  | 1006.1151 | 0.042792 | 30.5388 | 137.52 |  |  |  |  | 1350.3968 | 0.032001
|  |  | 30.6135 | 42.21 |  |  |  |  | 1353.6665 | 0.016392 | 30.6135 | 43.205 |  |  |  |  | 1008.5512 | 0.024497 | 30.6135 | 137.105 |  |  |  |  | 1353.6665 | 0.06351
|  |  | 30.6881 | 42.47 |  |  |  |  | 1356.9363 | 0.044155 | 30.6881 | 43.56 |  |  |  |  | 1010.9873 | 0.043438 | 30.6881 | 137.225 |  |  |  |  | 1356.9363 | 0.060818
|  |  | 30.7628 | 42.83 |  |  |  |  | 1360.206 | 0.094527 | 30.7628 | 43.925 |  |  |  |  | 1013.4235 | 0.063235 | 30.7628 | 138.04 |  |  |  |  | 1360.206 | 0.02923
|  |  | 30.8375 | 43.37 |  |  |  |  | 1363.4757 | 0.10672 | 30.8375 | 44.02 |  |  |  |  | 1015.8596 | 0.066473 | 30.8375 | 138.315 |  |  |  |  | 1363.4757 | 0.017069
|  |  | 30.9121 | 42.5 |  |  |  |  | 1366.7454 | 0.069364 | 30.9121 | 43.96 |  |  |  |  | 1018.2957 | 0.053734 | 30.9121 | 138.015 |  |  |  |  | 1366.7454 | 0.030272
|  |  | 30.9868 | 41.815 |  |  |  |  | 1370.0152 | 0.038348 | 30.9868 | 43.975 |  |  |  |  | 1020.7318 | 0.034843 | 30.9868 | 137.535 |  |  |  |  | 1370.0152 | 0.041738
|  |  | 31.0615 | 41.375 |  |  |  |  | 1373.2849 | 0.07593 | 31.0615 | 44.21 |  |  |  |  | 1023.1679 | 0.025921 | 31.0615 | 137.345 |  |  |  |  | 1373.2849 | 0.07408
|  |  | 31.1361 | 41.69 |  |  |  |  | 1376.5546 | 0.078472 | 31.1361 | 44.315 |  |  |  |  | 1025.604 | 0.029153 | 31.1361 | 137.29 |  |  |  |  | 1376.5546 | 0.097315
|  |  | 31.2108 | 41.275 |  |  |  |  | 1379.8243 | 0.039126 | 31.2108 | 44.26 |  |  |  |  | 1028.0401 | 0.032342 | 31.2108 | 137.025 |  |  |  |  | 1379.8243 | 0.092705
|  |  | 31.2855 | 42.385 |  |  |  |  | 1383.0941 | 0.036709 | 31.2855 | 44.14 |  |  |  |  | 1030.4763 | 0.038584 | 31.2855 | 136.685 |  |  |  |  | 1383.0941 | 0.071438
|  |  | 31.3601 | 42.775 |  |  |  |  | 1386.3638 | 0.057089 | 31.3601 | 44.115 |  |  |  |  | 1032.9124 | 0.050172 | 31.3601 | 136.26 |  |  |  |  | 1386.3638 | 0.050415
|  |  | 31.4348 | 42.215 |  |  |  |  | 1389.6335 | 0.043315 | 31.4348 | 44.075 |  |  |  |  | 1035.3485 | 0.058767 | 31.4348 | 135.885 |  |  |  |  | 1389.6335 | 0.021908
|  |  | 31.5095 | 42.255 |  |  |  |  | 1392.9032 | 0.046967 | 31.5095 | 43.765 |  |  |  |  | 1037.7846 | 0.058293 | 31.5095 | 135.66 |  |  |  |  | 1392.9032 | 0.025193
|  |  | 31.5841 | 42.415 |  |  |  |  | 1396.173 | 0.069248 | 31.5841 | 43.655 |  |  |  |  | 1040.2207 | 0.05387 | 31.5841 | 135.89 |  |  |  |  | 1396.173 | 0.064555
|  |  | 31.6588 | 41.83 |  |  |  |  | 1399.4427 | 0.060573 | 31.6588 | 43.46 |  |  |  |  | 1042.6568 | 0.060813 | 31.6588 | 136.56 |  |  |  |  | 1399.4427 | 0.078789
|  |  | 31.7335 | 43.46 |  |  |  |  | 1402.7124 | 0.043936 | 31.7335 | 43.13 |  |  |  |  | 1045.0929 | 0.079582 | 31.7335 | 136.55 |  |  |  |  | 1402.7124 | 0.065611
|  |  | 31.8081 | 43.09 |  |  |  |  | 1405.9821 | 0.059941 | 31.8081 | 43.235 |  |  |  |  | 1047.5291 | 0.093993 | 31.8081 | 136.43 |  |  |  |  | 1405.9821 | 0.046151
|  |  | 31.8828 | 42.435 |  |  |  |  | 1409.2519 | 0.056061 | 31.8828 | 43.665 |  |  |  |  | 1049.9652 | 0.092539 | 31.8828 | 136.465 |  |  |  |  | 1409.2519 | 0.031582
|  |  | 31.9575 | 41.75 |  |  |  |  | 1412.5216 | 0.02013 | 31.9575 | 44.535 |  |  |  |  | 1052.4013 | 0.072524 | 31.9575 | 136.295 |  |  |  |  | 1412.5216 | 0.0099693
|  |  | 32.0321 | 41.15 |  |  |  |  | 1415.7913 | 0.044366 | 32.0321 | 45.09 |  |  |  |  | 1054.8374 | 0.039321 | 32.0321 | 136.235 |  |  |  |  | 1415.7913 | 0.024628
|  |  | 32.1068 | 41.605 |  |  |  |  | 1419.061 | 0.064342 | 32.1068 | 45.485 |  |  |  |  | 1057.2735 | 0.01325 | 32.1068 | 135.93 |  |  |  |  | 1419.061 | 0.035035
|  |  | 32.1815 | 42.375 |  |  |  |  | 1422.3308 | 0.036122 | 32.1815 | 45.705 |  |  |  |  | 1059.7096 | 0.035149 | 32.1815 | 134.605 |  |  |  |  | 1422.3308 | 0.013249
|  |  | 32.2561 | 42.905 |  |  |  |  | 1425.6005 | 0.029387 | 32.2561 | 46.03 |  |  |  |  | 1062.1457 | 0.049749 | 32.2561 | 134.605 |  |  |  |  | 1425.6005 | 0.029581
|  |  | 32.3308 | 42.515 |  |  |  |  | 1428.8702 | 0.079174 | 32.3308 | 46.175 |  |  |  |  | 1064.5819 | 0.046003 | 32.3308 | 134.925 |  |  |  |  | 1428.8702 | 0.061389
|  |  | 32.4055 | 42.865 |  |  |  |  | 1432.14 | 0.082663 | 32.4055 | 46.355 |  |  |  |  | 1067.018 | 0.026736 | 32.4055 | 135.08 |  |  |  |  | 1432.14 | 0.05581
|  |  | 32.4801 | 42.925 |  |  |  |  | 1435.4097 | 0.039385 | 32.4801 | 46.615 |  |  |  |  | 1069.4541 | 0.0087463 | 32.4801 | 134.7 |  |  |  |  | 1435.4097 | 0.013001
|  |  | 32.5548 | 43.585 |  |  |  |  | 1438.6794 | 0.06354 | 32.5548 | 46.79 |  |  |  |  | 1071.8902 | 0.029657 | 32.5548 | 134.195 |  |  |  |  | 1438.6794 | 0.043117
|  |  | 32.6295 | 43.96 |  |  |  |  | 1441.9491 | 0.1166 | 32.6295 | 46.805 |  |  |  |  | 1074.3263 | 0.044433 | 32.6295 | 134.165 |  |  |  |  | 1441.9491 | 0.07481
|  |  | 32.7041 | 43.4 |  |  |  |  | 1445.2189 | 0.11968 | 32.7041 | 46.565 |  |  |  |  | 1076.7624 | 0.044487 | 32.7041 | 134.795 |  |  |  |  | 1445.2189 | 0.069309
|  |  | 32.7788 | 44.28 |  |  |  |  | 1448.4886 | 0.082075 | 32.7788 | 46.415 |  |  |  |  | 1079.1985 | 0.030835 | 32.7788 | 135.025 |  |  |  |  | 1448.4886 | 0.048177
|  |  | 32.8535 | 43.955 |  |  |  |  | 1451.7583 | 0.077852 | 32.8535 | 46.395 |  |  |  |  | 1081.6347 | 0.009543 | 32.8535 | 135.005 |  |  |  |  | 1451.7583 | 0.053643
|  |  | 32.9281 | 43.44 |  |  |  |  | 1455.028 | 0.10678 | 32.9281 | 46.2 |  |  |  |  | 1084.0708 | 0.010865 | 32.9281 | 134.295 |  |  |  |  | 1455.028 | 0.066062
|  |  | 33.0028 | 44.15 |  |  |  |  | 1458.2978 | 0.098612 | 33.0028 | 45.97 |  |  |  |  | 1086.5069 | 0.022551 | 33.0028 | 133.77 |  |  |  |  | 1458.2978 | 0.064507
|  |  | 33.0775 | 42.98 |  |  |  |  | 1461.5675 | 0.049998 | 33.0775 | 45.765 |  |  |  |  | 1088.943 | 0.021297 | 33.0775 | 133.775 |  |  |  |  | 1461.5675 | 0.049169
|  |  | 33.1521 | 42.485 |  |  |  |  | 1464.8372 | 0.0095574 | 33.1521 | 45.265 |  |  |  |  | 1091.3791 | 0.010851 | 33.1521 | 134.155 |  |  |  |  | 1464.8372 | 0.016696
|  |  | 33.2268 | 42.94 |  |  |  |  | 1468.1069 | 0.045348 | 33.2268 | 44.985 |  |  |  |  | 1093.8152 | 0.02365 | 33.2268 | 134.515 |  |  |  |  | 1468.1069 | 0.031772
|  |  | 33.3015 | 42.07 |  |  |  |  | 1471.3767 | 0.051487 | 33.3015 | 45.03 |  |  |  |  | 1096.2513 | 0.049618 | 33.3015 | 134.945 |  |  |  |  | 1471.3767 | 0.071814
|  |  | 33.3761 | 43.315 |  |  |  |  | 1474.6464 | 0.043746 | 33.3761 | 46.035 |  |  |  |  | 1098.6875 | 0.072191 | 33.3761 | 134.405 |  |  |  |  | 1474.6464 | 0.081765
|  |  | 33.4508 | 43.92 |  |  |  |  | 1477.9161 | 0.045588 | 33.4508 | 46.605 |  |  |  |  | 1101.1236 | 0.085517 | 33.4508 | 134.05 |  |  |  |  | 1477.9161 | 0.056802
|  |  | 33.5255 | 42.815 |  |  |  |  | 1481.1858 | 0.060792 | 33.5255 | 46.955 |  |  |  |  | 1103.5597 | 0.087542 | 33.5255 | 133.785 |  |  |  |  | 1481.1858 | 0.019944
|  |  | 33.6002 | 43.2 |  |  |  |  | 1484.4556 | 0.073491 | 33.6002 | 47.205 |  |  |  |  | 1105.9958 | 0.080381 | 33.6002 | 134.24 |  |  |  |  | 1484.4556 | 0.018487
|  |  | 33.6748 | 42.55 |  |  |  |  | 1487.7253 | 0.065981 | 33.6748 | 47.33 |  |  |  |  | 1108.4319 | 0.070623 | 33.6748 | 133.825 |  |  |  |  | 1487.7253 | 0.012122
|  |  | 33.7495 | 44.125 |  |  |  |  | 1490.995 | 0.036273 | 33.7495 | 47.385 |  |  |  |  | 1110.868 | 0.06688 | 33.7495 | 132.93 |  |  |  |  | 1490.995 | 0.018095
|  |  | 33.8242 | 44.995 |  |  |  |  | 1494.2647 | 0.017147 | 33.8242 | 47.31 |  |  |  |  | 1113.3041 | 0.070839 | 33.8242 | 132.62 |  |  |  |  | 1494.2647 | 0.039009
|  |  | 33.8988 | 45.12 |  |  |  |  | 1497.5345 | 0.029365 | 33.8988 | 47.19 |  |  |  |  | 1115.7403 | 0.074586 | 33.8988 | 133.345 |  |  |  |  | 1497.5345 | 0.034709
|  |  | 33.9735 | 43.75 |  |  |  |  | 1500.8042 | 0.022772 | 33.9735 | 47.045 |  |  |  |  | 1118.1764 | 0.070546 | 33.9735 | 133.94 |  |  |  |  | 1500.8042 | 0.03381
|  |  | 34.0482 | 42.55 |  |  |  |  | 1504.0739 | 0.020715 | 34.0482 | 46.81 |  |  |  |  | 1120.6125 | 0.056552 | 34.0482 | 134.405 |  |  |  |  | 1504.0739 | 0.06973
|  |  | 34.1228 | 42.07 |  |  |  |  | 1507.3436 | 0.030186 | 34.1228 | 46.545 |  |  |  |  | 1123.0486 | 0.035249 | 34.1228 | 134.8 |  |  |  |  | 1507.3436 | 0.091369
|  |  | 34.1975 | 42.515 |  |  |  |  | 1510.6134 | 0.029881 | 34.1975 | 46.25 |  |  |  |  | 1125.4847 | 0.012366 | 34.1975 | 134.775 |  |  |  |  | 1510.6134 | 0.082847
|  |  | 34.2722 | 42.46 |  |  |  |  | 1513.8831 | 0.041413 | 34.2722 | 46.05 |  |  |  |  | 1127.9208 | 0.012117 | 34.2722 | 134.03 |  |  |  |  | 1513.8831 | 0.058009
|  |  | 34.3468 | 42.705 |  |  |  |  | 1517.1528 | 0.048508 | 34.3468 | 45.715 |  |  |  |  | 1130.3569 | 0.028475 | 34.3468 | 134.255 |  |  |  |  | 1517.1528 | 0.035495
|  |  | 34.4215 | 42.485 |  |  |  |  | 1520.4226 | 0.034264 | 34.4215 | 45.305 |  |  |  |  | 1132.793 | 0.042225 | 34.4215 | 134.845 |  |  |  |  | 1520.4226 | 0.015285
|  |  | 34.4962 | 42.455 |  |  |  |  | 1523.6923 | 0.057731 | 34.4962 | 44.99 |  |  |  |  | 1135.2292 | 0.053765 | 34.4962 | 134.85 |  |  |  |  | 1523.6923 | 0.026389
|  |  | 34.5708 | 41.865 |  |  |  |  | 1526.962 | 0.098858 | 34.5708 | 44.215 |  |  |  |  | 1137.6653 | 0.061166 | 34.5708 | 134.495 |  |  |  |  | 1526.962 | 0.043169
|  |  | 34.6455 | 42.61 |  |  |  |  | 1530.2317 | 0.103 | 34.6455 | 43.595 |  |  |  |  | 1140.1014 | 0.060696 | 34.6455 | 134.22 |  |  |  |  | 1530.2317 | 0.033582
|  |  | 34.7202 | 43.245 |  |  |  |  | 1533.5015 | 0.065886 | 34.7202 | 43.19 |  |  |  |  | 1142.5375 | 0.049958 | 34.7202 | 134.4 |  |  |  |  | 1533.5015 | 0.012782
|  |  | 34.7948 | 44.76 |  |  |  |  | 1536.7712 | 0.036406 | 34.7948 | 43.32 |  |  |  |  | 1144.9736 | 0.030479 | 34.7948 | 134.875 |  |  |  |  | 1536.7712 | 0.037162
|  |  | 34.8695 | 43.755 |  |  |  |  | 1540.0409 | 0.054168 | 34.8695 | 44.035 |  |  |  |  | 1147.4097 | 0.014677 | 34.8695 | 135.16 |  |  |  |  | 1540.0409 | 0.040471
|  |  | 34.9442 | 41.87 |  |  |  |  | 1543.3106 | 0.052215 | 34.9442 | 44.56 |  |  |  |  | 1149.8458 | 0.030027 | 34.9442 | 134.115 |  |  |  |  | 1543.3106 | 0.014408
|  |  | 35.0188 | 41.975 |  |  |  |  | 1546.5804 | 0.039606 | 35.0188 | 44.845 |  |  |  |  | 1152.282 | 0.047913 | 35.0188 | 133.645 |  |  |  |  | 1546.5804 | 0.044487
|  |  | 35.0935 | 41.84 |  |  |  |  | 1549.8501 | 0.045024 | 35.0935 | 44.71 |  |  |  |  | 1154.7181 | 0.056167 | 35.0935 | 133.64 |  |  |  |  | 1549.8501 | 0.075432
|  |  | 35.1682 | 42.475 |  |  |  |  | 1553.1198 | 0.042853 | 35.1682 | 44.555 |  |  |  |  | 1157.1542 | 0.053391 | 35.1682 | 133.615 |  |  |  |  | 1553.1198 | 0.070087
|  |  | 35.2428 | 41.88 |  |  |  |  | 1556.3895 | 0.026027 | 35.2428 | 44.745 |  |  |  |  | 1159.5903 | 0.04229 | 35.2428 | 133.595 |  |  |  |  | 1556.3895 | 0.040054
|  |  | 35.3175 | 42.81 |  |  |  |  | 1559.6593 | 0.014919 | 35.3175 | 44.88 |  |  |  |  | 1162.0264 | 0.028093 | 35.3175 | 133.74 |  |  |  |  | 1559.6593 | 0.042691
|  |  | 35.3922 | 42.375 |  |  |  |  | 1562.929 | 0.016993 | 35.3922 | 44.84 |  |  |  |  | 1164.4625 | 0.016497 | 35.3922 | 133.665 |  |  |  |  | 1562.929 | 0.057666
|  |  | 35.4668 | 41.445 |  |  |  |  | 1566.1987 | 0.026515 | 35.4668 | 44.455 |  |  |  |  | 1166.8986 | 0.010798 | 35.4668 | 133.65 |  |  |  |  | 1566.1987 | 0.041775
|  |  | 35.5415 | 42.435 |  |  |  |  | 1569.4684 | 0.032247 | 35.5415 | 44.205 |  |  |  |  | 1169.3348 | 0.010635 | 35.5415 | 133.76 |  |  |  |  | 1569.4684 | 0.0056418
|  |  | 35.6162 | 42.22 |  |  |  |  | 1572.7382 | 0.034777 | 35.6162 | 44.085 |  |  |  |  | 1171.7709 | 0.015515 | 35.6162 | 133.78 |  |  |  |  | 1572.7382 | 0.045989
|  |  | 35.6908 | 43.435 |  |  |  |  | 1576.0079 | 0.054311 | 35.6908 | 43.91 |  |  |  |  | 1174.207 | 0.02134 | 35.6908 | 134.135 |  |  |  |  | 1576.0079 | 0.076303
|  |  | 35.7655 | 43.74 |  |  |  |  | 1579.2776 | 0.068907 | 35.7655 | 43.615 |  |  |  |  | 1176.6431 | 0.02267 | 35.7655 | 134.805 |  |  |  |  | 1579.2776 | 0.080909
|  |  | 35.8402 | 41.64 |  |  |  |  | 1582.5473 | 0.057966 | 35.8402 | 43.305 |  |  |  |  | 1179.0792 | 0.017014 | 35.8402 | 135.26 |  |  |  |  | 1582.5473 | 0.056851
|  |  | 35.9148 | 42.875 |  |  |  |  | 1585.8171 | 0.041845 | 35.9148 | 42.9 |  |  |  |  | 1181.5153 | 0.01111 | 35.9148 | 135.275 |  |  |  |  | 1585.8171 | 0.010726
|  |  | 35.9895 | 42.615 |  |  |  |  | 1589.0868 | 0.052395 | 35.9895 | 42.505 |  |  |  |  | 1183.9514 | 0.022059 | 35.9895 | 134.9 |  |  |  |  | 1589.0868 | 0.038531
|  |  | 36.0642 | 43.245 |  |  |  |  | 1592.3565 | 0.055522 | 36.0642 | 42.22 |  |  |  |  | 1186.3876 | 0.035169 | 36.0642 | 134.825 |  |  |  |  | 1592.3565 | 0.065958
|  |  | 36.1388 | 42.095 |  |  |  |  | 1595.6262 | 0.051074 | 36.1388 | 42.515 |  |  |  |  | 1188.8237 | 0.03995 | 36.1388 | 135.675 |  |  |  |  | 1595.6262 | 0.058726
|  |  | 36.2135 | 40.535 |  |  |  |  | 1598.896 | 0.057706 | 36.2135 | 43.27 |  |  |  |  | 1191.2598 | 0.033309 | 36.2135 | 135.74 |  |  |  |  | 1598.896 | 0.046309
|  |  | 36.2882 | 41.85 |  |  |  |  | 1602.1657 | 0.046265 | 36.2882 | 43.38 |  |  |  |  | 1193.6959 | 0.018088 | 36.2882 | 135.63 |  |  |  |  | 1602.1657 | 0.078202
|  |  | 36.3628 | 42.85 |  |  |  |  | 1605.4354 | 0.0023988 | 36.3628 | 43.42 |  |  |  |  | 1196.132 | 0.011473 | 36.3628 | 135.46 |  |  |  |  | 1605.4354 | 0.10231
|  |  | 36.4375 | 42.225 |  |  |  |  | 1608.7052 | 0.050161 | 36.4375 | 43.63 |  |  |  |  | 1198.5681 | 0.023537 | 36.4375 | 135.405 |  |  |  |  | 1608.7052 | 0.088646
|  |  | 36.5122 | 41.495 |  |  |  |  | 1611.9749 | 0.074395 | 36.5122 | 43.78 |  |  |  |  | 1201.0042 | 0.028012 | 36.5122 | 135.765 |  |  |  |  | 1611.9749 | 0.047645
|  |  | 36.5868 | 40.615 |  |  |  |  | 1615.2446 | 0.053872 | 36.5868 | 43.685 |  |  |  |  | 1203.4404 | 0.020447 | 36.5868 | 136.645 |  |  |  |  | 1615.2446 | 0.025817
|  |  | 36.6615 | 41.755 |  |  |  |  | 1618.5143 | 0.014862 | 36.6615 | 43.595 |  |  |  |  | 1205.8765 | 0.010725 | 36.6615 | 136.415 |  |  |  |  | 1618.5143 | 0.032262
|  |  | 36.7362 | 41.6 |  |  |  |  | 1621.7841 | 0.046923 | 36.7362 | 43.59 |  |  |  |  | 1208.3126 | 0.025702 | 36.7362 | 135.89 |  |  |  |  | 1621.7841 | 0.022058
|  |  | 36.8108 | 41.23 |  |  |  |  | 1625.0538 | 0.068718 | 36.8108 | 43.475 |  |  |  |  | 1210.7487 | 0.041963 | 36.8108 | 135.88 |  |  |  |  | 1625.0538 | 0.0526
|  |  | 36.8855 | 41.975 |  |  |  |  | 1628.3235 | 0.077374 | 36.8855 | 43.365 |  |  |  |  | 1213.1848 | 0.048294 | 36.8855 | 136.01 |  |  |  |  | 1628.3235 | 0.091209
|  |  | 36.9602 | 42.485 |  |  |  |  | 1631.5932 | 0.079273 | 36.9602 | 43.035 |  |  |  |  | 1215.6209 | 0.043211 | 36.9602 | 136.465 |  |  |  |  | 1631.5932 | 0.098866
|  |  | 37.0348 | 42.765 |  |  |  |  | 1634.863 | 0.059756 | 37.0348 | 42.755 |  |  |  |  | 1218.057 | 0.031089 | 37.0348 | 136.605 |  |  |  |  | 1634.863 | 0.068872
|  |  | 37.1095 | 41.905 |  |  |  |  | 1638.1327 | 0.017899 | 37.1095 | 42.26 |  |  |  |  | 1220.4932 | 0.021212 | 37.1095 | 136.59 |  |  |  |  | 1638.1327 | 0.020682
|  |  | 37.1842 | 40.315 |  |  |  |  | 1641.4024 | 0.028227 | 37.1842 | 41.78 |  |  |  |  | 1222.9293 | 0.021105 | 37.1842 | 136.61 |  |  |  |  | 1641.4024 | 0.027247
|  |  | 37.2588 | 40.305 |  |  |  |  | 1644.6721 | 0.049028 | 37.2588 | 41.385 |  |  |  |  | 1225.3654 | 0.024356 | 37.2588 | 136.845 |  |  |  |  | 1644.6721 | 0.037322
|  |  | 37.3335 | 41.91 |  |  |  |  | 1647.9419 | 0.041905 | 37.3335 | 41.06 |  |  |  |  | 1227.8015 | 0.026745 | 37.3335 | 137.635 |  |  |  |  | 1647.9419 | 0.016929
|  |  | 37.4082 | 41.2 |  |  |  |  | 1651.2116 | 0.021873 | 37.4082 | 40.64 |  |  |  |  | 1230.2376 | 0.029582 | 37.4082 | 138.37 |  |  |  |  | 1651.2116 | 0.018067
|  |  | 37.4828 | 40.68 |  |  |  |  | 1654.4813 | 0.019244 | 37.4828 | 40.715 |  |  |  |  | 1232.6737 | 0.032026 | 37.4828 | 138.57 |  |  |  |  | 1654.4813 | 0.036915
|  |  | 37.5575 | 41.08 |  |  |  |  | 1657.751 | 0.03493 | 37.5575 | 41.135 |  |  |  |  | 1235.1098 | 0.032418 | 37.5575 | 138.34 |  |  |  |  | 1657.751 | 0.028648
|  |  | 37.6322 | 38.95 |  |  |  |  | 1661.0208 | 0.0433 | 37.6322 | 41.87 |  |  |  |  | 1237.546 | 0.032634 | 37.6322 | 138.05 |  |  |  |  | 1661.0208 | 0.0054932
|  |  | 37.7068 | 40.74 |  |  |  |  | 1664.2905 | 0.045788 | 37.7068 | 42.345 |  |  |  |  | 1239.9821 | 0.035312 | 37.7068 | 137.91 |  |  |  |  | 1664.2905 | 0.028145
|  |  | 37.7815 | 40.19 |  |  |  |  | 1667.5602 | 0.065125 | 37.7815 | 42.7 |  |  |  |  | 1242.4182 | 0.037302 | 37.7815 | 137.39 |  |  |  |  | 1667.5602 | 0.039478
|  |  | 37.8562 | 41.27 |  |  |  |  | 1670.8299 | 0.084926 | 37.8562 | 42.87 |  |  |  |  | 1244.8543 | 0.032807 | 37.8562 | 137.23 |  |  |  |  | 1670.8299 | 0.051373
|  |  | 37.9308 | 40.815 |  |  |  |  | 1674.0997 | 0.075624 | 37.9308 | 42.93 |  |  |  |  | 1247.2904 | 0.019784 | 37.9308 | 137.935 |  |  |  |  | 1674.0997 | 0.075513
|  |  | 38.0055 | 42.115 |  |  |  |  | 1677.3694 | 0.053043 | 38.0055 | 43.09 |  |  |  |  | 1249.7265 | 0.0073057 | 38.0055 | 138.04 |  |  |  |  | 1677.3694 | 0.08392
|  |  | 38.0802 | 41.2 |  |  |  |  | 1680.6391 | 0.079445 | 38.0802 | 43.345 |  |  |  |  | 1252.1626 | 0.02256 | 38.0802 | 137.52 |  |  |  |  | 1680.6391 | 0.058369
|  |  | 38.1548 | 41.24 |  |  |  |  | 1683.9088 | 0.10793 | 38.1548 | 43.355 |  |  |  |  | 1254.5988 | 0.035526 | 38.1548 | 137.34 |  |  |  |  | 1683.9088 | 0.024143
|  |  | 38.2295 | 42.105 |  |  |  |  | 1687.1786 | 0.099804 | 38.2295 | 43.345 |  |  |  |  | 1257.0349 | 0.038504 | 38.2295 | 137.885 |  |  |  |  | 1687.1786 | 0.059027
|  |  | 38.3042 | 41.93 |  |  |  |  | 1690.4483 | 0.066723 | 38.3042 | 43.235 |  |  |  |  | 1259.471 | 0.032822 | 38.3042 | 138.045 |  |  |  |  | 1690.4483 | 0.078833
|  |  | 38.3788 | 41.22 |  |  |  |  | 1693.718 | 0.036959 | 38.3788 | 43.17 |  |  |  |  | 1261.9071 | 0.026567 | 38.3788 | 137.77 |  |  |  |  | 1693.718 | 0.058978
|  |  | 38.4535 | 41.7 |  |  |  |  | 1696.9878 | 0.021947 | 38.4535 | 43.11 |  |  |  |  | 1264.3432 | 0.028197 | 38.4535 | 137.79 |  |  |  |  | 1696.9878 | 0.017667
|  |  | 38.5282 | 41.44 |  |  |  |  | 1700.2575 | 0.031784 | 38.5282 | 42.895 |  |  |  |  | 1266.7793 | 0.03193 | 38.5282 | 137.9 |  |  |  |  | 1700.2575 | 0.039862
|  |  | 38.6028 | 41.905 |  |  |  |  | 1703.5272 | 0.044922 | 38.6028 | 42.66 |  |  |  |  | 1269.2154 | 0.029742 | 38.6028 | 137.535 |  |  |  |  | 1703.5272 | 0.06237
|  |  | 38.6775 | 42.62 |  |  |  |  | 1706.7969 | 0.037606 | 38.6775 | 42.545 |  |  |  |  | 1271.6516 | 0.020649 | 38.6775 | 137.8 |  |  |  |  | 1706.7969 | 0.0576
|  |  | 38.7522 | 42.665 |  |  |  |  | 1710.0667 | 0.022718 | 38.7522 | 42.285 |  |  |  |  | 1274.0877 | 0.0085015 | 38.7522 | 137.865 |  |  |  |  | 1710.0667 | 0.035705
|  |  | 38.8268 | 43.2 |  |  |  |  | 1713.3364 | 0.025224 | 38.8268 | 42.175 |  |  |  |  | 1276.5238 | 0.0081278 | 38.8268 | 137.62 |  |  |  |  | 1713.3364 | 0.012756
|  |  | 38.9015 | 40.665 |  |  |  |  | 1716.6061 | 0.016277 | 38.9015 | 42.54 |  |  |  |  | 1278.9599 | 0.017778 | 38.9015 | 137.99 |  |  |  |  | 1716.6061 | 0.0084265
|  |  | 38.9762 | 41.01 |  |  |  |  | 1719.8758 | 0.022265 | 38.9762 | 43.64 |  |  |  |  | 1281.396 | 0.025484 | 38.9762 | 138.025 |  |  |  |  | 1719.8758 | 0.010308
|  |  | 39.0508 | 41.62 |  |  |  |  | 1723.1456 | 0.044974 | 39.0508 | 44.18 |  |  |  |  | 1283.8321 | 0.03036 | 39.0508 | 137.82 |  |  |  |  | 1723.1456 | 0.005961
|  |  | 39.1255 | 43.61 |  |  |  |  | 1726.4153 | 0.040555 | 39.1255 | 44.525 |  |  |  |  | 1286.2682 | 0.030933 | 39.1255 | 137.76 |  |  |  |  | 1726.4153 | 0.023734
|  |  | 39.2002 | 41.305 |  |  |  |  | 1729.685 | 0.020487 | 39.2002 | 44.915 |  |  |  |  | 1288.7044 | 0.025936 | 39.2002 | 137.8 |  |  |  |  | 1729.685 | 0.042305
|  |  | 39.2748 | 41.33 |  |  |  |  | 1732.9547 | 0.042764 | 39.2748 | 45.35 |  |  |  |  | 1291.1405 | 0.016639 | 39.2748 | 137.42 |  |  |  |  | 1732.9547 | 0.047133
|  |  | 39.3495 | 41.64 |  |  |  |  | 1736.2245 | 0.051622 | 39.3495 | 45.43 |  |  |  |  | 1293.5766 | 0.010018 | 39.3495 | 136.41 |  |  |  |  | 1736.2245 | 0.034437
|  |  | 39.4242 | 41.92 |  |  |  |  | 1739.4942 | 0.031853 | 39.4242 | 45.515 |  |  |  |  | 1296.0127 | 0.01429 | 39.4242 | 136.17 |  |  |  |  | 1739.4942 | 0.015681
|  |  | 39.4988 | 42.18 |  |  |  |  | 1742.7639 | 0.0441 | 39.4988 | 45.715 |  |  |  |  | 1298.4488 | 0.018663 | 39.4988 | 136.115 |  |  |  |  | 1742.7639 | 0.018841
|  |  | 39.5735 | 41.595 |  |  |  |  | 1746.0336 | 0.074792 | 39.5735 | 45.82 |  |  |  |  | 1300.8849 | 0.01906 | 39.5735 | 135.96 |  |  |  |  | 1746.0336 | 0.02275
|  |  | 39.6482 | 41.21 |  |  |  |  | 1749.3034 | 0.078304 | 39.6482 | 45.875 |  |  |  |  | 1303.321 | 0.017893 | 39.6482 | 136.19 |  |  |  |  | 1749.3034 | 0.01437
|  |  | 39.7228 | 40.935 |  |  |  |  | 1752.5731 | 0.057709 | 39.7228 | 45.89 |  |  |  |  | 1305.7571 | 0.016983 | 39.7228 | 136.27 |  |  |  |  | 1752.5731 | 0.0115
|  |  | 39.7975 | 42.055 |  |  |  |  | 1755.8428 | 0.037041 | 39.7975 | 45.97 |  |  |  |  | 1308.1933 | 0.014244 | 39.7975 | 136.255 |  |  |  |  | 1755.8428 | 0.025122
|  |  | 39.8722 | 40.83 |  |  |  |  | 1759.1125 | 0.039301 | 39.8722 | 46.03 |  |  |  |  | 1310.6294 | 0.010175 | 39.8722 | 135.44 |  |  |  |  | 1759.1125 | 0.034925
|  |  | 39.9468 | 42.28 |  |  |  |  | 1762.3823 | 0.046784 | 39.9468 | 45.865 |  |  |  |  | 1313.0655 | 0.015162 | 39.9468 | 135.115 |  |  |  |  | 1762.3823 | 0.035435
|  |  | 40.0215 | 41.845 |  |  |  |  | 1765.652 | 0.035808 | 40.0215 | 45.545 |  |  |  |  | 1315.5016 | 0.025232 | 40.0215 | 135.15 |  |  |  |  | 1765.652 | 0.02615
|  |  | 40.0962 | 42.01 |  |  |  |  | 1768.9217 | 0.0091171 | 40.0962 | 45.32 |  |  |  |  | 1317.9377 | 0.030578 | 40.0962 | 135.27 |  |  |  |  | 1768.9217 | 0.019714
|  |  | 40.1708 | 42.735 |  |  |  |  | 1772.1914 | 0.040239 | 40.1708 | 45.19 |  |  |  |  | 1320.3738 | 0.027284 | 40.1708 | 135.16 |  |  |  |  | 1772.1914 | 0.030741
|  |  | 40.2455 | 42.86 |  |  |  |  | 1775.4612 | 0.054741 | 40.2455 | 45.345 |  |  |  |  | 1322.8099 | 0.016774 | 40.2455 | 135.59 |  |  |  |  | 1775.4612 | 0.039955
|  |  | 40.3202 | 42.935 |  |  |  |  | 1778.7309 | 0.027591 | 40.3202 | 46 |  |  |  |  | 1325.2461 | 0.011185 | 40.3202 | 135.8 |  |  |  |  | 1778.7309 | 0.040202
|  |  | 40.3948 | 43.955 |  |  |  |  | 1782.0006 | 0.033816 | 40.3948 | 47.36 |  |  |  |  | 1327.6822 | 0.020333 | 40.3948 | 135.38 |  |  |  |  | 1782.0006 | 0.035395
|  |  | 40.4695 | 43.56 |  |  |  |  | 1785.2704 | 0.088908 | 40.4695 | 47.94 |  |  |  |  | 1330.1183 | 0.025789 | 40.4695 | 135.265 |  |  |  |  | 1785.2704 | 0.030529
|  |  | 40.5442 | 42.94 |  |  |  |  | 1788.5401 | 0.10757 | 40.5442 | 48.19 |  |  |  |  | 1332.5544 | 0.022402 | 40.5442 | 135.275 |  |  |  |  | 1788.5401 | 0.032657
|  |  | 40.6188 | 43.345 |  |  |  |  | 1791.8098 | 0.081661 | 40.6188 | 48.46 |  |  |  |  | 1334.9905 | 0.012535 | 40.6188 | 134.785 |  |  |  |  | 1791.8098 | 0.040125
|  |  | 40.6935 | 42.885 |  |  |  |  | 1795.0795 | 0.033887 | 40.6935 | 48.79 |  |  |  |  | 1337.4266 | 0.0058073 | 40.6935 | 134.045 |  |  |  |  | 1795.0795 | 0.036832
|  |  | 40.7682 | 41.62 |  |  |  |  | 1798.3493 | 0.013214 | 40.7682 | 49.085 |  |  |  |  | 1339.8627 | 0.0097206 | 40.7682 | 133.845 |  |  |  |  | 1798.3493 | 0.014572
|  |  | 40.8428 | 41.845 |  |  |  |  | 1801.619 | 0.026662 | 40.8428 | 49.16 |  |  |  |  | 1342.2989 | 0.0079302 | 40.8428 | 133.47 |  |  |  |  | 1801.619 | 0.021269
|  |  | 40.9175 | 43.955 |  |  |  |  | 1804.8887 | 0.035737 | 40.9175 | 49.065 |  |  |  |  | 1344.735 | 0.002525 | 40.9175 | 133.37 |  |  |  |  | 1804.8887 | 0.048833
|  |  | 40.9922 | 44.51 |  |  |  |  | 1808.1584 | 0.041091 | 40.9922 | 49.02 |  |  |  |  | 1347.1711 | 0.01555 | 40.9922 | 133.575 |  |  |  |  | 1808.1584 | 0.053979
|  |  | 41.0668 | 44.395 |  |  |  |  | 1811.4282 | 0.028836 | 41.0668 | 48.94 |  |  |  |  | 1349.6072 | 0.026341 | 41.0668 | 133.48 |  |  |  |  | 1811.4282 | 0.036117
|  |  | 41.1415 | 43.56 |  |  |  |  | 1814.6979 | 0.025907 | 41.1415 | 48.815 |  |  |  |  | 1352.0433 | 0.029053 | 41.1415 | 133.535 |  |  |  |  | 1814.6979 | 0.020318
|  |  | 41.2162 | 42.665 |  |  |  |  | 1817.9676 | 0.048697 | 41.2162 | 48.605 |  |  |  |  | 1354.4794 | 0.022172 | 41.2162 | 133.035 |  |  |  |  | 1817.9676 | 0.031444
|  |  | 41.2909 | 42.61 |  |  |  |  | 1821.2373 | 0.050637 | 41.2909 | 48.365 |  |  |  |  | 1356.9155 | 0.012838 | 41.2909 | 132.33 |  |  |  |  | 1821.2373 | 0.034508
|  |  | 41.3655 | 42.85 |  |  |  |  | 1824.5071 | 0.042369 | 41.3655 | 48.075 |  |  |  |  | 1359.3517 | 0.021312 | 41.3655 | 131.98 |  |  |  |  | 1824.5071 | 0.027827
|  |  | 41.4402 | 43.99 |  |  |  |  | 1827.7768 | 0.071616 | 41.4402 | 47.92 |  |  |  |  | 1361.7878 | 0.034312 | 41.4402 | 132 |  |  |  |  | 1827.7768 | 0.024308
|  |  | 41.5149 | 43.17 |  |  |  |  | 1831.0465 | 0.099899 | 41.5149 | 48.02 |  |  |  |  | 1364.2239 | 0.03971 | 41.5149 | 131.595 |  |  |  |  | 1831.0465 | 0.024879
|  |  | 41.5895 | 44.445 |  |  |  |  | 1834.3162 | 0.091718 | 41.5895 | 48.345 |  |  |  |  | 1366.66 | 0.035495 | 41.5895 | 131.16 |  |  |  |  | 1834.3162 | 0.028832
|  |  | 41.6642 | 43.785 |  |  |  |  | 1837.586 | 0.046795 | 41.6642 | 48.69 |  |  |  |  | 1369.0961 | 0.024254 | 41.6642 | 130.635 |  |  |  |  | 1837.586 | 0.03005
|  |  | 41.7389 | 43.77 |  |  |  |  | 1840.8557 | 0.018418 | 41.7389 | 48.94 |  |  |  |  | 1371.5322 | 0.011026 | 41.7389 | 130.865 |  |  |  |  | 1840.8557 | 0.015881
|  |  | 41.8135 | 43.39 |  |  |  |  | 1844.1254 | 0.058339 | 41.8135 | 49.26 |  |  |  |  | 1373.9683 | 0.0010989 | 41.8135 | 131.105 |  |  |  |  | 1844.1254 | 0.015264
|  |  | 41.8882 | 43.14 |  |  |  |  | 1847.3951 | 0.063437 | 41.8882 | 49.415 |  |  |  |  | 1376.4045 | 0.0039886 | 41.8882 | 130.99 |  |  |  |  | 1847.3951 | 0.044574
|  |  | 41.9629 | 44.04 |  |  |  |  | 1850.6649 | 0.032852 | 41.9629 | 49.415 |  |  |  |  | 1378.8406 | 0.003642 | 41.9629 | 130.745 |  |  |  |  | 1850.6649 | 0.054066
|  |  | 42.0375 | 43.5 |  |  |  |  | 1853.9346 | 0.023848 | 42.0375 | 49.3 |  |  |  |  | 1381.2767 | 0.0038164 | 42.0375 | 130.54 |  |  |  |  | 1853.9346 | 0.03552
|  |  | 42.1122 | 43.505 |  |  |  |  | 1857.2043 | 0.059245 | 42.1122 | 49.265 |  |  |  |  | 1383.7128 | 0.0063999 | 42.1122 | 130.665 |  |  |  |  | 1857.2043 | 0.00023037
|  |  | 42.1869 | 43.73 |  |  |  |  | 1860.474 | 0.067271 | 42.1869 | 49.22 |  |  |  |  | 1386.1489 | 0.0069523 | 42.1869 | 130.59 |  |  |  |  | 1860.474 | 0.029922
|  |  | 42.2615 | 43.9 |  |  |  |  | 1863.7438 | 0.047521 | 42.2615 | 48.99 |  |  |  |  | 1388.585 | 0.0050436 | 42.2615 | 130.42 |  |  |  |  | 1863.7438 | 0.037132
|  |  | 42.3362 | 44.45 |  |  |  |  | 1867.0135 | 0.02647 | 42.3362 | 48.89 |  |  |  |  | 1391.0211 | 0.0036562 | 42.3362 | 130.535 |  |  |  |  | 1867.0135 | 0.02335
|  |  | 42.4109 | 43.075 |  |  |  |  | 1870.2832 | 0.034978 | 42.4109 | 48.715 |  |  |  |  | 1393.4573 | 0.0049968 | 42.4109 | 130.725 |  |  |  |  | 1870.2832 | 0.010953
|  |  | 42.4855 | 42.94 |  |  |  |  | 1873.553 | 0.0507 | 42.4855 | 48.49 |  |  |  |  | 1395.8934 | 0.0073467 | 42.4855 | 131.05 |  |  |  |  | 1873.553 | 0.0090423
|  |  | 42.5602 | 43.23 |  |  |  |  | 1876.8227 | 0.07066 | 42.5602 | 48.055 |  |  |  |  | 1398.3295 | 0.012766 | 42.5602 | 131.105 |  |  |  |  | 1876.8227 | 0.02083
|  |  | 42.6349 | 45.1 |  |  |  |  | 1880.0924 | 0.083855 | 42.6349 | 47.34 |  |  |  |  | 1400.7656 | 0.021145 | 42.6349 | 130.725 |  |  |  |  | 1880.0924 | 0.050241
|  |  | 42.7095 | 45.83 |  |  |  |  | 1883.3621 | 0.073487 | 42.7095 | 46.945 |  |  |  |  | 1403.2017 | 0.029567 | 42.7095 | 130.71 |  |  |  |  | 1883.3621 | 0.065695
|  |  | 42.7842 | 45.105 |  |  |  |  | 1886.6319 | 0.041984 | 42.7842 | 46.575 |  |  |  |  | 1405.6378 | 0.035269 | 42.7842 | 131.14 |  |  |  |  | 1886.6319 | 0.05082
|  |  | 42.8589 | 42.795 |  |  |  |  | 1889.9016 | 0.012636 | 42.8589 | 45.85 |  |  |  |  | 1408.0739 | 0.036848 | 42.8589 | 131.15 |  |  |  |  | 1889.9016 | 0.011909
|  |  | 42.9335 | 42.725 |  |  |  |  | 1893.1713 | 0.0062894 | 42.9335 | 45.615 |  |  |  |  | 1410.5101 | 0.034573 | 42.9335 | 131.05 |  |  |  |  | 1893.1713 | 0.032867
|  |  | 43.0082 | 42.945 |  |  |  |  | 1896.441 | 0.031172 | 43.0082 | 45.67 |  |  |  |  | 1412.9462 | 0.029966 | 43.0082 | 130.805 |  |  |  |  | 1896.441 | 0.052927
|  |  | 43.0829 | 43.01 |  |  |  |  | 1899.7108 | 0.068981 | 43.0829 | 45.975 |  |  |  |  | 1415.3823 | 0.024883 | 43.0829 | 130.69 |  |  |  |  | 1899.7108 | 0.045677
|  |  | 43.1575 | 43.04 |  |  |  |  | 1902.9805 | 0.086656 | 43.1575 | 46.22 |  |  |  |  | 1417.8184 | 0.020259 | 43.1575 | 131.085 |  |  |  |  | 1902.9805 | 0.035223
|  |  | 43.2322 | 43.635 |  |  |  |  | 1906.2502 | 0.067213 | 43.2322 | 46.445 |  |  |  |  | 1420.2545 | 0.015254 | 43.2322 | 131.13 |  |  |  |  | 1906.2502 | 0.038732
|  |  | 43.3069 | 42.925 |  |  |  |  | 1909.5199 | 0.031609 | 43.3069 | 46.635 |  |  |  |  | 1422.6906 | 0.0081306 | 43.3069 | 130.76 |  |  |  |  | 1909.5199 | 0.026641
|  |  | 43.3815 | 43.395 |  |  |  |  | 1912.7897 | 0.045176 | 43.3815 | 46.615 |  |  |  |  | 1425.1267 | 0.0025832 | 43.3815 | 130.5 |  |  |  |  | 1912.7897 | 0.011796
|  |  | 43.4562 | 42.13 |  |  |  |  | 1916.0594 | 0.055797 | 43.4562 | 46.5 |  |  |  |  | 1427.5629 | 0.014199 | 43.4562 | 130.455 |  |  |  |  | 1916.0594 | 0.046136
|  |  | 43.5309 | 42.515 |  |  |  |  | 1919.3291 | 0.051153 | 43.5309 | 46.48 |  |  |  |  | 1429.999 | 0.0251 | 43.5309 | 130.205 |  |  |  |  | 1919.3291 | 0.056855
|  |  | 43.6055 | 42.805 |  |  |  |  | 1922.5988 | 0.072235 | 43.6055 | 46.44 |  |  |  |  | 1432.4351 | 0.031636 | 43.6055 | 130.135 |  |  |  |  | 1922.5988 | 0.031015
|  |  | 43.6802 | 42.45 |  |  |  |  | 1925.8686 | 0.090834 | 43.6802 | 46.21 |  |  |  |  | 1434.8712 | 0.03151 | 43.6802 | 130.6 |  |  |  |  | 1925.8686 | 0.022219
|  |  | 43.7549 | 42.245 |  |  |  |  | 1929.1383 | 0.074744 | 43.7549 | 45.985 |  |  |  |  | 1437.3073 | 0.024692 | 43.7549 | 131.345 |  |  |  |  | 1929.1383 | 0.068968
|  |  | 43.8295 | 43.26 |  |  |  |  | 1932.408 | 0.030647 | 43.8295 | 45.785 |  |  |  |  | 1439.7434 | 0.013577 | 43.8295 | 131.43 |  |  |  |  | 1932.408 | 0.089105
|  |  | 43.9042 | 43.57 |  |  |  |  | 1935.6777 | 0.0196 | 43.9042 | 45.495 |  |  |  |  | 1442.1795 | 0.0022467 | 43.9042 | 131.675 |  |  |  |  | 1935.6777 | 0.078851
|  |  | 43.9789 | 41.335 |  |  |  |  | 1938.9475 | 0.044073 | 43.9789 | 45.07 |  |  |  |  | 1444.6157 | 0.0078309 | 43.9789 | 131.795 |  |  |  |  | 1938.9475 | 0.05734
|  |  | 44.0535 | 42.575 |  |  |  |  | 1942.2172 | 0.046529 | 44.0535 | 44.73 |  |  |  |  | 1447.0518 | 0.014369 | 44.0535 | 132.29 |  |  |  |  | 1942.2172 | 0.041513
|  |  | 44.1282 | 42.82 |  |  |  |  | 1945.4869 | 0.039453 | 44.1282 | 44.445 |  |  |  |  | 1449.4879 | 0.020058 | 44.1282 | 133.065 |  |  |  |  | 1945.4869 | 0.029371
|  |  | 44.2029 | 42.83 |  |  |  |  | 1948.7566 | 0.032597 | 44.2029 | 44.135 |  |  |  |  | 1451.924 | 0.02567 | 44.2029 | 133.585 |  |  |  |  | 1948.7566 | 0.051935
|  |  | 44.2775 | 43.045 |  |  |  |  | 1952.0264 | 0.029657 | 44.2775 | 43.515 |  |  |  |  | 1454.3601 | 0.029293 | 44.2775 | 133.16 |  |  |  |  | 1952.0264 | 0.091161
|  |  | 44.3522 | 43.165 |  |  |  |  | 1955.2961 | 0.030657 | 44.3522 | 43.03 |  |  |  |  | 1456.7962 | 0.029056 | 44.3522 | 133.22 |  |  |  |  | 1955.2961 | 0.11016
|  |  | 44.4269 | 43.195 |  |  |  |  | 1958.5658 | 0.026706 | 44.4269 | 42.855 |  |  |  |  | 1459.2323 | 0.02555 | 44.4269 | 133.195 |  |  |  |  | 1958.5658 | 0.09687
|  |  | 44.5015 | 42.04 |  |  |  |  | 1961.8355 | 0.025105 | 44.5015 | 43.44 |  |  |  |  | 1461.6685 | 0.021411 | 44.5015 | 133.16 |  |  |  |  | 1961.8355 | 0.060809
|  |  | 44.5762 | 41.325 |  |  |  |  | 1965.1053 | 0.036719 | 44.5762 | 44.035 |  |  |  |  | 1464.1046 | 0.017745 | 44.5762 | 133.965 |  |  |  |  | 1965.1053 | 0.029021
|  |  | 44.6509 | 42.58 |  |  |  |  | 1968.375 | 0.043319 | 44.6509 | 44.28 |  |  |  |  | 1466.5407 | 0.012192 | 44.6509 | 134.405 |  |  |  |  | 1968.375 | 0.033553
|  |  | 44.7255 | 43.375 |  |  |  |  | 1971.6447 | 0.038382 | 44.7255 | 44.355 |  |  |  |  | 1468.9768 | 0.0033882 | 44.7255 | 134.455 |  |  |  |  | 1971.6447 | 0.044452
|  |  | 44.8002 | 44.105 |  |  |  |  | 1974.9145 | 0.027829 | 44.8002 | 44.49 |  |  |  |  | 1471.4129 | 0.0091583 | 44.8002 | 134.485 |  |  |  |  | 1974.9145 | 0.050406
|  |  | 44.8749 | 42.715 |  |  |  |  | 1978.1842 | 0.015605 | 44.8749 | 44.455 |  |  |  |  | 1473.849 | 0.018572 | 44.8749 | 134.35 |  |  |  |  | 1978.1842 | 0.058051
|  |  | 44.9495 | 41.345 |  |  |  |  | 1981.4539 | 0.01946 | 44.9495 | 44.395 |  |  |  |  | 1476.2851 | 0.022382 | 44.9495 | 134.175 |  |  |  |  | 1981.4539 | 0.064865
|  |  | 45.0242 | 42.405 |  |  |  |  | 1984.7236 | 0.038061 | 45.0242 | 44.36 |  |  |  |  | 1478.7212 | 0.020205 | 45.0242 | 134.42 |  |  |  |  | 1984.7236 | 0.06014
|  |  | 45.0989 | 41.885 |  |  |  |  | 1987.9934 | 0.045262 | 45.0989 | 44.225 |  |  |  |  | 1481.1574 | 0.017918 | 45.0989 | 135.095 |  |  |  |  | 1987.9934 | 0.037804
|  |  | 45.1735 | 40.72 |  |  |  |  | 1991.2631 | 0.03287 | 45.1735 | 43.79 |  |  |  |  | 1483.5935 | 0.021733 | 45.1735 | 135.92 |  |  |  |  | 1991.2631 | 0.0055499
|  |  | 45.2482 | 41.295 |  |  |  |  | 1994.5328 | 0.0058713 | 45.2482 | 43.47 |  |  |  |  | 1486.0296 | 0.025147 | 45.2482 | 136.255 |  |  |  |  | 1994.5328 | 0.023583
|  |  | 45.3229 | 42.02 |  |  |  |  | 1997.8025 | 0.036339 | 45.3229 | 43.155 |  |  |  |  | 1488.4657 | 0.021216 | 45.3229 | 136.53 |  |  |  |  | 1997.8025 | 0.02912
|  |  | 45.3975 | 41.525 |  |  |  |  | 2001.0723 | 0.073568 | 45.3975 | 42.645 |  |  |  |  | 1490.9018 | 0.0098211 | 45.3975 | 137.1 |  |  |  |  | 2001.0723 | 0.011204
|  |  | 45.4722 | 40.99 |  |  |  |  | 2004.342 | 0.095735 | 45.4722 | 42.235 |  |  |  |  | 1493.3379 | 0.0099083 | 45.4722 | 137.41 |  |  |  |  | 2004.342 | 0.015988
|  |  | 45.5469 | 41.935 |  |  |  |  | 2007.6117 | 0.09189 | 45.5469 | 41.875 |  |  |  |  | 1495.774 | 0.02265 | 45.5469 | 137.69 |  |  |  |  | 2007.6117 | 0.03115
|  |  | 45.6215 | 41.73 |  |  |  |  | 2010.8814 | 0.065183 | 45.6215 | 41.525 |  |  |  |  | 1498.2102 | 0.02779 | 45.6215 | 138.3 |  |  |  |  | 2010.8814 | 0.031656
|  |  | 45.6962 | 41.83 |  |  |  |  | 2014.1512 | 0.03647 | 45.6962 | 41.08 |  |  |  |  | 1500.6463 | 0.022876 | 45.6962 | 138.345 |  |  |  |  | 2014.1512 | 0.045803
|  |  | 45.7709 | 42.24 |  |  |  |  | 2017.4209 | 0.025179 | 45.7709 | 40.7 |  |  |  |  | 1503.0824 | 0.011667 | 45.7709 | 138.68 |  |  |  |  | 2017.4209 | 0.067154
|  |  | 45.8455 | 42.535 |  |  |  |  | 2020.6906 | 0.018432 | 45.8455 | 41.205 |  |  |  |  | 1505.5185 | 0.0065737 | 45.8455 | 139.345 |  |  |  |  | 2020.6906 | 0.067105
|  |  | 45.9202 | 42.28 |  |  |  |  | 2023.9603 | 0.025375 | 45.9202 | 41.86 |  |  |  |  | 1507.9546 | 0.0083683 | 45.9202 | 139.12 |  |  |  |  | 2023.9603 | 0.044601
|  |  | 45.9949 | 41.305 |  |  |  |  | 2027.2301 | 0.035655 | 45.9949 | 42.395 |  |  |  |  | 1510.3907 | 0.00070877 | 45.9949 | 138.89 |  |  |  |  | 2027.2301 | 0.034935
|  |  | 46.0695 | 41.56 |  |  |  |  | 2030.4998 | 0.037967 | 46.0695 | 42.835 |  |  |  |  | 1512.8268 | 0.020371 | 46.0695 | 138.695 |  |  |  |  | 2030.4998 | 0.044968
|  |  | 46.1442 | 41.265 |  |  |  |  | 2033.7695 | 0.060319 | 46.1442 | 42.99 |  |  |  |  | 1515.263 | 0.045206 | 46.1442 | 138.725 |  |  |  |  | 2033.7695 | 0.035134
|  |  | 46.2189 | 41.405 |  |  |  |  | 2037.0392 | 0.092569 | 46.2189 | 43.17 |  |  |  |  | 1517.6991 | 0.065164 | 46.2189 | 138.65 |  |  |  |  | 2037.0392 | 0.007258
|  |  | 46.2935 | 43.17 |  |  |  |  | 2040.309 | 0.1047 | 46.2935 | 43.365 |  |  |  |  | 1520.1352 | 0.071629 | 46.2935 | 138.525 |  |  |  |  | 2040.309 | 0.018183
|  |  | 46.3682 | 43.305 |  |  |  |  | 2043.5787 | 0.0885 | 46.3682 | 43.385 |  |  |  |  | 1522.5713 | 0.061191 | 46.3682 | 138.65 |  |  |  |  | 2043.5787 | 0.023066
|  |  | 46.4429 | 42.73 |  |  |  |  | 2046.8484 | 0.054348 | 46.4429 | 43.36 |  |  |  |  | 1525.0074 | 0.038296 | 46.4429 | 138.795 |  |  |  |  | 2046.8484 | 0.027435
|  |  | 46.5175 | 42.925 |  |  |  |  | 2050.1181 | 0.015496 | 46.5175 | 43.29 |  |  |  |  | 1527.4435 | 0.02051 | 46.5175 | 138.675 |  |  |  |  | 2050.1181 | 0.048528
|  |  | 46.5922 | 40.65 |  |  |  |  | 2053.3879 | 0.022775 | 46.5922 | 43.065 |  |  |  |  | 1529.8796 | 0.030079 | 46.5922 | 138.705 |  |  |  |  | 2053.3879 | 0.05459
|  |  | 46.6669 | 41.475 |  |  |  |  | 2056.6576 | 0.053108 | 46.6669 | 43.01 |  |  |  |  | 1532.3158 | 0.036778 | 46.6669 | 138.745 |  |  |  |  | 2056.6576 | 0.033865
|  |  | 46.7415 | 42.135 |  |  |  |  | 2059.9273 | 0.066091 | 46.7415 | 42.985 |  |  |  |  | 1534.7519 | 0.02757 | 46.7415 | 138.19 |  |  |  |  | 2059.9273 | 0.0063525
|  |  | 46.8162 | 41.185 |  |  |  |  | 2063.1971 | 0.061722 | 46.8162 | 42.94 |  |  |  |  | 1537.188 | 0.005443 | 46.8162 | 137.965 |  |  |  |  | 2063.1971 | 0.026468
|  |  | 46.8909 | 40.525 |  |  |  |  | 2066.4668 | 0.057774 | 46.8909 | 42.69 |  |  |  |  | 1539.6241 | 0.020521 | 46.8909 | 138.19 |  |  |  |  | 2066.4668 | 0.025998
|  |  | 46.9655 | 41.12 |  |  |  |  | 2069.7365 | 0.058666 | 46.9655 | 42.445 |  |  |  |  | 1542.0602 | 0.039534 | 46.9655 | 138.535 |  |  |  |  | 2069.7365 | 0.016697
|  |  | 47.0402 | 43.47 |  |  |  |  | 2073.0062 | 0.048281 | 47.0402 | 42.37 |  |  |  |  | 1544.4963 | 0.044208 | 47.0402 | 138.87 |  |  |  |  | 2073.0062 | 0.027329
|  |  | 47.1149 | 43.8 |  |  |  |  | 2076.276 | 0.051386 | 47.1149 | 42.715 |  |  |  |  | 1546.9324 | 0.033252 | 47.1149 | 138.795 |  |  |  |  | 2076.276 | 0.033408
|  |  | 47.1895 | 44.3 |  |  |  |  | 2079.5457 | 0.090808 | 47.1895 | 43.1 |  |  |  |  | 1549.3686 | 0.01353 | 47.1895 | 138.395 |  |  |  |  | 2079.5457 | 0.054464
|  |  | 47.2642 | 42.72 |  |  |  |  | 2082.8154 | 0.12107 | 47.2642 | 43.57 |  |  |  |  | 1551.8047 | 0.019377 | 47.2642 | 137.58 |  |  |  |  | 2082.8154 | 0.090395
|  |  | 47.3389 | 42.615 |  |  |  |  | 2086.0851 | 0.1182 | 47.3389 | 43.955 |  |  |  |  | 1554.2408 | 0.037303 | 47.3389 | 137.595 |  |  |  |  | 2086.0851 | 0.10624
|  |  | 47.4135 | 41.705 |  |  |  |  | 2089.3549 | 0.085465 | 47.4135 | 44.145 |  |  |  |  | 1556.6769 | 0.04471 | 47.4135 | 137.745 |  |  |  |  | 2089.3549 | 0.085927
|  |  | 47.4882 | 42.395 |  |  |  |  | 2092.6246 | 0.045955 | 47.4882 | 44.465 |  |  |  |  | 1559.113 | 0.039734 | 47.4882 | 138.42 |  |  |  |  | 2092.6246 | 0.051316
|  |  | 47.5629 | 42.76 |  |  |  |  | 2095.8943 | 0.027886 | 47.5629 | 44.71 |  |  |  |  | 1561.5491 | 0.025541 | 47.5629 | 138.245 |  |  |  |  | 2095.8943 | 0.062201
|  |  | 47.6375 | 42.005 |  |  |  |  | 2099.164 | 0.025029 | 47.6375 | 44.615 |  |  |  |  | 1563.9852 | 0.010398 | 47.6375 | 137.405 |  |  |  |  | 2099.164 | 0.080935
|  |  | 47.7122 | 42.465 |  |  |  |  | 2102.4338 | 0.025382 | 47.7122 | 44.535 |  |  |  |  | 1566.4214 | 0.019333 | 47.7122 | 136.87 |  |  |  |  | 2102.4338 | 0.066298
|  |  | 47.7869 | 41.885 |  |  |  |  | 2105.7035 | 0.031563 | 47.7869 | 44.655 |  |  |  |  | 1568.8575 | 0.035506 | 47.7869 | 136.685 |  |  |  |  | 2105.7035 | 0.023437
|  |  | 47.8615 | 42.27 |  |  |  |  | 2108.9732 | 0.025303 | 47.8615 | 44.77 |  |  |  |  | 1571.2936 | 0.046769 | 47.8615 | 136.565 |  |  |  |  | 2108.9732 | 0.026685
|  |  | 47.9362 | 41.4 |  |  |  |  | 2112.2429 | 0.0048162 | 47.9362 | 44.73 |  |  |  |  | 1573.7297 | 0.049978 | 47.9362 | 136.185 |  |  |  |  | 2112.2429 | 0.061035
|  |  | 48.0109 | 41.695 |  |  |  |  | 2115.5127 | 0.030637 | 48.0109 | 44.57 |  |  |  |  | 1576.1658 | 0.043775 | 48.0109 | 135.72 |  |  |  |  | 2115.5127 | 0.068491
|  |  | 48.0855 | 41.415 |  |  |  |  | 2118.7824 | 0.047413 | 48.0855 | 44.43 |  |  |  |  | 1578.6019 | 0.029316 | 48.0855 | 135.7 |  |  |  |  | 2118.7824 | 0.050665
|  |  | 48.1602 | 40.97 |  |  |  |  | 2122.0521 | 0.040455 | 48.1602 | 44.31 |  |  |  |  | 1581.038 | 0.010485 | 48.1602 | 136.6 |  |  |  |  | 2122.0521 | 0.023691
|  |  | 48.2349 | 41.87 |  |  |  |  | 2125.3218 | 0.02337 | 48.2349 | 44.09 |  |  |  |  | 1583.4742 | 0.0073434 | 48.2349 | 136.855 |  |  |  |  | 2125.3218 | 0.019385
|  |  | 48.3095 | 41.585 |  |  |  |  | 2128.5916 | 0.040907 | 48.3095 | 43.895 |  |  |  |  | 1585.9103 | 0.019509 | 48.3095 | 136.24 |  |  |  |  | 2128.5916 | 0.023984
|  |  | 48.3842 | 41.56 |  |  |  |  | 2131.8613 | 0.064996 | 48.3842 | 43.71 |  |  |  |  | 1588.3464 | 0.024189 | 48.3842 | 136.32 |  |  |  |  | 2131.8613 | 0.034393
|  |  | 48.4589 | 42.495 |  |  |  |  | 2135.131 | 0.077227 | 48.4589 | 43.835 |  |  |  |  | 1590.7825 | 0.023316 | 48.4589 | 136.11 |  |  |  |  | 2135.131 | 0.058117
|  |  | 48.5336 | 42.145 |  |  |  |  | 2138.4007 | 0.080947 | 48.5336 | 44.24 |  |  |  |  | 1593.2186 | 0.02114 | 48.5336 | 136.005 |  |  |  |  | 2138.4007 | 0.067102
|  |  | 48.6082 | 42.425 |  |  |  |  | 2141.6705 | 0.078923 | 48.6082 | 44.705 |  |  |  |  | 1595.6547 | 0.019627 | 48.6082 | 136.41 |  |  |  |  | 2141.6705 | 0.04518
|  |  | 48.6829 | 43.59 |  |  |  |  | 2144.9402 | 0.067517 | 48.6829 | 45.42 |  |  |  |  | 1598.0908 | 0.016596 | 48.6829 | 136.68 |  |  |  |  | 2144.9402 | 0.00086701
|  |  | 48.7575 | 41.595 |  |  |  |  | 2148.2099 | 0.045227 | 48.7575 | 45.785 |  |  |  |  | 1600.527 | 0.010955 | 48.7575 | 136.56 |  |  |  |  | 2148.2099 | 0.042222
|  |  | 48.8322 | 42.245 |  |  |  |  | 2151.4797 | 0.021347 | 48.8322 | 46.045 |  |  |  |  | 1602.9631 | 0.0077176 | 48.8322 | 135.845 |  |  |  |  | 2151.4797 | 0.057142
|  |  | 48.9069 | 42.515 |  |  |  |  | 2154.7494 | 0.023548 | 48.9069 | 46.5 |  |  |  |  | 1605.3992 | 0.011295 | 48.9069 | 134.6 |  |  |  |  | 2154.7494 | 0.038693
|  |  | 48.9816 | 43.845 |  |  |  |  | 2158.0191 | 0.034908 | 48.9816 | 46.67 |  |  |  |  | 1607.8353 | 0.012943 | 48.9816 | 134.46 |  |  |  |  | 2158.0191 | 0.013495
|  |  | 49.0562 | 43.45 |  |  |  |  | 2161.2888 | 0.038602 | 49.0562 | 46.845 |  |  |  |  | 1610.2714 | 0.0095272 | 49.0562 | 134.68 |  |  |  |  | 2161.2888 | 0.032722
|  |  | 49.1309 | 43.315 |  |  |  |  | 2164.5586 | 0.043237 | 49.1309 | 47.135 |  |  |  |  | 1612.7075 | 0.0083351 | 49.1309 | 135.56 |  |  |  |  | 2164.5586 | 0.036152
|  |  | 49.2056 | 43.21 |  |  |  |  | 2167.8283 | 0.051748 | 49.2056 | 47.175 |  |  |  |  | 1615.1436 | 0.018585 | 49.2056 | 135.23 |  |  |  |  | 2167.8283 | 0.019408
|  |  | 49.2802 | 43.985 |  |  |  |  | 2171.098 | 0.055225 | 49.2802 | 47.2 |  |  |  |  | 1617.5798 | 0.030144 | 49.2802 | 134.565 |  |  |  |  | 2171.098 | 0.0025061
|  |  | 49.3549 | 43.185 |  |  |  |  | 2174.3677 | 0.052972 | 49.3549 | 47.145 |  |  |  |  | 1620.0159 | 0.038195 | 49.3549 | 133.73 |  |  |  |  | 2174.3677 | 0.019666
|  |  | 49.4296 | 43.235 |  |  |  |  | 2177.6375 | 0.057804 | 49.4296 | 46.965 |  |  |  |  | 1622.452 | 0.04079 | 49.4296 | 133.595 |  |  |  |  | 2177.6375 | 0.04054
|  |  | 49.5042 | 44.03 |  |  |  |  | 2180.9072 | 0.07209 | 49.5042 | 46.77 |  |  |  |  | 1624.8881 | 0.037706 | 49.5042 | 133.835 |  |  |  |  | 2180.9072 | 0.056331
|  |  | 49.5789 | 44.095 |  |  |  |  | 2184.1769 | 0.080736 | 49.5789 | 46.485 |  |  |  |  | 1627.3242 | 0.030326 | 49.5789 | 134.43 |  |  |  |  | 2184.1769 | 0.049476
|  |  | 49.6536 | 42.3 |  |  |  |  | 2187.4466 | 0.074043 | 49.6536 | 46.285 |  |  |  |  | 1629.7603 | 0.021875 | 49.6536 | 134.295 |  |  |  |  | 2187.4466 | 0.017506
|  |  | 49.7282 | 42.515 |  |  |  |  | 2190.7164 | 0.05483 | 49.7282 | 46.06 |  |  |  |  | 1632.1964 | 0.017221 | 49.7282 | 133.94 |  |  |  |  | 2190.7164 | 0.027991
|  |  | 49.8029 | 41.975 |  |  |  |  | 2193.9861 | 0.04022 | 49.8029 | 45.85 |  |  |  |  | 1634.6325 | 0.017423 | 49.8029 | 133.725 |  |  |  |  | 2193.9861 | 0.053305
|  |  | 49.8776 | 42.15 |  |  |  |  | 2197.2558 | 0.047531 | 49.8776 | 45.62 |  |  |  |  | 1637.0687 | 0.017066 | 49.8776 | 133.715 |  |  |  |  | 2197.2558 | 0.050549
|  |  | 49.9522 | 42.49 |  |  |  |  | 2200.5255 | 0.056403 | 49.9522 | 45.63 |  |  |  |  | 1639.5048 | 0.013332 | 49.9522 | 134.045 |  |  |  |  | 2200.5255 | 0.036644
|  |  | 50.0269 | 43.38 |  |  |  |  | 2203.7953 | 0.047543 | 50.0269 | 46.455 |  |  |  |  | 1641.9409 | 0.0077099 | 50.0269 | 134.08 |  |  |  |  | 2203.7953 | 0.039519
|  |  | 50.1016 | 44.05 |  |  |  |  | 2207.065 | 0.018359 | 50.1016 | 47.065 |  |  |  |  | 1644.377 | 0.0035698 | 50.1016 | 133.96 |  |  |  |  | 2207.065 | 0.041089
|  |  | 50.1762 | 43.86 |  |  |  |  | 2210.3347 | 0.021729 | 50.1762 | 47.53 |  |  |  |  | 1646.8131 | 0.0036888 | 50.1762 | 133.975 |  |  |  |  | 2210.3347 | 0.052539
|  |  | 50.2509 | 42.835 |  |  |  |  | 2213.6044 | 0.052083 | 50.2509 | 47.8 |  |  |  |  | 1649.2492 | 0.01139 | 50.2509 | 133.935 |  |  |  |  | 2213.6044 | 0.083809
|  |  | 50.3256 | 42.21 |  |  |  |  | 2216.8742 | 0.058242 | 50.3256 | 47.87 |  |  |  |  | 1651.6853 | 0.024132 | 50.3256 | 134 |  |  |  |  | 2216.8742 | 0.096736
|  |  | 50.4002 | 43.09 |  |  |  |  | 2220.1439 | 0.0394 | 50.4002 | 47.79 |  |  |  |  | 1654.1215 | 0.037257 | 50.4002 | 133.66 |  |  |  |  | 2220.1439 | 0.071768
|  |  | 50.4749 | 43.315 |  |  |  |  | 2223.4136 | 0.016232 | 50.4749 | 47.745 |  |  |  |  | 1656.5576 | 0.045405 | 50.4749 | 133.64 |  |  |  |  | 2223.4136 | 0.02779
|  |  | 50.5496 | 42.38 |  |  |  |  | 2226.6833 | 0.022497 | 50.5496 | 47.665 |  |  |  |  | 1658.9937 | 0.044718 | 50.5496 | 133.765 |  |  |  |  | 2226.6833 | 0.044647
|  |  | 50.6242 | 43.26 |  |  |  |  | 2229.9531 | 0.027386 | 50.6242 | 47.545 |  |  |  |  | 1661.4298 | 0.034748 | 50.6242 | 133.82 |  |  |  |  | 2229.9531 | 0.062751
|  |  | 50.6989 | 43.155 |  |  |  |  | 2233.2228 | 0.03366 | 50.6989 | 47.385 |  |  |  |  | 1663.8659 | 0.019077 | 50.6989 | 133.92 |  |  |  |  | 2233.2228 | 0.053073
|  |  | 50.7736 | 44.325 |  |  |  |  | 2236.4925 | 0.042499 | 50.7736 | 47.16 |  |  |  |  | 1666.302 | 0.0066429 | 50.7736 | 133.85 |  |  |  |  | 2236.4925 | 0.043322
|  |  | 50.8482 | 42.5 |  |  |  |  | 2239.7623 | 0.041258 | 50.8482 | 46.92 |  |  |  |  | 1668.7381 | 0.011294 | 50.8482 | 133.76 |  |  |  |  | 2239.7623 | 0.038894
|  |  | 50.9229 | 42.64 |  |  |  |  | 2243.032 | 0.04062 | 50.9229 | 46.595 |  |  |  |  | 1671.1743 | 0.012824 | 50.9229 | 134.18 |  |  |  |  | 2243.032 | 0.011885
|  |  | 50.9976 | 43.32 |  |  |  |  | 2246.3017 | 0.047215 | 50.9976 | 46.055 |  |  |  |  | 1673.6104 | 0.0086644 | 50.9976 | 134.35 |  |  |  |  | 2246.3017 | 0.042676
|  |  | 51.0722 | 43.285 |  |  |  |  | 2249.5714 | 0.039163 | 51.0722 | 45.66 |  |  |  |  | 1676.0465 | 0.011875 | 51.0722 | 134.72 |  |  |  |  | 2249.5714 | 0.087604
|  |  | 51.1469 | 42.32 |  |  |  |  | 2252.8412 | 0.023327 | 51.1469 | 45.095 |  |  |  |  | 1678.4826 | 0.020307 | 51.1469 | 134.785 |  |  |  |  | 2252.8412 | 0.09702
|  |  | 51.2216 | 43.175 |  |  |  |  | 2256.1109 | 0.044522 | 51.2216 | 44.125 |  |  |  |  | 1680.9187 | 0.023319 | 51.2216 | 134.46 |  |  |  |  | 2256.1109 | 0.062738
|  |  | 51.2962 | 42.895 |  |  |  |  | 2259.3806 | 0.054744 | 51.2962 | 43.805 |  |  |  |  | 1683.3548 | 0.018083 | 51.2962 | 134.135 |  |  |  |  | 2259.3806 | 0.013611
|  |  | 51.3709 | 43.25 |  |  |  |  | 2262.6503 | 0.031906 | 51.3709 | 43.47 |  |  |  |  | 1685.7909 | 0.0075359 | 51.3709 | 133.72 |  |  |  |  | 2262.6503 | 0.051459
|  |  | 51.4456 | 43.55 |  |  |  |  | 2265.9201 | 0.024348 | 51.4456 | 43.565 |  |  |  |  | 1688.2271 | 0.011796 | 51.4456 | 133.685 |  |  |  |  | 2265.9201 | 0.06667
|  |  | 51.5202 | 43.07 |  |  |  |  | 2269.1898 | 0.056585 | 51.5202 | 44.135 |  |  |  |  | 1690.6632 | 0.023023 | 51.5202 | 133.85 |  |  |  |  | 2269.1898 | 0.045782
|  |  | 51.5949 | 42.675 |  |  |  |  | 2272.4595 | 0.057661 | 51.5949 | 44.59 |  |  |  |  | 1693.0993 | 0.028486 | 51.5949 | 133.89 |  |  |  |  | 2272.4595 | 0.028282
|  |  | 51.6696 | 43.435 |  |  |  |  | 2275.7292 | 0.028862 | 51.6696 | 44.85 |  |  |  |  | 1695.5354 | 0.026975 | 51.6696 | 134.09 |  |  |  |  | 2275.7292 | 0.049978
|  |  | 51.7442 | 43.365 |  |  |  |  | 2278.999 | 0.026887 | 51.7442 | 44.935 |  |  |  |  | 1697.9715 | 0.020358 | 51.7442 | 134.475 |  |  |  |  | 2278.999 | 0.054443
|  |  | 51.8189 | 43.295 |  |  |  |  | 2282.2687 | 0.041958 | 51.8189 | 45.01 |  |  |  |  | 1700.4076 | 0.011542 | 51.8189 | 134.88 |  |  |  |  | 2282.2687 | 0.033434
|  |  | 51.8936 | 43.755 |  |  |  |  | 2285.5384 | 0.030883 | 51.8936 | 45.145 |  |  |  |  | 1702.8437 | 0.0026881 | 51.8936 | 134.89 |  |  |  |  | 2285.5384 | 0.0086874
|  |  | 51.9682 | 41.5 |  |  |  |  | 2288.8081 | 0.035156 | 51.9682 | 45.125 |  |  |  |  | 1705.2799 | 0.0066483 | 51.9682 | 134.68 |  |  |  |  | 2288.8081 | 0.0207
|  |  | 52.0429 | 41.545 |  |  |  |  | 2292.0779 | 0.062294 | 52.0429 | 44.935 |  |  |  |  | 1707.716 | 0.014667 | 52.0429 | 134.44 |  |  |  |  | 2292.0779 | 0.023491
|  |  | 52.1176 | 42.305 |  |  |  |  | 2295.3476 | 0.069357 | 52.1176 | 44.71 |  |  |  |  | 1710.1521 | 0.020519 | 52.1176 | 133.745 |  |  |  |  | 2295.3476 | 0.026714
|  |  | 52.1922 | 42.435 |  |  |  |  | 2298.6173 | 0.053557 | 52.1922 | 44.545 |  |  |  |  | 1712.5882 | 0.022803 | 52.1922 | 133.91 |  |  |  |  | 2298.6173 | 0.037347
|  |  | 52.2669 | 42.14 |  |  |  |  | 2301.887 | 0.032379 | 52.2669 | 44.435 |  |  |  |  | 1715.0243 | 0.021884 | 52.2669 | 134.86 |  |  |  |  | 2301.887 | 0.036231
|  |  | 52.3416 | 41.275 |  |  |  |  | 2305.1568 | 0.025825 | 52.3416 | 44.275 |  |  |  |  | 1717.4604 | 0.020151 | 52.3416 | 135.235 |  |  |  |  | 2305.1568 | 0.021201
|  |  | 52.4162 | 40.725 |  |  |  |  | 2308.4265 | 0.036009 | 52.4162 | 43.92 |  |  |  |  | 1719.8965 | 0.019235 | 52.4162 | 134.71 |  |  |  |  | 2308.4265 | 0.025883
|  |  | 52.4909 | 41.915 |  |  |  |  | 2311.6962 | 0.04968 | 52.4909 | 43.485 |  |  |  |  | 1722.3327 | 0.0173 | 52.4909 | 134.61 |  |  |  |  | 2311.6962 | 0.042477
|  |  | 52.5656 | 43.28 |  |  |  |  | 2314.9659 | 0.053575 | 52.5656 | 42.995 |  |  |  |  | 1724.7688 | 0.011962 | 52.5656 | 135 |  |  |  |  | 2314.9659 | 0.042085
|  |  | 52.6402 | 43.76 |  |  |  |  | 2318.2357 | 0.04264 | 52.6402 | 42.405 |  |  |  |  | 1727.2049 | 0.0034633 | 52.6402 | 134.83 |  |  |  |  | 2318.2357 | 0.024606
|  |  | 52.7149 | 40.92 |  |  |  |  | 2321.5054 | 0.027276 | 52.7149 | 42.295 |  |  |  |  | 1729.641 | 0.0056162 | 52.7149 | 135.06 |  |  |  |  | 2321.5054 | 0.004445
|  |  | 52.7896 | 41.76 |  |  |  |  | 2324.7751 | 0.029563 | 52.7896 | 42.565 |  |  |  |  | 1732.0771 | 0.011795 | 52.7896 | 135.91 |  |  |  |  | 2324.7751 | 0.019381
|  |  | 52.8642 | 41.12 |  |  |  |  | 2328.0449 | 0.039697 | 52.8642 | 43.16 |  |  |  |  | 1734.5132 | 0.01316 | 52.8642 | 136.04 |  |  |  |  | 2328.0449 | 0.036317
|  |  | 52.9389 | 42.84 |  |  |  |  | 2331.3146 | 0.042145 | 52.9389 | 43.4 |  |  |  |  | 1736.9493 | 0.010739 | 52.9389 | 135.84 |  |  |  |  | 2331.3146 | 0.050117
|  |  | 53.0136 | 41.92 |  |  |  |  | 2334.5843 | 0.033982 | 53.0136 | 43.575 |  |  |  |  | 1739.3855 | 0.0094589 | 53.0136 | 135.675 |  |  |  |  | 2334.5843 | 0.049551
|  |  | 53.0882 | 43.09 |  |  |  |  | 2337.854 | 0.019128 | 53.0882 | 43.795 |  |  |  |  | 1741.8216 | 0.011841 | 53.0882 | 135.795 |  |  |  |  | 2337.854 | 0.031974
|  |  | 53.1629 | 43.405 |  |  |  |  | 2341.1238 | 0.033197 | 53.1629 | 43.815 |  |  |  |  | 1744.2577 | 0.013092 | 53.1629 | 135.755 |  |  |  |  | 2341.1238 | 0.04307
|  |  | 53.2376 | 42.53 |  |  |  |  | 2344.3935 | 0.064022 | 53.2376 | 43.86 |  |  |  |  | 1746.6938 | 0.01106 | 53.2376 | 136.175 |  |  |  |  | 2344.3935 | 0.080097
|  |  | 53.3122 | 41.83 |  |  |  |  | 2347.6632 | 0.08015 | 53.3122 | 43.815 |  |  |  |  | 1749.1299 | 0.0074218 | 53.3122 | 136.365 |  |  |  |  | 2347.6632 | 0.097472
|  |  | 53.3869 | 41.05 |  |  |  |  | 2350.9329 | 0.068222 | 53.3869 | 43.585 |  |  |  |  | 1751.566 | 0.0054053 | 53.3869 | 135.945 |  |  |  |  | 2350.9329 | 0.083923
|  |  | 53.4616 | 41.09 |  |  |  |  | 2354.2027 | 0.032713 | 53.4616 | 43.46 |  |  |  |  | 1754.0021 | 0.0057301 | 53.4616 | 135.895 |  |  |  |  | 2354.2027 | 0.050139
|  |  | 53.5362 | 41.1 |  |  |  |  | 2357.4724 | 0.017761 | 53.5362 | 43.21 |  |  |  |  | 1756.4383 | 0.0094065 | 53.5362 | 136.54 |  |  |  |  | 2357.4724 | 0.017775
|  |  | 53.6109 | 41.925 |  |  |  |  | 2360.7421 | 0.0381 | 53.6109 | 43.055 |  |  |  |  | 1758.8744 | 0.016335 | 53.6109 | 136.88 |  |  |  |  | 2360.7421 | 0.016336
|  |  | 53.6856 | 42.72 |  |  |  |  | 2364.0118 | 0.042506 | 53.6856 | 42.895 |  |  |  |  | 1761.3105 | 0.022422 | 53.6856 | 136.715 |  |  |  |  | 2364.0118 | 0.037642
|  |  | 53.7602 | 42.17 |  |  |  |  | 2367.2816 | 0.055009 | 53.7602 | 42.63 |  |  |  |  | 1763.7466 | 0.023685 | 53.7602 | 136.545 |  |  |  |  | 2367.2816 | 0.056259
|  |  | 53.8349 | 41.135 |  |  |  |  | 2370.5513 | 0.070229 | 53.8349 | 42.22 |  |  |  |  | 1766.1827 | 0.018362 | 53.8349 | 136.725 |  |  |  |  | 2370.5513 | 0.055679
|  |  | 53.9096 | 41.66 |  |  |  |  | 2373.821 | 0.06658 | 53.9096 | 41.905 |  |  |  |  | 1768.6188 | 0.0080349 | 53.9096 | 136.91 |  |  |  |  | 2373.821 | 0.030695
|  |  | 53.9842 | 41.885 |  |  |  |  | 2377.0907 | 0.065945 | 53.9842 | 41.45 |  |  |  |  | 1771.0549 | 0.0054386 | 53.9842 | 136.53 |  |  |  |  | 2377.0907 | 0.011547
|  |  | 54.0589 | 42.2 |  |  |  |  | 2380.3605 | 0.086547 | 54.0589 | 40.99 |  |  |  |  | 1773.4911 | 0.013215 | 54.0589 | 136.855 |  |  |  |  | 2380.3605 | 0.029973
|  |  | 54.1336 | 42.575 |  |  |  |  | 2383.6302 | 0.085284 | 54.1336 | 41.005 |  |  |  |  | 1775.9272 | 0.014583 | 54.1336 | 137.72 |  |  |  |  | 2383.6302 | 0.018417
|  |  | 54.2082 | 41.49 |  |  |  |  | 2386.8999 | 0.042301 | 54.2082 | 41.31 |  |  |  |  | 1778.3633 | 0.0091874 | 54.2082 | 138.72 |  |  |  |  | 2386.8999 | 0.02279
|  |  | 54.2829 | 40.995 |  |  |  |  | 2390.1696 | 0.030099 | 54.2829 | 42.055 |  |  |  |  | 1780.7994 | 0.0048407 | 54.2829 | 138.595 |  |  |  |  | 2390.1696 | 0.068098
|  |  | 54.3576 | 41.51 |  |  |  |  | 2393.4394 | 0.074237 | 54.3576 | 42.33 |  |  |  |  | 1783.2355 | 0.011648 | 54.3576 | 138.365 |  |  |  |  | 2393.4394 | 0.088312
|  |  | 54.4322 | 40.62 |  |  |  |  | 2396.7091 | 0.072189 | 54.4322 | 42.54 |  |  |  |  | 1785.6716 | 0.014665 | 54.4322 | 138.05 |  |  |  |  | 2396.7091 | 0.068542
|  |  | 54.5069 | 40.41 |  |  |  |  | 2399.9788 | 0.03659 | 54.5069 | 42.835 |  |  |  |  | 1788.1077 | 0.010776 | 54.5069 | 138.485 |  |  |  |  | 2399.9788 | 0.020237
|  |  | 54.5816 | 41.105 |  |  |  |  | 2403.2485 | 0.050955 | 54.5816 | 43.065 |  |  |  |  | 1790.5439 | 0.0062089 | 54.5816 | 138.865 |  |  |  |  | 2403.2485 | 0.036137
|  |  | 54.6562 | 41.26 |  |  |  |  | 2406.5183 | 0.069612 | 54.6562 | 43.255 |  |  |  |  | 1792.98 | 0.014226 | 54.6562 | 138.01 |  |  |  |  | 2406.5183 | 0.062885
|  |  | 54.7309 | 42.855 |  |  |  |  | 2409.788 | 0.045181 | 54.7309 | 43.435 |  |  |  |  | 1795.4161 | 0.02098 | 54.7309 | 137.035 |  |  |  |  | 2409.788 | 0.055069
|  |  | 54.8056 | 43.085 |  |  |  |  | 2413.0577 | 0.012377 | 54.8056 | 43.49 |  |  |  |  | 1797.8522 | 0.020519 | 54.8056 | 136.7 |  |  |  |  | 2413.0577 | 0.025271
|  |  | 54.8802 | 41.765 |  |  |  |  | 2416.3275 | 0.048864 | 54.8802 | 43.535 |  |  |  |  | 1800.2883 | 0.013244 | 54.8802 | 137.335 |  |  |  |  | 2416.3275 | 0.027639
|  |  | 54.9549 | 42.085 |  |  |  |  | 2419.5972 | 0.052838 | 54.9549 | 43.275 |  |  |  |  | 1802.7244 | 0.0074932 | 54.9549 | 138.055 |  |  |  |  | 2419.5972 | 0.053501
|  |  | 55.0296 | 42.065 |  |  |  |  | 2422.8669 | 0.027639 | 55.0296 | 43.125 |  |  |  |  | 1805.1605 | 0.013687 | 55.0296 | 138.495 |  |  |  |  | 2422.8669 | 0.066475
|  |  | 55.1042 | 41.72 |  |  |  |  | 2426.1366 | 0.029109 | 55.1042 | 43.045 |  |  |  |  | 1807.5966 | 0.01748 | 55.1042 | 138.59 |  |  |  |  | 2426.1366 | 0.064793
|  |  | 55.1789 | 41.595 |  |  |  |  | 2429.4064 | 0.043316 | 55.1789 | 43.17 |  |  |  |  | 1810.0328 | 0.015467 | 55.1789 | 138.245 |  |  |  |  | 2429.4064 | 0.049586
|  |  | 55.2536 | 41.795 |  |  |  |  | 2432.6761 | 0.034341 | 55.2536 | 43.05 |  |  |  |  | 1812.4689 | 0.014041 | 55.2536 | 137.875 |  |  |  |  | 2432.6761 | 0.028142
|  |  | 55.3282 | 41.485 |  |  |  |  | 2435.9458 | 0.020295 | 55.3282 | 42.675 |  |  |  |  | 1814.905 | 0.019255 | 55.3282 | 137.685 |  |  |  |  | 2435.9458 | 0.025172
|  |  | 55.4029 | 42.145 |  |  |  |  | 2439.2155 | 0.022932 | 55.4029 | 42.285 |  |  |  |  | 1817.3411 | 0.023275 | 55.4029 | 137.805 |  |  |  |  | 2439.2155 | 0.039615
|  |  | 55.4776 | 41.23 |  |  |  |  | 2442.4853 | 0.018396 | 55.4776 | 42.335 |  |  |  |  | 1819.7772 | 0.020094 | 55.4776 | 138.45 |  |  |  |  | 2442.4853 | 0.049823
|  |  | 55.5522 | 40.905 |  |  |  |  | 2445.755 | 0.010798 | 55.5522 | 43.115 |  |  |  |  | 1822.2133 | 0.0096628 | 55.5522 | 138.46 |  |  |  |  | 2445.755 | 0.05743
|  |  | 55.6269 | 42.475 |  |  |  |  | 2449.0247 | 0.025125 | 55.6269 | 43.98 |  |  |  |  | 1824.6494 | 0.0052728 | 55.6269 | 138.115 |  |  |  |  | 2449.0247 | 0.067155
|  |  | 55.7016 | 42.005 |  |  |  |  | 2452.2944 | 0.043359 | 55.7016 | 44.77 |  |  |  |  | 1827.0856 | 0.016556 | 55.7016 | 138.055 |  |  |  |  | 2452.2944 | 0.075544
|  |  | 55.7762 | 42.575 |  |  |  |  | 2455.5642 | 0.063005 | 55.7762 | 45.06 |  |  |  |  | 1829.5217 | 0.021436 | 55.7762 | 137.865 |  |  |  |  | 2455.5642 | 0.072312
|  |  | 55.8509 | 42.66 |  |  |  |  | 2458.8339 | 0.076845 | 55.8509 | 45.255 |  |  |  |  | 1831.9578 | 0.019027 | 55.8509 | 137.3 |  |  |  |  | 2458.8339 | 0.050153
|  |  | 55.9256 | 42.575 |  |  |  |  | 2462.1036 | 0.078226 | 55.9256 | 45.445 |  |  |  |  | 1834.3939 | 0.015091 | 55.9256 | 136.38 |  |  |  |  | 2462.1036 | 0.020181
|  |  | 56.0002 | 42.05 |  |  |  |  | 2465.3733 | 0.077958 | 56.0002 | 45.585 |  |  |  |  | 1836.83 | 0.018329 | 56.0002 | 135.97 |  |  |  |  | 2465.3733 | 0.038674
|  |  | 56.0749 | 41.88 |  |  |  |  | 2468.6431 | 0.091858 | 56.0749 | 45.75 |  |  |  |  | 1839.2661 | 0.023495 | 56.0749 | 136.105 |  |  |  |  | 2468.6431 | 0.056812
|  |  | 56.1496 | 41.6 |  |  |  |  | 2471.9128 | 0.10519 | 56.1496 | 45.975 |  |  |  |  | 1841.7022 | 0.023795 | 56.1496 | 136.41 |  |  |  |  | 2471.9128 | 0.046716
|  |  | 56.2243 | 42.745 |  |  |  |  | 2475.1825 | 0.095401 | 56.2243 | 45.955 |  |  |  |  | 1844.1384 | 0.019212 | 56.2243 | 136.24 |  |  |  |  | 2475.1825 | 0.019981
|  |  | 56.2989 | 42.315 |  |  |  |  | 2478.4522 | 0.058828 | 56.2989 | 45.965 |  |  |  |  | 1846.5745 | 0.01417 | 56.2989 | 136.225 |  |  |  |  | 2478.4522 | 0.029508
|  |  | 56.3736 | 42.23 |  |  |  |  | 2481.722 | 0.015506 | 56.3736 | 45.97 |  |  |  |  | 1849.0106 | 0.012346 | 56.3736 | 136.35 |  |  |  |  | 2481.722 | 0.040553
|  |  | 56.4482 | 42.49 |  |  |  |  | 2484.9917 | 0.031344 | 56.4482 | 45.965 |  |  |  |  | 1851.4467 | 0.010925 | 56.4482 | 136.78 |  |  |  |  | 2484.9917 | 0.026204
|  |  | 56.5229 | 43.635 |  |  |  |  | 2488.2614 | 0.039301 | 56.5229 | 45.95 |  |  |  |  | 1853.8828 | 0.009549 | 56.5229 | 136.355 |  |  |  |  | 2488.2614 | 0.010595
|  |  | 56.5976 | 42.79 |  |  |  |  | 2491.5311 | 0.029342 | 56.5976 | 45.945 |  |  |  |  | 1856.3189 | 0.013218 | 56.5976 | 135.795 |  |  |  |  | 2491.5311 | 0.032292
|  |  | 56.6723 | 43.395 |  |  |  |  | 2494.8009 | 0.039659 | 56.6723 | 45.71 |  |  |  |  | 1858.755 | 0.017704 | 56.6723 | 135.305 |  |  |  |  | 2494.8009 | 0.037246
|  |  | 56.7469 | 42.175 |  |  |  |  | 2498.0706 | 0.054229 | 56.7469 | 45.465 |  |  |  |  | 1861.1912 | 0.017058 | 56.7469 | 135.255 |  |  |  |  | 2498.0706 | 0.026054
|  |  | 56.8216 | 41.745 |  |  |  |  | 2501.3403 | 0.048325 | 56.8216 | 45.45 |  |  |  |  | 1863.6273 | 0.0098567 | 56.8216 | 135.655 |  |  |  |  | 2501.3403 | 0.022091
|  |  | 56.8963 | 40.13 |  |  |  |  | 2504.6101 | 0.029666 | 56.8963 | 45.6 |  |  |  |  | 1866.0634 | 0.0015307 | 56.8963 | 136.05 |  |  |  |  | 2504.6101 | 0.026251
|  |  | 56.9709 | 41.365 |  |  |  |  | 2507.8798 | 0.023792 | 56.9709 | 46.175 |  |  |  |  | 1868.4995 | 0.011454 | 56.9709 | 136.005 |  |  |  |  | 2507.8798 | 0.021616
|  |  | 57.0456 | 41.655 |  |  |  |  | 2511.1495 | 0.028188 | 57.0456 | 47.32 |  |  |  |  | 1870.9356 | 0.015622 | 57.0456 | 135.815 |  |  |  |  | 2511.1495 | 0.023287
|  |  | 57.1203 | 41.21 |  |  |  |  | 2514.4192 | 0.036483 | 57.1203 | 47.89 |  |  |  |  | 1873.3717 | 0.012368 | 57.1203 | 135.64 |  |  |  |  | 2514.4192 | 0.035079
|  |  | 57.1949 | 41.5 |  |  |  |  | 2517.689 | 0.05129 | 57.1949 | 48.305 |  |  |  |  | 1875.8078 | 0.0094213 | 57.1949 | 135.025 |  |  |  |  | 2517.689 | 0.038188
|  |  | 57.2696 | 42.715 |  |  |  |  | 2520.9587 | 0.058912 | 57.2696 | 48.615 |  |  |  |  | 1878.244 | 0.018376 | 57.2696 | 134.685 |  |  |  |  | 2520.9587 | 0.030134
|  |  | 57.3443 | 43.155 |  |  |  |  | 2524.2284 | 0.059406 | 57.3443 | 48.845 |  |  |  |  | 1880.6801 | 0.026616 | 57.3443 | 134.47 |  |  |  |  | 2524.2284 | 0.019744
|  |  | 57.4189 | 43.455 |  |  |  |  | 2527.4981 | 0.069722 | 57.4189 | 49 |  |  |  |  | 1883.1162 | 0.027799 | 57.4189 | 134.17 |  |  |  |  | 2527.4981 | 0.016557
|  |  | 57.4936 | 43.88 |  |  |  |  | 2530.7679 | 0.085555 | 57.4936 | 49.175 |  |  |  |  | 1885.5523 | 0.023645 | 57.4936 | 134.235 |  |  |  |  | 2530.7679 | 0.016132
|  |  | 57.5683 | 42.96 |  |  |  |  | 2534.0376 | 0.087004 | 57.5683 | 49.27 |  |  |  |  | 1887.9884 | 0.023094 | 57.5683 | 134.69 |  |  |  |  | 2534.0376 | 0.0087351
|  |  | 57.6429 | 42.355 |  |  |  |  | 2537.3073 | 0.067038 | 57.6429 | 49.25 |  |  |  |  | 1890.4245 | 0.028565 | 57.6429 | 134.6 |  |  |  |  | 2537.3073 | 0.011096
|  |  | 57.7176 | 42.61 |  |  |  |  | 2540.577 | 0.032212 | 57.7176 | 49.175 |  |  |  |  | 1892.8606 | 0.030258 | 57.7176 | 134.13 |  |  |  |  | 2540.577 | 0.040213
|  |  | 57.7923 | 44.05 |  |  |  |  | 2543.8468 | 0.0064841 | 57.7923 | 48.855 |  |  |  |  | 1895.2968 | 0.022889 | 57.7923 | 133.73 |  |  |  |  | 2543.8468 | 0.068403
|  |  | 57.8669 | 43.55 |  |  |  |  | 2547.1165 | 0.032763 | 57.8669 | 48.63 |  |  |  |  | 1897.7329 | 0.012303 | 57.8669 | 132.995 |  |  |  |  | 2547.1165 | 0.081399
|  |  | 57.9416 | 43.31 |  |  |  |  | 2550.3862 | 0.041242 | 57.9416 | 48.455 |  |  |  |  | 1900.169 | 0.021979 | 57.9416 | 132.51 |  |  |  |  | 2550.3862 | 0.070947
|  |  | 58.0163 | 42.28 |  |  |  |  | 2553.6559 | 0.031743 | 58.0163 | 48.225 |  |  |  |  | 1902.6051 | 0.036244 | 58.0163 | 132.29 |  |  |  |  | 2553.6559 | 0.042738
|  |  | 58.0909 | 43.805 |  |  |  |  | 2556.9257 | 0.02597 | 58.0909 | 48.02 |  |  |  |  | 1905.0412 | 0.041098 | 58.0909 | 132.65 |  |  |  |  | 2556.9257 | 0.025201
|  |  | 58.1656 | 44.3 |  |  |  |  | 2560.1954 | 0.039026 | 58.1656 | 47.89 |  |  |  |  | 1907.4773 | 0.033853 | 58.1656 | 132.875 |  |  |  |  | 2560.1954 | 0.035674
|  |  | 58.2403 | 44.475 |  |  |  |  | 2563.4651 | 0.039809 | 58.2403 | 48.295 |  |  |  |  | 1909.9134 | 0.017395 | 58.2403 | 132.955 |  |  |  |  | 2563.4651 | 0.033621
|  |  | 58.3149 | 44.105 |  |  |  |  | 2566.7348 | 0.018642 | 58.3149 | 48.735 |  |  |  |  | 1912.3496 | 0.0029555 | 58.3149 | 132.485 |  |  |  |  | 2566.7348 | 0.015215
|  |  | 58.3896 | 43.3 |  |  |  |  | 2570.0046 | 0.025237 | 58.3896 | 49.175 |  |  |  |  | 1914.7857 | 0.017697 | 58.3896 | 131.845 |  |  |  |  | 2570.0046 | 0.010921
|  |  | 58.4643 | 42.61 |  |  |  |  | 2573.2743 | 0.051744 | 58.4643 | 49.435 |  |  |  |  | 1917.2218 | 0.024672 | 58.4643 | 131.38 |  |  |  |  | 2573.2743 | 0.0151
|  |  | 58.5389 | 42.61 |  |  |  |  | 2576.544 | 0.051646 | 58.5389 | 49.545 |  |  |  |  | 1919.6579 | 0.023 | 58.5389 | 130.73 |  |  |  |  | 2576.544 | 0.0052435
|  |  | 58.6136 | 43.475 |  |  |  |  | 2579.8137 | 0.022888 | 58.6136 | 49.605 |  |  |  |  | 1922.094 | 0.016571 | 58.6136 | 130.88 |  |  |  |  | 2579.8137 | 0.033988
|  |  | 58.6883 | 44.6 |  |  |  |  | 2583.0835 | 0.01811 | 58.6883 | 49.56 |  |  |  |  | 1924.5301 | 0.011642 | 58.6883 | 131.165 |  |  |  |  | 2583.0835 | 0.059046
|  |  | 58.7629 | 44.59 |  |  |  |  | 2586.3532 | 0.044396 | 58.7629 | 49.645 |  |  |  |  | 1926.9662 | 0.010405 | 58.7629 | 131.52 |  |  |  |  | 2586.3532 | 0.063342
|  |  | 58.8376 | 43.7 |  |  |  |  | 2589.6229 | 0.059287 | 58.8376 | 49.665 |  |  |  |  | 1929.4024 | 0.0085385 | 58.8376 | 131.54 |  |  |  |  | 2589.6229 | 0.046383
|  |  | 58.9123 | 43.26 |  |  |  |  | 2592.8927 | 0.094021 | 58.9123 | 49.34 |  |  |  |  | 1931.8385 | 0.0063512 | 58.9123 | 131.41 |  |  |  |  | 2592.8927 | 0.022042
|  |  | 58.9869 | 43.945 |  |  |  |  | 2596.1624 | 0.13879 | 58.9869 | 49.065 |  |  |  |  | 1934.2746 | 0.0078484 | 58.9869 | 131.28 |  |  |  |  | 2596.1624 | 0.0098854
|  |  | 59.0616 | 44.485 |  |  |  |  | 2599.4321 | 0.15708 | 59.0616 | 48.825 |  |  |  |  | 1936.7107 | 0.0093888 | 59.0616 | 131.245 |  |  |  |  | 2599.4321 | 0.029484
|  |  | 59.1363 | 42.44 |  |  |  |  | 2602.7018 | 0.13266 | 59.1363 | 48.555 |  |  |  |  | 1939.1468 | 0.0087563 | 59.1363 | 130.985 |  |  |  |  | 2602.7018 | 0.059554
|  |  | 59.2109 | 42.775 |  |  |  |  | 2605.9716 | 0.078774 | 59.2109 | 48.135 |  |  |  |  | 1941.5829 | 0.0095663 | 59.2109 | 131.075 |  |  |  |  | 2605.9716 | 0.081406
|  |  | 59.2856 | 42.37 |  |  |  |  | 2609.2413 | 0.036917 | 59.2856 | 47.755 |  |  |  |  | 1944.019 | 0.012859 | 59.2856 | 131.975 |  |  |  |  | 2609.2413 | 0.080535
|  |  | 59.3603 | 42.49 |  |  |  |  | 2612.511 | 0.050706 | 59.3603 | 47.455 |  |  |  |  | 1946.4552 | 0.013753 | 59.3603 | 131.655 |  |  |  |  | 2612.511 | 0.059184
|  |  | 59.4349 | 42.46 |  |  |  |  | 2615.7807 | 0.06981 | 59.4349 | 46.885 |  |  |  |  | 1948.8913 | 0.0097238 | 59.4349 | 131.365 |  |  |  |  | 2615.7807 | 0.039197
|  |  | 59.5096 | 43.27 |  |  |  |  | 2619.0505 | 0.085118 | 59.5096 | 46.355 |  |  |  |  | 1951.3274 | 0.0052512 | 59.5096 | 131.765 |  |  |  |  | 2619.0505 | 0.035756
|  |  | 59.5843 | 43.45 |  |  |  |  | 2622.3202 | 0.094108 | 59.5843 | 45.785 |  |  |  |  | 1953.7635 | 0.012287 | 59.5843 | 132.015 |  |  |  |  | 2622.3202 | 0.027432
|  |  | 59.6589 | 43.41 |  |  |  |  | 2625.5899 | 0.084834 | 59.6589 | 45.77 |  |  |  |  | 1956.1996 | 0.018804 | 59.6589 | 131.925 |  |  |  |  | 2625.5899 | 0.010048
|  |  | 59.7336 | 43.12 |  |  |  |  | 2628.8596 | 0.055222 | 59.7336 | 46.18 |  |  |  |  | 1958.6357 | 0.019371 | 59.7336 | 131.415 |  |  |  |  | 2628.8596 | 0.020369
|  |  | 59.8083 | 43.47 |  |  |  |  | 2632.1294 | 0.027335 | 59.8083 | 46.38 |  |  |  |  | 1961.0718 | 0.013888 | 59.8083 | 131.26 |  |  |  |  | 2632.1294 | 0.03073
|  |  | 59.8829 | 43.905 |  |  |  |  | 2635.3991 | 0.043244 | 59.8829 | 46.535 |  |  |  |  | 1963.508 | 0.0070632 | 59.8829 | 131.325 |  |  |  |  | 2635.3991 | 0.02583
|  |  | 59.9576 | 44.31 |  |  |  |  | 2638.6688 | 0.057631 | 59.9576 | 46.65 |  |  |  |  | 1965.9441 | 0.0085457 | 59.9576 | 131.585 |  |  |  |  | 2638.6688 | 0.0098765
|  |  | 60.0323 | 43.5 |  |  |  |  | 2641.9385 | 0.059949 | 60.0323 | 46.765 |  |  |  |  | 1968.3802 | 0.010305 | 60.0323 | 131.48 |  |  |  |  | 2641.9385 | 0.018486
|  |  | 60.1069 | 41.865 |  |  |  |  | 2645.2083 | 0.078066 | 60.1069 | 46.72 |  |  |  |  | 1970.8163 | 0.0056226 | 60.1069 | 131.28 |  |  |  |  | 2645.2083 | 0.034933
|  |  | 60.1816 | 43.3 |  |  |  |  | 2648.478 | 0.10881 | 60.1816 | 46.54 |  |  |  |  | 1973.2524 | 0.0045395 | 60.1816 | 131.165 |  |  |  |  | 2648.478 | 0.033991
|  |  | 60.2563 | 42.745 |  |  |  |  | 2651.7477 | 0.11966 | 60.2563 | 46.38 |  |  |  |  | 1975.6885 | 0.015953 | 60.2563 | 130.88 |  |  |  |  | 2651.7477 | 0.012228
|  |  | 60.3309 | 41.82 |  |  |  |  | 2655.0174 | 0.093876 | 60.3309 | 46.31 |  |  |  |  | 1978.1246 | 0.023506 | 60.3309 | 130.78 |  |  |  |  | 2655.0174 | 0.018594
|  |  | 60.4056 | 41.655 |  |  |  |  | 2658.2872 | 0.038753 | 60.4056 | 46.155 |  |  |  |  | 1980.5607 | 0.02356 | 60.4056 | 131.45 |  |  |  |  | 2658.2872 | 0.035409
|  |  | 60.4803 | 43.29 |  |  |  |  | 2661.5569 | 0.024893 | 60.4803 | 46 |  |  |  |  | 1982.9969 | 0.0156 | 60.4803 | 131.915 |  |  |  |  | 2661.5569 | 0.024172
|  |  | 60.5549 | 43.4 |  |  |  |  | 2664.8266 | 0.061563 | 60.5549 | 45.715 |  |  |  |  | 1985.433 | 0.0031896 | 60.5549 | 131.875 |  |  |  |  | 2664.8266 | 0.036744
|  |  | 60.6296 | 42.92 |  |  |  |  | 2668.0963 | 0.061168 | 60.6296 | 45.43 |  |  |  |  | 1987.8691 | 0.012109 | 60.6296 | 131.985 |  |  |  |  | 2668.0963 | 0.085067
|  |  | 60.7043 | 42.925 |  |  |  |  | 2671.3661 | 0.041623 | 60.7043 | 45.05 |  |  |  |  | 1990.3052 | 0.021765 | 60.7043 | 132.675 |  |  |  |  | 2671.3661 | 0.10976
|  |  | 60.7789 | 43.91 |  |  |  |  | 2674.6358 | 0.063343 | 60.7789 | 44.615 |  |  |  |  | 1992.7413 | 0.024727 | 60.7789 | 133.055 |  |  |  |  | 2674.6358 | 0.091616
|  |  | 60.8536 | 44.025 |  |  |  |  | 2677.9055 | 0.09273 | 60.8536 | 44.355 |  |  |  |  | 1995.1774 | 0.021337 | 60.8536 | 133.165 |  |  |  |  | 2677.9055 | 0.042075
|  |  | 60.9283 | 42.29 |  |  |  |  | 2681.1753 | 0.090045 | 60.9283 | 43.905 |  |  |  |  | 1997.6135 | 0.01475 | 60.9283 | 132.925 |  |  |  |  | 2681.1753 | 0.024475
|  |  | 61.0029 | 42.99 |  |  |  |  | 2684.445 | 0.055216 | 61.0029 | 43.265 |  |  |  |  | 2000.0497 | 0.0094958 | 61.0029 | 132.885 |  |  |  |  | 2684.445 | 0.045599
|  |  | 61.0776 | 43.6 |  |  |  |  | 2687.7147 | 0.011044 | 61.0776 | 43.19 |  |  |  |  | 2002.4858 | 0.0075886 | 61.0776 | 133.145 |  |  |  |  | 2687.7147 | 0.033844
|  |  | 61.1523 | 43.59 |  |  |  |  | 2690.9844 | 0.027417 | 61.1523 | 43.535 |  |  |  |  | 2004.9219 | 0.0059417 | 61.1523 | 133.715 |  |  |  |  | 2690.9844 | 0.014437
|  |  | 61.2269 | 42.295 |  |  |  |  | 2694.2542 | 0.040503 | 61.2269 | 44.075 |  |  |  |  | 2007.358 | 0.0046134 | 61.2269 | 134.375 |  |  |  |  | 2694.2542 | 0.035538
|  |  | 61.3016 | 41.605 |  |  |  |  | 2697.5239 | 0.042824 | 61.3016 | 44.565 |  |  |  |  | 2009.7941 | 0.0069436 | 61.3016 | 134.94 |  |  |  |  | 2697.5239 | 0.042373
|  |  | 61.3763 | 41.28 |  |  |  |  | 2700.7936 | 0.051856 | 61.3763 | 44.6 |  |  |  |  | 2012.2302 | 0.0096563 | 61.3763 | 134.895 |  |  |  |  | 2700.7936 | 0.035183
|  |  | 61.4509 | 41.865 |  |  |  |  | 2704.0633 | 0.074765 | 61.4509 | 44.625 |  |  |  |  | 2014.6663 | 0.010657 | 61.4509 | 135.015 |  |  |  |  | 2704.0633 | 0.03044
|  |  | 61.5256 | 42.43 |  |  |  |  | 2707.3331 | 0.0993 | 61.5256 | 44.55 |  |  |  |  | 2017.1025 | 0.010865 | 61.5256 | 135.185 |  |  |  |  | 2707.3331 | 0.021834
|  |  | 61.6003 | 42.32 |  |  |  |  | 2710.6028 | 0.10329 | 61.6003 | 44.56 |  |  |  |  | 2019.5386 | 0.011904 | 61.6003 | 135.085 |  |  |  |  | 2710.6028 | 0.0024472
|  |  | 61.6749 | 43.24 |  |  |  |  | 2713.8725 | 0.076493 | 61.6749 | 44.57 |  |  |  |  | 2021.9747 | 0.013371 | 61.6749 | 135.445 |  |  |  |  | 2713.8725 | 0.024823
|  |  | 61.7496 | 42.54 |  |  |  |  | 2717.1422 | 0.045062 | 61.7496 | 44.315 |  |  |  |  | 2024.4108 | 0.01375 | 61.7496 | 136.075 |  |  |  |  | 2717.1422 | 0.045804
|  |  | 61.8243 | 42.42 |  |  |  |  | 2720.412 | 0.055385 | 61.8243 | 44.085 |  |  |  |  | 2026.8469 | 0.012359 | 61.8243 | 136.205 |  |  |  |  | 2720.412 | 0.063219
|  |  | 61.8989 | 41.71 |  |  |  |  | 2723.6817 | 0.054573 | 61.8989 | 43.86 |  |  |  |  | 2029.283 | 0.0093569 | 61.8989 | 136.515 |  |  |  |  | 2723.6817 | 0.072993
|  |  | 61.9736 | 42.105 |  |  |  |  | 2726.9514 | 0.017688 | 61.9736 | 43.605 |  |  |  |  | 2031.7191 | 0.005137 | 61.9736 | 136.53 |  |  |  |  | 2726.9514 | 0.062834
|  |  | 62.0483 | 41.38 |  |  |  |  | 2730.2211 | 0.03319 | 62.0483 | 43.285 |  |  |  |  | 2034.1553 | 0.00083699 | 62.0483 | 137.26 |  |  |  |  | 2730.2211 | 0.035727
|  |  | 62.1229 | 40.53 |  |  |  |  | 2733.4909 | 0.05891 | 62.1229 | 42.965 |  |  |  |  | 2036.5914 | 0.0050211 | 62.1229 | 137.59 |  |  |  |  | 2733.4909 | 0.037156
|  |  | 62.1976 | 41.295 |  |  |  |  | 2736.7606 | 0.04198 | 62.1976 | 42.385 |  |  |  |  | 2039.0275 | 0.0089556 | 62.1976 | 137.5 |  |  |  |  | 2736.7606 | 0.051318
|  |  | 62.2723 | 41.91 |  |  |  |  | 2740.0303 | 0.011783 | 62.2723 | 41.95 |  |  |  |  | 2041.4636 | 0.01157 | 62.2723 | 137.01 |  |  |  |  | 2740.0303 | 0.03481
|  |  | 62.3469 | 43.855 |  |  |  |  | 2743.3 | 0.046989 | 62.3469 | 41.66 |  |  |  |  | 2043.8997 | 0.014062 | 62.3469 | 137.665 |  |  |  |  | 2743.3 | 0.018944
|  |  | 62.4216 | 43.49 |  |  |  |  | 2746.5698 | 0.055303 | 62.4216 | 41.155 |  |  |  |  | 2046.3358 | 0.017838 | 62.4216 | 138.28 |  |  |  |  | 2746.5698 | 0.056705
|  |  | 62.4963 | 41.975 |  |  |  |  | 2749.8395 | 0.049005 | 62.4963 | 41.265 |  |  |  |  | 2048.7719 | 0.021901 | 62.4963 | 138.555 |  |  |  |  | 2749.8395 | 0.070411
|  |  | 62.5709 | 41.73 |  |  |  |  | 2753.1092 | 0.063591 | 62.5709 | 41.815 |  |  |  |  | 2051.2081 | 0.023775 | 62.5709 | 138.985 |  |  |  |  | 2753.1092 | 0.05613
|  |  | 62.6456 | 42.66 |  |  |  |  | 2756.3789 | 0.066663 | 62.6456 | 42.545 |  |  |  |  | 2053.6442 | 0.021836 | 62.6456 | 139.16 |  |  |  |  | 2756.3789 | 0.050231
|  |  | 62.7203 | 41.905 |  |  |  |  | 2759.6487 | 0.038869 | 62.7203 | 42.835 |  |  |  |  | 2056.0803 | 0.016432 | 62.7203 | 138.91 |  |  |  |  | 2759.6487 | 0.059821
|  |  | 62.7949 | 42.205 |  |  |  |  | 2762.9184 | 0.028162 | 62.7949 | 43.15 |  |  |  |  | 2058.5164 | 0.009973 | 62.7949 | 138.68 |  |  |  |  | 2762.9184 | 0.047368
|  |  | 62.8696 | 41.14 |  |  |  |  | 2766.1881 | 0.049862 | 62.8696 | 43.365 |  |  |  |  | 2060.9525 | 0.0061717 | 62.8696 | 138.695 |  |  |  |  | 2766.1881 | 0.018839
|  |  | 62.9443 | 41.495 |  |  |  |  | 2769.4579 | 0.044737 | 62.9443 | 43.475 |  |  |  |  | 2063.3886 | 0.0051526 | 62.9443 | 139.16 |  |  |  |  | 2769.4579 | 0.041097
|  |  | 63.0189 | 40.58 |  |  |  |  | 2772.7276 | 0.030843 | 63.0189 | 43.685 |  |  |  |  | 2065.8247 | 0.0024304 | 63.0189 | 139.53 |  |  |  |  | 2772.7276 | 0.058282
|  |  | 63.0936 | 41.22 |  |  |  |  | 2775.9973 | 0.031474 | 63.0936 | 43.95 |  |  |  |  | 2068.2609 | 0.0028682 | 63.0936 | 139.62 |  |  |  |  | 2775.9973 | 0.047443
|  |  | 63.1683 | 41.885 |  |  |  |  | 2779.267 | 0.013436 | 63.1683 | 43.945 |  |  |  |  | 2070.697 | 0.0083651 | 63.1683 | 139.735 |  |  |  |  | 2779.267 | 0.026129
|  |  | 63.2429 | 42.54 |  |  |  |  | 2782.5368 | 0.034942 | 63.2429 | 43.83 |  |  |  |  | 2073.1331 | 0.011586 | 63.2429 | 138.92 |  |  |  |  | 2782.5368 | 0.029853
|  |  | 63.3176 | 42.33 |  |  |  |  | 2785.8065 | 0.074352 | 63.3176 | 43.605 |  |  |  |  | 2075.5692 | 0.010944 | 63.3176 | 138.83 |  |  |  |  | 2785.8065 | 0.040517
|  |  | 63.3923 | 42.01 |  |  |  |  | 2789.0762 | 0.081095 | 63.3923 | 43.19 |  |  |  |  | 2078.0053 | 0.0069764 | 63.3923 | 139.035 |  |  |  |  | 2789.0762 | 0.047728
|  |  | 63.4669 | 42.3 |  |  |  |  | 2792.3459 | 0.052444 | 63.4669 | 43.075 |  |  |  |  | 2080.4414 | 0.0046106 | 63.4669 | 139.23 |  |  |  |  | 2792.3459 | 0.058757
|  |  | 63.5416 | 42.22 |  |  |  |  | 2795.6157 | 0.015167 | 63.5416 | 42.98 |  |  |  |  | 2082.8775 | 0.0085545 | 63.5416 | 138.94 |  |  |  |  | 2795.6157 | 0.06746
|  |  | 63.6163 | 42.695 |  |  |  |  | 2798.8854 | 0.0064045 | 63.6163 | 42.81 |  |  |  |  | 2085.3137 | 0.011522 | 63.6163 | 138.79 |  |  |  |  | 2798.8854 | 0.062489
|  |  | 63.6909 | 41.44 |  |  |  |  | 2802.1551 | 0.013208 | 63.6909 | 42.705 |  |  |  |  | 2087.7498 | 0.011481 | 63.6909 | 138.91 |  |  |  |  | 2802.1551 | 0.039255
|  |  | 63.7656 | 41.265 |  |  |  |  | 2805.4248 | 0.048246 | 63.7656 | 42.855 |  |  |  |  | 2090.1859 | 0.0096883 | 63.7656 | 138.235 |  |  |  |  | 2805.4248 | 0.018386
|  |  | 63.8403 | 41.52 |  |  |  |  | 2808.6946 | 0.070073 | 63.8403 | 43.49 |  |  |  |  | 2092.622 | 0.0091856 | 63.8403 | 137.755 |  |  |  |  | 2808.6946 | 0.046461
|  |  | 63.915 | 41.395 |  |  |  |  | 2811.9643 | 0.061205 | 63.915 | 43.995 |  |  |  |  | 2095.0581 | 0.010942 | 63.915 | 137.67 |  |  |  |  | 2811.9643 | 0.063935
|  |  | 63.9896 | 41.88 |  |  |  |  | 2815.234 | 0.031218 | 63.9896 | 44.37 |  |  |  |  | 2097.4942 | 0.012604 | 63.9896 | 137.335 |  |  |  |  | 2815.234 | 0.056545
|  |  | 64.0643 | 41.29 |  |  |  |  | 2818.5037 | 0.012141 | 64.0643 | 44.635 |  |  |  |  | 2099.9303 | 0.012627 | 64.0643 | 136.92 |  |  |  |  | 2818.5037 | 0.041909
|  |  | 64.139 | 42.075 |  |  |  |  | 2821.7735 | 0.0090756 | 64.139 | 44.725 |  |  |  |  | 2102.3665 | 0.011026 | 64.139 | 136.695 |  |  |  |  | 2821.7735 | 0.044322
|  |  | 64.2136 | 42.83 |  |  |  |  | 2825.0432 | 0.009247 | 64.2136 | 44.97 |  |  |  |  | 2104.8026 | 0.010036 | 64.2136 | 137.045 |  |  |  |  | 2825.0432 | 0.042496
|  |  | 64.2883 | 43.645 |  |  |  |  | 2828.3129 | 0.02977 | 64.2883 | 45.29 |  |  |  |  | 2107.2387 | 0.013582 | 64.2883 | 137 |  |  |  |  | 2828.3129 | 0.027418
|  |  | 64.363 | 43.49 |  |  |  |  | 2831.5826 | 0.038193 | 64.363 | 45.43 |  |  |  |  | 2109.6748 | 0.020282 | 64.363 | 136.295 |  |  |  |  | 2831.5826 | 0.031761
|  |  | 64.4376 | 42.03 |  |  |  |  | 2834.8524 | 0.041597 | 64.4376 | 45.42 |  |  |  |  | 2112.1109 | 0.026202 | 64.4376 | 136.12 |  |  |  |  | 2834.8524 | 0.041742
|  |  | 64.5123 | 40.685 |  |  |  |  | 2838.1221 | 0.053917 | 64.5123 | 45.2 |  |  |  |  | 2114.547 | 0.028702 | 64.5123 | 136.145 |  |  |  |  | 2838.1221 | 0.030141
|  |  | 64.587 | 42.155 |  |  |  |  | 2841.3918 | 0.064954 | 64.587 | 45.075 |  |  |  |  | 2116.9831 | 0.026804 | 64.587 | 135.725 |  |  |  |  | 2841.3918 | 0.0035795
|  |  | 64.6616 | 43.31 |  |  |  |  | 2844.6615 | 0.065243 | 64.6616 | 45.015 |  |  |  |  | 2119.4193 | 0.021343 | 64.6616 | 135.345 |  |  |  |  | 2844.6615 | 0.01778
|  |  | 64.7363 | 44.02 |  |  |  |  | 2847.9313 | 0.053397 | 64.7363 | 44.94 |  |  |  |  | 2121.8554 | 0.014504 | 64.7363 | 135.4 |  |  |  |  | 2847.9313 | 0.022685
|  |  | 64.811 | 42.245 |  |  |  |  | 2851.201 | 0.02941 | 64.811 | 44.89 |  |  |  |  | 2124.2915 | 0.0086142 | 64.811 | 135.93 |  |  |  |  | 2851.201 | 0.031862
|  |  | 64.8856 | 41.695 |  |  |  |  | 2854.4707 | 0.00042933 | 64.8856 | 44.805 |  |  |  |  | 2126.7276 | 0.0049864 | 64.8856 | 136.29 |  |  |  |  | 2854.4707 | 0.054616
|  |  | 64.9603 | 43.025 |  |  |  |  | 2857.7405 | 0.019067 | 64.9603 | 44.69 |  |  |  |  | 2129.1637 | 0.0050488 | 64.9603 | 136.435 |  |  |  |  | 2857.7405 | 0.066421
|  |  | 65.035 | 42.495 |  |  |  |  | 2861.0102 | 0.022797 | 65.035 | 44.57 |  |  |  |  | 2131.5998 | 0.0080822 | 65.035 | 136.435 |  |  |  |  | 2861.0102 | 0.058541
|  |  | 65.1096 | 42.495 |  |  |  |  | 2864.2799 | 0.037076 | 65.1096 | 44.505 |  |  |  |  | 2134.0359 | 0.0105 | 65.1096 | 136.45 |  |  |  |  | 2864.2799 | 0.039863
|  |  | 65.1843 | 42.495 |  |  |  |  | 2867.5496 | 0.059647 | 65.1843 | 44.68 |  |  |  |  | 2136.4721 | 0.010416 | 65.1843 | 136.42 |  |  |  |  | 2867.5496 | 0.028327
|  |  | 65.259 | 42.575 |  |  |  |  | 2870.8194 | 0.064537 | 65.259 | 45.38 |  |  |  |  | 2138.9082 | 0.008558 | 65.259 | 135.965 |  |  |  |  | 2870.8194 | 0.03343
|  |  | 65.3336 | 43.37 |  |  |  |  | 2874.0891 | 0.050833 | 65.3336 | 45.895 |  |  |  |  | 2141.3443 | 0.0088659 | 65.3336 | 134.985 |  |  |  |  | 2874.0891 | 0.038631
|  |  | 65.4083 | 43.645 |  |  |  |  | 2877.3588 | 0.042212 | 65.4083 | 46.38 |  |  |  |  | 2143.7804 | 0.011574 | 65.4083 | 134.7 |  |  |  |  | 2877.3588 | 0.030299
|  |  | 65.483 | 43.54 |  |  |  |  | 2880.6285 | 0.052516 | 65.483 | 46.96 |  |  |  |  | 2146.2165 | 0.012022 | 65.483 | 135.135 |  |  |  |  | 2880.6285 | 0.021425
|  |  | 65.5576 | 43.005 |  |  |  |  | 2883.8983 | 0.075597 | 65.5576 | 47.185 |  |  |  |  | 2148.6526 | 0.0079929 | 65.5576 | 134.91 |  |  |  |  | 2883.8983 | 0.043032
|  |  | 65.6323 | 42.325 |  |  |  |  | 2887.168 | 0.097688 | 65.6323 | 47.255 |  |  |  |  | 2151.0887 | 0.0029811 | 65.6323 | 134.745 |  |  |  |  | 2887.168 | 0.056269
|  |  | 65.707 | 43.24 |  |  |  |  | 2890.4377 | 0.091416 | 65.707 | 47.4 |  |  |  |  | 2153.5248 | 0.010835 | 65.707 | 134.74 |  |  |  |  | 2890.4377 | 0.044332
|  |  | 65.7816 | 42.84 |  |  |  |  | 2893.7074 | 0.044387 | 65.7816 | 47.705 |  |  |  |  | 2155.961 | 0.018158 | 65.7816 | 134.205 |  |  |  |  | 2893.7074 | 0.025921
|  |  | 65.8563 | 43.455 |  |  |  |  | 2896.9772 | 0.029741 | 65.8563 | 47.99 |  |  |  |  | 2158.3971 | 0.020816 | 65.8563 | 134.13 |  |  |  |  | 2896.9772 | 0.036257
|  |  | 65.931 | 43.8 |  |  |  |  | 2900.2469 | 0.085989 | 65.931 | 48.01 |  |  |  |  | 2160.8332 | 0.01816 | 65.931 | 134.385 |  |  |  |  | 2900.2469 | 0.036965
|  |  | 66.0056 | 43.395 |  |  |  |  | 2903.5166 | 0.10449 | 66.0056 | 47.87 |  |  |  |  | 2163.2693 | 0.012003 | 66.0056 | 134.62 |  |  |  |  | 2903.5166 | 0.020471
|  |  | 66.0803 | 43.265 |  |  |  |  | 2906.7863 | 0.09325 | 66.0803 | 47.71 |  |  |  |  | 2165.7054 | 0.0060719 | 66.0803 | 134.325 |  |  |  |  | 2906.7863 | 0.029886
|  |  | 66.155 | 43.555 |  |  |  |  | 2910.0561 | 0.083579 | 66.155 | 47.53 |  |  |  |  | 2168.1415 | 0.0043377 | 66.155 | 134.25 |  |  |  |  | 2910.0561 | 0.03644
|  |  | 66.2296 | 42.12 |  |  |  |  | 2913.3258 | 0.074168 | 66.2296 | 47.385 |  |  |  |  | 2170.5776 | 0.0041348 | 66.2296 | 134.075 |  |  |  |  | 2913.3258 | 0.013817
|  |  | 66.3043 | 41.985 |  |  |  |  | 2916.5955 | 0.044465 | 66.3043 | 47.28 |  |  |  |  | 2173.0138 | 0.0053724 | 66.3043 | 133.66 |  |  |  |  | 2916.5955 | 0.025647
|  |  | 66.379 | 42.14 |  |  |  |  | 2919.8652 | 0.0040188 | 66.379 | 46.91 |  |  |  |  | 2175.4499 | 0.0099477 | 66.379 | 133.735 |  |  |  |  | 2919.8652 | 0.051215
|  |  | 66.4536 | 42.855 |  |  |  |  | 2923.135 | 0.023475 | 66.4536 | 46.525 |  |  |  |  | 2177.886 | 0.013958 | 66.4536 | 133.64 |  |  |  |  | 2923.135 | 0.041449
|  |  | 66.5283 | 42.695 |  |  |  |  | 2926.4047 | 0.027118 | 66.5283 | 46.31 |  |  |  |  | 2180.3221 | 0.014885 | 66.5283 | 133.93 |  |  |  |  | 2926.4047 | 0.0064519
|  |  | 66.603 | 42.97 |  |  |  |  | 2929.6744 | 0.037444 | 66.603 | 46.665 |  |  |  |  | 2182.7582 | 0.012575 | 66.603 | 134.18 |  |  |  |  | 2929.6744 | 0.049581
|  |  | 66.6776 | 43.38 |  |  |  |  | 2932.9441 | 0.055486 | 66.6776 | 47.335 |  |  |  |  | 2185.1943 | 0.0091087 | 66.6776 | 133.71 |  |  |  |  | 2932.9441 | 0.071161
|  |  | 66.7523 | 42.61 |  |  |  |  | 2936.2139 | 0.055647 | 66.7523 | 47.905 |  |  |  |  | 2187.6304 | 0.0074664 | 66.7523 | 133.235 |  |  |  |  | 2936.2139 | 0.053482
|  |  | 66.827 | 42.635 |  |  |  |  | 2939.4836 | 0.049466 | 66.827 | 48.385 |  |  |  |  | 2190.0666 | 0.0079121 | 66.827 | 133.26 |  |  |  |  | 2939.4836 | 0.014523
|  |  | 66.9016 | 43.62 |  |  |  |  | 2942.7533 | 0.060912 | 66.9016 | 48.575 |  |  |  |  | 2192.5027 | 0.0097522 | 66.9016 | 133.78 |  |  |  |  | 2942.7533 | 0.03847
|  |  | 66.9763 | 44.125 |  |  |  |  | 2946.0231 | 0.06356 | 66.9763 | 48.62 |  |  |  |  | 2194.9388 | 0.013574 | 66.9763 | 134.14 |  |  |  |  | 2946.0231 | 0.056952
|  |  | 67.051 | 44.31 |  |  |  |  | 2949.2928 | 0.036305 | 67.051 | 48.56 |  |  |  |  | 2197.3749 | 0.017528 | 67.051 | 134.15 |  |  |  |  | 2949.2928 | 0.050587
|  |  | 67.1256 | 43.64 |  |  |  |  | 2952.5625 | 0.015889 | 67.1256 | 48.51 |  |  |  |  | 2199.811 | 0.018705 | 67.1256 | 133.57 |  |  |  |  | 2952.5625 | 0.035657
|  |  | 67.2003 | 42.78 |  |  |  |  | 2955.8322 | 0.061419 | 67.2003 | 48.5 |  |  |  |  | 2202.2471 | 0.015589 | 67.2003 | 133.245 |  |  |  |  | 2955.8322 | 0.030767
|  |  | 67.275 | 43.24 |  |  |  |  | 2959.102 | 0.083383 | 67.275 | 48.41 |  |  |  |  | 2204.6832 | 0.0088593 | 67.275 | 132.785 |  |  |  |  | 2959.102 | 0.040426
|  |  | 67.3496 | 43.4 |  |  |  |  | 2962.3717 | 0.072864 | 67.3496 | 48.01 |  |  |  |  | 2207.1194 | 0.0014562 | 67.3496 | 133.505 |  |  |  |  | 2962.3717 | 0.052485
|  |  | 67.4243 | 43.615 |  |  |  |  | 2965.6414 | 0.038211 | 67.4243 | 47.73 |  |  |  |  | 2209.5555 | 0.0053608 | 67.4243 | 134.36 |  |  |  |  | 2965.6414 | 0.050465
|  |  | 67.499 | 42.825 |  |  |  |  | 2968.9111 | 0.011324 | 67.499 | 47.47 |  |  |  |  | 2211.9916 | 0.0075568 | 67.499 | 134.28 |  |  |  |  | 2968.9111 | 0.030673
|  |  | 67.5736 | 42.775 |  |  |  |  | 2972.1809 | 0.05027 | 67.5736 | 47.205 |  |  |  |  | 2214.4277 | 0.0068425 | 67.5736 | 134.295 |  |  |  |  | 2972.1809 | 0.018172
|  |  | 67.6483 | 43.065 |  |  |  |  | 2975.4506 | 0.083161 | 67.6483 | 46.93 |  |  |  |  | 2216.8638 | 0.0082642 | 67.6483 | 134.675 |  |  |  |  | 2975.4506 | 0.026186
|  |  | 67.723 | 44.115 |  |  |  |  | 2978.7203 | 0.10029 | 67.723 | 46.465 |  |  |  |  | 2219.2999 | 0.012605 | 67.723 | 134.925 |  |  |  |  | 2978.7203 | 0.015345
|  |  | 67.7976 | 43.515 |  |  |  |  | 2981.99 | 0.094696 | 67.7976 | 45.81 |  |  |  |  | 2221.736 | 0.015336 | 67.7976 | 135.015 |  |  |  |  | 2981.99 | 0.017499
|  |  | 67.8723 | 42.195 |  |  |  |  | 2985.2598 | 0.068582 | 67.8723 | 45.385 |  |  |  |  | 2224.1722 | 0.014365 | 67.8723 | 134.495 |  |  |  |  | 2985.2598 | 0.04371
|  |  | 67.947 | 42.605 |  |  |  |  | 2988.5295 | 0.046264 | 67.947 | 44.805 |  |  |  |  | 2226.6083 | 0.010086 | 67.947 | 134.17 |  |  |  |  | 2988.5295 | 0.046386
|  |  | 68.0216 | 44.255 |  |  |  |  | 2991.7992 | 0.062134 | 68.0216 | 44.025 |  |  |  |  | 2229.0444 | 0.0044497 | 68.0216 | 133.93 |  |  |  |  | 2991.7992 | 0.020129
|  |  | 68.0963 | 43.475 |  |  |  |  | 2995.0689 | 0.084841 | 68.0963 | 44 |  |  |  |  | 2231.4805 | 0.00080343 | 68.0963 | 134.425 |  |  |  |  | 2995.0689 | 0.0206
|  |  | 68.171 | 42.345 |  |  |  |  | 2998.3387 | 0.090686 | 68.171 | 44.455 |  |  |  |  | 2233.9166 | 0.0042261 | 68.171 | 135.025 |  |  |  |  | 2998.3387 | 0.051667
|  |  | 68.2456 | 42.445 |  |  |  |  | 3001.6084 | 0.079466 | 68.2456 | 45.355 |  |  |  |  | 2236.3527 | 0.0075876 | 68.2456 | 134.765 |  |  |  |  | 3001.6084 | 0.058636
|  |  | 68.3203 | 43.285 |  |  |  |  | 3004.8781 | 0.062462 | 68.3203 | 45.62 |  |  |  |  | 2238.7888 | 0.01116 | 68.3203 | 133.95 |  |  |  |  | 3004.8781 | 0.046452
|  |  | 68.395 | 44.195 |  |  |  |  | 3008.1478 | 0.053032 | 68.395 | 45.72 |  |  |  |  | 2241.225 | 0.013568 | 68.395 | 133.565 |  |  |  |  | 3008.1478 | 0.037303
|  |  | 68.4696 | 43.085 |  |  |  |  | 3011.4176 | 0.055101 | 68.4696 | 45.735 |  |  |  |  | 2243.6611 | 0.013335 | 68.4696 | 133.485 |  |  |  |  | 3011.4176 | 0.038698
|  |  | 68.5443 | 43.215 |  |  |  |  | 3014.6873 | 0.066007 | 68.5443 | 45.645 |  |  |  |  | 2246.0972 | 0.010409 | 68.5443 | 133.81 |  |  |  |  | 3014.6873 | 0.036191
|  |  | 68.619 | 42.78 |  |  |  |  | 3017.957 | 0.072194 | 68.619 | 45.575 |  |  |  |  | 2248.5333 | 0.0068636 | 68.619 | 133.935 |  |  |  |  | 3017.957 | 0.033546
|  |  | 68.6936 | 42.335 |  |  |  |  | 3021.2267 | 0.053773 | 68.6936 | 45.465 |  |  |  |  | 2250.9694 | 0.006534 | 68.6936 | 134.23 |  |  |  |  | 3021.2267 | 0.040853
|  |  | 68.7683 | 42.335 |  |  |  |  | 3024.4965 | 0.016623 | 68.7683 | 45.365 |  |  |  |  | 2253.4055 | 0.008468 | 68.7683 | 134.21 |  |  |  |  | 3024.4965 | 0.051943
|  |  | 68.843 | 43.11 |  |  |  |  | 3027.7662 | 0.067601 | 68.843 | 45.235 |  |  |  |  | 2255.8416 | 0.0096402 | 68.843 | 133.73 |  |  |  |  | 3027.7662 | 0.058089
|  |  | 68.9176 | 43.015 |  |  |  |  | 3031.0359 | 0.11646 | 68.9176 | 44.965 |  |  |  |  | 2258.2778 | 0.0098218 | 68.9176 | 133.75 |  |  |  |  | 3031.0359 | 0.052614
|  |  | 68.9923 | 42.305 |  |  |  |  | 3034.3057 | 0.1262 | 68.9923 | 44.67 |  |  |  |  | 2260.7139 | 0.0090294 | 68.9923 | 134.3 |  |  |  |  | 3034.3057 | 0.034468
|  |  | 69.067 | 42.215 |  |  |  |  | 3037.5754 | 0.097014 | 69.067 | 44.355 |  |  |  |  | 2263.15 | 0.006514 | 69.067 | 134.565 |  |  |  |  | 3037.5754 | 0.023104
|  |  | 69.1416 | 43.725 |  |  |  |  | 3040.8451 | 0.057743 | 69.1416 | 44.085 |  |  |  |  | 2265.5861 | 0.0020177 | 69.1416 | 134.915 |  |  |  |  | 3040.8451 | 0.035481
|  |  | 69.2163 | 42.655 |  |  |  |  | 3044.1148 | 0.044936 | 69.2163 | 43.645 |  |  |  |  | 2268.0222 | 0.0038033 | 69.2163 | 135.24 |  |  |  |  | 3044.1148 | 0.036352
|  |  | 69.291 | 42.73 |  |  |  |  | 3047.3846 | 0.037988 | 69.291 | 43.07 |  |  |  |  | 2270.4583 | 0.0080912 | 69.291 | 134.925 |  |  |  |  | 3047.3846 | 0.016434
|  |  | 69.3656 | 42.54 |  |  |  |  | 3050.6543 | 0.021715 | 69.3656 | 42.84 |  |  |  |  | 2272.8944 | 0.0089762 | 69.3656 | 134.365 |  |  |  |  | 3050.6543 | 0.015995
|  |  | 69.4403 | 42.1 |  |  |  |  | 3053.924 | 0.047848 | 69.4403 | 43.255 |  |  |  |  | 2275.3306 | 0.0057545 | 69.4403 | 134.12 |  |  |  |  | 3053.924 | 0.035583
|  |  | 69.515 | 42.355 |  |  |  |  | 3057.1937 | 0.084699 | 69.515 | 43.67 |  |  |  |  | 2277.7667 | 0.0037825 | 69.515 | 134.67 |  |  |  |  | 3057.1937 | 0.02992
|  |  | 69.5896 | 42.275 |  |  |  |  | 3060.4635 | 0.10908 | 69.5896 | 44.05 |  |  |  |  | 2280.2028 | 0.010232 | 69.5896 | 135.325 |  |  |  |  | 3060.4635 | 0.0046277
|  |  | 69.6643 | 42.425 |  |  |  |  | 3063.7332 | 0.11625 | 69.6643 | 44.28 |  |  |  |  | 2282.6389 | 0.015071 | 69.6643 | 135.86 |  |  |  |  | 3063.7332 | 0.047645
|  |  | 69.739 | 43.205 |  |  |  |  | 3067.0029 | 0.10359 | 69.739 | 44.38 |  |  |  |  | 2285.075 | 0.015278 | 69.739 | 136.41 |  |  |  |  | 3067.0029 | 0.083554
|  |  | 69.8136 | 42.88 |  |  |  |  | 3070.2726 | 0.071369 | 69.8136 | 44.415 |  |  |  |  | 2287.5111 | 0.010768 | 69.8136 | 137.255 |  |  |  |  | 3070.2726 | 0.094928
|  |  | 69.8883 | 43.16 |  |  |  |  | 3073.5424 | 0.028939 | 69.8883 | 44.23 |  |  |  |  | 2289.9472 | 0.0038374 | 69.8883 | 136.935 |  |  |  |  | 3073.5424 | 0.078501
|  |  | 69.963 | 41.6 |  |  |  |  | 3076.8121 | 0.0071459 | 69.963 | 44.035 |  |  |  |  | 2292.3834 | 0.0026615 | 69.963 | 136.505 |  |  |  |  | 3076.8121 | 0.043056
|  |  | 70.0376 | 40.75 |  |  |  |  | 3080.0818 | 0.019899 | 70.0376 | 43.93 |  |  |  |  | 2294.8195 | 0.0048283 | 70.0376 | 136.765 |  |  |  |  | 3080.0818 | 0.011712
|  |  | 70.1123 | 42.825 |  |  |  |  | 3083.3515 | 0.021453 | 70.1123 | 43.675 |  |  |  |  | 2297.2556 | 0.0040936 | 70.1123 | 136.725 |  |  |  |  | 3083.3515 | 0.034223
|  |  | 70.187 | 42.365 |  |  |  |  | 3086.6213 | 0.045777 | 70.187 | 43.62 |  |  |  |  | 2299.6917 | 0.008605 | 70.187 | 136.395 |  |  |  |  | 3086.6213 | 0.045256
|  |  | 70.2617 | 42.88 |  |  |  |  | 3089.891 | 0.071788 | 70.2617 | 43.505 |  |  |  |  | 2302.1278 | 0.015679 | 70.2617 | 136.305 |  |  |  |  | 3089.891 | 0.03338
|  |  | 70.3363 | 41.985 |  |  |  |  | 3093.1607 | 0.080624 | 70.3363 | 43.12 |  |  |  |  | 2304.5639 | 0.019558 | 70.3363 | 136.605 |  |  |  |  | 3093.1607 | 0.010735
|  |  | 70.411 | 42.01 |  |  |  |  | 3096.4304 | 0.070472 | 70.411 | 43.03 |  |  |  |  | 2307 | 0.017483 | 70.411 | 137.11 |  |  |  |  | 3096.4304 | 0.028116
|  |  | 70.4856 | 41.43 |  |  |  |  | 3099.7002 | 0.046121 | 70.4856 | 42.785 |  |  |  |  | 2309.4362 | 0.0096837 | 70.4856 | 137.505 |  |  |  |  | 3099.7002 | 0.042095
|  |  | 70.5603 | 40.16 |  |  |  |  | 3102.9699 | 0.015615 | 70.5603 | 42.08 |  |  |  |  | 2311.8723 | 0.0035575 | 70.5603 | 137.52 |  |  |  |  | 3102.9699 | 0.037646
|  |  | 70.635 | 40.395 |  |  |  |  | 3106.2396 | 0.031592 | 70.635 | 41.69 |  |  |  |  | 2314.3084 | 0.012064 | 70.635 | 137.285 |  |  |  |  | 3106.2396 | 0.022279
|  |  | 70.7096 | 41.11 |  |  |  |  | 3109.5093 | 0.051564 | 70.7096 | 41.355 |  |  |  |  | 2316.7445 | 0.016549 | 70.7096 | 137.395 |  |  |  |  | 3109.5093 | 0.0091275
|  |  | 70.7843 | 39.925 |  |  |  |  | 3112.7791 | 0.042524 | 70.7843 | 41.335 |  |  |  |  | 2319.1806 | 0.014221 | 70.7843 | 138.2 |  |  |  |  | 3112.7791 | 0.020286
|  |  | 70.859 | 40.61 |  |  |  |  | 3116.0488 | 0.011703 | 70.859 | 41.715 |  |  |  |  | 2321.6167 | 0.0065497 | 70.859 | 138.705 |  |  |  |  | 3116.0488 | 0.040532
|  |  | 70.9336 | 41.995 |  |  |  |  | 3119.3185 | 0.043778 | 70.9336 | 42.21 |  |  |  |  | 2324.0528 | 0.0061718 | 70.9336 | 138.68 |  |  |  |  | 3119.3185 | 0.050765
|  |  | 71.0083 | 42.62 |  |  |  |  | 3122.5883 | 0.075661 | 71.0083 | 42.78 |  |  |  |  | 2326.4889 | 0.014268 | 71.0083 | 138.36 |  |  |  |  | 3122.5883 | 0.039564
|  |  | 71.083 | 42.895 |  |  |  |  | 3125.858 | 0.084057 | 71.083 | 42.99 |  |  |  |  | 2328.9251 | 0.017933 | 71.083 | 138.35 |  |  |  |  | 3125.858 | 0.012138
|  |  | 71.1577 | 43.305 |  |  |  |  | 3129.1277 | 0.072657 | 71.1577 | 43.255 |  |  |  |  | 2331.3612 | 0.015981 | 71.1577 | 138.37 |  |  |  |  | 3129.1277 | 0.024141
|  |  | 71.2323 | 41.555 |  |  |  |  | 3132.3974 | 0.057917 | 71.2323 | 43.44 |  |  |  |  | 2333.7973 | 0.010095 | 71.2323 | 138.145 |  |  |  |  | 3132.3974 | 0.036106
|  |  | 71.307 | 40.95 |  |  |  |  | 3135.6672 | 0.070582 | 71.307 | 43.56 |  |  |  |  | 2336.2334 | 0.0033664 | 71.307 | 137.43 |  |  |  |  | 3135.6672 | 0.025675
|  |  | 71.3817 | 41.21 |  |  |  |  | 3138.9369 | 0.098268 | 71.3817 | 43.6 |  |  |  |  | 2338.6695 | 0.0019979 | 71.3817 | 136.715 |  |  |  |  | 3138.9369 | 0.018892
|  |  | 71.4563 | 43.225 |  |  |  |  | 3142.2066 | 0.10248 | 71.4563 | 43.655 |  |  |  |  | 2341.1056 | 0.0055877 | 71.4563 | 136.58 |  |  |  |  | 3142.2066 | 0.030954
|  |  | 71.531 | 43.36 |  |  |  |  | 3145.4763 | 0.069237 | 71.531 | 43.72 |  |  |  |  | 2343.5417 | 0.0097856 | 71.531 | 137.175 |  |  |  |  | 3145.4763 | 0.032759
|  |  | 71.6057 | 42.71 |  |  |  |  | 3148.7461 | 0.015883 | 71.6057 | 43.79 |  |  |  |  | 2345.9779 | 0.014484 | 71.6057 | 137.59 |  |  |  |  | 3148.7461 | 0.038341
|  |  | 71.6803 | 41.905 |  |  |  |  | 3152.0158 | 0.037829 | 71.6803 | 43.785 |  |  |  |  | 2348.414 | 0.017475 | 71.6803 | 137.645 |  |  |  |  | 3152.0158 | 0.050627
|  |  | 71.755 | 41.75 |  |  |  |  | 3155.2855 | 0.051259 | 71.755 | 43.535 |  |  |  |  | 2350.8501 | 0.017169 | 71.755 | 137.595 |  |  |  |  | 3155.2855 | 0.045844
|  |  | 71.8297 | 41.86 |  |  |  |  | 3158.5552 | 0.031437 | 71.8297 | 43.435 |  |  |  |  | 2353.2862 | 0.013886 | 71.8297 | 137.39 |  |  |  |  | 3158.5552 | 0.02405
|  |  | 71.9043 | 42.43 |  |  |  |  | 3161.825 | 0.031311 | 71.9043 | 43.37 |  |  |  |  | 2355.7223 | 0.0097537 | 71.9043 | 137.625 |  |  |  |  | 3161.825 | 0.029687
|  |  | 71.979 | 41.785 |  |  |  |  | 3165.0947 | 0.059863 | 71.979 | 43.075 |  |  |  |  | 2358.1584 | 0.0072133 | 71.979 | 137.925 |  |  |  |  | 3165.0947 | 0.042216
|  |  | 72.0537 | 41.345 |  |  |  |  | 3168.3644 | 0.0681 | 72.0537 | 42.75 |  |  |  |  | 2360.5945 | 0.0062931 | 72.0537 | 138.105 |  |  |  |  | 3168.3644 | 0.032378
|  |  | 72.1283 | 40.665 |  |  |  |  | 3171.6341 | 0.053057 | 72.1283 | 42.64 |  |  |  |  | 2363.0307 | 0.0062444 | 72.1283 | 138.11 |  |  |  |  | 3171.6341 | 0.022912
|  |  | 72.203 | 40.34 |  |  |  |  | 3174.9039 | 0.027094 | 72.203 | 43.54 |  |  |  |  | 2365.4668 | 0.0077095 | 72.203 | 137.8 |  |  |  |  | 3174.9039 | 0.032291
|  |  | 72.2777 | 42.8 |  |  |  |  | 3178.1736 | 0.018014 | 72.2777 | 44.375 |  |  |  |  | 2367.9029 | 0.0090037 | 72.2777 | 137.63 |  |  |  |  | 3178.1736 | 0.03222
|  |  | 72.3523 | 42.21 |  |  |  |  | 3181.4433 | 0.039017 | 72.3523 | 44.98 |  |  |  |  | 2370.339 | 0.0077563 | 72.3523 | 137.59 |  |  |  |  | 3181.4433 | 0.044786
|  |  | 72.427 | 41.68 |  |  |  |  | 3184.713 | 0.064334 | 72.427 | 45.385 |  |  |  |  | 2372.7751 | 0.0034617 | 72.427 | 137.42 |  |  |  |  | 3184.713 | 0.06931
|  |  | 72.5017 | 42.835 |  |  |  |  | 3187.9828 | 0.084636 | 72.5017 | 45.81 |  |  |  |  | 2375.2112 | 0.0033982 | 72.5017 | 137.53 |  |  |  |  | 3187.9828 | 0.072587
|  |  | 72.5763 | 41.695 |  |  |  |  | 3191.2525 | 0.084691 | 72.5763 | 45.87 |  |  |  |  | 2377.6473 | 0.0089057 | 72.5763 | 137.395 |  |  |  |  | 3191.2525 | 0.047976
|  |  | 72.651 | 41.405 |  |  |  |  | 3194.5222 | 0.057004 | 72.651 | 45.87 |  |  |  |  | 2380.0835 | 0.011655 | 72.651 | 137.18 |  |  |  |  | 3194.5222 | 0.02118
|  |  | 72.7257 | 41.735 |  |  |  |  | 3197.7919 | 0.023923 | 72.7257 | 45.97 |  |  |  |  | 2382.5196 | 0.010629 | 72.7257 | 136.955 |  |  |  |  | 3197.7919 | 0.028013
|  |  | 72.8003 | 41.345 |  |  |  |  | 3201.0617 | 0.043264 | 72.8003 | 46.105 |  |  |  |  | 2384.9557 | 0.0068404 | 72.8003 | 136.57 |  |  |  |  | 3201.0617 | 0.033315
|  |  | 72.875 | 41.305 |  |  |  |  | 3204.3314 | 0.052554 | 72.875 | 46.165 |  |  |  |  | 2387.3918 | 0.0029719 | 72.875 | 136.35 |  |  |  |  | 3204.3314 | 0.045759
|  |  | 72.9497 | 41.975 |  |  |  |  | 3207.6011 | 0.039138 | 72.9497 | 46.125 |  |  |  |  | 2389.8279 | 0.0022911 | 72.9497 | 136.89 |  |  |  |  | 3207.6011 | 0.0577
|  |  | 73.0243 | 42.445 |  |  |  |  | 3210.8708 | 0.043337 | 73.0243 | 46.175 |  |  |  |  | 2392.264 | 0.0031408 | 73.0243 | 137.645 |  |  |  |  | 3210.8708 | 0.045535
|  |  | 73.099 | 42.09 |  |  |  |  | 3214.1406 | 0.063051 | 73.099 | 46.405 |  |  |  |  | 2394.7001 | 0.0064358 | 73.099 | 137.685 |  |  |  |  | 3214.1406 | 0.010849
|  |  | 73.1737 | 41.995 |  |  |  |  | 3217.4103 | 0.069408 | 73.1737 | 46.52 |  |  |  |  | 2397.1363 | 0.010976 | 73.1737 | 136.655 |  |  |  |  | 3217.4103 | 0.024402
|  |  | 73.2483 | 41.68 |  |  |  |  | 3220.68 | 0.065988 | 73.2483 | 46.375 |  |  |  |  | 2399.5724 | 0.013649 | 73.2483 | 136.105 |  |  |  |  | 3220.68 | 0.035045
|  |  | 73.323 | 41.075 |  |  |  |  | 3223.9498 | 0.062884 | 73.323 | 46.095 |  |  |  |  | 2402.0085 | 0.012571 | 73.323 | 135.94 |  |  |  |  | 3223.9498 | 0.02123
|  |  | 73.3977 | 41.545 |  |  |  |  | 3227.2195 | 0.070799 | 73.3977 | 45.875 |  |  |  |  | 2404.4446 | 0.0081784 | 73.3977 | 136.22 |  |  |  |  | 3227.2195 | 0.026002
|  |  | 73.4723 | 41.895 |  |  |  |  | 3230.4892 | 0.096072 | 73.4723 | 45.785 |  |  |  |  | 2406.8807 | 0.0049155 | 73.4723 | 136.155 |  |  |  |  | 3230.4892 | 0.035028
|  |  | 73.547 | 41.545 |  |  |  |  | 3233.7589 | 0.11949 | 73.547 | 45.84 |  |  |  |  | 2409.3168 | 0.007362 | 73.547 | 136.245 |  |  |  |  | 3233.7589 | 0.01717
|  |  | 73.6217 | 42.38 |  |  |  |  | 3237.0287 | 0.12094 | 73.6217 | 46.655 |  |  |  |  | 2411.7529 | 0.0084637 | 73.6217 | 136.65 |  |  |  |  | 3237.0287 | 0.023273
|  |  | 73.6963 | 42.34 |  |  |  |  | 3240.2984 | 0.098036 | 73.6963 | 47.35 |  |  |  |  | 2414.1891 | 0.006969 | 73.6963 | 136.425 |  |  |  |  | 3240.2984 | 0.04813
|  |  | 73.771 | 42.235 |  |  |  |  | 3243.5681 | 0.063279 | 73.771 | 47.87 |  |  |  |  | 2416.6252 | 0.0086889 | 73.771 | 135.995 |  |  |  |  | 3243.5681 | 0.041205
|  |  | 73.8457 | 42.615 |  |  |  |  | 3246.8378 | 0.035137 | 73.8457 | 48.67 |  |  |  |  | 2419.0613 | 0.013707 | 73.8457 | 134.54 |  |  |  |  | 3246.8378 | 0.0055499
|  |  | 73.9203 | 43.175 |  |  |  |  | 3250.1076 | 0.028356 | 73.9203 | 49.14 |  |  |  |  | 2421.4974 | 0.016024 | 73.9203 | 134.49 |  |  |  |  | 3250.1076 | 0.035739
|  |  | 73.995 | 43.68 |  |  |  |  | 3253.3773 | 0.026612 | 73.995 | 49.33 |  |  |  |  | 2423.9335 | 0.013073 | 73.995 | 134.77 |  |  |  |  | 3253.3773 | 0.05158
|  |  | 74.0697 | 42.3 |  |  |  |  | 3256.647 | 0.024084 | 74.0697 | 49.37 |  |  |  |  | 2426.3696 | 0.0059333 | 74.0697 | 134.605 |  |  |  |  | 3256.647 | 0.034516
|  |  | 74.1443 | 42.67 |  |  |  |  | 3259.9167 | 0.045494 | 74.1443 | 49.435 |  |  |  |  | 2428.8057 | 0.0057397 | 74.1443 | 134.005 |  |  |  |  | 3259.9167 | 0.011498
|  |  | 74.219 | 43.24 |  |  |  |  | 3263.1865 | 0.072729 | 74.219 | 49.39 |  |  |  |  | 2431.2419 | 0.011935 | 74.219 | 133.735 |  |  |  |  | 3263.1865 | 0.029233
|  |  | 74.2937 | 43.48 |  |  |  |  | 3266.4562 | 0.08425 | 74.2937 | 49.275 |  |  |  |  | 2433.678 | 0.013122 | 74.2937 | 133.565 |  |  |  |  | 3266.4562 | 0.028928
|  |  | 74.3683 | 43.615 |  |  |  |  | 3269.7259 | 0.070585 | 74.3683 | 49.32 |  |  |  |  | 2436.1141 | 0.0088114 | 74.3683 | 133.52 |  |  |  |  | 3269.7259 | 0.021835
|  |  | 74.443 | 43.01 |  |  |  |  | 3272.9956 | 0.033944 | 74.443 | 49.295 |  |  |  |  | 2438.5502 | 0.0066141 | 74.443 | 133.745 |  |  |  |  | 3272.9956 | 0.027507
|  |  | 74.5177 | 43.88 |  |  |  |  | 3276.2654 | 0.015621 | 74.5177 | 49.135 |  |  |  |  | 2440.9863 | 0.013402 | 74.5177 | 134 |  |  |  |  | 3276.2654 | 0.017226
|  |  | 74.5923 | 42.865 |  |  |  |  | 3279.5351 | 0.048644 | 74.5923 | 49.005 |  |  |  |  | 2443.4224 | 0.017791 | 74.5923 | 133.71 |  |  |  |  | 3279.5351 | 0.015312
|  |  | 74.667 | 43.835 |  |  |  |  | 3282.8048 | 0.054442 | 74.667 | 48.74 |  |  |  |  | 2445.8585 | 0.01601 | 74.667 | 133.025 |  |  |  |  | 3282.8048 | 0.045227
|  |  | 74.7417 | 43.095 |  |  |  |  | 3286.0745 | 0.034457 | 74.7417 | 48.235 |  |  |  |  | 2448.2947 | 0.0094969 | 74.7417 | 132.35 |  |  |  |  | 3286.0745 | 0.054099
|  |  | 74.8163 | 44.79 |  |  |  |  | 3289.3443 | 0.020218 | 74.8163 | 48.21 |  |  |  |  | 2450.7308 | 0.0088781 | 74.8163 | 132.78 |  |  |  |  | 3289.3443 | 0.051082
|  |  | 74.891 | 43.105 |  |  |  |  | 3292.614 | 0.031402 | 74.891 | 48.465 |  |  |  |  | 2453.1669 | 0.016557 | 74.891 | 132.82 |  |  |  |  | 3292.614 | 0.067818
|  |  | 74.9657 | 42.175 |  |  |  |  | 3295.8837 | 0.04602 | 74.9657 | 49.1 |  |  |  |  | 2455.603 | 0.021392 | 74.9657 | 132.24 |  |  |  |  | 3295.8837 | 0.079326
|  |  | 75.0403 | 42.76 |  |  |  |  | 3299.1534 | 0.070137 | 75.0403 | 49.415 |  |  |  |  | 2458.0391 | 0.021326 | 75.0403 | 131.92 |  |  |  |  | 3299.1534 | 0.057867
|  |  | 75.115 | 43.425 |  |  |  |  | 3302.4232 | 0.083639 | 75.115 | 49.63 |  |  |  |  | 2460.4752 | 0.017972 | 75.115 | 131.92 |  |  |  |  | 3302.4232 | 0.032979
|  |  | 75.1897 | 43.28 |  |  |  |  | 3305.6929 | 0.069677 | 75.1897 | 49.77 |  |  |  |  | 2462.9113 | 0.013654 | 75.1897 | 132.62 |  |  |  |  | 3305.6929 | 0.067986
|  |  | 75.2643 | 43.115 |  |  |  |  | 3308.9626 | 0.036962 | 75.2643 | 49.895 |  |  |  |  | 2465.3475 | 0.0084651 | 75.2643 | 132.745 |  |  |  |  | 3308.9626 | 0.086095
|  |  | 75.339 | 42.58 |  |  |  |  | 3312.2324 | 0.026197 | 75.339 | 49.87 |  |  |  |  | 2467.7836 | 0.0014057 | 75.339 | 132.285 |  |  |  |  | 3312.2324 | 0.064308
|  |  | 75.4137 | 43.1 |  |  |  |  | 3315.5021 | 0.052813 | 75.4137 | 49.85 |  |  |  |  | 2470.2197 | 0.0091188 | 75.4137 | 131.605 |  |  |  |  | 3315.5021 | 0.024124
|  |  | 75.4883 | 42.69 |  |  |  |  | 3318.7718 | 0.072992 | 75.4883 | 49.71 |  |  |  |  | 2472.6558 | 0.018184 | 75.4883 | 131.445 |  |  |  |  | 3318.7718 | 0.034543
|  |  | 75.563 | 43.105 |  |  |  |  | 3322.0415 | 0.06519 | 75.563 | 49.47 |  |  |  |  | 2475.0919 | 0.022903 | 75.563 | 131.25 |  |  |  |  | 3322.0415 | 0.046678
|  |  | 75.6377 | 43.27 |  |  |  |  | 3325.3113 | 0.022016 | 75.6377 | 49.24 |  |  |  |  | 2477.528 | 0.020948 | 75.6377 | 131.205 |  |  |  |  | 3325.3113 | 0.033558
|  |  | 75.7123 | 43.405 |  |  |  |  | 3328.581 | 0.039573 | 75.7123 | 48.995 |  |  |  |  | 2479.9641 | 0.013279 | 75.7123 | 131.45 |  |  |  |  | 3328.581 | 0.018784
|  |  | 75.787 | 44.075 |  |  |  |  | 3331.8507 | 0.069873 | 75.787 | 48.835 |  |  |  |  | 2482.4003 | 0.0086035 | 75.787 | 132 |  |  |  |  | 3331.8507 | 0.028525
|  |  | 75.8617 | 43.495 |  |  |  |  | 3335.1204 | 0.043571 | 75.8617 | 48.58 |  |  |  |  | 2484.8364 | 0.016027 | 75.8617 | 132.425 |  |  |  |  | 3335.1204 | 0.03603
|  |  | 75.9363 | 43.875 |  |  |  |  | 3338.3902 | 0.025579 | 75.9363 | 47.93 |  |  |  |  | 2487.2725 | 0.021808 | 75.9363 | 132.28 |  |  |  |  | 3338.3902 | 0.037652
|  |  | 76.011 | 42.12 |  |  |  |  | 3341.6599 | 0.08565 | 76.011 | 47.43 |  |  |  |  | 2489.7086 | 0.021374 | 76.011 | 132.285 |  |  |  |  | 3341.6599 | 0.037032
|  |  | 76.0857 | 43.865 |  |  |  |  | 3344.9296 | 0.087916 | 76.0857 | 46.91 |  |  |  |  | 2492.1447 | 0.015566 | 76.0857 | 132.345 |  |  |  |  | 3344.9296 | 0.032753
|  |  | 76.1603 | 44.66 |  |  |  |  | 3348.1993 | 0.041473 | 76.1603 | 46.4 |  |  |  |  | 2494.5808 | 0.0077698 | 76.1603 | 132.22 |  |  |  |  | 3348.1993 | 0.023162
|  |  | 76.235 | 43.85 |  |  |  |  | 3351.4691 | 0.080447 | 76.235 | 46.1 |  |  |  |  | 2497.0169 | 0.0025281 | 76.235 | 132.125 |  |  |  |  | 3351.4691 | 0.01731
|  |  | 76.3097 | 43.23 |  |  |  |  | 3354.7388 | 0.119 | 76.3097 | 45.85 |  |  |  |  | 2499.453 | 0.0046472 | 76.3097 | 132.17 |  |  |  |  | 3354.7388 | 0.034749
|  |  | 76.3843 | 43.41 |  |  |  |  | 3358.0085 | 0.10638 | 76.3843 | 46.025 |  |  |  |  | 2501.8892 | 0.010008 | 76.3843 | 132.185 |  |  |  |  | 3358.0085 | 0.051707
|  |  | 76.459 | 43.8 |  |  |  |  | 3361.2782 | 0.10993 | 76.459 | 46.4 |  |  |  |  | 2504.3253 | 0.017371 | 76.459 | 131.8 |  |  |  |  | 3361.2782 | 0.049798
|  |  | 76.5337 | 42.64 |  |  |  |  | 3364.548 | 0.13246 | 76.5337 | 46.695 |  |  |  |  | 2506.7614 | 0.023551 | 76.5337 | 131.385 |  |  |  |  | 3364.548 | 0.026708
|  |  | 76.6083 | 42.28 |  |  |  |  | 3367.8177 | 0.095936 | 76.6083 | 46.79 |  |  |  |  | 2509.1975 | 0.025369 | 76.6083 | 130.865 |  |  |  |  | 3367.8177 | 0.023531
|  |  | 76.683 | 42.275 |  |  |  |  | 3371.0874 | 0.043993 | 76.683 | 46.84 |  |  |  |  | 2511.6336 | 0.021778 | 76.683 | 131.045 |  |  |  |  | 3371.0874 | 0.050591
|  |  | 76.7577 | 41.915 |  |  |  |  | 3374.3571 | 0.12583 | 76.7577 | 46.885 |  |  |  |  | 2514.0697 | 0.014655 | 76.7577 | 131.575 |  |  |  |  | 3374.3571 | 0.058119
|  |  | 76.8323 | 42.035 |  |  |  |  | 3377.6269 | 0.15784 | 76.8323 | 46.965 |  |  |  |  | 2516.5058 | 0.0089771 | 76.8323 | 131.47 |  |  |  |  | 3377.6269 | 0.043903
|  |  | 76.907 | 43.055 |  |  |  |  | 3380.8966 | 0.11107 | 76.907 | 46.88 |  |  |  |  | 2518.942 | 0.0090549 | 76.907 | 131.17 |  |  |  |  | 3380.8966 | 0.028881
|  |  | 76.9817 | 42.8 |  |  |  |  | 3384.1663 | 0.070204 | 76.9817 | 46.69 |  |  |  |  | 2521.3781 | 0.0094405 | 76.9817 | 131.05 |  |  |  |  | 3384.1663 | 0.03474
|  |  | 77.0563 | 41.76 |  |  |  |  | 3387.436 | 0.098435 | 77.0563 | 46.395 |  |  |  |  | 2523.8142 | 0.0077881 | 77.0563 | 131.445 |  |  |  |  | 3387.436 | 0.03895
|  |  | 77.131 | 42.305 |  |  |  |  | 3390.7058 | 0.083678 | 77.131 | 46.315 |  |  |  |  | 2526.2503 | 0.0081121 | 77.131 | 131.965 |  |  |  |  | 3390.7058 | 0.031791
|  |  | 77.2057 | 42.24 |  |  |  |  | 3393.9755 | 0.066229 | 77.2057 | 46.19 |  |  |  |  | 2528.6864 | 0.01074 | 77.2057 | 131.995 |  |  |  |  | 3393.9755 | 0.018597
|  |  | 77.2803 | 43.955 |  |  |  |  | 3397.2452 | 0.12101 | 77.2803 | 45.9 |  |  |  |  | 2531.1225 | 0.011039 | 77.2803 | 132.675 |  |  |  |  | 3397.2452 | 0.015413
|  |  | 77.355 | 43.99 |  |  |  |  | 3400.515 | 0.14111 | 77.355 | 45.425 |  |  |  |  | 2533.5586 | 0.0072851 | 77.355 | 133.16 |  |  |  |  | 3400.515 | 0.031813
|  |  | 77.4297 | 44.045 |  |  |  |  | 3403.7847 | 0.10202 | 77.4297 | 45.12 |  |  |  |  | 2535.9948 | 0.0014065 | 77.4297 | 133.755 |  |  |  |  | 3403.7847 | 0.042381
|  |  | 77.5043 | 43.505 |  |  |  |  | 3407.0544 | 0.059782 | 77.5043 | 44.76 |  |  |  |  | 2538.4309 | 0.0067696 | 77.5043 | 133.47 |  |  |  |  | 3407.0544 | 0.034957
|  |  | 77.579 | 41.82 |  |  |  |  | 3410.3241 | 0.078526 | 77.579 | 44.155 |  |  |  |  | 2540.867 | 0.011124 | 77.579 | 133.465 |  |  |  |  | 3410.3241 | 0.026378
|  |  | 77.6537 | 41.865 |  |  |  |  | 3413.5939 | 0.086905 | 77.6537 | 43.71 |  |  |  |  | 2543.3031 | 0.011308 | 77.6537 | 133.925 |  |  |  |  | 3413.5939 | 0.055718
|  |  | 77.7283 | 42.435 |  |  |  |  | 3416.8636 | 0.069655 | 77.7283 | 43.37 |  |  |  |  | 2545.7392 | 0.0076868 | 77.7283 | 133.625 |  |  |  |  | 3416.8636 | 0.08023
|  |  | 77.803 | 41.385 |  |  |  |  | 3420.1333 | 0.048519 | 77.803 | 43.565 |  |  |  |  | 2548.1753 | 0.0037795 | 77.803 | 133.36 |  |  |  |  | 3420.1333 | 0.07362
|  |  | 77.8777 | 41.475 |  |  |  |  | 3423.403 | 0.026666 | 77.8777 | 44.255 |  |  |  |  | 2550.6114 | 0.0069514 | 77.8777 | 133.56 |  |  |  |  | 3423.403 | 0.038309
|  |  | 77.9523 | 41.85 |  |  |  |  | 3426.6728 | 0.0079116 | 77.9523 | 44.725 |  |  |  |  | 2553.0476 | 0.010915 | 77.9523 | 133.835 |  |  |  |  | 3426.6728 | 0.039243
|  |  | 78.027 | 41.72 |  |  |  |  | 3429.9425 | 0.039591 | 78.027 | 44.99 |  |  |  |  | 2555.4837 | 0.012658 | 78.027 | 134.88 |  |  |  |  | 3429.9425 | 0.073632
|  |  | 78.1017 | 41.605 |  |  |  |  | 3433.2122 | 0.064263 | 78.1017 | 45.01 |  |  |  |  | 2557.9198 | 0.012178 | 78.1017 | 135.61 |  |  |  |  | 3433.2122 | 0.076972
|  |  | 78.1763 | 42.99 |  |  |  |  | 3436.4819 | 0.075135 | 78.1763 | 44.92 |  |  |  |  | 2560.3559 | 0.009939 | 78.1763 | 135.78 |  |  |  |  | 3436.4819 | 0.047478
|  |  | 78.251 | 42.82 |  |  |  |  | 3439.7517 | 0.074885 | 78.251 | 44.825 |  |  |  |  | 2562.792 | 0.0066131 | 78.251 | 135.925 |  |  |  |  | 3439.7517 | 0.0085274
|  |  | 78.3257 | 42.82 |  |  |  |  | 3443.0214 | 0.067987 | 78.3257 | 44.84 |  |  |  |  | 2565.2281 | 0.0043865 | 78.3257 | 136.16 |  |  |  |  | 3443.0214 | 0.014645
|  |  | 78.4004 | 42.275 |  |  |  |  | 3446.2911 | 0.056284 | 78.4004 | 44.805 |  |  |  |  | 2567.6642 | 0.0058474 | 78.4004 | 135.99 |  |  |  |  | 3446.2911 | 0.0079523
|  |  | 78.475 | 41.885 |  |  |  |  | 3449.5608 | 0.043283 | 78.475 | 44.56 |  |  |  |  | 2570.1004 | 0.0070906 | 78.475 | 136.385 |  |  |  |  | 3449.5608 | 0.025053
|  |  | 78.5497 | 41.835 |  |  |  |  | 3452.8306 | 0.032739 | 78.5497 | 44.245 |  |  |  |  | 2572.5365 | 0.0057857 | 78.5497 | 136.545 |  |  |  |  | 3452.8306 | 0.045287
|  |  | 78.6244 | 42.615 |  |  |  |  | 3456.1003 | 0.012516 | 78.6244 | 43.82 |  |  |  |  | 2574.9726 | 0.0047424 | 78.6244 | 136.305 |  |  |  |  | 3456.1003 | 0.037659
|  |  | 78.699 | 43.255 |  |  |  |  | 3459.37 | 0.029006 | 78.699 | 43.345 |  |  |  |  | 2577.4087 | 0.0095193 | 78.699 | 136.365 |  |  |  |  | 3459.37 | 0.018594
|  |  | 78.7737 | 42.86 |  |  |  |  | 3462.6397 | 0.078699 | 78.7737 | 43.145 |  |  |  |  | 2579.8448 | 0.015058 | 78.7737 | 137.245 |  |  |  |  | 3462.6397 | 0.049605
|  |  | 78.8484 | 43.06 |  |  |  |  | 3465.9095 | 0.10906 | 78.8484 | 42.75 |  |  |  |  | 2582.2809 | 0.017738 | 78.8484 | 137.95 |  |  |  |  | 3465.9095 | 0.067232
|  |  | 78.923 | 41.855 |  |  |  |  | 3469.1792 | 0.10306 | 78.923 | 41.76 |  |  |  |  | 2584.717 | 0.016362 | 78.923 | 137.965 |  |  |  |  | 3469.1792 | 0.055757
|  |  | 78.9977 | 42 |  |  |  |  | 3472.4489 | 0.07969 | 78.9977 | 41.315 |  |  |  |  | 2587.1532 | 0.011787 | 78.9977 | 137.69 |  |  |  |  | 3472.4489 | 0.046203
|  |  | 79.0723 | 42.24 |  |  |  |  | 3475.7186 | 0.078469 | 79.0723 | 41.24 |  |  |  |  | 2589.5893 | 0.0070703 | 79.0723 | 138.215 |  |  |  |  | 3475.7186 | 0.055037
|  |  | 79.147 | 41.735 |  |  |  |  | 3478.9884 | 0.077184 | 79.147 | 41.325 |  |  |  |  | 2592.0254 | 0.0066709 | 79.147 | 138.725 |  |  |  |  | 3478.9884 | 0.04027
|  |  | 79.2217 | 42.535 |  |  |  |  | 3482.2581 | 0.064499 | 79.2217 | 42.22 |  |  |  |  | 2594.4615 | 0.0081226 | 79.2217 | 139.27 |  |  |  |  | 3482.2581 | 0.0065655
|  |  | 79.2964 | 42.45 |  |  |  |  | 3485.5278 | 0.096428 | 79.2964 | 42.7 |  |  |  |  | 2596.8976 | 0.0077955 | 79.2964 | 139.595 |  |  |  |  | 3485.5278 | 0.059254
|  |  | 79.371 | 42.86 |  |  |  |  | 3488.7976 | 0.13586 | 79.371 | 43.105 |  |  |  |  | 2599.3337 | 0.0062854 | 79.371 | 138.985 |  |  |  |  | 3488.7976 | 0.08326
|  |  | 79.4457 | 42.66 |  |  |  |  | 3492.0673 | 0.12974 | 79.4457 | 43.505 |  |  |  |  | 2601.7698 | 0.0058436 | 79.4457 | 138.375 |  |  |  |  | 3492.0673 | 0.062569
|  |  | 79.5204 | 42.01 |  |  |  |  | 3495.337 | 0.074766 | 79.5204 | 43.675 |  |  |  |  | 2604.206 | 0.0071373 | 79.5204 | 138.1 |  |  |  |  | 3495.337 | 0.020674
|  |  | 79.595 | 42.16 |  |  |  |  | 3498.6067 | 0.013359 | 79.595 | 43.73 |  |  |  |  | 2606.6421 | 0.0091878 | 79.595 | 138.565 |  |  |  |  | 3498.6067 | 0.049623
|  |  | 79.6697 | 41.6 |  |  |  |  | 3501.8765 | 0.0608 | 79.6697 | 43.64 |  |  |  |  | 2609.0782 | 0.011572 | 79.6697 | 138.825 |  |  |  |  | 3501.8765 | 0.065921
|  |  | 79.7444 | 42.31 |  |  |  |  | 3505.1462 | 0.082462 | 79.7444 | 43.69 |  |  |  |  | 2611.5143 | 0.013457 | 79.7444 | 138.88 |  |  |  |  | 3505.1462 | 0.043022
|  |  | 79.819 | 42.615 |  |  |  |  | 3508.4159 | 0.073846 | 79.819 | 43.74 |  |  |  |  | 2613.9504 | 0.013601 | 79.819 | 138.805 |  |  |  |  | 3508.4159 | 0.0017894
|  |  | 79.8937 | 42.03 |  |  |  |  | 3511.6856 | 0.05678 | 79.8937 | 43.845 |  |  |  |  | 2616.3865 | 0.011361 | 79.8937 | 138.35 |  |  |  |  | 3511.6856 | 0.031911
|  |  | 79.9684 | 41.69 |  |  |  |  | 3514.9554 | 0.063995 | 79.9684 | 43.935 |  |  |  |  | 2618.8226 | 0.0072148 | 79.9684 | 137.995 |  |  |  |  | 3514.9554 | 0.032529
|  |  | 80.043 | 40.82 |  |  |  |  | 3518.2251 | 0.088871 | 80.043 | 43.77 |  |  |  |  | 2621.2588 | 0.0025176 | 80.043 | 138.1 |  |  |  |  | 3518.2251 | 0.0099504
|  |  | 80.1177 | 41.635 |  |  |  |  | 3521.4948 | 0.096324 | 80.1177 | 43.505 |  |  |  |  | 2623.6949 | 0.0012264 | 80.1177 | 138.76 |  |  |  |  | 3521.4948 | 0.033167
|  |  | 80.1924 | 42.18 |  |  |  |  | 3524.7645 | 0.071199 | 80.1924 | 43.34 |  |  |  |  | 2626.131 | 0.0030901 | 80.1924 | 138.81 |  |  |  |  | 3524.7645 | 0.050352
|  |  | 80.267 | 43.39 |  |  |  |  | 3528.0343 | 0.023383 | 80.267 | 43.15 |  |  |  |  | 2628.5671 | 0.0030502 | 80.267 | 138.535 |  |  |  |  | 3528.0343 | 0.042566
|  |  | 80.3417 | 42.46 |  |  |  |  | 3531.304 | 0.021643 | 80.3417 | 42.84 |  |  |  |  | 2631.0032 | 0.0019581 | 80.3417 | 138.005 |  |  |  |  | 3531.304 | 0.027103
|  |  | 80.4164 | 41.545 |  |  |  |  | 3534.5737 | 0.048266 | 80.4164 | 43.045 |  |  |  |  | 2633.4393 | 0.0030917 | 80.4164 | 137.96 |  |  |  |  | 3534.5737 | 0.038464
|  |  | 80.491 | 42.69 |  |  |  |  | 3537.8434 | 0.07458 | 80.491 | 43.535 |  |  |  |  | 2635.8754 | 0.0070312 | 80.491 | 138.1 |  |  |  |  | 3537.8434 | 0.052132
|  |  | 80.5657 | 42.545 |  |  |  |  | 3541.1132 | 0.11025 | 80.5657 | 43.95 |  |  |  |  | 2638.3116 | 0.011752 | 80.5657 | 138.165 |  |  |  |  | 3541.1132 | 0.052463
|  |  | 80.6404 | 42.93 |  |  |  |  | 3544.3829 | 0.12431 | 80.6404 | 44.5 |  |  |  |  | 2640.7477 | 0.015698 | 80.6404 | 137.74 |  |  |  |  | 3544.3829 | 0.045816
|  |  | 80.715 | 42.965 |  |  |  |  | 3547.6526 | 0.09553 | 80.715 | 44.88 |  |  |  |  | 2643.1838 | 0.017307 | 80.715 | 137.415 |  |  |  |  | 3547.6526 | 0.040445
|  |  | 80.7897 | 42.15 |  |  |  |  | 3550.9223 | 0.045606 | 80.7897 | 45.09 |  |  |  |  | 2645.6199 | 0.015723 | 80.7897 | 137.06 |  |  |  |  | 3550.9223 | 0.031874
|  |  | 80.8644 | 41.485 |  |  |  |  | 3554.1921 | 0.045792 | 80.8644 | 45.28 |  |  |  |  | 2648.056 | 0.011359 | 80.8644 | 137.195 |  |  |  |  | 3554.1921 | 0.011756
|  |  | 80.939 | 42.43 |  |  |  |  | 3557.4618 | 0.043194 | 80.939 | 45.325 |  |  |  |  | 2650.4921 | 0.0058828 | 80.939 | 137.15 |  |  |  |  | 3557.4618 | 0.021653
|  |  | 81.0137 | 42.355 |  |  |  |  | 3560.7315 | 0.017125 | 81.0137 | 45.405 |  |  |  |  | 2652.9282 | 0.0022639 | 81.0137 | 137.255 |  |  |  |  | 3560.7315 | 0.048949
|  |  | 81.0884 | 43.845 |  |  |  |  | 3564.0012 | 0.083796 | 81.0884 | 45.565 |  |  |  |  | 2655.3644 | 0.0030636 | 81.0884 | 136.53 |  |  |  |  | 3564.0012 | 0.054318
|  |  | 81.163 | 43.11 |  |  |  |  | 3567.271 | 0.12533 | 81.163 | 45.585 |  |  |  |  | 2657.8005 | 0.0031467 | 81.163 | 135.565 |  |  |  |  | 3567.271 | 0.028702
|  |  | 81.2377 | 42.58 |  |  |  |  | 3570.5407 | 0.11203 | 81.2377 | 45.43 |  |  |  |  | 2660.2366 | 0.0022254 | 81.2377 | 134.79 |  |  |  |  | 3570.5407 | 0.024006
|  |  | 81.3124 | 41.19 |  |  |  |  | 3573.8104 | 0.076343 | 81.3124 | 45.265 |  |  |  |  | 2662.6727 | 0.0019817 | 81.3124 | 135.21 |  |  |  |  | 3573.8104 | 0.068441
|  |  | 81.387 | 40.925 |  |  |  |  | 3577.0802 | 0.081223 | 81.387 | 45.195 |  |  |  |  | 2665.1088 | 0.0039169 | 81.387 | 136.17 |  |  |  |  | 3577.0802 | 0.084508
|  |  | 81.4617 | 43.015 |  |  |  |  | 3580.3499 | 0.075085 | 81.4617 | 45.105 |  |  |  |  | 2667.5449 | 0.0084775 | 81.4617 | 136.725 |  |  |  |  | 3580.3499 | 0.064368
|  |  | 81.5364 | 43.545 |  |  |  |  | 3583.6196 | 0.028746 | 81.5364 | 45.015 |  |  |  |  | 2669.981 | 0.014389 | 81.5364 | 136.525 |  |  |  |  | 3583.6196 | 0.043114
|  |  | 81.611 | 42.575 |  |  |  |  | 3586.8893 | 0.066403 | 81.611 | 44.9 |  |  |  |  | 2672.4171 | 0.01946 | 81.611 | 136.435 |  |  |  |  | 3586.8893 | 0.062596
|  |  | 81.6857 | 43.075 |  |  |  |  | 3590.1591 | 0.1018 | 81.6857 | 44.73 |  |  |  |  | 2674.8533 | 0.021819 | 81.6857 | 136.755 |  |  |  |  | 3590.1591 | 0.067424
|  |  | 81.7604 | 43.495 |  |  |  |  | 3593.4288 | 0.075407 | 81.7604 | 44.515 |  |  |  |  | 2677.2894 | 0.021166 | 81.7604 | 136.41 |  |  |  |  | 3593.4288 | 0.048106
|  |  | 81.835 | 42.035 |  |  |  |  | 3596.6985 | 0.041842 | 81.835 | 44.695 |  |  |  |  | 2679.7255 | 0.019132 | 81.835 | 135.89 |  |  |  |  | 3596.6985 | 0.058431
|  |  | 81.9097 | 41.94 |  |  |  |  | 3599.9682 | 0.088018 | 81.9097 | 45.25 |  |  |  |  | 2682.1616 | 0.017466 | 81.9097 | 135.755 |  |  |  |  | 3599.9682 | 0.081349
|  |  | 81.9844 | 42.545 |  |  |  |  | 3603.238 | 0.096573 | 81.9844 | 46.135 |  |  |  |  | 2684.5977 | 0.015186 | 81.9844 | 136.655 |  |  |  |  | 3603.238 | 0.069701
|  |  | 82.059 | 44.005 |  |  |  |  | 3606.5077 | 0.05191 | 82.059 | 46.545 |  |  |  |  | 2687.0338 | 0.0099596 | 82.059 | 136.64 |  |  |  |  | 3606.5077 | 0.03102
|  |  | 82.1337 | 41.825 |  |  |  |  | 3609.7774 | 0.0099177 | 82.1337 | 46.855 |  |  |  |  | 2689.4699 | 0.0030977 | 82.1337 | 135.88 |  |  |  |  | 3609.7774 | 0.040657
|  |  | 82.2084 | 42.545 |  |  |  |  | 3613.0471 | 0.023418 | 82.2084 | 47.17 |  |  |  |  | 2691.9061 | 0.010139 | 82.2084 | 135.02 |  |  |  |  | 3613.0471 | 0.057062
|  |  | 82.283 | 43.32 |  |  |  |  | 3616.3169 | 0.040116 | 82.283 | 47.515 |  |  |  |  | 2694.3422 | 0.017675 | 82.283 | 134.895 |  |  |  |  | 3616.3169 | 0.034604
|  |  | 82.3577 | 43.735 |  |  |  |  | 3619.5866 | 0.10041 | 82.3577 | 47.65 |  |  |  |  | 2696.7783 | 0.019619 | 82.3577 | 134.85 |  |  |  |  | 3619.5866 | 0.02206
|  |  | 82.4324 | 42.835 |  |  |  |  | 3622.8563 | 0.12571 | 82.4324 | 47.75 |  |  |  |  | 2699.2144 | 0.014756 | 82.4324 | 134.9 |  |  |  |  | 3622.8563 | 0.061431
|  |  | 82.507 | 42.455 |  |  |  |  | 3626.126 | 0.082626 | 82.507 | 47.9 |  |  |  |  | 2701.6505 | 0.0048928 | 82.507 | 134.915 |  |  |  |  | 3626.126 | 0.073654
|  |  | 82.5817 | 43.48 |  |  |  |  | 3629.3958 | 0.015439 | 82.5817 | 48.015 |  |  |  |  | 2704.0866 | 0.0060501 | 82.5817 | 134.73 |  |  |  |  | 3629.3958 | 0.058846
|  |  | 82.6564 | 44.88 |  |  |  |  | 3632.6655 | 0.067247 | 82.6564 | 48.08 |  |  |  |  | 2706.5227 | 0.013285 | 82.6564 | 134.035 |  |  |  |  | 3632.6655 | 0.038408
|  |  | 82.731 | 44.41 |  |  |  |  | 3635.9352 | 0.04906 | 82.731 | 47.94 |  |  |  |  | 2708.9589 | 0.01402 | 82.731 | 134.105 |  |  |  |  | 3635.9352 | 0.021069
|  |  | 82.8057 | 44.05 |  |  |  |  | 3639.2049 | 0.04583 | 82.8057 | 47.84 |  |  |  |  | 2711.395 | 0.0091939 | 82.8057 | 134.16 |  |  |  |  | 3639.2049 | 0.023305
|  |  | 82.8804 | 43.2 |  |  |  |  | 3642.4747 | 0.15187 | 82.8804 | 47.57 |  |  |  |  | 2713.8311 | 0.0093218 | 82.8804 | 134.395 |  |  |  |  | 3642.4747 | 0.055171
|  |  | 82.955 | 43.585 |  |  |  |  | 3645.7444 | 0.19748 | 82.955 | 47.4 |  |  |  |  | 2716.2672 | 0.017509 | 82.955 | 134.39 |  |  |  |  | 3645.7444 | 0.072595
|  |  | 83.0297 | 43.145 |  |  |  |  | 3649.0141 | 0.1554 | 83.0297 | 47.235 |  |  |  |  | 2718.7033 | 0.02249 | 83.0297 | 134.005 |  |  |  |  | 3649.0141 | 0.064892
|  |  | 83.1044 | 42.82 |  |  |  |  | 3652.2838 | 0.075315 | 83.1044 | 46.865 |  |  |  |  | 2721.1394 | 0.020806 | 83.1044 | 134.06 |  |  |  |  | 3652.2838 | 0.049951
|  |  | 83.179 | 42.94 |  |  |  |  | 3655.5536 | 0.067996 | 83.179 | 46.615 |  |  |  |  | 2723.5755 | 0.013263 | 83.179 | 134.555 |  |  |  |  | 3655.5536 | 0.044796
|  |  | 83.2537 | 43.57 |  |  |  |  | 3658.8233 | 0.068892 | 83.2537 | 46.585 |  |  |  |  | 2726.0117 | 0.0044488 | 83.2537 | 134.715 |  |  |  |  | 3658.8233 | 0.028789
|  |  | 83.3284 | 44.25 |  |  |  |  | 3662.093 | 0.035352 | 83.3284 | 47.12 |  |  |  |  | 2728.4478 | 0.0073372 | 83.3284 | 134.755 |  |  |  |  | 3662.093 | 0.0053287
|  |  | 83.403 | 42.965 |  |  |  |  | 3665.3628 | 0.034344 | 83.403 | 48.06 |  |  |  |  | 2730.8839 | 0.010806 | 83.403 | 134.485 |  |  |  |  | 3665.3628 | 0.036448
|  |  | 83.4777 | 42.96 |  |  |  |  | 3668.6325 | 0.011481 | 83.4777 | 48.355 |  |  |  |  | 2733.32 | 0.0094643 | 83.4777 | 133.88 |  |  |  |  | 3668.6325 | 0.040868
|  |  | 83.5524 | 42.635 |  |  |  |  | 3671.9022 | 0.062415 | 83.5524 | 48.64 |  |  |  |  | 2735.7561 | 0.0062754 | 83.5524 | 133.835 |  |  |  |  | 3671.9022 | 0.015585
|  |  | 83.627 | 42.97 |  |  |  |  | 3675.1719 | 0.13265 | 83.627 | 48.95 |  |  |  |  | 2738.1922 | 0.00672 | 83.627 | 133.91 |  |  |  |  | 3675.1719 | 0.023634
|  |  | 83.7017 | 43.745 |  |  |  |  | 3678.4417 | 0.14155 | 83.7017 | 48.995 |  |  |  |  | 2740.6283 | 0.0078347 | 83.7017 | 133.885 |  |  |  |  | 3678.4417 | 0.038733
|  |  | 83.7764 | 43.535 |  |  |  |  | 3681.7114 | 0.076654 | 83.7764 | 48.9 |  |  |  |  | 2743.0645 | 0.0057844 | 83.7764 | 133.82 |  |  |  |  | 3681.7114 | 0.019013
|  |  | 83.851 | 42.92 |  |  |  |  | 3684.9811 | 0.02808 | 83.851 | 48.74 |  |  |  |  | 2745.5006 | 0.0030211 | 83.851 | 133.85 |  |  |  |  | 3684.9811 | 0.023146
|  |  | 83.9257 | 42.885 |  |  |  |  | 3688.2508 | 0.081906 | 83.9257 | 48.685 |  |  |  |  | 2747.9367 | 0.0059965 | 83.9257 | 133.48 |  |  |  |  | 3688.2508 | 0.055836
|  |  | 84.0004 | 42.95 |  |  |  |  | 3691.5206 | 0.075409 | 84.0004 | 48.585 |  |  |  |  | 2750.3728 | 0.0076272 | 84.0004 | 132.815 |  |  |  |  | 3691.5206 | 0.054784
|  |  | 84.075 | 42.9 |  |  |  |  | 3694.7903 | 0.060395 | 84.075 | 48.305 |  |  |  |  | 2752.8089 | 0.0052205 | 84.075 | 132.68 |  |  |  |  | 3694.7903 | 0.028949
|  |  | 84.1497 | 43.15 |  |  |  |  | 3698.06 | 0.086618 | 84.1497 | 48.145 |  |  |  |  | 2755.245 | 0.0048757 | 84.1497 | 132.81 |  |  |  |  | 3698.06 | 0.041648
|  |  | 84.2244 | 43.35 |  |  |  |  | 3701.3297 | 0.077981 | 84.2244 | 47.87 |  |  |  |  | 2757.6811 | 0.011704 | 84.2244 | 133.84 |  |  |  |  | 3701.3297 | 0.054656
|  |  | 84.299 | 43.295 |  |  |  |  | 3704.5995 | 0.036452 | 84.299 | 47.385 |  |  |  |  | 2760.1173 | 0.017001 | 84.299 | 134.2 |  |  |  |  | 3704.5995 | 0.030726
|  |  | 84.3737 | 43.005 |  |  |  |  | 3707.8692 | 0.053386 | 84.3737 | 46.54 |  |  |  |  | 2762.5534 | 0.018161 | 84.3737 | 134.215 |  |  |  |  | 3707.8692 | 0.025599
|  |  | 84.4484 | 43.38 |  |  |  |  | 3711.1389 | 0.060925 | 84.4484 | 46.045 |  |  |  |  | 2764.9895 | 0.016506 | 84.4484 | 134.455 |  |  |  |  | 3711.1389 | 0.067137
|  |  | 84.523 | 43.665 |  |  |  |  | 3714.4086 | 0.027895 | 84.523 | 45.655 |  |  |  |  | 2767.4256 | 0.016418 | 84.523 | 134.545 |  |  |  |  | 3714.4086 | 0.079405
|  |  | 84.5977 | 42.145 |  |  |  |  | 3717.6784 | 0.0522 | 84.5977 | 45.04 |  |  |  |  | 2769.8617 | 0.018768 | 84.5977 | 134.725 |  |  |  |  | 3717.6784 | 0.063487
|  |  | 84.6724 | 43.24 |  |  |  |  | 3720.9481 | 0.08427 | 84.6724 | 44.66 |  |  |  |  | 2772.2978 | 0.019171 | 84.6724 | 134.645 |  |  |  |  | 3720.9481 | 0.045714
|  |  | 84.747 | 42.445 |  |  |  |  | 3724.2178 | 0.076045 | 84.747 | 44.52 |  |  |  |  | 2774.7339 | 0.015168 | 84.747 | 134.65 |  |  |  |  | 3724.2178 | 0.03561
|  |  | 84.8217 | 42.915 |  |  |  |  | 3727.4875 | 0.041324 | 84.8217 | 44.965 |  |  |  |  | 2777.1701 | 0.0084701 | 84.8217 | 135.05 |  |  |  |  | 3727.4875 | 0.014284
|  |  | 84.8964 | 42.95 |  |  |  |  | 3730.7573 | 0.039616 | 84.8964 | 45.395 |  |  |  |  | 2779.6062 | 0.006264 | 84.8964 | 135.19 |  |  |  |  | 3730.7573 | 0.023189
|  |  | 84.971 | 43.285 |  |  |  |  | 3734.027 | 0.088701 | 84.971 | 45.785 |  |  |  |  | 2782.0423 | 0.0089394 | 84.971 | 134.66 |  |  |  |  | 3734.027 | 0.047168
|  |  | 85.0457 | 43.275 |  |  |  |  | 3737.2967 | 0.12777 | 85.0457 | 46.055 |  |  |  |  | 2784.4784 | 0.0079794 | 85.0457 | 134.3 |  |  |  |  | 3737.2967 | 0.049514
|  |  | 85.1204 | 41.345 |  |  |  |  | 3740.5664 | 0.12936 | 85.1204 | 45.995 |  |  |  |  | 2786.9145 | 0.0031833 | 85.1204 | 134.105 |  |  |  |  | 3740.5664 | 0.036307
|  |  | 85.195 | 43.575 |  |  |  |  | 3743.8362 | 0.089574 | 85.195 | 45.915 |  |  |  |  | 2789.3506 | 0.0052108 | 85.195 | 133.9 |  |  |  |  | 3743.8362 | 0.027292
|  |  | 85.2697 | 43.2 |  |  |  |  | 3747.1059 | 0.035687 | 85.2697 | 45.99 |  |  |  |  | 2791.7867 | 0.0095531 | 85.2697 | 134.375 |  |  |  |  | 3747.1059 | 0.021326
|  |  | 85.3444 | 42.78 |  |  |  |  | 3750.3756 | 0.034627 | 85.3444 | 45.985 |  |  |  |  | 2794.2229 | 0.0090127 | 85.3444 | 134.515 |  |  |  |  | 3750.3756 | 0.0064958
|  |  | 85.419 | 43.19 |  |  |  |  | 3753.6454 | 0.057519 | 85.419 | 45.89 |  |  |  |  | 2796.659 | 0.0035709 | 85.419 | 133.93 |  |  |  |  | 3753.6454 | 0.02085
|  |  | 85.4937 | 40.88 |  |  |  |  | 3756.9151 | 0.067042 | 85.4937 | 45.555 |  |  |  |  | 2799.0951 | 0.0061888 | 85.4937 | 134.1 |  |  |  |  | 3756.9151 | 0.039217
|  |  | 85.5684 | 41.91 |  |  |  |  | 3760.1848 | 0.057705 | 85.5684 | 45.31 |  |  |  |  | 2801.5312 | 0.013314 | 85.5684 | 134.15 |  |  |  |  | 3760.1848 | 0.045503
|  |  | 85.6431 | 41.23 |  |  |  |  | 3763.4545 | 0.036052 | 85.6431 | 45.07 |  |  |  |  | 2803.9673 | 0.015625 | 85.6431 | 134.135 |  |  |  |  | 3763.4545 | 0.037965
|  |  | 85.7177 | 41.345 |  |  |  |  | 3766.7243 | 0.052523 | 85.7177 | 44.65 |  |  |  |  | 2806.4034 | 0.011694 | 85.7177 | 134.18 |  |  |  |  | 3766.7243 | 0.022501
|  |  | 85.7924 | 41.72 |  |  |  |  | 3769.994 | 0.077241 | 85.7924 | 44.305 |  |  |  |  | 2808.8395 | 0.0041218 | 85.7924 | 134.285 |  |  |  |  | 3769.994 | 0.026136
|  |  | 85.8671 | 42.655 |  |  |  |  | 3773.2637 | 0.076913 | 85.8671 | 44.055 |  |  |  |  | 2811.2757 | 0.0078538 | 85.8671 | 134.365 |  |  |  |  | 3773.2637 | 0.039703
|  |  | 85.9417 | 42.84 |  |  |  |  | 3776.5334 | 0.058642 | 85.9417 | 43.575 |  |  |  |  | 2813.7118 | 0.0134 | 85.9417 | 134.58 |  |  |  |  | 3776.5334 | 0.035918
|  |  | 86.0164 | 42.465 |  |  |  |  | 3779.8032 | 0.031291 | 86.0164 | 43.33 |  |  |  |  | 2816.1479 | 0.012868 | 86.0164 | 134.155 |  |  |  |  | 3779.8032 | 0.017381
|  |  | 86.091 | 41.995 |  |  |  |  | 3783.0729 | 0.0090349 | 86.091 | 43.455 |  |  |  |  | 2818.584 | 0.0063668 | 86.091 | 133.88 |  |  |  |  | 3783.0729 | 0.024164
|  |  | 86.1657 | 40.285 |  |  |  |  | 3786.3426 | 0.050977 | 86.1657 | 43.915 |  |  |  |  | 2821.0201 | 0.0034829 | 86.1657 | 134.615 |  |  |  |  | 3786.3426 | 0.034409
|  |  | 86.2404 | 41.475 |  |  |  |  | 3789.6123 | 0.073653 | 86.2404 | 44.23 |  |  |  |  | 2823.4562 | 0.01133 | 86.2404 | 135.09 |  |  |  |  | 3789.6123 | 0.026022
|  |  | 86.315 | 42.655 |  |  |  |  | 3792.8821 | 0.070298 | 86.315 | 44.375 |  |  |  |  | 2825.8923 | 0.014339 | 86.315 | 135.07 |  |  |  |  | 3792.8821 | 0.027518
|  |  | 86.3897 | 43.075 |  |  |  |  | 3796.1518 | 0.06297 | 86.3897 | 44.51 |  |  |  |  | 2828.3285 | 0.011798 | 86.3897 | 135.17 |  |  |  |  | 3796.1518 | 0.044963
|  |  | 86.4644 | 42.085 |  |  |  |  | 3799.4215 | 0.058153 | 86.4644 | 44.605 |  |  |  |  | 2830.7646 | 0.0078307 | 86.4644 | 135.465 |  |  |  |  | 3799.4215 | 0.048238
|  |  | 86.5391 | 41.525 |  |  |  |  | 3802.6912 | 0.043251 | 86.5391 | 44.68 |  |  |  |  | 2833.2007 | 0.009676 | 86.5391 | 136.12 |  |  |  |  | 3802.6912 | 0.041452
|  |  | 86.6137 | 42.53 |  |  |  |  | 3805.961 | 0.049182 | 86.6137 | 44.56 |  |  |  |  | 2835.6368 | 0.012017 | 86.6137 | 136.555 |  |  |  |  | 3805.961 | 0.039519
|  |  | 86.6884 | 41.42 |  |  |  |  | 3809.2307 | 0.056962 | 86.6884 | 44.3 |  |  |  |  | 2838.0729 | 0.010102 | 86.6884 | 136.74 |  |  |  |  | 3809.2307 | 0.030263
|  |  | 86.7631 | 40.595 |  |  |  |  | 3812.5004 | 0.031431 | 86.7631 | 44.105 |  |  |  |  | 2840.509 | 0.0077595 | 86.7631 | 136.325 |  |  |  |  | 3812.5004 | 0.0034443
|  |  | 86.8377 | 40.945 |  |  |  |  | 3815.7701 | 0.059712 | 86.8377 | 43.975 |  |  |  |  | 2842.9451 | 0.012177 | 86.8377 | 136.32 |  |  |  |  | 3815.7701 | 0.029206
|  |  | 86.9124 | 41.075 |  |  |  |  | 3819.0399 | 0.10836 | 86.9124 | 43.79 |  |  |  |  | 2845.3812 | 0.016114 | 86.9124 | 136.475 |  |  |  |  | 3819.0399 | 0.043065
|  |  | 86.9871 | 41.66 |  |  |  |  | 3822.3096 | 0.10732 | 86.9871 | 43.41 |  |  |  |  | 2847.8174 | 0.014035 | 86.9871 | 136.015 |  |  |  |  | 3822.3096 | 0.029706
|  |  | 87.0617 | 43.205 |  |  |  |  | 3825.5793 | 0.069032 | 87.0617 | 42.975 |  |  |  |  | 2850.2535 | 0.0058576 | 87.0617 | 136.155 |  |  |  |  | 3825.5793 | 0.025316
|  |  | 87.1364 | 41.935 |  |  |  |  | 3828.849 | 0.067405 | 87.1364 | 42.825 |  |  |  |  | 2852.6896 | 0.0075391 | 87.1364 | 136.89 |  |  |  |  | 3828.849 | 0.047209
|  |  | 87.2111 | 40.86 |  |  |  |  | 3832.1188 | 0.072226 | 87.2111 | 42.595 |  |  |  |  | 2855.1257 | 0.017535 | 87.2111 | 137.345 |  |  |  |  | 3832.1188 | 0.043372
|  |  | 87.2857 | 39.91 |  |  |  |  | 3835.3885 | 0.06268 | 87.2857 | 42.24 |  |  |  |  | 2857.5618 | 0.021647 | 87.2857 | 137.28 |  |  |  |  | 3835.3885 | 0.011152
|  |  | 87.3604 | 40.65 |  |  |  |  | 3838.6582 | 0.13285 | 87.3604 | 41.81 |  |  |  |  | 2859.9979 | 0.017821 | 87.3604 | 137.305 |  |  |  |  | 3838.6582 | 0.042143
|  |  | 87.4351 | 41.005 |  |  |  |  | 3841.928 | 0.21425 | 87.4351 | 41.88 |  |  |  |  | 2862.434 | 0.0077882 | 87.4351 | 138.175 |  |  |  |  | 3841.928 | 0.073795
|  |  | 87.5097 | 41.075 |  |  |  |  | 3845.1977 | 0.25461 | 87.5097 | 42.125 |  |  |  |  | 2864.8702 | 0.0039586 | 87.5097 | 138.77 |  |  |  |  | 3845.1977 | 0.079236
|  |  | 87.5844 | 42.525 |  |  |  |  | 3848.4674 | 0.25355 | 87.5844 | 42.585 |  |  |  |  | 2867.3063 | 0.012105 | 87.5844 | 138.275 |  |  |  |  | 3848.4674 | 0.07191
|  |  | 87.6591 | 42.385 |  |  |  |  | 3851.7371 | 0.22176 | 87.6591 | 42.925 |  |  |  |  | 2869.7424 | 0.013225 | 87.6591 | 137.86 |  |  |  |  | 3851.7371 | 0.065283
|  |  | 87.7337 | 41.68 |  |  |  |  | 3855.0069 | 0.15808 | 87.7337 | 43.17 |  |  |  |  | 2872.1785 | 0.0072406 | 87.7337 | 137.885 |  |  |  |  | 3855.0069 | 0.046719
|  |  | 87.8084 | 41.86 |  |  |  |  | 3858.2766 | 0.070412 | 87.8084 | 43.465 |  |  |  |  | 2874.6146 | 0.0040842 | 87.8084 | 138.13 |  |  |  |  | 3858.2766 | 0.022908
|  |  | 87.8831 | 42.18 |  |  |  |  | 3861.5463 | 0.013876 | 87.8831 | 43.64 |  |  |  |  | 2877.0507 | 0.013021 | 87.8831 | 137.865 |  |  |  |  | 3861.5463 | 0.053575
|  |  | 87.9577 | 42.04 |  |  |  |  | 3864.816 | 0.058644 | 87.9577 | 43.835 |  |  |  |  | 2879.4868 | 0.017206 | 87.9577 | 137.765 |  |  |  |  | 3864.816 | 0.076547
|  |  | 88.0324 | 41.525 |  |  |  |  | 3868.0858 | 0.068917 | 88.0324 | 44.025 |  |  |  |  | 2881.923 | 0.014653 | 88.0324 | 137.58 |  |  |  |  | 3868.0858 | 0.064597
|  |  | 88.1071 | 41.595 |  |  |  |  | 3871.3555 | 0.065548 | 88.1071 | 44.045 |  |  |  |  | 2884.3591 | 0.0081814 | 88.1071 | 137.32 |  |  |  |  | 3871.3555 | 0.026896
|  |  | 88.1817 | 41.315 |  |  |  |  | 3874.6252 | 0.054275 | 88.1817 | 43.98 |  |  |  |  | 2886.7952 | 0.0088094 | 88.1817 | 137.505 |  |  |  |  | 3874.6252 | 0.017907
|  |  | 88.2564 | 41.83 |  |  |  |  | 3877.8949 | 0.036958 | 88.2564 | 44.025 |  |  |  |  | 2889.2313 | 0.014124 | 88.2564 | 137.655 |  |  |  |  | 3877.8949 | 0.033616
|  |  | 88.3311 | 41.84 |  |  |  |  | 3881.1647 | 0.045146 | 88.3311 | 44.105 |  |  |  |  | 2891.6674 | 0.014744 | 88.3311 | 137.835 |  |  |  |  | 3881.1647 | 0.031922
|  |  | 88.4057 | 41.285 |  |  |  |  | 3884.4344 | 0.052741 | 88.4057 | 44.06 |  |  |  |  | 2894.1035 | 0.010441 | 88.4057 | 137.575 |  |  |  |  | 3884.4344 | 0.043592
|  |  | 88.4804 | 41.075 |  |  |  |  | 3887.7041 | 0.032459 | 88.4804 | 43.905 |  |  |  |  | 2896.5396 | 0.0079613 | 88.4804 | 137.485 |  |  |  |  | 3887.7041 | 0.056893
|  |  | 88.5551 | 40.535 |  |  |  |  | 3890.9738 | 0.0042916 | 88.5551 | 43.625 |  |  |  |  | 2898.9758 | 0.012538 | 88.5551 | 137.48 |  |  |  |  | 3890.9738 | 0.047135
|  |  | 88.6297 | 42.29 |  |  |  |  | 3894.2436 | 0.031013 | 88.6297 | 43.405 |  |  |  |  | 2901.4119 | 0.015439 | 88.6297 | 137.515 |  |  |  |  | 3894.2436 | 0.01786
|  |  | 88.7044 | 41.615 |  |  |  |  | 3897.5133 | 0.052659 | 88.7044 | 43.36 |  |  |  |  | 2903.848 | 0.013459 | 88.7044 | 137.675 |  |  |  |  | 3897.5133 | 0.0129
|  |  | 88.7791 | 40.7 |  |  |  |  | 3900.783 | 0.091466 | 88.7791 | 43.415 |  |  |  |  | 2906.2841 | 0.008302 | 88.7791 | 137.87 |  |  |  |  | 3900.783 | 0.024386
|  |  | 88.8537 | 41.2 |  |  |  |  | 3904.0527 | 0.12757 | 88.8537 | 43.745 |  |  |  |  | 2908.7202 | 0.0055849 | 88.8537 | 138.06 |  |  |  |  | 3904.0527 | 0.030538
|  |  | 88.9284 | 41.27 |  |  |  |  | 3907.3225 | 0.12706 | 88.9284 | 44.725 |  |  |  |  | 2911.1563 | 0.0068989 | 88.9284 | 137.75 |  |  |  |  | 3907.3225 | 0.047084
|  |  | 89.0031 | 41.14 |  |  |  |  | 3910.5922 | 0.079756 | 89.0031 | 45.23 |  |  |  |  | 2913.5924 | 0.0056634 | 89.0031 | 137.67 |  |  |  |  | 3910.5922 | 0.050477
|  |  | 89.0777 | 41.21 |  |  |  |  | 3913.8619 | 0.021937 | 89.0777 | 45.525 |  |  |  |  | 2916.0286 | 0.0012915 | 89.0777 | 137.705 |  |  |  |  | 3913.8619 | 0.026091
|  |  | 89.1524 | 40.595 |  |  |  |  | 3917.1316 | 0.071853 | 89.1524 | 45.675 |  |  |  |  | 2918.4647 | 0.0052289 | 89.1524 | 137.15 |  |  |  |  | 3917.1316 | 0.018658
|  |  | 89.2271 | 40.635 |  |  |  |  | 3920.4014 | 0.098651 | 89.2271 | 46.005 |  |  |  |  | 2920.9008 | 0.009253 | 89.2271 | 136.64 |  |  |  |  | 3920.4014 | 0.051331
|  |  | 89.3017 | 40.515 |  |  |  |  | 3923.6711 | 0.079185 | 89.3017 | 46.38 |  |  |  |  | 2923.3369 | 0.0096631 | 89.3017 | 136.32 |  |  |  |  | 3923.6711 | 0.054974
|  |  | 89.3764 | 42.645 |  |  |  |  | 3926.9408 | 0.034118 | 89.3764 | 46.565 |  |  |  |  | 2925.773 | 0.0073017 | 89.3764 | 135.73 |  |  |  |  | 3926.9408 | 0.030792
|  |  | 89.4511 | 42.36 |  |  |  |  | 3930.2106 | 0.018388 | 89.4511 | 46.62 |  |  |  |  | 2928.2091 | 0.0062327 | 89.4511 | 135.665 |  |  |  |  | 3930.2106 | 0.016751
|  |  | 89.5257 | 41.82 |  |  |  |  | 3933.4803 | 0.030561 | 89.5257 | 46.715 |  |  |  |  | 2930.6452 | 0.0073754 | 89.5257 | 135.765 |  |  |  |  | 3933.4803 | 0.026436
|  |  | 89.6004 | 41.845 |  |  |  |  | 3936.75 | 0.020601 | 89.6004 | 46.82 |  |  |  |  | 2933.0814 | 0.0061689 | 89.6004 | 135.785 |  |  |  |  | 3936.75 | 0.022447
|  |  | 89.6751 | 40.785 |  |  |  |  | 3940.0197 | 0.024396 | 89.6751 | 46.78 |  |  |  |  | 2935.5175 | 0.0013828 | 89.6751 | 136.22 |  |  |  |  | 3940.0197 | 0.039605
|  |  | 89.7497 | 42.295 |  |  |  |  | 3943.2895 | 0.061533 | 89.7497 | 46.655 |  |  |  |  | 2937.9536 | 0.0060872 | 89.7497 | 136.14 |  |  |  |  | 3943.2895 | 0.056236
|  |  | 89.8244 | 42.98 |  |  |  |  | 3946.5592 | 0.097617 | 89.8244 | 46.605 |  |  |  |  | 2940.3897 | 0.012239 | 89.8244 | 136.17 |  |  |  |  | 3946.5592 | 0.044928
|  |  | 89.8991 | 42.275 |  |  |  |  | 3949.8289 | 0.10948 | 89.8991 | 46.495 |  |  |  |  | 2942.8258 | 0.014746 | 89.8991 | 136.395 |  |  |  |  | 3949.8289 | 0.023706
|  |  | 89.9737 | 41.775 |  |  |  |  | 3953.0986 | 0.087151 | 89.9737 | 46.36 |  |  |  |  | 2945.2619 | 0.012731 | 89.9737 | 136.315 |  |  |  |  | 3953.0986 | 0.046374
|  |  | 90.0484 | 43.18 |  |  |  |  | 3956.3684 | 0.066408 | 90.0484 | 46.305 |  |  |  |  | 2947.698 | 0.0077548 | 90.0484 | 136.69 |  |  |  |  | 3956.3684 | 0.0581
|  |  | 90.1231 | 42.9 |  |  |  |  | 3959.6381 | 0.093453 | 90.1231 | 46.28 |  |  |  |  | 2950.1342 | 0.0034823 | 90.1231 | 136.515 |  |  |  |  | 3959.6381 | 0.043764
|  |  | 90.1977 | 41.62 |  |  |  |  | 3962.9078 | 0.10987 | 90.1977 | 46.395 |  |  |  |  | 2952.5703 | 0.0027966 | 90.1977 | 136.18 |  |  |  |  | 3962.9078 | 0.032124
|  |  | 90.2724 | 41.785 |  |  |  |  | 3966.1775 | 0.08765 | 90.2724 | 46.66 |  |  |  |  | 2955.0064 | 0.0026035 | 90.2724 | 135.62 |  |  |  |  | 3966.1775 | 0.034585
|  |  | 90.3471 | 42.495 |  |  |  |  | 3969.4473 | 0.047088 | 90.3471 | 47.81 |  |  |  |  | 2957.4425 | 0.0070472 | 90.3471 | 135.595 |  |  |  |  | 3969.4473 | 0.025389
|  |  | 90.4217 | 41.295 |  |  |  |  | 3972.717 | 0.035614 | 90.4217 | 48.49 |  |  |  |  | 2959.8786 | 0.012825 | 90.4217 | 135.785 |  |  |  |  | 3972.717 | 0.040878
|  |  | 90.4964 | 42.105 |  |  |  |  | 3975.9867 | 0.051178 | 90.4964 | 48.99 |  |  |  |  | 2962.3147 | 0.015548 | 90.4964 | 135.625 |  |  |  |  | 3975.9867 | 0.068178
|  |  | 90.5711 | 42.12 |  |  |  |  | 3979.2564 | 0.058709 | 90.5711 | 49.285 |  |  |  |  | 2964.7508 | 0.012792 | 90.5711 | 135.29 |  |  |  |  | 3979.2564 | 0.070708
|  |  | 90.6457 | 42.295 |  |  |  |  | 3982.5262 | 0.051225 | 90.6457 | 49.54 |  |  |  |  | 2967.187 | 0.0047405 | 90.6457 | 134.91 |  |  |  |  | 3982.5262 | 0.049105
|  |  | 90.7204 | 42.545 |  |  |  |  | 3985.7959 | 0.039265 | 90.7204 | 49.675 |  |  |  |  | 2969.6231 | 0.0058363 | 90.7204 | 134.38 |  |  |  |  | 3985.7959 | 0.029883
|  |  | 90.7951 | 41.79 |  |  |  |  | 3989.0656 | 0.045083 | 90.7951 | 49.895 |  |  |  |  | 2972.0592 | 0.014762 | 90.7951 | 134.345 |  |  |  |  | 3989.0656 | 0.03009
|  |  | 90.8697 | 42.45 |  |  |  |  | 3992.3353 | 0.053783 | 90.8697 | 49.965 |  |  |  |  | 2974.4953 | 0.018463 | 90.8697 | 134.355 |  |  |  |  | 3992.3353 | 0.043828
|  |  | 90.9444 | 42.675 |  |  |  |  | 3995.6051 | 0.081606 | 90.9444 | 49.905 |  |  |  |  | 2976.9314 | 0.015649 | 90.9444 | 134.19 |  |  |  |  | 3995.6051 | 0.068349
|  |  | 91.0191 | 42.77 |  |  |  |  | 3998.8748 | 0.12694 | 91.0191 | 49.745 |  |  |  |  | 2979.3675 | 0.0082207 | 91.0191 | 134.185 |  |  |  |  | 3998.8748 | 0.07741
|  |  | 91.0937 | 43.48 | [6 empty] | 91.0937 | 49.705 |  |  |  |  | 2981.8036 | 0.0048111 | 91.0937 | 134.18
|  |  | 91.1684 | 42.67 | [6 empty] | 91.1684 | 49.57 |  |  |  |  | 2984.2398 | 0.0099596 | 91.1684 | 133.985
|  |  | 91.2431 | 43.045 | [6 empty] | 91.2431 | 49.315 |  |  |  |  | 2986.6759 | 0.011384 | 91.2431 | 133.32
|  |  | 91.3177 | 42.425 | [6 empty] | 91.3177 | 48.815 |  |  |  |  | 2989.112 | 0.0088206 | 91.3177 | 132.895
|  |  | 91.3924 | 41.79 | [6 empty] | 91.3924 | 48.54 |  |  |  |  | 2991.5481 | 0.007574 | 91.3924 | 132.875
|  |  | 91.4671 | 42.93 | [6 empty] | 91.4671 | 48.48 |  |  |  |  | 2993.9842 | 0.010029 | 91.4671 | 133.115
|  |  | 91.5417 | 43.46 | [6 empty] | 91.5417 | 48.79 |  |  |  |  | 2996.4203 | 0.010173 | 91.5417 | 133.615
|  |  | 91.6164 | 44.075 | [6 empty] | 91.6164 | 49.14 |  |  |  |  | 2998.8564 | 0.0058155 | 91.6164 | 133.585
|  |  | 91.6911 | 43.16 | [6 empty] | 91.6911 | 49.575 |  |  |  |  | 3001.2926 | 0.0038496 | 91.6911 | 133.285
|  |  | 91.7657 | 43.47 | [6 empty] | 91.7657 | 50.075 |  |  |  |  | 3003.7287 | 0.010872 | 91.7657 | 132.675
|  |  | 91.8404 | 43.155 | [6 empty] | 91.8404 | 50.22 |  |  |  |  | 3006.1648 | 0.014819 | 91.8404 | 132.05
|  |  | 91.9151 | 43.545 | [6 empty] | 91.9151 | 50.245 |  |  |  |  | 3008.6009 | 0.013374 | 91.9151 | 132.395
|  |  | 91.9897 | 43.53 | [6 empty] | 91.9897 | 50.175 |  |  |  |  | 3011.037 | 0.0076613 | 91.9897 | 132.9
|  |  | 92.0644 | 42.93 | [6 empty] | 92.0644 | 50.095 |  |  |  |  | 3013.4731 | 0.0051371 | 92.0644 | 132.615
|  |  | 92.1391 | 43.44 | [6 empty] | 92.1391 | 50.135 |  |  |  |  | 3015.9092 | 0.010187 | 92.1391 | 131.495
|  |  | 92.2137 | 43.71 | [6 empty] | 92.2137 | 50.165 |  |  |  |  | 3018.3453 | 0.012679 | 92.2137 | 131.015
|  |  | 92.2884 | 42.83 | [6 empty] | 92.2884 | 49.965 |  |  |  |  | 3020.7815 | 0.011384 | 92.2884 | 130.785
|  |  | 92.3631 | 42.555 | [6 empty] | 92.3631 | 49.525 |  |  |  |  | 3023.2176 | 0.0089147 | 92.3631 | 130.9
|  |  | 92.4377 | 42.055 | [6 empty] | 92.4377 | 49.07 |  |  |  |  | 3025.6537 | 0.0082898 | 92.4377 | 131.455
|  |  | 92.5124 | 42.73 | [6 empty] | 92.5124 | 48.62 |  |  |  |  | 3028.0898 | 0.007715 | 92.5124 | 131.88
|  |  | 92.5871 | 43.03 | [6 empty] | 92.5871 | 48.31 |  |  |  |  | 3030.5259 | 0.0045965 | 92.5871 | 131.73
|  |  | 92.6617 | 43.57 | [6 empty] | 92.6617 | 47.87 |  |  |  |  | 3032.962 | 0.0021057 | 92.6617 | 131.83
|  |  | 92.7364 | 42.925 | [6 empty] | 92.7364 | 47.295 |  |  |  |  | 3035.3981 | 0.0072193 | 92.7364 | 132.12
|  |  | 92.8111 | 43.15 | [6 empty] | 92.8111 | 46.87 |  |  |  |  | 3037.8343 | 0.011031 | 92.8111 | 132.145
|  |  | 92.8857 | 43.055 | [6 empty] | 92.8857 | 46.47 |  |  |  |  | 3040.2704 | 0.012616 | 92.8857 | 131.985
|  |  | 92.9604 | 43.05 | [6 empty] | 92.9604 | 46.18 |  |  |  |  | 3042.7065 | 0.013129 | 92.9604 | 131.99
|  |  | 93.0351 | 42.89 | [6 empty] | 93.0351 | 46.425 |  |  |  |  | 3045.1426 | 0.013083 | 93.0351 | 131.955
|  |  | 93.1097 | 42.39 | [6 empty] | 93.1097 | 46.81 |  |  |  |  | 3047.5787 | 0.010949 | 93.1097 | 131.91
|  |  | 93.1844 | 43.415 | [6 empty] | 93.1844 | 47.04 |  |  |  |  | 3050.0148 | 0.0053745 | 93.1844 | 131.985
|  |  | 93.2591 | 42.985 | [6 empty] | 93.2591 | 47.12 |  |  |  |  | 3052.4509 | 0.0028437 | 93.2591 | 132.055
|  |  | 93.3337 | 42.265 | [6 empty] | 93.3337 | 47.085 |  |  |  |  | 3054.8871 | 0.010851 | 93.3337 | 131.705
|  |  | 93.4084 | 43.18 | [6 empty] | 93.4084 | 47.115 |  |  |  |  | 3057.3232 | 0.015255 | 93.4084 | 131.22
|  |  | 93.4831 | 43.87 | [6 empty] | 93.4831 | 47.18 |  |  |  |  | 3059.7593 | 0.01395 | 93.4831 | 130.99
|  |  | 93.5577 | 43.03 | [6 empty] | 93.5577 | 47.095 |  |  |  |  | 3062.1954 | 0.0077582 | 93.5577 | 131.7
|  |  | 93.6324 | 42.025 | [6 empty] | 93.6324 | 46.92 |  |  |  |  | 3064.6315 | 0.0060423 | 93.6324 | 131.86
|  |  | 93.7071 | 42.38 | [6 empty] | 93.7071 | 46.73 |  |  |  |  | 3067.0676 | 0.012965 | 93.7071 | 131.795
|  |  | 93.7818 | 44.08 | [6 empty] | 93.7818 | 46.53 |  |  |  |  | 3069.5037 | 0.016332 | 93.7818 | 131.68
|  |  | 93.8564 | 43.015 | [6 empty] | 93.8564 | 46.34 |  |  |  |  | 3071.9399 | 0.014313 | 93.8564 | 131.58
|  |  | 93.9311 | 43.12 | [6 empty] | 93.9311 | 45.925 |  |  |  |  | 3074.376 | 0.010112 | 93.9311 | 131.67
|  |  | 94.0058 | 43.425 | [6 empty] | 94.0058 | 45.635 |  |  |  |  | 3076.8121 | 0.010549 | 94.0058 | 132.465
|  |  | 94.0804 | 43.355 | [6 empty] | 94.0804 | 45.41 |  |  |  |  | 3079.2482 | 0.013049 | 94.0804 | 133.08
|  |  | 94.1551 | 41.78 | [6 empty] | 94.1551 | 44.9 |  |  |  |  | 3081.6843 | 0.011594 | 94.1551 | 133.465
|  |  | 94.2298 | 41.31 | [6 empty] | 94.2298 | 44.335 |  |  |  |  | 3084.1204 | 0.0065815 | 94.2298 | 133.23
|  |  | 94.3044 | 41.3 | [6 empty] | 94.3044 | 43.925 |  |  |  |  | 3086.5565 | 0.0074939 | 94.3044 | 133.035
|  |  | 94.3791 | 42.17 | [6 empty] | 94.3791 | 43.755 |  |  |  |  | 3088.9927 | 0.012964 | 94.3791 | 133.02
|  |  | 94.4538 | 42.395 | [6 empty] | 94.4538 | 44.06 |  |  |  |  | 3091.4288 | 0.014411 | 94.4538 | 133.375
|  |  | 94.5284 | 42.935 | [6 empty] | 94.5284 | 44.495 |  |  |  |  | 3093.8649 | 0.011408 | 94.5284 | 133.815
|  |  | 94.6031 | 42.725 | [6 empty] | 94.6031 | 44.895 |  |  |  |  | 3096.301 | 0.0098804 | 94.6031 | 134.34
|  |  | 94.6778 | 43.255 | [6 empty] | 94.6778 | 45.15 |  |  |  |  | 3098.7371 | 0.01401 | 94.6778 | 134.925
|  |  | 94.7524 | 42.06 | [6 empty] | 94.7524 | 45.265 |  |  |  |  | 3101.1732 | 0.016502 | 94.7524 | 134.735
|  |  | 94.8271 | 42.47 | [6 empty] | 94.8271 | 45.295 |  |  |  |  | 3103.6093 | 0.013444 | 94.8271 | 134.83
|  |  | 94.9018 | 41.67 | [6 empty] | 94.9018 | 45.245 |  |  |  |  | 3106.0455 | 0.0057973 | 94.9018 | 135.61
|  |  | 94.9764 | 40.625 | [6 empty] | 94.9764 | 45.225 |  |  |  |  | 3108.4816 | 0.0047775 | 94.9764 | 136.065
|  |  | 95.0511 | 40.325 | [6 empty] | 95.0511 | 44.925 |  |  |  |  | 3110.9177 | 0.01059 | 95.0511 | 135.99
|  |  | 95.1258 | 42.18 | [6 empty] | 95.1258 | 44.605 |  |  |  |  | 3113.3538 | 0.010783 | 95.1258 | 135.595
|  |  | 95.2004 | 41.935 | [6 empty] | 95.2004 | 44.32 |  |  |  |  | 3115.7899 | 0.0052207 | 95.2004 | 135.865
|  |  | 95.2751 | 42.25 | [6 empty] | 95.2751 | 44.17 |  |  |  |  | 3118.226 | 0.0053406 | 95.2751 | 136.28
|  |  | 95.3498 | 41.52 | [6 empty] | 95.3498 | 43.91 |  |  |  |  | 3120.6621 | 0.013135 | 95.3498 | 137.21
|  |  | 95.4244 | 41.58 | [6 empty] | 95.4244 | 43.505 |  |  |  |  | 3123.0983 | 0.016545 | 95.4244 | 137.65
|  |  | 95.4991 | 40.97 | [6 empty] | 95.4991 | 43.01 |  |  |  |  | 3125.5344 | 0.014218 | 95.4991 | 137.78
|  |  | 95.5738 | 40.965 | [6 empty] | 95.5738 | 42.405 |  |  |  |  | 3127.9705 | 0.0092832 | 95.5738 | 138.245
|  |  | 95.6484 | 41.525 | [6 empty] | 95.6484 | 42 |  |  |  |  | 3130.4066 | 0.0096864 | 95.6484 | 138.395
|  |  | 95.7231 | 41.32 | [6 empty] | 95.7231 | 41.635 |  |  |  |  | 3132.8427 | 0.013052 | 95.7231 | 138.005
|  |  | 95.7978 | 42.05 | [6 empty] | 95.7978 | 41.57 |  |  |  |  | 3135.2788 | 0.012553 | 95.7978 | 137.385
|  |  | 95.8724 | 43.015 | [6 empty] | 95.8724 | 42.05 |  |  |  |  | 3137.7149 | 0.0077614 | 95.8724 | 137.67
|  |  | 95.9471 | 43.48 | [6 empty] | 95.9471 | 42.725 |  |  |  |  | 3140.1511 | 0.0026599 | 95.9471 | 138.555
|  |  | 96.0218 | 42.535 | [6 empty] | 96.0218 | 43.52 |  |  |  |  | 3142.5872 | 0.0044034 | 96.0218 | 139.085
|  |  | 96.0964 | 42.21 | [6 empty] | 96.0964 | 43.745 |  |  |  |  | 3145.0233 | 0.0046083 | 96.0964 | 139.205
|  |  | 96.1711 | 41.135 | [6 empty] | 96.1711 | 43.845 |  |  |  |  | 3147.4594 | 0.0069786 | 96.1711 | 138.79
|  |  | 96.2458 | 41.765 | [6 empty] | 96.2458 | 43.87 |  |  |  |  | 3149.8955 | 0.014956 | 96.2458 | 138.38
|  |  | 96.3204 | 41.71 | [6 empty] | 96.3204 | 43.85 |  |  |  |  | 3152.3316 | 0.022048 | 96.3204 | 138.565
|  |  | 96.3951 | 41.705 | [6 empty] | 96.3951 | 43.975 |  |  |  |  | 3154.7677 | 0.023982 | 96.3951 | 138.49
|  |  | 96.4698 | 42.5 | [6 empty] | 96.4698 | 44.095 |  |  |  |  | 3157.2039 | 0.019532 | 96.4698 | 138.735
|  |  | 96.5444 | 42.62 | [6 empty] | 96.5444 | 44.02 |  |  |  |  | 3159.64 | 0.011619 | 96.5444 | 139.375
|  |  | 96.6191 | 41.8 | [6 empty] | 96.6191 | 43.925 |  |  |  |  | 3162.0761 | 0.0097196 | 96.6191 | 139.44
|  |  | 96.6938 | 41.55 | [6 empty] | 96.6938 | 43.79 |  |  |  |  | 3164.5122 | 0.014086 | 96.6938 | 138.985
|  |  | 96.7684 | 42.54 | [6 empty] | 96.7684 | 43.64 |  |  |  |  | 3166.9483 | 0.014366 | 96.7684 | 138.23
|  |  | 96.8431 | 41.58 | [6 empty] | 96.8431 | 43.48 |  |  |  |  | 3169.3844 | 0.0089536 | 96.8431 | 138.095
|  |  | 96.9178 | 42.535 | [6 empty] | 96.9178 | 43.18 |  |  |  |  | 3171.8205 | 0.0012501 | 96.9178 | 138.27
|  |  | 96.9924 | 42.875 | [6 empty] | 96.9924 | 43.25 |  |  |  |  | 3174.2567 | 0.0070492 | 96.9924 | 138.42
|  |  | 97.0671 | 42.175 | [6 empty] | 97.0671 | 43.52 |  |  |  |  | 3176.6928 | 0.0098128 | 97.0671 | 138.64
|  |  | 97.1418 | 41.1 | [6 empty] | 97.1418 | 43.94 |  |  |  |  | 3179.1289 | 0.0075496 | 97.1418 | 138.775
|  |  | 97.2164 | 40.66 | [6 empty] | 97.2164 | 44.41 |  |  |  |  | 3181.565 | 0.0056812 | 97.2164 | 138.38
|  |  | 97.2911 | 41.98 | [6 empty] | 97.2911 | 44.855 |  |  |  |  | 3184.0011 | 0.010237 | 97.2911 | 138.075
|  |  | 97.3658 | 42.725 | [6 empty] | 97.3658 | 45.155 |  |  |  |  | 3186.4372 | 0.013188 | 97.3658 | 137.715
|  |  | 97.4404 | 42.795 | [6 empty] | 97.4404 | 45.5 |  |  |  |  | 3188.8733 | 0.011146 | 97.4404 | 137.055
|  |  | 97.5151 | 42.935 | [6 empty] | 97.5151 | 45.595 |  |  |  |  | 3191.3094 | 0.0049315 | 97.5151 | 137.015
|  |  | 97.5898 | 42.505 | [6 empty] | 97.5898 | 45.59 |  |  |  |  | 3193.7456 | 0.002466 | 97.5898 | 137.055
|  |  | 97.6644 | 41.72 | [6 empty] | 97.6644 | 45.6 |  |  |  |  | 3196.1817 | 0.0068281 | 97.6644 | 136.855
|  |  | 97.7391 | 41.465 | [6 empty] | 97.7391 | 45.655 |  |  |  |  | 3198.6178 | 0.006328 | 97.7391 | 136.24
|  |  | 97.8138 | 40.165 | [6 empty] | 97.8138 | 45.73 |  |  |  |  | 3201.0539 | 0.0041179 | 97.8138 | 135.285
|  |  | 97.8884 | 41.745 | [6 empty] | 97.8884 | 45.745 |  |  |  |  | 3203.49 | 0.0099363 | 97.8884 | 135.36
|  |  | 97.9631 | 42.135 | [6 empty] | 97.9631 | 45.545 |  |  |  |  | 3205.9261 | 0.015882 | 97.9631 | 136.055
|  |  | 98.0378 | 42.295 | [6 empty] | 98.0378 | 45.44 |  |  |  |  | 3208.3622 | 0.017068 | 98.0378 | 136.155
|  |  | 98.1124 | 42.3 | [6 empty] | 98.1124 | 45.4 |  |  |  |  | 3210.7984 | 0.012665 | 98.1124 | 136.27
|  |  | 98.1871 | 42.54 | [6 empty] | 98.1871 | 45.405 |  |  |  |  | 3213.2345 | 0.0059099 | 98.1871 | 136.295
|  |  | 98.2618 | 42.85 | [6 empty] | 98.2618 | 45.31 |  |  |  |  | 3215.6706 | 0.0078012 | 98.2618 | 136.295
|  |  | 98.3364 | 42 | [6 empty] | 98.3364 | 44.845 |  |  |  |  | 3218.1067 | 0.012395 | 98.3364 | 136.27
|  |  | 98.4111 | 41.62 | [6 empty] | 98.4111 | 44.745 |  |  |  |  | 3220.5428 | 0.01281 | 98.4111 | 136.245
|  |  | 98.4858 | 42.415 | [6 empty] | 98.4858 | 44.97 |  |  |  |  | 3222.9789 | 0.01034 | 98.4858 | 136.175
|  |  | 98.5604 | 42.695 | [6 empty] | 98.5604 | 45.44 |  |  |  |  | 3225.415 | 0.0091564 | 98.5604 | 135.85
|  |  | 98.6351 | 41.805 | [6 empty] | 98.6351 | 46.08 |  |  |  |  | 3227.8512 | 0.0093426 | 98.6351 | 135.69
|  |  | 98.7098 | 42.13 | [6 empty] | 98.7098 | 46.57 |  |  |  |  | 3230.2873 | 0.0068652 | 98.7098 | 135.915
|  |  | 98.7844 | 42.645 | [6 empty] | 98.7844 | 46.98 |  |  |  |  | 3232.7234 | 0.0019591 | 98.7844 | 135.865
|  |  | 98.8591 | 41.395 | [6 empty] | 98.8591 | 47.505 |  |  |  |  | 3235.1595 | 0.0072319 | 98.8591 | 135.415
|  |  | 98.9338 | 42.31 | [6 empty] | 98.9338 | 47.81 |  |  |  |  | 3237.5956 | 0.011976 | 98.9338 | 135.26
|  |  | 99.0084 | 42.48 | [6 empty] | 99.0084 | 48.02 |  |  |  |  | 3240.0317 | 0.011503 | 99.0084 | 135.075
|  |  | 99.0831 | 42.61 | [6 empty] | 99.0831 | 48.3 |  |  |  |  | 3242.4678 | 0.0056827 | 99.0831 | 135.24
|  |  | 99.1578 | 42.485 | [6 empty] | 99.1578 | 48.31 |  |  |  |  | 3244.904 | 0.0047761 | 99.1578 | 135.215
|  |  | 99.2324 | 42.61 | [6 empty] | 99.2324 | 48.29 |  |  |  |  | 3247.3401 | 0.012356 | 99.2324 | 135.035
|  |  | 99.3071 | 43.035 | [6 empty] | 99.3071 | 48.175 |  |  |  |  | 3249.7762 | 0.015281 | 99.3071 | 134.905
|  |  | 99.3818 | 42.735 | [6 empty] | 99.3818 | 48.13 |  |  |  |  | 3252.2123 | 0.01166 | 99.3818 | 134.755
|  |  | 99.4564 | 43.535 | [6 empty] | 99.4564 | 48.05 |  |  |  |  | 3254.6484 | 0.0030804 | 99.4564 | 134.875
|  |  | 99.5311 | 44.205 | [6 empty] | 99.5311 | 47.88 |  |  |  |  | 3257.0845 | 0.0067373 | 99.5311 | 134.665
|  |  | 99.6058 | 43.22 | [6 empty] | 99.6058 | 47.645 |  |  |  |  | 3259.5206 | 0.012928 | 99.6058 | 133.975
|  |  | 99.6804 | 43.285 | [6 empty] | 99.6804 | 47.335 |  |  |  |  | 3261.9568 | 0.012928 | 99.6804 | 133.805
|  |  | 99.7551 | 44.555 | [6 empty] | 99.7551 | 46.915 |  |  |  |  | 3264.3929 | 0.0075523 | 99.7551 | 134.115
|  |  | 99.8298 | 43.36 | [6 empty] | 99.8298 | 46.79 |  |  |  |  | 3266.829 | 0.0059964 | 99.8298 | 134.77
|  |  | 99.9044 | 43.02 | [6 empty] | 99.9044 | 46.75 |  |  |  |  | 3269.2651 | 0.011775 | 99.9044 | 135.285
|  |  | 99.9791 | 43.615 | [6 empty] | 99.9791 | 47.34 |  |  |  |  | 3271.7012 | 0.013376 | 99.9791 | 135.035
|  |  | 100.0538 | 43.715 | [6 empty] | 100.0538 | 47.96 |  |  |  |  | 3274.1373 | 0.0094231 | 100.0538 | 134.77
|  |  | 100.1284 | 43.025 | [6 empty] | 100.1284 | 48.46 |  |  |  |  | 3276.5734 | 0.0073839 | 100.1284 | 134.895
|  |  | 100.2031 | 43.805 | [6 empty] | 100.2031 | 48.795 |  |  |  |  | 3279.0096 | 0.013775 | 100.2031 | 134.7
|  |  | 100.2778 | 43.545 | [6 empty] | 100.2778 | 48.965 |  |  |  |  | 3281.4457 | 0.017472 | 100.2778 | 134.27
|  |  | 100.3524 | 43.445 | [6 empty] | 100.3524 | 48.94 |  |  |  |  | 3283.8818 | 0.014242 | 100.3524 | 133.63
|  |  | 100.4271 | 43.475 | [6 empty] | 100.4271 | 48.915 |  |  |  |  | 3286.3179 | 0.0048356 | 100.4271 | 133.1
|  |  | 100.5018 | 42.61 | [6 empty] | 100.5018 | 48.92 |  |  |  |  | 3288.754 | 0.0068 | 100.5018 | 133.36
|  |  | 100.5764 | 42.58 | [6 empty] | 100.5764 | 48.77 |  |  |  |  | 3291.1901 | 0.015498 | 100.5764 | 133.34
|  |  | 100.6511 | 43.855 | [6 empty] | 100.6511 | 48.6 |  |  |  |  | 3293.6262 | 0.017399 | 100.6511 | 132.92
|  |  | 100.7258 | 43.38 | [6 empty] | 100.7258 | 48.305 |  |  |  |  | 3296.0624 | 0.011746 | 100.7258 | 132.575
|  |  | 100.8004 | 42.49 | [6 empty] | 100.8004 | 48.08 |  |  |  |  | 3298.4985 | 0.0018628 | 100.8004 | 132.84
|  |  | 100.8751 | 43.425 | [6 empty] | 100.8751 | 47.925 |  |  |  |  | 3300.9346 | 0.0098295 | 100.8751 | 133.545
|  |  | 100.9498 | 43.645 | [6 empty] | 100.9498 | 47.45 |  |  |  |  | 3303.3707 | 0.015932 | 100.9498 | 134.05
|  |  | 101.0245 | 44.225 | [6 empty] | 101.0245 | 47.005 |  |  |  |  | 3305.8068 | 0.015074 | 101.0245 | 134.165
|  |  | 101.0991 | 43.25 | [6 empty] | 101.0991 | 46.585 |  |  |  |  | 3308.2429 | 0.0081344 | 101.0991 | 134.31
|  |  | 101.1738 | 42.95 | [6 empty] | 101.1738 | 45.94 |  |  |  |  | 3310.679 | 0.0028801 | 101.1738 | 135.07
|  |  | 101.2485 | 43.21 | [6 empty] | 101.2485 | 45.36 |  |  |  |  | 3313.1152 | 0.010355 | 101.2485 | 135.55
|  |  | 101.3231 | 43.495 | [6 empty] | 101.3231 | 44.855 |  |  |  |  | 3315.5513 | 0.01407 | 101.3231 | 135.305
|  |  | 101.3978 | 44.39 | [6 empty] | 101.3978 | 44.445 |  |  |  |  | 3317.9874 | 0.013646 | 101.3978 | 134.695
|  |  | 101.4725 | 43.63 | [6 empty] | 101.4725 | 44.82 |  |  |  |  | 3320.4235 | 0.012421 | 101.4725 | 134.705
|  |  | 101.5471 | 41.88 | [6 empty] | 101.5471 | 45.395 |  |  |  |  | 3322.8596 | 0.013139 | 101.5471 | 134.9
|  |  | 101.6218 | 42.195 | [6 empty] | 101.6218 | 45.845 |  |  |  |  | 3325.2957 | 0.01425 | 101.6218 | 134.29
|  |  | 101.6964 | 41.945 | [6 empty] | 101.6964 | 45.94 |  |  |  |  | 3327.7318 | 0.015119 | 101.6964 | 134.32
|  |  | 101.7711 | 43.57 | [6 empty] | 101.7711 | 46.095 |  |  |  |  | 3330.168 | 0.017313 | 101.7711 | 134.43
|  |  | 101.8458 | 43.995 | [6 empty] | 101.8458 | 46.08 |  |  |  |  | 3332.6041 | 0.019636 | 101.8458 | 133.865
|  |  | 101.9205 | 43.32 | [6 empty] | 101.9205 | 46.03 |  |  |  |  | 3335.0402 | 0.018665 | 101.9205 | 133.59
|  |  | 101.9951 | 43.665 | [6 empty] | 101.9951 | 46.06 |  |  |  |  | 3337.4763 | 0.012884 | 101.9951 | 133.68
|  |  | 102.0698 | 42.025 | [6 empty] | 102.0698 | 45.93 |  |  |  |  | 3339.9124 | 0.0051423 | 102.0698 | 133.925
|  |  | 102.1445 | 42.47 | [6 empty] | 102.1445 | 45.765 |  |  |  |  | 3342.3485 | 0.0083193 | 102.1445 | 134.315
|  |  | 102.2191 | 41.785 | [6 empty] | 102.2191 | 45.535 |  |  |  |  | 3344.7846 | 0.012888 | 102.2191 | 134.4
|  |  | 102.2938 | 42.115 | [6 empty] | 102.2938 | 45.295 |  |  |  |  | 3347.2208 | 0.011226 | 102.2938 | 134.19
|  |  | 102.3685 | 43.875 | [6 empty] | 102.3685 | 44.99 |  |  |  |  | 3349.6569 | 0.0036103 | 102.3685 | 134.275
|  |  | 102.4431 | 42.88 | [6 empty] | 102.4431 | 44.705 |  |  |  |  | 3352.093 | 0.0070636 | 102.4431 | 134.16
|  |  | 102.5178 | 42.81 | [6 empty] | 102.5178 | 44.31 |  |  |  |  | 3354.5291 | 0.015807 | 102.5178 | 133.895
|  |  | 102.5925 | 41.935 | [6 empty] | 102.5925 | 43.59 |  |  |  |  | 3356.9652 | 0.019313 | 102.5925 | 134.125
|  |  | 102.6671 | 42.12 | [6 empty] | 102.6671 | 43.365 |  |  |  |  | 3359.4013 | 0.016585 | 102.6671 | 134.08
|  |  | 102.7418 | 43.615 | [6 empty] | 102.7418 | 43.645 |  |  |  |  | 3361.8374 | 0.0095953 | 102.7418 | 134.435
|  |  | 102.8165 | 44.23 | [6 empty] | 102.8165 | 44.225 |  |  |  |  | 3364.2735 | 0.0022979 | 102.8165 | 134.835
|  |  | 102.8911 | 41.825 | [6 empty] | 102.8911 | 44.41 |  |  |  |  | 3366.7097 | 0.002897 | 102.8911 | 135.03
|  |  | 102.9658 | 42.615 | [6 empty] | 102.9658 | 44.6 |  |  |  |  | 3369.1458 | 0.0039685 | 102.9658 | 135.44
|  |  | 103.0405 | 42.73 | [6 empty] | 103.0405 | 44.61 |  |  |  |  | 3371.5819 | 0.0060162 | 103.0405 | 135.605
|  |  | 103.1151 | 42.74 | [6 empty] | 103.1151 | 44.645 |  |  |  |  | 3374.018 | 0.009653 | 103.1151 | 135.645
|  |  | 103.1898 | 42.86 | [6 empty] | 103.1898 | 44.725 |  |  |  |  | 3376.4541 | 0.01034 | 103.1898 | 136.06
|  |  | 103.2645 | 43.505 | [6 empty] | 103.2645 | 44.72 |  |  |  |  | 3378.8902 | 0.0056922 | 103.2645 | 136.46
|  |  | 103.3391 | 42.13 | [6 empty] | 103.3391 | 44.55 |  |  |  |  | 3381.3263 | 0.0035372 | 103.3391 | 136.66
|  |  | 103.4138 | 41.6 | [6 empty] | 103.4138 | 44.37 |  |  |  |  | 3383.7625 | 0.013714 | 103.4138 | 136.35
|  |  | 103.4885 | 40.97 | [6 empty] | 103.4885 | 44.19 |  |  |  |  | 3386.1986 | 0.020424 | 103.4885 | 136.12
|  |  | 103.5631 | 41.495 | [6 empty] | 103.5631 | 43.92 |  |  |  |  | 3388.6347 | 0.020425 | 103.5631 | 136.075
|  |  | 103.6378 | 41.37 | [6 empty] | 103.6378 | 43.69 |  |  |  |  | 3391.0708 | 0.013486 | 103.6378 | 136.49
|  |  | 103.7125 | 41.845 | [6 empty] | 103.7125 | 43.285 |  |  |  |  | 3393.5069 | 0.0034142 | 103.7125 | 137.04
|  |  | 103.7871 | 41.83 | [6 empty] | 103.7871 | 43.095 |  |  |  |  | 3395.943 | 0.0086224 | 103.7871 | 137.16
|  |  | 103.8618 | 41.46 | [6 empty] | 103.8618 | 42.875 |  |  |  |  | 3398.3791 | 0.013469 | 103.8618 | 137.055
|  |  | 103.9365 | 41.825 | [6 empty] | 103.9365 | 42.27 |  |  |  |  | 3400.8153 | 0.011243 | 103.9365 | 137.4
|  |  | 104.0111 | 42.37 | [6 empty] | 104.0111 | 41.9 |  |  |  |  | 3403.2514 | 0.0038214 | 104.0111 | 137.605
|  |  | 104.0858 | 41.805 | [6 empty] | 104.0858 | 41.705 |  |  |  |  | 3405.6875 | 0.007039 | 104.0858 | 137.325
|  |  | 104.1605 | 40.56 | [6 empty] | 104.1605 | 42.065 |  |  |  |  | 3408.1236 | 0.012946 | 104.1605 | 137.585
|  |  | 104.2351 | 40.655 | [6 empty] | 104.2351 | 42.575 |  |  |  |  | 3410.5597 | 0.012425 | 104.2351 | 137.9
|  |  | 104.3098 | 41.395 | [6 empty] | 104.3098 | 43 |  |  |  |  | 3412.9958 | 0.0055978 | 104.3098 | 137.935
|  |  | 104.3845 | 41.745 | [6 empty] | 104.3845 | 43.24 |  |  |  |  | 3415.4319 | 0.0053901 | 104.3845 | 138.36
|  |  | 104.4591 | 41.895 | [6 empty] | 104.4591 | 43.425 |  |  |  |  | 3417.8681 | 0.0135 | 104.4591 | 138.845
|  |  | 104.5338 | 39.88 | [6 empty] | 104.5338 | 43.505 |  |  |  |  | 3420.3042 | 0.015856 | 104.5338 | 138.62
|  |  | 104.6085 | 40.455 | [6 empty] | 104.6085 | 43.765 |  |  |  |  | 3422.7403 | 0.01114 | 104.6085 | 137.905
|  |  | 104.6831 | 41.065 | [6 empty] | 104.6831 | 44.095 |  |  |  |  | 3425.1764 | 0.0036821 | 104.6831 | 137.35
|  |  | 104.7578 | 42.165 | [6 empty] | 104.7578 | 44.355 |  |  |  |  | 3427.6125 | 0.010066 | 104.7578 | 137.12
|  |  | 104.8325 | 42.575 | [6 empty] | 104.8325 | 44.35 |  |  |  |  | 3430.0486 | 0.015005 | 104.8325 | 136.935
|  |  | 104.9071 | 43.725 | [6 empty] | 104.9071 | 44.21 |  |  |  |  | 3432.4847 | 0.012491 | 104.9071 | 137.365
|  |  | 104.9818 | 43.21 | [6 empty] | 104.9818 | 44.07 |  |  |  |  | 3434.9209 | 0.003158 | 104.9818 | 137.595
|  |  | 105.0565 | 43.34 | [6 empty] | 105.0565 | 44.005 |  |  |  |  | 3437.357 | 0.0094897 | 105.0565 | 137.03
|  |  | 105.1311 | 42.525 | [6 empty] | 105.1311 | 44.005 |  |  |  |  | 3439.7931 | 0.019796 | 105.1311 | 136.79
|  |  | 105.2058 | 42.71 | [6 empty] | 105.2058 | 43.755 |  |  |  |  | 3442.2292 | 0.023912 | 105.2058 | 136.71
|  |  | 105.2805 | 43.035 | [6 empty] | 105.2805 | 43.44 |  |  |  |  | 3444.6653 | 0.020797 | 105.2805 | 136.905
|  |  | 105.3551 | 40.905 | [6 empty] | 105.3551 | 43.305 |  |  |  |  | 3447.1014 | 0.01292 | 105.3551 | 137.425
|  |  | 105.4298 | 41.705 | [6 empty] | 105.4298 | 43.285 |  |  |  |  | 3449.5375 | 0.0056861 | 105.4298 | 138.01
|  |  | 105.5045 | 42.53 | [6 empty] | 105.5045 | 43.525 |  |  |  |  | 3451.9737 | 0.0048284 | 105.5045 | 138.125
|  |  | 105.5791 | 42.62 | [6 empty] | 105.5791 | 44.395 |  |  |  |  | 3454.4098 | 0.0034011 | 105.5791 | 137.915
|  |  | 105.6538 | 42.885 | [6 empty] | 105.6538 | 45 |  |  |  |  | 3456.8459 | 0.0023373 | 105.6538 | 137.705
|  |  | 105.7285 | 43.01 | [6 empty] | 105.7285 | 45.475 |  |  |  |  | 3459.282 | 0.0084858 | 105.7285 | 137.435
|  |  | 105.8031 | 42.835 | [6 empty] | 105.8031 | 45.88 |  |  |  |  | 3461.7181 | 0.011835 | 105.8031 | 137.01
|  |  | 105.8778 | 42.825 | [6 empty] | 105.8778 | 46.11 |  |  |  |  | 3464.1542 | 0.010138 | 105.8778 | 136.555
|  |  | 105.9525 | 42.86 | [6 empty] | 105.9525 | 46.39 |  |  |  |  | 3466.5903 | 0.0062404 | 105.9525 | 136.67
|  |  | 106.0271 | 42.05 | [6 empty] | 106.0271 | 46.64 |  |  |  |  | 3469.0265 | 0.010519 | 106.0271 | 136.7
|  |  | 106.1018 | 42.405 | [6 empty] | 106.1018 | 46.675 |  |  |  |  | 3471.4626 | 0.016538 | 106.1018 | 136.925
|  |  | 106.1765 | 42.66 | [6 empty] | 106.1765 | 46.695 |  |  |  |  | 3473.8987 | 0.017628 | 106.1765 | 136.76
|  |  | 106.2511 | 42.825 | [6 empty] | 106.2511 | 46.86 |  |  |  |  | 3476.3348 | 0.012896 | 106.2511 | 136.11
|  |  | 106.3258 | 42.62 | [6 empty] | 106.3258 | 46.825 |  |  |  |  | 3478.7709 | 0.0048923 | 106.3258 | 135.855
|  |  | 106.4005 | 41.68 | [6 empty] | 106.4005 | 46.79 |  |  |  |  | 3481.207 | 0.0025812 | 106.4005 | 136.085
|  |  | 106.4751 | 42.935 | [6 empty] | 106.4751 | 46.73 |  |  |  |  | 3483.6431 | 0.0045501 | 106.4751 | 136.335
|  |  | 106.5498 | 42.175 | [6 empty] | 106.5498 | 46.625 |  |  |  |  | 3486.0793 | 0.00091847 | 106.5498 | 136.13
|  |  | 106.6245 | 41.655 | [6 empty] | 106.6245 | 46.475 |  |  |  |  | 3488.5154 | 0.0093803 | 106.6245 | 136.085
|  |  | 106.6991 | 42.3 | [6 empty] | 106.6991 | 46.23 |  |  |  |  | 3490.9515 | 0.019378 | 106.6991 | 135.98
|  |  | 106.7738 | 41.055 | [6 empty] | 106.7738 | 46.17 |  |  |  |  | 3493.3876 | 0.025387 | 106.7738 | 135.915
|  |  | 106.8485 | 40.585 | [6 empty] | 106.8485 | 46.26 |  |  |  |  | 3495.8237 | 0.024426 | 106.8485 | 135.865
|  |  | 106.9231 | 41.26 | [6 empty] | 106.9231 | 46.96 |  |  |  |  | 3498.2598 | 0.016662 | 106.9231 | 135.63
|  |  | 106.9978 | 41.745 | [6 empty] | 106.9978 | 47.715 |  |  |  |  | 3500.6959 | 0.0052836 | 106.9978 | 135.67
|  |  | 107.0725 | 42.065 | [6 empty] | 107.0725 | 48.365 |  |  |  |  | 3503.1321 | 0.005327 | 107.0725 | 135.8
|  |  | 107.1471 | 42.15 | [6 empty] | 107.1471 | 48.89 |  |  |  |  | 3505.5682 | 0.011293 | 107.1471 | 135.77
|  |  | 107.2218 | 42.62 | [6 empty] | 107.2218 | 49.45 |  |  |  |  | 3508.0043 | 0.012224 | 107.2218 | 135.235
|  |  | 107.2965 | 42.44 | [6 empty] | 107.2965 | 49.735 |  |  |  |  | 3510.4404 | 0.011557 | 107.2965 | 134.73
|  |  | 107.3711 | 43.25 | [6 empty] | 107.3711 | 50.065 |  |  |  |  | 3512.8765 | 0.012973 | 107.3711 | 134.345
|  |  | 107.4458 | 42.29 | [6 empty] | 107.4458 | 50.165 |  |  |  |  | 3515.3126 | 0.013872 | 107.4458 | 134.395
|  |  | 107.5205 | 42.58 | [6 empty] | 107.5205 | 50.005 |  |  |  |  | 3517.7487 | 0.011104 | 107.5205 | 134.27
|  |  | 107.5951 | 43.32 | [6 empty] | 107.5951 | 49.89 |  |  |  |  | 3520.1849 | 0.0061671 | 107.5951 | 134.115
|  |  | 107.6698 | 43.305 | [6 empty] | 107.6698 | 49.9 |  |  |  |  | 3522.621 | 0.0071527 | 107.6698 | 133.97
|  |  | 107.7445 | 44.23 | [6 empty] | 107.7445 | 49.92 |  |  |  |  | 3525.0571 | 0.010444 | 107.7445 | 133.79
|  |  | 107.8191 | 43.265 | [6 empty] | 107.8191 | 49.79 |  |  |  |  | 3527.4932 | 0.0092748 | 107.8191 | 133.78
|  |  | 107.8938 | 43.2 | [6 empty] | 107.8938 | 49.445 |  |  |  |  | 3529.9293 | 0.0038024 | 107.8938 | 134.07
|  |  | 107.9685 | 43.05 | [6 empty] | 107.9685 | 49.15 |  |  |  |  | 3532.3654 | 0.0052712 | 107.9685 | 134.255
|  |  | 108.0431 | 42.515 | [6 empty] | 108.0431 | 48.745 |  |  |  |  | 3534.8015 | 0.010578 | 108.0431 | 133.62
|  |  | 108.1178 | 42.61 | [6 empty] | 108.1178 | 48.41 |  |  |  |  | 3537.2376 | 0.010989 | 108.1178 | 133.155
|  |  | 108.1925 | 43.655 | [6 empty] | 108.1925 | 48.71 |  |  |  |  | 3539.6738 | 0.0064271 | 108.1925 | 133.33
|  |  | 108.2672 | 43.71 | [6 empty] | 108.2672 | 49.3 |  |  |  |  | 3542.1099 | 0.0051261 | 108.2672 | 133.395
|  |  | 108.3418 | 43.91 | [6 empty] | 108.3418 | 49.89 |  |  |  |  | 3544.546 | 0.010751 | 108.3418 | 133.035
|  |  | 108.4165 | 43.91 | [6 empty] | 108.4165 | 50.08 |  |  |  |  | 3546.9821 | 0.0121 | 108.4165 | 132.7
|  |  | 108.4911 | 43.19 | [6 empty] | 108.4911 | 50.225 |  |  |  |  | 3549.4182 | 0.0069019 | 108.4911 | 132.04
|  |  | 108.5658 | 42.95 | [6 empty] | 108.5658 | 50.445 |  |  |  |  | 3551.8543 | 0.003676 | 108.5658 | 132.09
|  |  | 108.6405 | 43.585 | [6 empty] | 108.6405 | 50.485 |  |  |  |  | 3554.2904 | 0.014363 | 108.6405 | 132.29
|  |  | 108.7151 | 43.45 | [6 empty] | 108.7151 | 50.415 |  |  |  |  | 3556.7266 | 0.021216 | 108.7151 | 132.86
|  |  | 108.7898 | 43.375 | [6 empty] | 108.7898 | 50.17 |  |  |  |  | 3559.1627 | 0.021259 | 108.7898 | 132.5
|  |  | 108.8645 | 42.44 | [6 empty] | 108.8645 | 50.05 |  |  |  |  | 3561.5988 | 0.015228 | 108.8645 | 132.015
|  |  | 108.9392 | 42.995 | [6 empty] | 108.9392 | 49.895 |  |  |  |  | 3564.0349 | 0.0090539 | 108.9392 | 131.51
|  |  | 109.0138 | 43.345 | [6 empty] | 109.0138 | 49.595 |  |  |  |  | 3566.471 | 0.011318 | 109.0138 | 131.31
|  |  | 109.0885 | 43.015 | [6 empty] | 109.0885 | 49.44 |  |  |  |  | 3568.9071 | 0.013954 | 109.0885 | 131.985
|  |  | 109.1632 | 43.55 | [6 empty] | 109.1632 | 49.17 |  |  |  |  | 3571.3432 | 0.011894 | 109.1632 | 132.605
|  |  | 109.2378 | 43.035 | [6 empty] | 109.2378 | 48.85 |  |  |  |  | 3573.7794 | 0.0087081 | 109.2378 | 131.76
|  |  | 109.3125 | 43.285 | [6 empty] | 109.3125 | 48.335 |  |  |  |  | 3576.2155 | 0.010929 | 109.3125 | 131.33
|  |  | 109.3872 | 43.08 | [6 empty] | 109.3872 | 47.86 |  |  |  |  | 3578.6516 | 0.013501 | 109.3872 | 131.865
|  |  | 109.4618 | 43.475 | [6 empty] | 109.4618 | 47.26 |  |  |  |  | 3581.0877 | 0.011694 | 109.4618 | 133.005
|  |  | 109.5365 | 43.475 | [6 empty] | 109.5365 | 46.48 |  |  |  |  | 3583.5238 | 0.0071204 | 109.5365 | 133.15
|  |  | 109.6112 | 43.12 | [6 empty] | 109.6112 | 46.38 |  |  |  |  | 3585.9599 | 0.0074524 | 109.6112 | 132.775
|  |  | 109.6858 | 42.8 | [6 empty] | 109.6858 | 46.55 |  |  |  |  | 3588.396 | 0.010299 | 109.6858 | 132.1
|  |  | 109.7605 | 42.155 | [6 empty] | 109.7605 | 47.055 |  |  |  |  | 3590.8322 | 0.0090029 | 109.7605 | 131.97
|  |  | 109.8352 | 42.12 | [6 empty] | 109.8352 | 47.36 |  |  |  |  | 3593.2683 | 0.0056461 | 109.8352 | 131.62
|  |  | 109.9098 | 42.195 | [6 empty] | 109.9098 | 47.415 |  |  |  |  | 3595.7044 | 0.009764 | 109.9098 | 132.08
|  |  | 109.9845 | 44 | [6 empty] | 109.9845 | 47.26 |  |  |  |  | 3598.1405 | 0.014433 | 109.9845 | 132.51
|  |  | 110.0592 | 42.945 | [6 empty] | 110.0592 | 47.27 |  |  |  |  | 3600.5766 | 0.013239 | 110.0592 | 132.215
|  |  | 110.1338 | 42.26 | [6 empty] | 110.1338 | 47.49 |  |  |  |  | 3603.0127 | 0.0053989 | 110.1338 | 131.895
|  |  | 110.2085 | 42.095 | [6 empty] | 110.2085 | 47.625 |  |  |  |  | 3605.4488 | 0.0061349 | 110.2085 | 131.72
|  |  | 110.2832 | 42.615 | [6 empty] | 110.2832 | 47.39 |  |  |  |  | 3607.885 | 0.016174 | 110.2832 | 131.53
|  |  | 110.3578 | 43.11 | [6 empty] | 110.3578 | 47.21 |  |  |  |  | 3610.3211 | 0.019873 | 110.3578 | 132.125
|  |  | 110.4325 | 43.33 | [6 empty] | 110.4325 | 46.975 |  |  |  |  | 3612.7572 | 0.015106 | 110.4325 | 132.33
|  |  | 110.5072 | 42.26 | [6 empty] | 110.5072 | 46.715 |  |  |  |  | 3615.1933 | 0.004701 | 110.5072 | 132.26
|  |  | 110.5818 | 43.535 | [6 empty] | 110.5818 | 46.51 |  |  |  |  | 3617.6294 | 0.011867 | 110.5818 | 132.155
|  |  | 110.6565 | 43.735 | [6 empty] | 110.6565 | 46.32 |  |  |  |  | 3620.0655 | 0.02189 | 110.6565 | 132.075
|  |  | 110.7312 | 42.43 | [6 empty] | 110.7312 | 45.985 |  |  |  |  | 3622.5016 | 0.024379 | 110.7312 | 132.175
|  |  | 110.8058 | 42.16 | [6 empty] | 110.8058 | 45.445 |  |  |  |  | 3624.9378 | 0.018801 | 110.8058 | 132.11
|  |  | 110.8805 | 42.27 | [6 empty] | 110.8805 | 44.6 |  |  |  |  | 3627.3739 | 0.0097691 | 110.8805 | 132.905
|  |  | 110.9552 | 42.54 | [6 empty] | 110.9552 | 44.17 |  |  |  |  | 3629.81 | 0.0096063 | 110.9552 | 133.97
|  |  | 111.0298 | 43.865 | [6 empty] | 111.0298 | 43.955 |  |  |  |  | 3632.2461 | 0.01373 | 111.0298 | 134.535
|  |  | 111.1045 | 43.93 | [6 empty] | 111.1045 | 44.29 |  |  |  |  | 3634.6822 | 0.011738 | 111.1045 | 134.705
|  |  | 111.1792 | 42.695 | [6 empty] | 111.1792 | 44.685 |  |  |  |  | 3637.1183 | 0.0039128 | 111.1792 | 134.845
|  |  | 111.2538 | 43.055 | [6 empty] | 111.2538 | 45.01 |  |  |  |  | 3639.5544 | 0.006426 | 111.2538 | 135.055
|  |  | 111.3285 | 42.32 | [6 empty] | 111.3285 | 45.43 |  |  |  |  | 3641.9906 | 0.013373 | 111.3285 | 134.725
|  |  | 111.4032 | 41.31 | [6 empty] | 111.4032 | 45.585 |  |  |  |  | 3644.4267 | 0.014175 | 111.4032 | 134.065
|  |  | 111.4778 | 41.46 | [6 empty] | 111.4778 | 45.665 |  |  |  |  | 3646.8628 | 0.0087128 | 111.4778 | 134
|  |  | 111.5525 | 42.65 | [6 empty] | 111.5525 | 45.51 |  |  |  |  | 3649.2989 | 0.00093927 | 111.5525 | 134.185
|  |  | 111.6272 | 41.99 | [6 empty] | 111.6272 | 45.42 |  |  |  |  | 3651.735 | 0.0076251 | 111.6272 | 134.665
|  |  | 111.7018 | 41.865 | [6 empty] | 111.7018 | 45.35 |  |  |  |  | 3654.1711 | 0.010152 | 111.7018 | 135.385
|  |  | 111.7765 | 40.785 | [6 empty] | 111.7765 | 45.175 |  |  |  |  | 3656.6072 | 0.0067301 | 111.7765 | 135.76
|  |  | 111.8512 | 41.475 | [6 empty] | 111.8512 | 44.94 |  |  |  |  | 3659.0434 | 0.0015874 | 111.8512 | 136.44
|  |  | 111.9258 | 42.17 | [6 empty] | 111.9258 | 44.635 |  |  |  |  | 3661.4795 | 0.0084115 | 111.9258 | 137.12
|  |  | 112.0005 | 42.82 | [6 empty] | 112.0005 | 44.205 |  |  |  |  | 3663.9156 | 0.011964 | 112.0005 | 137.495
|  |  | 112.0752 | 42.29 | [6 empty] | 112.0752 | 43.935 |  |  |  |  | 3666.3517 | 0.0099194 | 112.0752 | 137.165
|  |  | 112.1498 | 41.33 | [6 empty] | 112.1498 | 43.64 |  |  |  |  | 3668.7878 | 0.0034129 | 112.1498 | 137.1
|  |  | 112.2245 | 42.16 | [6 empty] | 112.2245 | 43.25 |  |  |  |  | 3671.2239 | 0.00422 | 112.2245 | 137.32
|  |  | 112.2992 | 42.08 | [6 empty] | 112.2992 | 42.63 |  |  |  |  | 3673.66 | 0.0094393 | 112.2992 | 137.665
|  |  | 112.3738 | 41.04 | [6 empty] | 112.3738 | 42.235 |  |  |  |  | 3676.0962 | 0.011037 | 112.3738 | 138.175
|  |  | 112.4485 | 41.45 | [6 empty] | 112.4485 | 42.05 |  |  |  |  | 3678.5323 | 0.012074 | 112.4485 | 138.545
|  |  | 112.5232 | 41.21 | [6 empty] | 112.5232 | 42.69 |  |  |  |  | 3680.9684 | 0.0158 | 112.5232 | 138.58
|  |  | 112.5978 | 40.93 | [6 empty] | 112.5978 | 43.41 |  |  |  |  | 3683.4045 | 0.018955 | 112.5978 | 138.85
|  |  | 112.6725 | 42.7 | [6 empty] | 112.6725 | 43.89 |  |  |  |  | 3685.8406 | 0.017628 | 112.6725 | 139.225
|  |  | 112.7472 | 42.03 | [6 empty] | 112.7472 | 44.275 |  |  |  |  | 3688.2767 | 0.011712 | 112.7472 | 139.37
|  |  | 112.8218 | 42.12 | [6 empty] | 112.8218 | 44.465 |  |  |  |  | 3690.7128 | 0.0080628 | 112.8218 | 139.26
|  |  | 112.8965 | 43.01 | [6 empty] | 112.8965 | 44.51 |  |  |  |  | 3693.149 | 0.013429 | 112.8965 | 138.775
|  |  | 112.9712 | 43.095 | [6 empty] | 112.9712 | 44.535 |  |  |  |  | 3695.5851 | 0.017294 | 112.9712 | 138.365
|  |  | 113.0458 | 42.17 | [6 empty] | 113.0458 | 44.535 |  |  |  |  | 3698.0212 | 0.015823 | 113.0458 | 138.45
|  |  | 113.1205 | 42.05 | [6 empty] | 113.1205 | 44.53 |  |  |  |  | 3700.4573 | 0.010421 | 113.1205 | 139.03
|  |  | 113.1952 | 42.175 | [6 empty] | 113.1952 | 44.63 |  |  |  |  | 3702.8934 | 0.0058375 | 113.1952 | 139.12
|  |  | 113.2698 | 41.58 | [6 empty] | 113.2698 | 44.67 |  |  |  |  | 3705.3295 | 0.0068639 | 113.2698 | 138.82
|  |  | 113.3445 | 42.215 | [6 empty] | 113.3445 | 44.585 |  |  |  |  | 3707.7656 | 0.0097973 | 113.3445 | 138.975
|  |  | 113.4192 | 41.15 | [6 empty] | 113.4192 | 44.295 |  |  |  |  | 3710.2017 | 0.014066 | 113.4192 | 139.21
|  |  | 113.4938 | 41.415 | [6 empty] | 113.4938 | 44.125 |  |  |  |  | 3712.6379 | 0.018475 | 113.4938 | 138.705
|  |  | 113.5685 | 41.1 | [6 empty] | 113.5685 | 43.895 |  |  |  |  | 3715.074 | 0.019453 | 113.5685 | 138.165
|  |  | 113.6432 | 41.635 | [6 empty] | 113.6432 | 43.76 |  |  |  |  | 3717.5101 | 0.014853 | 113.6432 | 138.2
|  |  | 113.7178 | 41.945 | [6 empty] | 113.7178 | 44.065 |  |  |  |  | 3719.9462 | 0.0065213 | 113.7178 | 138.195
|  |  | 113.7925 | 41.52 | [6 empty] | 113.7925 | 44.465 |  |  |  |  | 3722.3823 | 0.0074236 | 113.7925 | 137.9
|  |  | 113.8672 | 41.865 | [6 empty] | 113.8672 | 45.17 |  |  |  |  | 3724.8184 | 0.012493 | 113.8672 | 137.135
|  |  | 113.9418 | 41.62 | [6 empty] | 113.9418 | 45.46 |  |  |  |  | 3727.2545 | 0.010157 | 113.9418 | 136.975
|  |  | 114.0165 | 41.855 | [6 empty] | 114.0165 | 45.605 |  |  |  |  | 3729.6907 | 0.0027116 | 114.0165 | 137.29
|  |  | 114.0912 | 43.23 | [6 empty] | 114.0912 | 45.805 |  |  |  |  | 3732.1268 | 0.016427 | 114.0912 | 137.68
|  |  | 114.1658 | 43.9 | [6 empty] | 114.1658 | 45.83 |  |  |  |  | 3734.5629 | 0.030104 | 114.1658 | 137.44
|  |  | 114.2405 | 43.86 | [6 empty] | 114.2405 | 45.83 |  |  |  |  | 3736.999 | 0.036322 | 114.2405 | 137.005
|  |  | 114.3152 | 43.5 | [6 empty] | 114.3152 | 45.925 |  |  |  |  | 3739.4351 | 0.03223 | 114.3152 | 136.11
|  |  | 114.3898 | 41.91 | [6 empty] | 114.3898 | 46.155 |  |  |  |  | 3741.8712 | 0.020365 | 114.3898 | 135.935
|  |  | 114.4645 | 42.24 | [6 empty] | 114.4645 | 46.195 |  |  |  |  | 3744.3073 | 0.01363 | 114.4645 | 136.1
|  |  | 114.5392 | 41.375 | [6 empty] | 114.5392 | 46.12 |  |  |  |  | 3746.7435 | 0.022456 | 114.5392 | 135.93
|  |  | 114.6138 | 41.895 | [6 empty] | 114.6138 | 45.795 |  |  |  |  | 3749.1796 | 0.028579 | 114.6138 | 135.85
|  |  | 114.6885 | 41.65 | [6 empty] | 114.6885 | 45.735 |  |  |  |  | 3751.6157 | 0.02628 | 114.6885 | 136.45
|  |  | 114.7632 | 40.82 | [6 empty] | 114.7632 | 45.69 |  |  |  |  | 3754.0518 | 0.01785 | 114.7632 | 136.515
|  |  | 114.8378 | 41.8 | [6 empty] | 114.8378 | 45.52 |  |  |  |  | 3756.4879 | 0.0093219 | 114.8378 | 135.9
|  |  | 114.9125 | 43.43 | [6 empty] | 114.9125 | 45.38 |  |  |  |  | 3758.924 | 0.007437 | 114.9125 | 135.905
|  |  | 114.9872 | 42.965 | [6 empty] | 114.9872 | 45.19 |  |  |  |  | 3761.3601 | 0.010503 | 114.9872 | 136.38
|  |  | 115.0618 | 42.495 | [6 empty] | 115.0618 | 45.055 |  |  |  |  | 3763.7963 | 0.017222 | 115.0618 | 136.47
|  |  | 115.1365 | 41.855 | [6 empty] | 115.1365 | 45.19 |  |  |  |  | 3766.2324 | 0.025355 | 115.1365 | 136.315
|  |  | 115.2112 | 41.805 | [6 empty] | 115.2112 | 45.565 |  |  |  |  | 3768.6685 | 0.02949 | 115.2112 | 135.995
|  |  | 115.2859 | 42.505 | [6 empty] | 115.2859 | 46.47 |  |  |  |  | 3771.1046 | 0.025947 | 115.2859 | 136.115
|  |  | 115.3605 | 42.22 | [6 empty] | 115.3605 | 47.015 |  |  |  |  | 3773.5407 | 0.014993 | 115.3605 | 135.91
|  |  | 115.4352 | 42.615 | [6 empty] | 115.4352 | 47.46 |  |  |  |  | 3775.9768 | 0.0024199 | 115.4352 | 134.86
|  |  | 115.5098 | 43.09 | [6 empty] | 115.5098 | 47.855 |  |  |  |  | 3778.4129 | 0.011958 | 115.5098 | 134.8
|  |  | 115.5845 | 42.58 | [6 empty] | 115.5845 | 47.91 |  |  |  |  | 3780.8491 | 0.016571 | 115.5845 | 135.15
|  |  | 115.6592 | 43.195 | [6 empty] | 115.6592 | 47.985 |  |  |  |  | 3783.2852 | 0.012428 | 115.6592 | 135.615
|  |  | 115.7338 | 42.09 | [6 empty] | 115.7338 | 48.27 |  |  |  |  | 3785.7213 | 0.0036581 | 115.7338 | 135.72
|  |  | 115.8085 | 42.385 | [6 empty] | 115.8085 | 48.445 |  |  |  |  | 3788.1574 | 0.01151 | 115.8085 | 135.755
|  |  | 115.8832 | 43.28 | [6 empty] | 115.8832 | 48.635 |  |  |  |  | 3790.5935 | 0.019465 | 115.8832 | 135.485
|  |  | 115.9578 | 43.925 | [6 empty] | 115.9578 | 48.585 |  |  |  |  | 3793.0296 | 0.021206 | 115.9578 | 135.05
|  |  | 116.0325 | 43.89 | [6 empty] | 116.0325 | 48.45 |  |  |  |  | 3795.4657 | 0.018769 | 116.0325 | 134.97
|  |  | 116.1072 | 43.615 | [6 empty] | 116.1072 | 48.355 |  |  |  |  | 3797.9019 | 0.01778 | 116.1072 | 134.855
|  |  | 116.1819 | 43.7 | [6 empty] | 116.1819 | 48.26 |  |  |  |  | 3800.338 | 0.019025 | 116.1819 | 134.76
|  |  | 116.2565 | 43.33 | [6 empty] | 116.2565 | 48.055 |  |  |  |  | 3802.7741 | 0.017385 | 116.2565 | 134.41
|  |  | 116.3312 | 43.36 | [6 empty] | 116.3312 | 47.9 |  |  |  |  | 3805.2102 | 0.010891 | 116.3312 | 133.765
|  |  | 116.4059 | 42.88 | [6 empty] | 116.4059 | 47.65 |  |  |  |  | 3807.6463 | 0.0021072 | 116.4059 | 134.21
|  |  | 116.4805 | 43.18 | [6 empty] | 116.4805 | 47.235 |  |  |  |  | 3810.0824 | 0.0045368 | 116.4805 | 134.645
|  |  | 116.5552 | 43.06 | [6 empty] | 116.5552 | 47.165 |  |  |  |  | 3812.5185 | 0.0054927 | 116.5552 | 134.705
|  |  | 116.6299 | 42.88 | [6 empty] | 116.6299 | 47.57 |  |  |  |  | 3814.9547 | 0.0049159 | 116.6299 | 134.305
|  |  | 116.7045 | 41.72 | [6 empty] | 116.7045 | 48.335 |  |  |  |  | 3817.3908 | 0.013057 | 116.7045 | 134.495
|  |  | 116.7792 | 43.355 | [6 empty] | 116.7792 | 48.76 |  |  |  |  | 3819.8269 | 0.020547 | 116.7792 | 134.6
|  |  | 116.8539 | 43.57 | [6 empty] | 116.8539 | 49.06 |  |  |  |  | 3822.263 | 0.02102 | 116.8539 | 134.265
|  |  | 116.9285 | 43.5 | [6 empty] | 116.9285 | 49.3 |  |  |  |  | 3824.6991 | 0.012363 | 116.9285 | 134.06
|  |  | 117.0032 | 42.86 | [6 empty] | 117.0032 | 49.32 |  |  |  |  | 3827.1352 | 0.0055122 | 117.0032 | 133.655
|  |  | 117.0779 | 42.665 | [6 empty] | 117.0779 | 49.265 |  |  |  |  | 3829.5713 | 0.021745 | 117.0779 | 133.58
|  |  | 117.1525 | 43.88 | [6 empty] | 117.1525 | 49.12 |  |  |  |  | 3832.0075 | 0.033342 | 117.1525 | 133.68
|  |  | 117.2272 | 44.47 | [6 empty] | 117.2272 | 49.085 |  |  |  |  | 3834.4436 | 0.034782 | 117.2272 | 134.065
|  |  | 117.3019 | 43.67 | [6 empty] | 117.3019 | 48.995 |  |  |  |  | 3836.8797 | 0.025738 | 117.3019 | 134.595
|  |  | 117.3765 | 43.315 | [6 empty] | 117.3765 | 48.72 |  |  |  |  | 3839.3158 | 0.010786 | 117.3765 | 134.345
|  |  | 117.4512 | 43.08 | [6 empty] | 117.4512 | 48.4 |  |  |  |  | 3841.7519 | 0.0066074 | 117.4512 | 134.11
|  |  | 117.5259 | 43.835 | [6 empty] | 117.5259 | 48.095 |  |  |  |  | 3844.188 | 0.013855 | 117.5259 | 133.99
|  |  | 117.6005 | 43.905 | [6 empty] | 117.6005 | 47.64 |  |  |  |  | 3846.6241 | 0.012359 | 117.6005 | 133.93
|  |  | 117.6752 | 42.935 | [6 empty] | 117.6752 | 47.21 |  |  |  |  | 3849.0603 | 0.0040304 | 117.6752 | 134.215
|  |  | 117.7499 | 41.29 | [6 empty] | 117.7499 | 46.705 |  |  |  |  | 3851.4964 | 0.0092357 | 117.7499 | 135.005
|  |  | 117.8245 | 42.705 | [6 empty] | 117.8245 | 46.01 |  |  |  |  | 3853.9325 | 0.016323 | 117.8245 | 135.14
|  |  | 117.8992 | 43.62 | [6 empty] | 117.8992 | 45.6 |  |  |  |  | 3856.3686 | 0.015249 | 117.8992 | 134.75
|  |  | 117.9739 | 43.565 | [6 empty] | 117.9739 | 45.14 |  |  |  |  | 3858.8047 | 0.0065817 | 117.9739 | 134.185
|  |  | 118.0485 | 42.32 | [6 empty] | 118.0485 | 44.69 |  |  |  |  | 3861.2408 | 0.0098092 | 118.0485 | 134.47
|  |  | 118.1232 | 43.065 | [6 empty] | 118.1232 | 45.02 |  |  |  |  | 3863.6769 | 0.021078 | 118.1232 | 134.93
|  |  | 118.1979 | 42.545 | [6 empty] | 118.1979 | 45.67 |  |  |  |  | 3866.1131 | 0.025264 | 118.1979 | 135.28
|  |  | 118.2725 | 43.13 | [6 empty] | 118.2725 | 46.12 |  |  |  |  | 3868.5492 | 0.020095 | 118.2725 | 135.39
|  |  | 118.3472 | 42.68 | [6 empty] | 118.3472 | 46.325 |  |  |  |  | 3870.9853 | 0.0087851 | 118.3472 | 135.275
|  |  | 118.4219 | 41.76 | [6 empty] | 118.4219 | 46.34 |  |  |  |  | 3873.4214 | 0.010522 | 118.4219 | 134.81
|  |  | 118.4965 | 43.235 | [6 empty] | 118.4965 | 46.445 |  |  |  |  | 3875.8575 | 0.020957 | 118.4965 | 134.36
|  |  | 118.5712 | 43.1 | [6 empty] | 118.5712 | 46.435 |  |  |  |  | 3878.2936 | 0.025397 | 118.5712 | 133.985
|  |  | 118.6459 | 42.55 | [6 empty] | 118.6459 | 46.285 |  |  |  |  | 3880.7297 | 0.023558 | 118.6459 | 133.79
|  |  | 118.7205 | 42.71 | [6 empty] | 118.7205 | 46.215 |  |  |  |  | 3883.1658 | 0.018602 | 118.7205 | 133.61
|  |  | 118.7952 | 41.765 | [6 empty] | 118.7952 | 46.03 |  |  |  |  | 3885.602 | 0.013282 | 118.7952 | 133.655
|  |  | 118.8699 | 42.58 | [6 empty] | 118.8699 | 45.775 |  |  |  |  | 3888.0381 | 0.0069759 | 118.8699 | 133.875
|  |  | 118.9445 | 42.315 | [6 empty] | 118.9445 | 45.42 |  |  |  |  | 3890.4742 | 0.0016164 | 118.9445 | 133.73
|  |  | 119.0192 | 41.945 | [6 empty] | 119.0192 | 44.865 |  |  |  |  | 3892.9103 | 0.010855 | 119.0192 | 133.845
|  |  | 119.0939 | 41.465 | [6 empty] | 119.0939 | 44.6 |  |  |  |  | 3895.3464 | 0.016919 | 119.0939 | 134.23
|  |  | 119.1685 | 41.95 | [6 empty] | 119.1685 | 44.21 |  |  |  |  | 3897.7825 | 0.016537 | 119.1685 | 134.005
|  |  | 119.2432 | 42.99 | [6 empty] | 119.2432 | 43.525 |  |  |  |  | 3900.2186 | 0.0092348 | 119.2432 | 133.77
|  |  | 119.3179 | 42.965 | [6 empty] | 119.3179 | 43.345 |  |  |  |  | 3902.6548 | 0.0023496 | 119.3179 | 134.095
|  |  | 119.3925 | 43.265 | [6 empty] | 119.3925 | 43.55 |  |  |  |  | 3905.0909 | 0.012716 | 119.3925 | 134.39
|  |  | 119.4672 | 42.84 | [6 empty] | 119.4672 | 44.115 |  |  |  |  | 3907.527 | 0.01783 | 119.4672 | 134.665
|  |  | 119.5419 | 41.82 | [6 empty] | 119.5419 | 44.42 |  |  |  |  | 3909.9631 | 0.01616 | 119.5419 | 134.91
|  |  | 119.6165 | 42.17 | [6 empty] | 119.6165 | 44.645 |  |  |  |  | 3912.3992 | 0.011733 | 119.6165 | 135.39
|  |  | 119.6912 | 42.04 | [6 empty] | 119.6912 | 44.89 |  |  |  |  | 3914.8353 | 0.013204 | 119.6912 | 135.65
|  |  | 119.7659 | 41.6 | [6 empty] | 119.7659 | 44.89 |  |  |  |  | 3917.2714 | 0.016811 | 119.7659 | 135.58
|  |  | 119.8405 | 41.58 | [6 empty] | 119.8405 | 44.805 |  |  |  |  | 3919.7076 | 0.015461 | 119.8405 | 135.745
|  |  | 119.9152 | 41.83 | [6 empty] | 119.9152 | 44.68 |  |  |  |  | 3922.1437 | 0.010374 | 119.9152 | 136.22
|  |  | 119.9899 | 41.22 | [6 empty] | 119.9899 | 44.59 |  |  |  |  | 3924.5798 | 0.01144 | 119.9899 | 136.695
|  |  | 120.0645 | 42.41 | [6 empty] | 120.0645 | 44.55 |  |  |  |  | 3927.0159 | 0.016505 | 120.0645 | 136.585
|  |  | 120.1392 | 42.32 | [6 empty] | 120.1392 | 44.44 |  |  |  |  | 3929.452 | 0.015935 | 120.1392 | 136.645
|  |  | 120.2139 | 42.53 | [6 empty] | 120.2139 | 44.22 |  |  |  |  | 3931.8881 | 0.0083543 | 120.2139 | 136.59
|  |  | 120.2885 | 41.25 | [6 empty] | 120.2885 | 44.055 |  |  |  |  | 3934.3242 | 0.0078424 | 120.2885 | 136.175
|  |  | 120.3632 | 41.465 | [6 empty] | 120.3632 | 43.765 |  |  |  |  | 3936.7604 | 0.018548 | 120.3632 | 136.4
|  |  | 120.4379 | 41.875 | [6 empty] | 120.4379 | 43.395 |  |  |  |  | 3939.1965 | 0.023314 | 120.4379 | 137.065
|  |  | 120.5125 | 41.75 | [6 empty] | 120.5125 | 43.065 |  |  |  |  | 3941.6326 | 0.019004 | 120.5125 | 137.56
|  |  | 120.5872 | 40.74 | [6 empty] | 120.5872 | 42.685 |  |  |  |  | 3944.0687 | 0.0085271 | 120.5872 | 137.185
|  |  | 120.6619 | 41.095 | [6 empty] | 120.6619 | 42.02 |  |  |  |  | 3946.5048 | 0.01051 | 120.6619 | 136.77
|  |  | 120.7365 | 41.005 | [6 empty] | 120.7365 | 41.86 |  |  |  |  | 3948.9409 | 0.019494 | 120.7365 | 137.405
|  |  | 120.8112 | 40.74 | [6 empty] | 120.8112 | 42.06 |  |  |  |  | 3951.377 | 0.02074 | 120.8112 | 138.31
|  |  | 120.8859 | 43.005 | [6 empty] | 120.8859 | 42.66 |  |  |  |  | 3953.8132 | 0.01389 | 120.8859 | 138.645
|  |  | 120.9605 | 42.67 | [6 empty] | 120.9605 | 43.185 |  |  |  |  | 3956.2493 | 0.0083184 | 120.9605 | 138.45
|  |  | 121.0352 | 43.395 | [6 empty] | 121.0352 | 43.6 |  |  |  |  | 3958.6854 | 0.015466 | 121.0352 | 138.255
|  |  | 121.1099 | 40.055 | [6 empty] | 121.1099 | 43.91 |  |  |  |  | 3961.1215 | 0.019289 | 121.1099 | 138.175
|  |  | 121.1845 | 39.995 | [6 empty] | 121.1845 | 43.88 |  |  |  |  | 3963.5576 | 0.014756 | 121.1845 | 137.96
|  |  | 121.2592 | 42.67 | [6 empty] | 121.2592 | 44 |  |  |  |  | 3965.9937 | 0.0080676 | 121.2592 | 137.845
|  |  | 121.3339 | 42.195 | [6 empty] | 121.3339 | 43.965 |  |  |  |  | 3968.4298 | 0.016472 | 121.3339 | 137.99
|  |  | 121.4085 | 41.785 | [6 empty] | 121.4085 | 43.98 |  |  |  |  | 3970.866 | 0.025179 | 121.4085 | 138.045
|  |  | 121.4832 | 42.12 | [6 empty] | 121.4832 | 44.205 |  |  |  |  | 3973.3021 | 0.025621 | 121.4832 | 137.495
|  |  | 121.5579 | 41.49 | [6 empty] | 121.5579 | 44.375 |  |  |  |  | 3975.7382 | 0.017247 | 121.5579 | 137.29
|  |  | 121.6325 | 42.11 | [6 empty] | 121.6325 | 44.335 |  |  |  |  | 3978.1743 | 0.0046537 | 121.6325 | 137.23
|  |  | 121.7072 | 41.7 | [6 empty] | 121.7072 | 44.3 |  |  |  |  | 3980.6104 | 0.010477 | 121.7072 | 137.56
|  |  | 121.7819 | 42.485 | [6 empty] | 121.7819 | 44.165 |  |  |  |  | 3983.0465 | 0.017972 | 121.7819 | 137.315
|  |  | 121.8565 | 40.665 | [6 empty] | 121.8565 | 43.87 |  |  |  |  | 3985.4826 | 0.018868 | 121.8565 | 137.175
|  |  | 121.9312 | 42.11 | [6 empty] | 121.9312 | 43.66 |  |  |  |  | 3987.9188 | 0.016538 | 121.9312 | 138.075
|  |  | 122.0059 | 43.295 | [6 empty] | 122.0059 | 43.395 |  |  |  |  | 3990.3549 | 0.017265 | 122.0059 | 138.38
|  |  | 122.0805 | 44.02 | [6 empty] | 122.0805 | 43.165 |  |  |  |  | 3992.791 | 0.020503 | 122.0805 | 138.415
|  |  | 122.1552 | 40.565 | [6 empty] | 122.1552 | 43.285 |  |  |  |  | 3995.2271 | 0.021509 | 122.1552 | 138.17
|  |  | 122.2299 | 41.535 | [6 empty] | 122.2299 | 43.865 |  |  |  |  | 3997.6632 | 0.019414 | 122.2299 | 137.9
|  |  | 122.3045 | 40.455 | [6 empty] | 122.3045 | 45.14 | [6 empty] | 122.3045 | 137.53
|  |  | 122.3792 | 41.92 | [6 empty] | 122.3792 | 45.59 | [6 empty] | 122.3792 | 137.89
|  |  | 122.4539 | 42.76 | [6 empty] | 122.4539 | 45.91 | [6 empty] | 122.4539 | 138.255
|  |  | 122.5285 | 42.175 | [6 empty] | 122.5285 | 46.305 | [6 empty] | 122.5285 | 138.61
|  |  | 122.6032 | 42.88 | [6 empty] | 122.6032 | 46.66 | [6 empty] | 122.6032 | 137.995
|  |  | 122.6779 | 42.445 | [6 empty] | 122.6779 | 46.82 | [6 empty] | 122.6779 | 136.78
|  |  | 122.7525 | 43.54 | [6 empty] | 122.7525 | 46.655 | [6 empty] | 122.7525 | 136.17
|  |  | 122.8272 | 42.75 | [6 empty] | 122.8272 | 46.65 | [6 empty] | 122.8272 | 136.275
|  |  | 122.9019 | 42.175 | [6 empty] | 122.9019 | 46.825 | [6 empty] | 122.9019 | 136.44
|  |  | 122.9765 | 41.515 | [6 empty] | 122.9765 | 47.015 | [6 empty] | 122.9765 | 136.145
|  |  | 123.0512 | 40.615 | [6 empty] | 123.0512 | 46.95 | [6 empty] | 123.0512 | 136.265
|  |  | 123.1259 | 40.54 | [6 empty] | 123.1259 | 46.91 | [6 empty] | 123.1259 | 136.195
|  |  | 123.2005 | 41.49 | [6 empty] | 123.2005 | 46.91 | [6 empty] | 123.2005 | 136.28
|  |  | 123.2752 | 42.615 | [6 empty] | 123.2752 | 46.825 | [6 empty] | 123.2752 | 136.39
|  |  | 123.3499 | 41.33 | [6 empty] | 123.3499 | 46.74 | [6 empty] | 123.3499 | 136.38
|  |  | 123.4246 | 41.365 | [6 empty] | 123.4246 | 46.605 | [6 empty] | 123.4246 | 136.42
|  |  | 123.4992 | 41.885 | [6 empty] | 123.4992 | 46.425 | [6 empty] | 123.4992 | 136.58
|  |  | 123.5739 | 42.055 | [6 empty] | 123.5739 | 46.73 | [6 empty] | 123.5739 | 136.695
|  |  | 123.6486 | 41.985 | [6 empty] | 123.6486 | 47.515 | [6 empty] | 123.6486 | 136.72
|  |  | 123.7232 | 42.835 | [6 empty] | 123.7232 | 48.485 | [6 empty] | 123.7232 | 135.995
|  |  | 123.7979 | 42.73 | [6 empty] | 123.7979 | 48.865 | [6 empty] | 123.7979 | 134.79
|  |  | 123.8726 | 42.24 | [6 empty] | 123.8726 | 49.18 | [6 empty] | 123.8726 | 134.9
|  |  | 123.9472 | 42.325 | [6 empty] | 123.9472 | 49.675 | [6 empty] | 123.9472 | 135.315
|  |  | 124.0219 | 42.395 | [6 empty] | 124.0219 | 49.825 | [6 empty] | 124.0219 | 135.09
|  |  | 124.0965 | 43.54 | [6 empty] | 124.0965 | 49.975 | [6 empty] | 124.0965 | 134.685
|  |  | 124.1712 | 42.845 | [6 empty] | 124.1712 | 50.02 | [6 empty] | 124.1712 | 134.285
|  |  | 124.2459 | 43.02 | [6 empty] | 124.2459 | 50.05 | [6 empty] | 124.2459 | 134.645
|  |  | 124.3206 | 43.04 | [6 empty] | 124.3206 | 50.09 | [6 empty] | 124.3206 | 134.9
|  |  | 124.3952 | 43.565 | [6 empty] | 124.3952 | 50.08 | [6 empty] | 124.3952 | 134.9
|  |  | 124.4699 | 43.28 | [6 empty] | 124.4699 | 49.995 | [6 empty] | 124.4699 | 134.56
|  |  | 124.5446 | 43.115 | [6 empty] | 124.5446 | 49.79 | [6 empty] | 124.5446 | 133.995
|  |  | 124.6192 | 42.45 | [6 empty] | 124.6192 | 49.395 | [6 empty] | 124.6192 | 133.64
|  |  | 124.6939 | 42.77 | [6 empty] | 124.6939 | 49.105 | [6 empty] | 124.6939 | 133.29
|  |  | 124.7686 | 41.705 | [6 empty] | 124.7686 | 48.785 | [6 empty] | 124.7686 | 132.805
|  |  | 124.8432 | 42.23 | [6 empty] | 124.8432 | 48.855 | [6 empty] | 124.8432 | 132.395
|  |  | 124.9179 | 43.315 | [6 empty] | 124.9179 | 49.365 | [6 empty] | 124.9179 | 132.26
|  |  | 124.9926 | 43.67 | [6 empty] | 124.9926 | 49.92 | [6 empty] | 124.9926 | 132.42
|  |  | 125.0672 | 43.345 | [6 empty] | 125.0672 | 50.26 | [6 empty] | 125.0672 | 132.98
|  |  | 125.1419 | 42.765 | [6 empty] | 125.1419 | 50.285 | [6 empty] | 125.1419 | 132.905
|  |  | 125.2166 | 43.38 | [6 empty] | 125.2166 | 50.35 | [6 empty] | 125.2166 | 132.075
|  |  | 125.2912 | 43.995 | [6 empty] | 125.2912 | 50.38 | [6 empty] | 125.2912 | 132.035
|  |  | 125.3659 | 43.54 | [6 empty] | 125.3659 | 50.38 | [6 empty] | 125.3659 | 131.96
|  |  | 125.4406 | 43.615 | [6 empty] | 125.4406 | 50.39 | [6 empty] | 125.4406 | 131.73
|  |  | 125.5152 | 44.115 | [6 empty] | 125.5152 | 50.265 | [6 empty] | 125.5152 | 131.575
|  |  | 125.5899 | 43.725 | [6 empty] | 125.5899 | 50.105 | [6 empty] | 125.5899 | 131.855
|  |  | 125.6646 | 42.935 | [6 empty] | 125.6646 | 49.845 | [6 empty] | 125.6646 | 132.27
|  |  | 125.7392 | 42.55 | [6 empty] | 125.7392 | 49.3 | [6 empty] | 125.7392 | 132.75
|  |  | 125.8139 | 42.945 | [6 empty] | 125.8139 | 49.005 | [6 empty] | 125.8139 | 132.975
|  |  | 125.8886 | 43.265 | [6 empty] | 125.8886 | 48.745 | [6 empty] | 125.8886 | 132.575
|  |  | 125.9632 | 42.49 | [6 empty] | 125.9632 | 48.26 | [6 empty] | 125.9632 | 132.19
|  |  | 126.0379 | 43.525 | [6 empty] | 126.0379 | 47.745 | [6 empty] | 126.0379 | 132.38
|  |  | 126.1126 | 42.885 | [6 empty] | 126.1126 | 47.245 | [6 empty] | 126.1126 | 133.44
|  |  | 126.1872 | 42.01 | [6 empty] | 126.1872 | 46.8 | [6 empty] | 126.1872 | 133.585
|  |  | 126.2619 | 42.515 | [6 empty] | 126.2619 | 46.26 | [6 empty] | 126.2619 | 133.145
|  |  | 126.3366 | 42.575 | [6 empty] | 126.3366 | 46.315 | [6 empty] | 126.3366 | 131.83
|  |  | 126.4112 | 43.58 | [6 empty] | 126.4112 | 46.675 | [6 empty] | 126.4112 | 131.33
|  |  | 126.4859 | 42.77 | [6 empty] | 126.4859 | 47.105 | [6 empty] | 126.4859 | 131.225
|  |  | 126.5606 | 42.58 | [6 empty] | 126.5606 | 47.27 | [6 empty] | 126.5606 | 131.81
|  |  | 126.6352 | 43.385 | [6 empty] | 126.6352 | 47.27 | [6 empty] | 126.6352 | 132.14
|  |  | 126.7099 | 43.235 | [6 empty] | 126.7099 | 47.36 | [6 empty] | 126.7099 | 132.415
|  |  | 126.7846 | 43.375 | [6 empty] | 126.7846 | 47.44 | [6 empty] | 126.7846 | 132.645
|  |  | 126.8592 | 43.44 | [6 empty] | 126.8592 | 47.34 | [6 empty] | 126.8592 | 132.035
|  |  | 126.9339 | 43.81 | [6 empty] | 126.9339 | 47.165 | [6 empty] | 126.9339 | 131.62
|  |  | 127.0086 | 43.155 | [6 empty] | 127.0086 | 47.06 | [6 empty] | 127.0086 | 131.97
|  |  | 127.0832 | 42.58 | [6 empty] | 127.0832 | 46.95 | [6 empty] | 127.0832 | 132.15
|  |  | 127.1579 | 41.79 | [6 empty] | 127.1579 | 46.685 | [6 empty] | 127.1579 | 132.12
|  |  | 127.2326 | 43.74 | [6 empty] | 127.2326 | 46.395 | [6 empty] | 127.2326 | 132.05
|  |  | 127.3072 | 44.25 | [6 empty] | 127.3072 | 46.1 | [6 empty] | 127.3072 | 132.165
|  |  | 127.3819 | 43.405 | [6 empty] | 127.3819 | 45.485 | [6 empty] | 127.3819 | 132.195
|  |  | 127.4566 | 43.33 | [6 empty] | 127.4566 | 45.02 | [6 empty] | 127.4566 | 132.63
|  |  | 127.5312 | 42.47 | [6 empty] | 127.5312 | 44.75 | [6 empty] | 127.5312 | 133.265
|  |  | 127.6059 | 42.205 | [6 empty] | 127.6059 | 44.205 | [6 empty] | 127.6059 | 133.43
|  |  | 127.6806 | 41.185 | [6 empty] | 127.6806 | 43.895 | [6 empty] | 127.6806 | 133.81
|  |  | 127.7552 | 41.825 | [6 empty] | 127.7552 | 43.87 | [6 empty] | 127.7552 | 133.81
|  |  | 127.8299 | 41.885 | [6 empty] | 127.8299 | 44.645 | [6 empty] | 127.8299 | 134.025
|  |  | 127.9046 | 42.65 | [6 empty] | 127.9046 | 45.32 | [6 empty] | 127.9046 | 134.285
|  |  | 127.9792 | 43.28 | [6 empty] | 127.9792 | 45.66 | [6 empty] | 127.9792 | 134.58
|  |  | 128.0539 | 42.58 | [6 empty] | 128.0539 | 45.7 | [6 empty] | 128.0539 | 134.805
|  |  | 128.1286 | 42.44 | [6 empty] | 128.1286 | 45.61 | [6 empty] | 128.1286 | 135.07
|  |  | 128.2032 | 42.44 | [6 empty] | 128.2032 | 45.54 | [6 empty] | 128.2032 | 135.195
|  |  | 128.2779 | 42.26 | [6 empty] | 128.2779 | 45.56 | [6 empty] | 128.2779 | 135.075
|  |  | 128.3526 | 42.355 | [6 empty] | 128.3526 | 45.545 | [6 empty] | 128.3526 | 134.81
|  |  | 128.4272 | 41.945 | [6 empty] | 128.4272 | 45.34 | [6 empty] | 128.4272 | 135.12
|  |  | 128.5019 | 42.035 | [6 empty] | 128.5019 | 45.095 | [6 empty] | 128.5019 | 135.545
|  |  | 128.5766 | 41.32 | [6 empty] | 128.5766 | 44.72 | [6 empty] | 128.5766 | 136.2
|  |  | 128.6512 | 42.175 | [6 empty] | 128.6512 | 44.59 | [6 empty] | 128.6512 | 137.32
|  |  | 128.7259 | 41.895 | [6 empty] | 128.7259 | 44.32 | [6 empty] | 128.7259 | 137.78
|  |  | 128.8006 | 42.105 | [6 empty] | 128.8006 | 43.825 | [6 empty] | 128.8006 | 137.975
|  |  | 128.8752 | 41.42 | [6 empty] | 128.8752 | 43.545 | [6 empty] | 128.8752 | 137.65
|  |  | 128.9499 | 42.49 | [6 empty] | 128.9499 | 43.115 | [6 empty] | 128.9499 | 137.29
|  |  | 129.0246 | 42.41 | [6 empty] | 129.0246 | 42.49 | [6 empty] | 129.0246 | 137.515
|  |  | 129.0992 | 42.16 | [6 empty] | 129.0992 | 42.265 | [6 empty] | 129.0992 | 138.01
|  |  | 129.1739 | 40.79 | [6 empty] | 129.1739 | 42.615 | [6 empty] | 129.1739 | 138.415
|  |  | 129.2486 | 41.25 | [6 empty] | 129.2486 | 43.68 | [6 empty] | 129.2486 | 138.635
|  |  | 129.3232 | 43.035 | [6 empty] | 129.3232 | 44.09 | [6 empty] | 129.3232 | 137.9
|  |  | 129.3979 | 42.325 | [6 empty] | 129.3979 | 44.385 | [6 empty] | 129.3979 | 137.815
|  |  | 129.4726 | 42.29 | [6 empty] | 129.4726 | 44.53 | [6 empty] | 129.4726 | 138.44
|  |  | 129.5472 | 41.77 | [6 empty] | 129.5472 | 44.7 | [6 empty] | 129.5472 | 138.855
|  |  | 129.6219 | 42.46 | [6 empty] | 129.6219 | 44.92 | [6 empty] | 129.6219 | 138.42
|  |  | 129.6966 | 43.62 | [6 empty] | 129.6966 | 44.955 | [6 empty] | 129.6966 | 138.155
|  |  | 129.7712 | 42.49 | [6 empty] | 129.7712 | 44.805 | [6 empty] | 129.7712 | 137.98
|  |  | 129.8459 | 41.925 | [6 empty] | 129.8459 | 44.74 | [6 empty] | 129.8459 | 138.03
|  |  | 129.9206 | 41.66 | [6 empty] | 129.9206 | 44.74 | [6 empty] | 129.9206 | 138.18
|  |  | 129.9952 | 41.065 | [6 empty] | 129.9952 | 44.64 | [6 empty] | 129.9952 | 138.575
|  |  | 130.0699 | 41.31 | [6 empty] | 130.0699 | 44.495 | [6 empty] | 130.0699 | 138.895
|  |  | 130.1446 | 41.545 | [6 empty] | 130.1446 | 44.395 | [6 empty] | 130.1446 | 138.85
|  |  | 130.2193 | 40.89 | [6 empty] | 130.2193 | 44.12 | [6 empty] | 130.2193 | 138.785
|  |  | 130.2939 | 41.67 | [6 empty] | 130.2939 | 43.9 | [6 empty] | 130.2939 | 138.345
|  |  | 130.3686 | 41.4 | [6 empty] | 130.3686 | 43.89 | [6 empty] | 130.3686 | 138.095
|  |  | 130.4433 | 41.42 | [6 empty] | 130.4433 | 44.285 | [6 empty] | 130.4433 | 137.725
|  |  | 130.5179 | 41.67 | [6 empty] | 130.5179 | 44.735 | [6 empty] | 130.5179 | 137.315
|  |  | 130.5926 | 41.67 | [6 empty] | 130.5926 | 45.135 | [6 empty] | 130.5926 | 137.015
|  |  | 130.6673 | 41.65 | [6 empty] | 130.6673 | 45.59 | [6 empty] | 130.6673 | 137.625
|  |  | 130.7419 | 41.79 | [6 empty] | 130.7419 | 45.775 | [6 empty] | 130.7419 | 137.945
|  |  | 130.8166 | 41.32 | [6 empty] | 130.8166 | 45.845 | [6 empty] | 130.8166 | 137.675
|  |  | 130.8913 | 42.57 | [6 empty] | 130.8913 | 46 | [6 empty] | 130.8913 | 137.165
|  |  | 130.9659 | 42.655 | [6 empty] | 130.9659 | 46.15 | [6 empty] | 130.9659 | 136.81
|  |  | 131.0406 | 43.14 | [6 empty] | 131.0406 | 46.255 | [6 empty] | 131.0406 | 136.61
|  |  | 131.1152 | 43.255 | [6 empty] | 131.1152 | 46.37 | [6 empty] | 131.1152 | 136.89
|  |  | 131.1899 | 42.725 | [6 empty] | 131.1899 | 46.34 | [6 empty] | 131.1899 | 136.73
|  |  | 131.2646 | 43.42 | [6 empty] | 131.2646 | 46.29 | [6 empty] | 131.2646 | 136.6
|  |  | 131.3392 | 42.615 | [6 empty] | 131.3392 | 46.09 | [6 empty] | 131.3392 | 136.935
|  |  | 131.4139 | 42.615 | [6 empty] | 131.4139 | 46.055 | [6 empty] | 131.4139 | 137.02
|  |  | 131.4886 | 42.535 | [6 empty] | 131.4886 | 45.965 | [6 empty] | 131.4886 | 136.485
|  |  | 131.5632 | 41.955 | [6 empty] | 131.5632 | 45.73 | [6 empty] | 131.5632 | 135.825
|  |  | 131.6379 | 42.535 | [6 empty] | 131.6379 | 45.555 | [6 empty] | 131.6379 | 135.17
|  |  | 131.7126 | 43.15 | [6 empty] | 131.7126 | 45.31 | [6 empty] | 131.7126 | 135.41
|  |  | 131.7872 | 42.17 | [6 empty] | 131.7872 | 45.43 | [6 empty] | 131.7872 | 135.995
|  |  | 131.8619 | 41.505 | [6 empty] | 131.8619 | 46.09 | [6 empty] | 131.8619 | 136.51
|  |  | 131.9366 | 41.99 | [6 empty] | 131.9366 | 46.725 | [6 empty] | 131.9366 | 136.165
|  |  | 132.0113 | 41.98 | [6 empty] | 132.0113 | 47.385 | [6 empty] | 132.0113 | 135.495
|  |  | 132.0859 | 42.27 | [6 empty] | 132.0859 | 47.805 | [6 empty] | 132.0859 | 135.08
|  |  | 132.1606 | 42.64 | [6 empty] | 132.1606 | 48.145 | [6 empty] | 132.1606 | 134.905
|  |  | 132.2353 | 42.08 | [6 empty] | 132.2353 | 48.44 | [6 empty] | 132.2353 | 134.835
|  |  | 132.3099 | 42.48 | [6 empty] | 132.3099 | 48.715 | [6 empty] | 132.3099 | 134.49
|  |  | 132.3846 | 42.645 | [6 empty] | 132.3846 | 48.795 | [6 empty] | 132.3846 | 134.44
|  |  | 132.4593 | 42.32 | [6 empty] | 132.4593 | 48.97 | [6 empty] | 132.4593 | 134.785
|  |  | 132.5339 | 42.265 | [6 empty] | 132.5339 | 49.285 | [6 empty] | 132.5339 | 134.71
|  |  | 132.6086 | 42.58 | [6 empty] | 132.6086 | 49.34 | [6 empty] | 132.6086 | 134.275
|  |  | 132.6833 | 43.535 | [6 empty] | 132.6833 | 49.08 | [6 empty] | 132.6833 | 134.405
|  |  | 132.7579 | 42.675 | [6 empty] | 132.7579 | 48.7 | [6 empty] | 132.7579 | 135.105
|  |  | 132.8326 | 42.735 | [6 empty] | 132.8326 | 48.55 | [6 empty] | 132.8326 | 135.08
|  |  | 132.9073 | 43.22 | [6 empty] | 132.9073 | 48.4 | [6 empty] | 132.9073 | 134.615
|  |  | 132.9819 | 42.765 | [6 empty] | 132.9819 | 48.11 | [6 empty] | 132.9819 | 134.465
|  |  | 133.0566 | 42.61 | [6 empty] | 133.0566 | 47.895 | [6 empty] | 133.0566 | 134.215
|  |  | 133.1313 | 43.32 | [6 empty] | 133.1313 | 47.565 | [6 empty] | 133.1313 | 134.6
|  |  | 133.2059 | 43.94 | [6 empty] | 133.2059 | 47.215 | [6 empty] | 133.2059 | 134.93
|  |  | 133.2806 | 43.615 | [6 empty] | 133.2806 | 47.675 | [6 empty] | 133.2806 | 134.925
|  |  | 133.3553 | 44.53 | [6 empty] | 133.3553 | 48.38 | [6 empty] | 133.3553 | 135.165
|  |  | 133.4299 | 43.795 | [6 empty] | 133.4299 | 49.19 | [6 empty] | 133.4299 | 134.83
|  |  | 133.5046 | 43.915 | [6 empty] | 133.5046 | 49.43 | [6 empty] | 133.5046 | 133.745
|  |  | 133.5793 | 43.225 | [6 empty] | 133.5793 | 49.395 | [6 empty] | 133.5793 | 133.535
|  |  | 133.6539 | 42.93 | [6 empty] | 133.6539 | 49.46 | [6 empty] | 133.6539 | 133.47
|  |  | 133.7286 | 43.435 | [6 empty] | 133.7286 | 49.46 | [6 empty] | 133.7286 | 133.61
|  |  | 133.8033 | 43.65 | [6 empty] | 133.8033 | 49.5 | [6 empty] | 133.8033 | 133.54
|  |  | 133.8779 | 43.745 | [6 empty] | 133.8779 | 49.475 | [6 empty] | 133.8779 | 133.36
|  |  | 133.9526 | 43.535 | [6 empty] | 133.9526 | 49.36 | [6 empty] | 133.9526 | 133.13
|  |  | 134.0273 | 42.325 | [6 empty] | 134.0273 | 49.065 | [6 empty] | 134.0273 | 133.045
|  |  | 134.1019 | 42.36 | [6 empty] | 134.1019 | 48.485 | [6 empty] | 134.1019 | 133.485
|  |  | 134.1766 | 41.82 | [6 empty] | 134.1766 | 48.25 | [6 empty] | 134.1766 | 134.245
|  |  | 134.2513 | 42.075 | [6 empty] | 134.2513 | 48.025 | [6 empty] | 134.2513 | 134.605
|  |  | 134.3259 | 43.46 | [6 empty] | 134.3259 | 47.62 | [6 empty] | 134.3259 | 134.55
|  |  | 134.4006 | 43.785 | [6 empty] | 134.4006 | 47.265 | [6 empty] | 134.4006 | 134.955
|  |  | 134.4753 | 42.295 | [6 empty] | 134.4753 | 46.87 | [6 empty] | 134.4753 | 135.29
|  |  | 134.5499 | 42.58 | [6 empty] | 134.5499 | 46.065 | [6 empty] | 134.5499 | 134.955
|  |  | 134.6246 | 42.935 | [6 empty] | 134.6246 | 45.335 | [6 empty] | 134.6246 | 134.335
|  |  | 134.6993 | 42.84 | [6 empty] | 134.6993 | 44.84 | [6 empty] | 134.6993 | 134.255
|  |  | 134.7739 | 43.375 | [6 empty] | 134.7739 | 45.015 | [6 empty] | 134.7739 | 134.515
|  |  | 134.8486 | 42.95 | [6 empty] | 134.8486 | 45.915 | [6 empty] | 134.8486 | 134.63
|  |  | 134.9233 | 43.075 | [6 empty] | 134.9233 | 46.35 | [6 empty] | 134.9233 | 134.175
|  |  | 134.9979 | 43.05 | [6 empty] | 134.9979 | 46.655 | [6 empty] | 134.9979 | 133.87
|  |  | 135.0726 | 42.95 | [6 empty] | 135.0726 | 46.655 | [6 empty] | 135.0726 | 134.015
|  |  | 135.1473 | 44.04 | [6 empty] | 135.1473 | 46.6 | [6 empty] | 135.1473 | 133.885
|  |  | 135.2219 | 44.165 | [6 empty] | 135.2219 | 46.645 | [6 empty] | 135.2219 | 133.6
|  |  | 135.2966 | 43.225 | [6 empty] | 135.2966 | 46.54 | [6 empty] | 135.2966 | 133.395
|  |  | 135.3713 | 42.75 | [6 empty] | 135.3713 | 46.28 | [6 empty] | 135.3713 | 133.67
|  |  | 135.4459 | 43.59 | [6 empty] | 135.4459 | 46.125 | [6 empty] | 135.4459 | 133.92
|  |  | 135.5206 | 43.58 | [6 empty] | 135.5206 | 45.945 | [6 empty] | 135.5206 | 133.99
|  |  | 135.5953 | 42.795 | [6 empty] | 135.5953 | 45.8 | [6 empty] | 135.5953 | 134.225
|  |  | 135.6699 | 43.085 | [6 empty] | 135.6699 | 45.61 | [6 empty] | 135.6699 | 134.615
|  |  | 135.7446 | 43.065 | [6 empty] | 135.7446 | 45.305 | [6 empty] | 135.7446 | 135.025
|  |  | 135.8193 | 44.36 | [6 empty] | 135.8193 | 45.03 | [6 empty] | 135.8193 | 135.295
|  |  | 135.8939 | 42.735 | [6 empty] | 135.8939 | 44.575 | [6 empty] | 135.8939 | 135.155
|  |  | 135.9686 | 42.62 | [6 empty] | 135.9686 | 43.91 | [6 empty] | 135.9686 | 134.605
|  |  | 136.0433 | 42.77 | [6 empty] | 136.0433 | 43.97 | [6 empty] | 136.0433 | 134.42
|  |  | 136.1179 | 41.725 | [6 empty] | 136.1179 | 44.28 | [6 empty] | 136.1179 | 134.9
|  |  | 136.1926 | 42.155 | [6 empty] | 136.1926 | 44.775 | [6 empty] | 136.1926 | 135.31
|  |  | 136.2673 | 42.845 | [6 empty] | 136.2673 | 44.885 | [6 empty] | 136.2673 | 135.385
|  |  | 136.3419 | 42.85 | [6 empty] | 136.3419 | 45.025 | [6 empty] | 136.3419 | 135.45
|  |  | 136.4166 | 42.735 | [6 empty] | 136.4166 | 45.295 | [6 empty] | 136.4166 | 135.275
|  |  | 136.4913 | 40.995 | [6 empty] | 136.4913 | 45.2 | [6 empty] | 136.4913 | 135.18
|  |  | 136.5659 | 40.25 | [6 empty] | 136.5659 | 44.98 | [6 empty] | 136.5659 | 135.61
|  |  | 136.6406 | 40.385 | [6 empty] | 136.6406 | 44.97 | [6 empty] | 136.6406 | 136.335
|  |  | 136.7153 | 41.605 | [6 empty] | 136.7153 | 45.06 | [6 empty] | 136.7153 | 136.965
|  |  | 136.7899 | 42.445 | [6 empty] | 136.7899 | 44.91 | [6 empty] | 136.7899 | 136.88
|  |  | 136.8646 | 42.165 | [6 empty] | 136.8646 | 44.575 | [6 empty] | 136.8646 | 136.685
|  |  | 136.9393 | 41.87 | [6 empty] | 136.9393 | 44.42 | [6 empty] | 136.9393 | 136.48
|  |  | 137.0139 | 41.855 | [6 empty] | 137.0139 | 44.16 | [6 empty] | 137.0139 | 136.615
|  |  | 137.0886 | 41.57 | [6 empty] | 137.0886 | 43.645 | [6 empty] | 137.0886 | 136.69
|  |  | 137.1633 | 40.695 | [6 empty] | 137.1633 | 43.23 | [6 empty] | 137.1633 | 137.11
|  |  | 137.2379 | 40.74 | [6 empty] | 137.2379 | 42.845 | [6 empty] | 137.2379 | 137.09
|  |  | 137.3126 | 40.24 | [6 empty] | 137.3126 | 42.36 | [6 empty] | 137.3126 | 136.865
|  |  | 137.3873 | 41.05 | [6 empty] | 137.3873 | 42.235 | [6 empty] | 137.3873 | 137.145
|  |  | 137.4619 | 41.685 | [6 empty] | 137.4619 | 42.56 | [6 empty] | 137.4619 | 137.495
|  |  | 137.5366 | 41.99 | [6 empty] | 137.5366 | 43.02 | [6 empty] | 137.5366 | 137.455
|  |  | 137.6113 | 42.76 | [6 empty] | 137.6113 | 43.585 | [6 empty] | 137.6113 | 137.65
|  |  | 137.6859 | 41.925 | [6 empty] | 137.6859 | 43.775 | [6 empty] | 137.6859 | 138.14
|  |  | 137.7606 | 41.48 | [6 empty] | 137.7606 | 44.03 | [6 empty] | 137.7606 | 138.66
|  |  | 137.8353 | 42.13 | [6 empty] | 137.8353 | 44.31 | [6 empty] | 137.8353 | 138.575
|  |  | 137.9099 | 42.665 | [6 empty] | 137.9099 | 44.4 | [6 empty] | 137.9099 | 138.685
|  |  | 137.9846 | 43.31 | [6 empty] | 137.9846 | 44.385 | [6 empty] | 137.9846 | 138.67
|  |  | 138.0593 | 41.025 | [6 empty] | 138.0593 | 44.54 | [6 empty] | 138.0593 | 137.425
|  |  | 138.1339 | 41.145 | [6 empty] | 138.1339 | 44.61 | [6 empty] | 138.1339 | 136.825
|  |  | 138.2086 | 41.66 | [6 empty] | 138.2086 | 44.565 | [6 empty] | 138.2086 | 137.075
|  |  | 138.2833 | 41.83 | [6 empty] | 138.2833 | 44.525 | [6 empty] | 138.2833 | 138.305
|  |  | 138.358 | 42.715 | [6 empty] | 138.358 | 44.485 | [6 empty] | 138.358 | 138.79
|  |  | 138.4326 | 41.83 | [6 empty] | 138.4326 | 44.44 | [6 empty] | 138.4326 | 138.585
|  |  | 138.5073 | 41.89 | [6 empty] | 138.5073 | 44.265 | [6 empty] | 138.5073 | 137.9
|  |  | 138.582 | 42.06 | [6 empty] | 138.582 | 44 | [6 empty] | 138.582 | 137.74
|  |  | 138.6566 | 42.435 | [6 empty] | 138.6566 | 43.725 | [6 empty] | 138.6566 | 137.935
|  |  | 138.7313 | 40.62 | [6 empty] | 138.7313 | 43.59 | [6 empty] | 138.7313 | 137.76
|  |  | 138.806 | 40.165 | [6 empty] | 138.806 | 43.95 | [6 empty] | 138.806 | 137.735
|  |  | 138.8806 | 40.415 | [6 empty] | 138.8806 | 44.68 | [6 empty] | 138.8806 | 138.21
|  |  | 138.9553 | 41.41 | [6 empty] | 138.9553 | 45.5 | [6 empty] | 138.9553 | 138.59
|  |  | 139.03 | 41.84 | [6 empty] | 139.03 | 45.895 | [6 empty] | 139.03 | 138.685
|  |  | 139.1046 | 42.16 | [6 empty] | 139.1046 | 46.21 | [6 empty] | 139.1046 | 138.255
|  |  | 139.1793 | 42.37 | [6 empty] | 139.1793 | 46.57 | [6 empty] | 139.1793 | 137.52
|  |  | 139.254 | 41.695 | [6 empty] | 139.254 | 46.885 | [6 empty] | 139.254 | 136.5
|  |  | 139.3286 | 43.175 | [6 empty] | 139.3286 | 47.015 | [6 empty] | 139.3286 | 136.26
|  |  | 139.4033 | 43.085 | [6 empty] | 139.4033 | 47 | [6 empty] | 139.4033 | 135.97
|  |  | 139.478 | 41.975 | [6 empty] | 139.478 | 46.985 | [6 empty] | 139.478 | 135.795
|  |  | 139.5526 | 40.905 | [6 empty] | 139.5526 | 47.07 | [6 empty] | 139.5526 | 136.335
|  |  | 139.6273 | 40.165 | [6 empty] | 139.6273 | 47.165 | [6 empty] | 139.6273 | 136.7
|  |  | 139.702 | 40.475 | [6 empty] | 139.702 | 47.095 | [6 empty] | 139.702 | 136.4
|  |  | 139.7766 | 39.805 | [6 empty] | 139.7766 | 47.08 | [6 empty] | 139.7766 | 136.445
|  |  | 139.8513 | 41.25 | [6 empty] | 139.8513 | 47.06 | [6 empty] | 139.8513 | 136.245
|  |  | 139.9259 | 40.665 | [6 empty] | 139.9259 | 46.87 | [6 empty] | 139.9259 | 136.865
|  |  | 140.0006 | 40.21 | [6 empty] | 140.0006 | 46.755 | [6 empty] | 140.0006 | 137.36
|  |  | 140.0753 | 42.01 | [6 empty] | 140.0753 | 46.55 | [6 empty] | 140.0753 | 137.435
|  |  | 140.15 | 41.88 | [6 empty] | 140.15 | 46.5 | [6 empty] | 140.15 | 136.745
|  |  | 140.2246 | 41.9 | [6 empty] | 140.2246 | 47.01 | [6 empty] | 140.2246 | 136.195
|  |  | 140.2993 | 41.55 | [6 empty] | 140.2993 | 47.725 | [6 empty] | 140.2993 | 135.88
|  |  | 140.374 | 41.55 | [6 empty] | 140.374 | 48.735 | [6 empty] | 140.374 | 135.975
|  |  | 140.4486 | 42.88 | [6 empty] | 140.4486 | 49.195 | [6 empty] | 140.4486 | 136.01
|  |  | 140.5233 | 43.425 | [6 empty] | 140.5233 | 49.6 | [6 empty] | 140.5233 | 135.515
|  |  | 140.598 | 43.075 | [6 empty] | 140.598 | 49.935 | [6 empty] | 140.598 | 134.7
|  |  | 140.6726 | 43.925 | [6 empty] | 140.6726 | 50.075 | [6 empty] | 140.6726 | 134.435
|  |  | 140.7473 | 43.15 | [6 empty] | 140.7473 | 50.245 | [6 empty] | 140.7473 | 134.685
|  |  | 140.822 | 42.355 | [6 empty] | 140.822 | 50.515 | [6 empty] | 140.822 | 134.745
|  |  | 140.8966 | 41.935 | [6 empty] | 140.8966 | 50.57 | [6 empty] | 140.8966 | 134.56
|  |  | 140.9713 | 43.02 | [6 empty] | 140.9713 | 50.505 | [6 empty] | 140.9713 | 134.455
|  |  | 141.046 | 43.12 | [6 empty] | 141.046 | 50.23 | [6 empty] | 141.046 | 134.55
|  |  | 141.1206 | 43.96 | [6 empty] | 141.1206 | 49.975 | [6 empty] | 141.1206 | 134.425
|  |  | 141.1953 | 43.775 | [6 empty] | 141.1953 | 49.78 | [6 empty] | 141.1953 | 134.31
|  |  | 141.27 | 42.39 | [6 empty] | 141.27 | 49.605 | [6 empty] | 141.27 | 134.145
|  |  | 141.3446 | 42.21 | [6 empty] | 141.3446 | 49.3 | [6 empty] | 141.3446 | 133.975
|  |  | 141.4193 | 42.27 | [6 empty] | 141.4193 | 49.115 | [6 empty] | 141.4193 | 133.755
|  |  | 141.494 | 42.985 | [6 empty] | 141.494 | 49.15 | [6 empty] | 141.494 | 133.55
|  |  | 141.5686 | 43.73 | [6 empty] | 141.5686 | 49.81 | [6 empty] | 141.5686 | 133.78
|  |  | 141.6433 | 44.335 | [6 empty] | 141.6433 | 50.19 | [6 empty] | 141.6433 | 133.52
|  |  | 141.718 | 44.265 | [6 empty] | 141.718 | 50.505 | [6 empty] | 141.718 | 132.21
|  |  | 141.7926 | 43.725 | [6 empty] | 141.7926 | 50.655 | [6 empty] | 141.7926 | 131.875
|  |  | 141.8673 | 42.555 | [6 empty] | 141.8673 | 50.685 | [6 empty] | 141.8673 | 132.44
|  |  | 141.942 | 43.77 | [6 empty] | 141.942 | 50.725 | [6 empty] | 141.942 | 132.85
|  |  | 142.0166 | 44.44 | [6 empty] | 142.0166 | 50.68 | [6 empty] | 142.0166 | 132.4
|  |  | 142.0913 | 43.3 | [6 empty] | 142.0913 | 50.6 | [6 empty] | 142.0913 | 132.52
|  |  | 142.166 | 44.095 | [6 empty] | 142.166 | 50.46 | [6 empty] | 142.166 | 132.655
|  |  | 142.2406 | 44.265 | [6 empty] | 142.2406 | 50.28 | [6 empty] | 142.2406 | 132.29
|  |  | 142.3153 | 44.695 | [6 empty] | 142.3153 | 50.055 | [6 empty] | 142.3153 | 132.17
|  |  | 142.39 | 43.805 | [6 empty] | 142.39 | 49.855 | [6 empty] | 142.39 | 132.455
|  |  | 142.4646 | 43.7 | [6 empty] | 142.4646 | 49.43 | [6 empty] | 142.4646 | 133.24
|  |  | 142.5393 | 43.045 | [6 empty] | 142.5393 | 49.06 | [6 empty] | 142.5393 | 133.345
|  |  | 142.614 | 42.19 | [6 empty] | 142.614 | 48.68 | [6 empty] | 142.614 | 132.755
|  |  | 142.6886 | 42.235 | [6 empty] | 142.6886 | 48.13 | [6 empty] | 142.6886 | 132.405
|  |  | 142.7633 | 42.645 | [6 empty] | 142.7633 | 47.58 | [6 empty] | 142.7633 | 132.56
|  |  | 142.838 | 42.765 | [6 empty] | 142.838 | 47.05 | [6 empty] | 142.838 | 132.78
|  |  | 142.9126 | 42.52 | [6 empty] | 142.9126 | 46.44 | [6 empty] | 142.9126 | 133.06
|  |  | 142.9873 | 43.485 | [6 empty] | 142.9873 | 46.51 | [6 empty] | 142.9873 | 133.02
|  |  | 143.062 | 42.365 | [6 empty] | 143.062 | 46.9 | [6 empty] | 143.062 | 132.705
|  |  | 143.1366 | 43.04 | [6 empty] | 143.1366 | 47.32 | [6 empty] | 143.1366 | 132.17
|  |  | 143.2113 | 44.3 | [6 empty] | 143.2113 | 47.465 | [6 empty] | 143.2113 | 132.18
|  |  | 143.286 | 43.425 | [6 empty] | 143.286 | 47.51 | [6 empty] | 143.286 | 132.51
|  |  | 143.3606 | 42.45 | [6 empty] | 143.3606 | 47.505 | [6 empty] | 143.3606 | 132.935
|  |  | 143.4353 | 41.74 | [6 empty] | 143.4353 | 47.485 | [6 empty] | 143.4353 | 132.885
|  |  | 143.51 | 41.545 | [6 empty] | 143.51 | 47.455 | [6 empty] | 143.51 | 132.35
|  |  | 143.5846 | 42.185 | [6 empty] | 143.5846 | 47.2 | [6 empty] | 143.5846 | 131.88
|  |  | 143.6593 | 43.11 | [6 empty] | 143.6593 | 47.065 | [6 empty] | 143.6593 | 131.71
|  |  | 143.734 | 43.625 | [6 empty] | 143.734 | 46.95 | [6 empty] | 143.734 | 131.875
|  |  | 143.8086 | 42.37 | [6 empty] | 143.8086 | 46.695 | [6 empty] | 143.8086 | 132.385
|  |  | 143.8833 | 43.25 | [6 empty] | 143.8833 | 46.465 | [6 empty] | 143.8833 | 132.81
|  |  | 143.958 | 43.355 | [6 empty] | 143.958 | 46.19 | [6 empty] | 143.958 | 132.82
|  |  | 144.0326 | 42.47 | [6 empty] | 144.0326 | 45.985 | [6 empty] | 144.0326 | 132.76
|  |  | 144.1073 | 41.685 | [6 empty] | 144.1073 | 45.845 | [6 empty] | 144.1073 | 132.7
|  |  | 144.182 | 40.495 | [6 empty] | 144.182 | 45.49 | [6 empty] | 144.182 | 133.13
|  |  | 144.2566 | 40.595 | [6 empty] | 144.2566 | 44.945 | [6 empty] | 144.2566 | 133.405
|  |  | 144.3313 | 41.29 | [6 empty] | 144.3313 | 44.14 | [6 empty] | 144.3313 | 133.73
|  |  | 144.406 | 41.555 | [6 empty] | 144.406 | 44.035 | [6 empty] | 144.406 | 134.175
|  |  | 144.4806 | 41.325 | [6 empty] | 144.4806 | 44.41 | [6 empty] | 144.4806 | 134.31
|  |  | 144.5553 | 42.21 | [6 empty] | 144.5553 | 45.265 | [6 empty] | 144.5553 | 134.19
|  |  | 144.63 | 42.935 | [6 empty] | 144.63 | 45.665 | [6 empty] | 144.63 | 134.455
|  |  | 144.7047 | 42.86 | [6 empty] | 144.7047 | 45.845 | [6 empty] | 144.7047 | 134.805
|  |  | 144.7793 | 44.25 | [6 empty] | 144.7793 | 45.92 | [6 empty] | 144.7793 | 135.39
|  |  | 144.854 | 42.08 | [6 empty] | 144.854 | 45.935 | [6 empty] | 144.854 | 135.835
|  |  | 144.9286 | 42.785 | [6 empty] | 144.9286 | 45.885 | [6 empty] | 144.9286 | 135.6
|  |  | 145.0033 | 43.11 | [6 empty] | 145.0033 | 45.81 | [6 empty] | 145.0033 | 135.195
|  |  | 145.078 | 43.095 | [6 empty] | 145.078 | 45.61 | [6 empty] | 145.078 | 135.755
|  |  | 145.1526 | 43.39 | [6 empty] | 145.1526 | 45.32 | [6 empty] | 145.1526 | 136.57
|  |  | 145.2273 | 41.505 | [6 empty] | 145.2273 | 44.97 | [6 empty] | 145.2273 | 137.22
|  |  | 145.302 | 41.64 | [6 empty] | 145.302 | 44.78 | [6 empty] | 145.302 | 137.15
|  |  | 145.3766 | 42.13 | [6 empty] | 145.3766 | 44.51 | [6 empty] | 145.3766 | 137.055
|  |  | 145.4513 | 41.975 | [6 empty] | 145.4513 | 43.92 | [6 empty] | 145.4513 | 136.69
|  |  | 145.526 | 42.34 | [6 empty] | 145.526 | 43.56 | [6 empty] | 145.526 | 136.54
|  |  | 145.6006 | 42.495 | [6 empty] | 145.6006 | 43.15 | [6 empty] | 145.6006 | 136.73
|  |  | 145.6753 | 42.82 | [6 empty] | 145.6753 | 42.615 | [6 empty] | 145.6753 | 137.985
|  |  | 145.75 | 42.025 | [6 empty] | 145.75 | 42.345 | [6 empty] | 145.75 | 138.715
|  |  | 145.8246 | 41.105 | [6 empty] | 145.8246 | 42.585 | [6 empty] | 145.8246 | 138.725
|  |  | 145.8993 | 42.455 | [6 empty] | 145.8993 | 43.57 | [6 empty] | 145.8993 | 138.015
|  |  | 145.974 | 42.87 | [6 empty] | 145.974 | 44.12 | [6 empty] | 145.974 | 137.985
|  |  | 146.0486 | 41.935 | [6 empty] | 146.0486 | 44.395 | [6 empty] | 146.0486 | 138.605
|  |  | 146.1233 | 41.75 | [6 empty] | 146.1233 | 44.53 | [6 empty] | 146.1233 | 138.825
|  |  | 146.198 | 41.525 | [6 empty] | 146.198 | 44.73 | [6 empty] | 146.198 | 138.535
|  |  | 146.2726 | 41.785 | [6 empty] | 146.2726 | 44.895 | [6 empty] | 146.2726 | 138.23
|  |  | 146.3473 | 42.11 | [6 empty] | 146.3473 | 45.04 | [6 empty] | 146.3473 | 138.27
|  |  | 146.422 | 42.275 | [6 empty] | 146.422 | 45.115 | [6 empty] | 146.422 | 138.32
|  |  | 146.4967 | 42.165 | [6 empty] | 146.4967 | 45.055 | [6 empty] | 146.4967 | 138.19
|  |  | 146.5713 | 43.115 | [6 empty] | 146.5713 | 44.94 | [6 empty] | 146.5713 | 138.525
|  |  | 146.646 | 42.85 | [6 empty] | 146.646 | 44.895 | [6 empty] | 146.646 | 138.565
|  |  | 146.7207 | 43.2 | [6 empty] | 146.7207 | 44.805 | [6 empty] | 146.7207 | 138.485
|  |  | 146.7953 | 42.655 | [6 empty] | 146.7953 | 44.6 | [6 empty] | 146.7953 | 138.645
|  |  | 146.87 | 42.03 | [6 empty] | 146.87 | 44.29 | [6 empty] | 146.87 | 138.375
|  |  | 146.9447 | 43.265 | [6 empty] | 146.9447 | 44.155 | [6 empty] | 146.9447 | 138.01
|  |  | 147.0193 | 43.07 | [6 empty] | 147.0193 | 44.205 | [6 empty] | 147.0193 | 137.955
|  |  | 147.094 | 42.915 | [6 empty] | 147.094 | 44.53 | [6 empty] | 147.094 | 138.1
|  |  | 147.1687 | 42.36 | [6 empty] | 147.1687 | 44.815 | [6 empty] | 147.1687 | 138.075
|  |  | 147.2433 | 42.115 | [6 empty] | 147.2433 | 45.25 | [6 empty] | 147.2433 | 138.125
|  |  | 147.318 | 42.545 | [6 empty] | 147.318 | 45.69 | [6 empty] | 147.318 | 138.2
|  |  | 147.3927 | 42.8 | [6 empty] | 147.3927 | 45.905 | [6 empty] | 147.3927 | 138.24
|  |  | 147.4673 | 43.56 | [6 empty] | 147.4673 | 46.105 | [6 empty] | 147.4673 | 137.985
|  |  | 147.542 | 43.215 | [6 empty] | 147.542 | 46.415 | [6 empty] | 147.542 | 137.06
|  |  | 147.6167 | 42.345 | [6 empty] | 147.6167 | 46.475 | [6 empty] | 147.6167 | 136.78
|  |  | 147.6913 | 42.17 | [6 empty] | 147.6913 | 46.45 | [6 empty] | 147.6913 | 136.59
|  |  | 147.766 | 41.72 | [6 empty] | 147.766 | 46.24 | [6 empty] | 147.766 | 136.52
|  |  | 147.8407 | 43.445 | [6 empty] | 147.8407 | 46.225 | [6 empty] | 147.8407 | 135.955
|  |  | 147.9153 | 43.68 | [6 empty] | 147.9153 | 46.3 | [6 empty] | 147.9153 | 136.025
|  |  | 147.99 | 40.575 | [6 empty] | 147.99 | 46.11 | [6 empty] | 147.99 | 136.125
|  |  | 148.0647 | 41.075 | [6 empty] | 148.0647 | 45.97 | [6 empty] | 148.0647 | 136.065
|  |  | 148.1393 | 41.97 | [6 empty] | 148.1393 | 45.87 | [6 empty] | 148.1393 | 136.17
|  |  | 148.214 | 41.76 | [6 empty] | 148.214 | 45.72 | [6 empty] | 148.214 | 136.265
|  |  | 148.2887 | 41.825 | [6 empty] | 148.2887 | 45.465 | [6 empty] | 148.2887 | 136.44
|  |  | 148.3633 | 42.415 | [6 empty] | 148.3633 | 45.455 | [6 empty] | 148.3633 | 136.465
|  |  | 148.438 | 42.095 | [6 empty] | 148.438 | 45.45 | [6 empty] | 148.438 | 136.435
|  |  | 148.5127 | 43.595 | [6 empty] | 148.5127 | 45.905 | [6 empty] | 148.5127 | 136.32
|  |  | 148.5873 | 42.88 | [6 empty] | 148.5873 | 46.56 | [6 empty] | 148.5873 | 136.13
|  |  | 148.662 | 41.515 | [6 empty] | 148.662 | 47.16 | [6 empty] | 148.662 | 135.575
|  |  | 148.7367 | 41.705 | [6 empty] | 148.7367 | 47.69 | [6 empty] | 148.7367 | 134.91
|  |  | 148.8113 | 43.63 | [6 empty] | 148.8113 | 47.965 | [6 empty] | 148.8113 | 134.635
|  |  | 148.886 | 43.79 | [6 empty] | 148.886 | 48.215 | [6 empty] | 148.886 | 134.515
|  |  | 148.9607 | 42.81 | [6 empty] | 148.9607 | 48.48 | [6 empty] | 148.9607 | 134.59
|  |  | 149.0353 | 42.515 | [6 empty] | 149.0353 | 48.655 | [6 empty] | 149.0353 | 134.785
|  |  | 149.11 | 42.045 | [6 empty] | 149.11 | 48.895 | [6 empty] | 149.11 | 135.205
|  |  | 149.1847 | 43.02 | [6 empty] | 149.1847 | 49.005 | [6 empty] | 149.1847 | 134.73
|  |  | 149.2593 | 44.09 | [6 empty] | 149.2593 | 48.95 | [6 empty] | 149.2593 | 134.235
|  |  | 149.334 | 44.29 | [6 empty] | 149.334 | 48.86 | [6 empty] | 149.334 | 133.98
|  |  | 149.4087 | 43.34 | [6 empty] | 149.4087 | 48.91 | [6 empty] | 149.4087 | 134.145
|  |  | 149.4833 | 43.39 | [6 empty] | 149.4833 | 48.89 | [6 empty] | 149.4833 | 134.095
|  |  | 149.558 | 43.585 | [6 empty] | 149.558 | 48.645 | [6 empty] | 149.558 | 133.635
|  |  | 149.6327 | 41.765 | [6 empty] | 149.6327 | 48.215 | [6 empty] | 149.6327 | 133.79
|  |  | 149.7073 | 44.025 | [6 empty] | 149.7073 | 47.94 | [6 empty] | 149.7073 | 134.23
|  |  | 149.782 | 44.85 | [6 empty] | 149.782 | 47.75 | [6 empty] | 149.782 | 134.585
|  |  | 149.8567 | 44.085 | [6 empty] | 149.8567 | 47.56 | [6 empty] | 149.8567 | 134.34
|  |  | 149.9313 | 42.55 | [6 empty] | 149.9313 | 47.875 | [6 empty] | 149.9313 | 134.055
|  |  | 150.006 | 42.67 | [6 empty] | 150.006 | 48.505 | [6 empty] | 150.006 | 133.615
|  |  | 150.0807 | 43.65 | [6 empty] | 150.0807 | 49.195 | [6 empty] | 150.0807 | 133.615
|  |  | 150.1553 | 43.31 | [6 empty] | 150.1553 | 49.49 | [6 empty] | 150.1553 | 133.62
|  |  | 150.23 | 43.535 | [6 empty] | 150.23 | 49.61 | [6 empty] | 150.23 | 133.57
|  |  | 150.3047 | 44.205 | [6 empty] | 150.3047 | 49.735 | [6 empty] | 150.3047 | 133.435
|  |  | 150.3793 | 43.355 | [6 empty] | 150.3793 | 49.715 | [6 empty] | 150.3793 | 133.21
|  |  | 150.454 | 44.96 | [6 empty] | 150.454 | 49.655 | [6 empty] | 150.454 | 133.41
|  |  | 150.5287 | 44.315 | [6 empty] | 150.5287 | 49.61 | [6 empty] | 150.5287 | 133.325
|  |  | 150.6033 | 42.83 | [6 empty] | 150.6033 | 49.56 | [6 empty] | 150.6033 | 133.625
|  |  | 150.678 | 43.74 | [6 empty] | 150.678 | 49.345 | [6 empty] | 150.678 | 133.865
|  |  | 150.7527 | 43.89 | [6 empty] | 150.7527 | 48.94 | [6 empty] | 150.7527 | 133.785
|  |  | 150.8273 | 43.55 | [6 empty] | 150.8273 | 48.355 | [6 empty] | 150.8273 | 133.4
|  |  | 150.902 | 44.09 | [6 empty] | 150.902 | 48.055 | [6 empty] | 150.902 | 133.275
|  |  | 150.9767 | 42.135 | [6 empty] | 150.9767 | 47.715 | [6 empty] | 150.9767 | 134.085
|  |  | 151.0513 | 42.45 | [6 empty] | 151.0513 | 47.42 | [6 empty] | 151.0513 | 134.72
|  |  | 151.126 | 42.77 | [6 empty] | 151.126 | 46.92 | [6 empty] | 151.126 | 135.21
|  |  | 151.2007 | 43.735 | [6 empty] | 151.2007 | 46.34 | [6 empty] | 151.2007 | 135.245
|  |  | 151.2753 | 43.305 | [6 empty] | 151.2753 | 45.53 | [6 empty] | 151.2753 | 134.84
|  |  | 151.35 | 43.07 | [6 empty] | 151.35 | 45.13 | [6 empty] | 151.35 | 134.5
|  |  | 151.4247 | 42.265 | [6 empty] | 151.4247 | 45.275 | [6 empty] | 151.4247 | 135.285
|  |  | 151.4993 | 42.295 | [6 empty] | 151.4993 | 46.03 | [6 empty] | 151.4993 | 135.45
|  |  | 151.574 | 42.915 | [6 empty] | 151.574 | 46.37 | [6 empty] | 151.574 | 134.52
|  |  | 151.6487 | 43.055 | [6 empty] | 151.6487 | 46.565 | [6 empty] | 151.6487 | 133.82
|  |  | 151.7233 | 43.665 | [6 empty] | 151.7233 | 46.75 | [6 empty] | 151.7233 | 133.9
|  |  | 151.798 | 43.225 | [6 empty] | 151.798 | 46.77 | [6 empty] | 151.798 | 134.02
|  |  | 151.8727 | 43.205 | [6 empty] | 151.8727 | 46.695 | [6 empty] | 151.8727 | 134.325
|  |  | 151.9473 | 43.02 | [6 empty] | 151.9473 | 46.45 | [6 empty] | 151.9473 | 134.815
|  |  | 152.022 | 42.79 | [6 empty] | 152.022 | 46.385 | [6 empty] | 152.022 | 134.335
|  |  | 152.0967 | 42.69 | [6 empty] | 152.0967 | 46.4 | [6 empty] | 152.0967 | 133.49
|  |  | 152.1713 | 42.39 | [6 empty] | 152.1713 | 46.105 | [6 empty] | 152.1713 | 133.62
|  |  | 152.246 | 43.72 | [6 empty] | 152.246 | 45.755 | [6 empty] | 152.246 | 134.375
|  |  | 152.3207 | 42.5 | [6 empty] | 152.3207 | 45.53 | [6 empty] | 152.3207 | 134.785
|  |  | 152.3953 | 42.73 | [6 empty] | 152.3953 | 45.18 | [6 empty] | 152.3953 | 134.67
|  |  | 152.47 | 42.455 | [6 empty] | 152.47 | 44.84 | [6 empty] | 152.47 | 134.775
|  |  | 152.5447 | 42.35 | [6 empty] | 152.5447 | 44.48 | [6 empty] | 152.5447 | 134.995
|  |  | 152.6193 | 42.465 | [6 empty] | 152.6193 | 44.065 | [6 empty] | 152.6193 | 135.25
|  |  | 152.694 | 43.165 | [6 empty] | 152.694 | 44.255 | [6 empty] | 152.694 | 135.08
|  |  | 152.7687 | 42.65 | [6 empty] | 152.7687 | 44.615 | [6 empty] | 152.7687 | 134.835
|  |  | 152.8434 | 43.915 | [6 empty] | 152.8434 | 44.995 | [6 empty] | 152.8434 | 134.94
|  |  | 152.918 | 43.1 | [6 empty] | 152.918 | 45.245 | [6 empty] | 152.918 | 135.04
|  |  | 152.9927 | 42.685 | [6 empty] | 152.9927 | 45.34 | [6 empty] | 152.9927 | 135.31
|  |  | 153.0674 | 41.995 | [6 empty] | 153.0674 | 45.345 | [6 empty] | 153.0674 | 135.55
|  |  | 153.142 | 41.835 | [6 empty] | 153.142 | 45.385 | [6 empty] | 153.142 | 135.725
|  |  | 153.2167 | 42.22 | [6 empty] | 153.2167 | 45.45 | [6 empty] | 153.2167 | 136.12
|  |  | 153.2914 | 41.615 | [6 empty] | 153.2914 | 45.48 | [6 empty] | 153.2914 | 136.615
|  |  | 153.366 | 42.66 | [6 empty] | 153.366 | 45.155 | [6 empty] | 153.366 | 136.84
|  |  | 153.4407 | 41.795 | [6 empty] | 153.4407 | 44.94 | [6 empty] | 153.4407 | 136.93
|  |  | 153.5154 | 41.17 | [6 empty] | 153.5154 | 44.835 | [6 empty] | 153.5154 | 136.665
|  |  | 153.59 | 40.825 | [6 empty] | 153.59 | 44.655 | [6 empty] | 153.59 | 136.665
|  |  | 153.6647 | 41.925 | [6 empty] | 153.6647 | 44.4 | [6 empty] | 153.6647 | 136.77
|  |  | 153.7393 | 41.55 | [6 empty] | 153.7393 | 44.05 | [6 empty] | 153.7393 | 137.075
|  |  | 153.814 | 41.685 | [6 empty] | 153.814 | 43.585 | [6 empty] | 153.814 | 137.69
|  |  | 153.8887 | 39.61 | [6 empty] | 153.8887 | 43.155 | [6 empty] | 153.8887 | 138.09
|  |  | 153.9633 | 40.42 | [6 empty] | 153.9633 | 42.685 | [6 empty] | 153.9633 | 138.135
|  |  | 154.038 | 41.095 | [6 empty] | 154.038 | 42.475 | [6 empty] | 154.038 | 138.31
|  |  | 154.1127 | 41.35 | [6 empty] | 154.1127 | 42.71 | [6 empty] | 154.1127 | 138.32
|  |  | 154.1873 | 42.075 | [6 empty] | 154.1873 | 43.06 | [6 empty] | 154.1873 | 138.405
|  |  | 154.262 | 41.43 | [6 empty] | 154.262 | 43.48 | [6 empty] | 154.262 | 138.965
|  |  | 154.3367 | 40.915 | [6 empty] | 154.3367 | 43.96 | [6 empty] | 154.3367 | 138.855
|  |  | 154.4114 | 41.525 | [6 empty] | 154.4114 | 44.26 | [6 empty] | 154.4114 | 138.36
|  |  | 154.486 | 41.615 | [6 empty] | 154.486 | 44.445 | [6 empty] | 154.486 | 138.4
|  |  | 154.5607 | 40.695 | [6 empty] | 154.5607 | 44.81 | [6 empty] | 154.5607 | 138.51
|  |  | 154.6354 | 41.06 | [6 empty] | 154.6354 | 44.87 | [6 empty] | 154.6354 | 138.36
|  |  | 154.71 | 42.02 | [6 empty] | 154.71 | 44.82 | [6 empty] | 154.71 | 138.32
|  |  | 154.7847 | 41.445 | [6 empty] | 154.7847 | 44.845 | [6 empty] | 154.7847 | 137.965
|  |  | 154.8594 | 41.49 | [6 empty] | 154.8594 | 44.94 | [6 empty] | 154.8594 | 137.545
|  |  | 154.934 | 42.385 | [6 empty] | 154.934 | 45.055 | [6 empty] | 154.934 | 137.605
|  |  | 155.0087 | 42.67 | [6 empty] | 155.0087 | 44.95 | [6 empty] | 155.0087 | 137.775
|  |  | 155.0834 | 40.785 | [6 empty] | 155.0834 | 44.87 | [6 empty] | 155.0834 | 137.705
|  |  | 155.158 | 41.65 | [6 empty] | 155.158 | 44.735 | [6 empty] | 155.158 | 137.345
|  |  | 155.2327 | 41.685 | [6 empty] | 155.2327 | 44.33 | [6 empty] | 155.2327 | 137.03
|  |  | 155.3074 | 41.16 | [6 empty] | 155.3074 | 43.935 | [6 empty] | 155.3074 | 137.12
|  |  | 155.382 | 40.66 | [6 empty] | 155.382 | 43.695 | [6 empty] | 155.382 | 137.915
|  |  | 155.4567 | 40.545 | [6 empty] | 155.4567 | 43.825 | [6 empty] | 155.4567 | 138.38
|  |  | 155.5314 | 39.88 | [6 empty] | 155.5314 | 44.61 | [6 empty] | 155.5314 | 138.325
|  |  | 155.606 | 40.305 | [6 empty] | 155.606 | 45.395 | [6 empty] | 155.606 | 138.035
|  |  | 155.6807 | 40.54 | [6 empty] | 155.6807 | 46.1 | [6 empty] | 155.6807 | 137.76
|  |  | 155.7554 | 42.575 | [6 empty] | 155.7554 | 46.365 | [6 empty] | 155.7554 | 138.065
|  |  | 155.83 | 42.97 | [6 empty] | 155.83 | 46.675 | [6 empty] | 155.83 | 138.08
|  |  | 155.9047 | 40.67 | [6 empty] | 155.9047 | 46.905 | [6 empty] | 155.9047 | 137.35
|  |  | 155.9794 | 42.2 | [6 empty] | 155.9794 | 47.04 | [6 empty] | 155.9794 | 136.165
|  |  | 156.054 | 42.62 | [6 empty] | 156.054 | 47.15 | [6 empty] | 156.054 | 135.97
|  |  | 156.1287 | 42.315 | [6 empty] | 156.1287 | 47.4 | [6 empty] | 156.1287 | 136.08
|  |  | 156.2034 | 42.305 | [6 empty] | 156.2034 | 47.63 | [6 empty] | 156.2034 | 136.27
|  |  | 156.278 | 42.22 | [6 empty] | 156.278 | 47.605 | [6 empty] | 156.278 | 136.19
|  |  | 156.3527 | 41.865 | [6 empty] | 156.3527 | 47.545 | [6 empty] | 156.3527 | 136.17
|  |  | 156.4274 | 41.355 | [6 empty] | 156.4274 | 47.485 | [6 empty] | 156.4274 | 136.285
|  |  | 156.502 | 41.84 | [6 empty] | 156.502 | 47.405 | [6 empty] | 156.502 | 136.56
|  |  | 156.5767 | 41.76 | [6 empty] | 156.5767 | 47.275 | [6 empty] | 156.5767 | 136.55
|  |  | 156.6514 | 43.115 | [6 empty] | 156.6514 | 47.11 | [6 empty] | 156.6514 | 136.44
|  |  | 156.726 | 42.405 | [6 empty] | 156.726 | 47.065 | [6 empty] | 156.726 | 136.325
|  |  | 156.8007 | 42.27 | [6 empty] | 156.8007 | 47.045 | [6 empty] | 156.8007 | 136.515
|  |  | 156.8754 | 42.16 | [6 empty] | 156.8754 | 47.195 | [6 empty] | 156.8754 | 136.805
|  |  | 156.95 | 40.915 | [6 empty] | 156.95 | 47.825 | [6 empty] | 156.95 | 136.47
|  |  | 157.0247 | 40.745 | [6 empty] | 157.0247 | 48.57 | [6 empty] | 157.0247 | 135.985
|  |  | 157.0994 | 42.265 | [6 empty] | 157.0994 | 49.485 | [6 empty] | 157.0994 | 135.98
|  |  | 157.174 | 42.39 | [6 empty] | 157.174 | 49.86 | [6 empty] | 157.174 | 136.28
|  |  | 157.2487 | 41.85 | [6 empty] | 157.2487 | 50.035 | [6 empty] | 157.2487 | 136.06
|  |  | 157.3234 | 42.75 | [6 empty] | 157.3234 | 50.385 | [6 empty] | 157.3234 | 135.565
|  |  | 157.398 | 42.97 | [6 empty] | 157.398 | 50.55 | [6 empty] | 157.398 | 135.31
|  |  | 157.4727 | 42.32 | [6 empty] | 157.4727 | 50.695 | [6 empty] | 157.4727 | 135.555
|  |  | 157.5474 | 42.48 | [6 empty] | 157.5474 | 50.79 | [6 empty] | 157.5474 | 135.195
|  |  | 157.622 | 43.905 | [6 empty] | 157.622 | 50.62 | [6 empty] | 157.622 | 134.93
|  |  | 157.6967 | 43.305 | [6 empty] | 157.6967 | 50.545 | [6 empty] | 157.6967 | 134.42
|  |  | 157.7714 | 43.515 | [6 empty] | 157.7714 | 50.42 | [6 empty] | 157.7714 | 133.65
|  |  | 157.846 | 43.34 | [6 empty] | 157.846 | 50.22 | [6 empty] | 157.846 | 133.68
|  |  | 157.9207 | 42.97 | [6 empty] | 157.9207 | 49.955 | [6 empty] | 157.9207 | 134.005
|  |  | 157.9954 | 42.7 | [6 empty] | 157.9954 | 49.43 | [6 empty] | 157.9954 | 133.79
|  |  | 158.07 | 42.55 | [6 empty] | 158.07 | 49.245 | [6 empty] | 158.07 | 133.7
|  |  | 158.1447 | 43.48 | [6 empty] | 158.1447 | 49.205 | [6 empty] | 158.1447 | 133.645
|  |  | 158.2194 | 43.185 | [6 empty] | 158.2194 | 49.72 | [6 empty] | 158.2194 | 133.765
|  |  | 158.294 | 42.635 | [6 empty] | 158.294 | 50.085 | [6 empty] | 158.294 | 133.86
|  |  | 158.3687 | 43.515 | [6 empty] | 158.3687 | 50.42 | [6 empty] | 158.3687 | 133.545
|  |  | 158.4434 | 43.92 | [6 empty] | 158.4434 | 50.735 | [6 empty] | 158.4434 | 133.225
|  |  | 158.518 | 43.28 | [6 empty] | 158.518 | 50.84 | [6 empty] | 158.518 | 132.98
|  |  | 158.5927 | 43.135 | [6 empty] | 158.5927 | 50.905 | [6 empty] | 158.5927 | 132.5
|  |  | 158.6674 | 43.84 | [6 empty] | 158.6674 | 51.02 | [6 empty] | 158.6674 | 132.365
|  |  | 158.742 | 42.14 | [6 empty] | 158.742 | 51.025 | [6 empty] | 158.742 | 132.425
|  |  | 158.8167 | 43.51 | [6 empty] | 158.8167 | 50.885 | [6 empty] | 158.8167 | 132.1
|  |  | 158.8914 | 42.775 | [6 empty] | 158.8914 | 50.55 | [6 empty] | 158.8914 | 132.245
|  |  | 158.966 | 43.945 | [6 empty] | 158.966 | 50.38 | [6 empty] | 158.966 | 132.445
|  |  | 159.0407 | 44.285 | [6 empty] | 159.0407 | 50.18 | [6 empty] | 159.0407 | 132.28
|  |  | 159.1154 | 43.49 | [6 empty] | 159.1154 | 49.745 | [6 empty] | 159.1154 | 132.345
|  |  | 159.19 | 42.78 | [6 empty] | 159.19 | 49.43 | [6 empty] | 159.19 | 133.08
|  |  | 159.2647 | 43.03 | [6 empty] | 159.2647 | 48.975 | [6 empty] | 159.2647 | 133.49
|  |  | 159.3394 | 43.87 | [6 empty] | 159.3394 | 48.215 | [6 empty] | 159.3394 | 132.875
|  |  | 159.414 | 43.46 | [6 empty] | 159.414 | 47.64 | [6 empty] | 159.414 | 132.355
|  |  | 159.4887 | 42.55 | [6 empty] | 159.4887 | 47.19 | [6 empty] | 159.4887 | 132.59
|  |  | 159.5634 | 42.18 | [6 empty] | 159.5634 | 46.815 | [6 empty] | 159.5634 | 132.635
|  |  | 159.638 | 42.92 | [6 empty] | 159.638 | 46.86 | [6 empty] | 159.638 | 132.53
|  |  | 159.7127 | 43.765 | [6 empty] | 159.7127 | 47.08 | [6 empty] | 159.7127 | 132.545
|  |  | 159.7874 | 43.45 | [6 empty] | 159.7874 | 47.485 | [6 empty] | 159.7874 | 132.355
|  |  | 159.862 | 43.7 | [6 empty] | 159.862 | 47.67 | [6 empty] | 159.862 | 132.245
|  |  | 159.9367 | 43.395 | [6 empty] | 159.9367 | 47.735 | [6 empty] | 159.9367 | 132.65
|  |  | 160.0114 | 43.485 | [6 empty] | 160.0114 | 47.705 | [6 empty] | 160.0114 | 132.67
|  |  | 160.086 | 42.055 | [6 empty] | 160.086 | 47.725 | [6 empty] | 160.086 | 132.17
|  |  | 160.1607 | 42.55 | [6 empty] | 160.1607 | 47.72 | [6 empty] | 160.1607 | 131.69
|  |  | 160.2354 | 43.49 | [6 empty] | 160.2354 | 47.65 | [6 empty] | 160.2354 | 131.385
|  |  | 160.31 | 42.81 | [6 empty] | 160.31 | 47.505 | [6 empty] | 160.31 | 132.065
|  |  | 160.3847 | 43.16 | [6 empty] | 160.3847 | 47.33 | [6 empty] | 160.3847 | 132.645
|  |  | 160.4594 | 44.115 | [6 empty] | 160.4594 | 47.055 | [6 empty] | 160.4594 | 132.91
|  |  | 160.534 | 43.18 | [6 empty] | 160.534 | 46.7 | [6 empty] | 160.534 | 132.625
|  |  | 160.6087 | 43.05 | [6 empty] | 160.6087 | 46.5 | [6 empty] | 160.6087 | 132.595
|  |  | 160.6834 | 42.78 | [6 empty] | 160.6834 | 46.235 | [6 empty] | 160.6834 | 133.03
|  |  | 160.7581 | 42.685 | [6 empty] | 160.7581 | 45.69 | [6 empty] | 160.7581 | 133.325
|  |  | 160.8327 | 43.31 | [6 empty] | 160.8327 | 45.26 | [6 empty] | 160.8327 | 133.42
|  |  | 160.9074 | 43.425 | [6 empty] | 160.9074 | 44.9 | [6 empty] | 160.9074 | 133.165
|  |  | 160.9821 | 43.28 | [6 empty] | 160.9821 | 44.33 | [6 empty] | 160.9821 | 133.065
|  |  | 161.0567 | 43.535 | [6 empty] | 161.0567 | 44.175 | [6 empty] | 161.0567 | 133.615
|  |  | 161.1314 | 42.15 | [6 empty] | 161.1314 | 44.59 | [6 empty] | 161.1314 | 134.335
|  |  | 161.2061 | 43.775 | [6 empty] | 161.2061 | 45.565 | [6 empty] | 161.2061 | 134.3
|  |  | 161.2807 | 43.39 | [6 empty] | 161.2807 | 45.95 | [6 empty] | 161.2807 | 134.6
|  |  | 161.3554 | 42.545 | [6 empty] | 161.3554 | 46.015 | [6 empty] | 161.3554 | 135.405
|  |  | 161.4301 | 42.545 | [6 empty] | 161.4301 | 46.085 | [6 empty] | 161.4301 | 135.48
|  |  | 161.5047 | 41.38 | [6 empty] | 161.5047 | 46.135 | [6 empty] | 161.5047 | 135.505
|  |  | 161.5794 | 41.6 | [6 empty] | 161.5794 | 46.165 | [6 empty] | 161.5794 | 135.155
|  |  | 161.6541 | 40.98 | [6 empty] | 161.6541 | 45.935 | [6 empty] | 161.6541 | 134.895
|  |  | 161.7287 | 40.955 | [6 empty] | 161.7287 | 45.685 | [6 empty] | 161.7287 | 135.25
|  |  | 161.8034 | 41.945 | [6 empty] | 161.8034 | 45.475 | [6 empty] | 161.8034 | 135.72
|  |  | 161.8781 | 42.575 | [6 empty] | 161.8781 | 45.26 | [6 empty] | 161.8781 | 135.735
|  |  | 161.9527 | 42.66 | [6 empty] | 161.9527 | 45.205 | [6 empty] | 161.9527 | 136.1
|  |  | 162.0274 | 41.385 | [6 empty] | 162.0274 | 45.045 | [6 empty] | 162.0274 | 136.485
|  |  | 162.1021 | 40.93 | [6 empty] | 162.1021 | 44.31 | [6 empty] | 162.1021 | 137.53
|  |  | 162.1767 | 41.365 | [6 empty] | 162.1767 | 43.755 | [6 empty] | 162.1767 | 137.94
|  |  | 162.2514 | 40.525 | [6 empty] | 162.2514 | 43.305 | [6 empty] | 162.2514 | 138.005
|  |  | 162.326 | 41.675 | [6 empty] | 162.326 | 42.845 | [6 empty] | 162.326 | 137.815
|  |  | 162.4007 | 41.355 | [6 empty] | 162.4007 | 42.51 | [6 empty] | 162.4007 | 138.125
|  |  | 162.4754 | 42.535 | [6 empty] | 162.4754 | 42.735 | [6 empty] | 162.4754 | 138.885
|  |  | 162.5501 | 43.335 | [6 empty] | 162.5501 | 43.3 | [6 empty] | 162.5501 | 139.255
|  |  | 162.6247 | 42.575 | [6 empty] | 162.6247 | 44.11 | [6 empty] | 162.6247 | 139.25
|  |  | 162.6994 | 42.985 | [6 empty] | 162.6994 | 44.455 | [6 empty] | 162.6994 | 138.79
|  |  | 162.7741 | 44.21 | [6 empty] | 162.7741 | 44.76 | [6 empty] | 162.7741 | 138.55
|  |  | 162.8487 | 43.075 | [6 empty] | 162.8487 | 44.935 | [6 empty] | 162.8487 | 138.68
|  |  | 162.9234 | 42.815 | [6 empty] | 162.9234 | 44.97 | [6 empty] | 162.9234 | 138.935
|  |  | 162.9981 | 42.225 | [6 empty] | 162.9981 | 45.08 | [6 empty] | 162.9981 | 139.295
|  |  | 163.0727 | 42.915 | [6 empty] | 163.0727 | 45.22 | [6 empty] | 163.0727 | 138.775
|  |  | 163.1474 | 41.885 | [6 empty] | 163.1474 | 45.2 | [6 empty] | 163.1474 | 137.965
|  |  | 163.2221 | 42.335 | [6 empty] | 163.2221 | 45.19 | [6 empty] | 163.2221 | 137.925
|  |  | 163.2967 | 43.01 | [6 empty] | 163.2967 | 45.115 | [6 empty] | 163.2967 | 138.2
|  |  | 163.3714 | 42.965 | [6 empty] | 163.3714 | 45.055 | [6 empty] | 163.3714 | 138.41
|  |  | 163.4461 | 42.035 | [6 empty] | 163.4461 | 44.905 | [6 empty] | 163.4461 | 138.115
|  |  | 163.5207 | 41.32 | [6 empty] | 163.5207 | 44.405 | [6 empty] | 163.5207 | 138.105
|  |  | 163.5954 | 41.495 | [6 empty] | 163.5954 | 44.22 | [6 empty] | 163.5954 | 138.475
|  |  | 163.6701 | 41.995 | [6 empty] | 163.6701 | 44.2 | [6 empty] | 163.6701 | 138.75
|  |  | 163.7447 | 41.57 | [6 empty] | 163.7447 | 44.595 | [6 empty] | 163.7447 | 138.495
|  |  | 163.8194 | 41.505 | [6 empty] | 163.8194 | 45.005 | [6 empty] | 163.8194 | 137.935
|  |  | 163.8941 | 42.72 | [6 empty] | 163.8941 | 45.445 | [6 empty] | 163.8941 | 137.29
|  |  | 163.9687 | 42.99 | [6 empty] | 163.9687 | 45.88 | [6 empty] | 163.9687 | 137.165
|  |  | 164.0434 | 43.245 | [6 empty] | 164.0434 | 46.05 | [6 empty] | 164.0434 | 137.02
|  |  | 164.1181 | 43.68 | [6 empty] | 164.1181 | 46.255 | [6 empty] | 164.1181 | 137.475
|  |  | 164.1927 | 43.05 | [6 empty] | 164.1927 | 46.335 | [6 empty] | 164.1927 | 137.215
|  |  | 164.2674 | 42.785 | [6 empty] | 164.2674 | 46.46 | [6 empty] | 164.2674 | 137.01
|  |  | 164.3421 | 42.425 | [6 empty] | 164.3421 | 46.545 | [6 empty] | 164.3421 | 137.205
|  |  | 164.4167 | 41.55 | [6 empty] | 164.4167 | 46.57 | [6 empty] | 164.4167 | 136.89
|  |  | 164.4914 | 42.035 | [6 empty] | 164.4914 | 46.535 | [6 empty] | 164.4914 | 136.445
|  |  | 164.5661 | 42.105 | [6 empty] | 164.5661 | 46.485 | [6 empty] | 164.5661 | 136.035
|  |  | 164.6407 | 42.38 | [6 empty] | 164.6407 | 46.4 | [6 empty] | 164.6407 | 135.76
|  |  | 164.7154 | 41.935 | [6 empty] | 164.7154 | 46.29 | [6 empty] | 164.7154 | 135.83
|  |  | 164.7901 | 42.495 | [6 empty] | 164.7901 | 46.08 | [6 empty] | 164.7901 | 136.04
|  |  | 164.8647 | 42.375 | [6 empty] | 164.8647 | 45.9 | [6 empty] | 164.8647 | 136.735
|  |  | 164.9394 | 41.435 | [6 empty] | 164.9394 | 45.635 | [6 empty] | 164.9394 | 137.165
|  |  | 165.0141 | 41.64 | [6 empty] | 165.0141 | 45.505 | [6 empty] | 165.0141 | 137.195
|  |  | 165.0887 | 42.375 | [6 empty] | 165.0887 | 45.485 | [6 empty] | 165.0887 | 136.18
|  |  | 165.1634 | 43.34 | [6 empty] | 165.1634 | 45.845 | [6 empty] | 165.1634 | 135.5
|  |  | 165.2381 | 43.38 | [6 empty] | 165.2381 | 46.525 | [6 empty] | 165.2381 | 135.16
|  |  | 165.3127 | 43.32 | [6 empty] | 165.3127 | 47.025 | [6 empty] | 165.3127 | 135.235
|  |  | 165.3874 | 42.895 | [6 empty] | 165.3874 | 47.77 | [6 empty] | 165.3874 | 135.64
|  |  | 165.4621 | 43.02 | [6 empty] | 165.4621 | 48.055 | [6 empty] | 165.4621 | 136.17
|  |  | 165.5367 | 42.105 | [6 empty] | 165.5367 | 48.26 | [6 empty] | 165.5367 | 136.23
|  |  | 165.6114 | 43.56 | [6 empty] | 165.6114 | 48.6 | [6 empty] | 165.6114 | 135.725
|  |  | 165.6861 | 43.455 | [6 empty] | 165.6861 | 48.795 | [6 empty] | 165.6861 | 135.425
|  |  | 165.7607 | 42.74 | [6 empty] | 165.7607 | 48.885 | [6 empty] | 165.7607 | 135.07
|  |  | 165.8354 | 43.455 | [6 empty] | 165.8354 | 48.815 | [6 empty] | 165.8354 | 134.97
|  |  | 165.9101 | 43.285 | [6 empty] | 165.9101 | 48.85 | [6 empty] | 165.9101 | 135.03
|  |  | 165.9847 | 42.99 | [6 empty] | 165.9847 | 48.96 | [6 empty] | 165.9847 | 135.29
|  |  | 166.0594 | 42.77 | [6 empty] | 166.0594 | 49 | [6 empty] | 166.0594 | 135.175
|  |  | 166.1341 | 43.57 | [6 empty] | 166.1341 | 48.92 | [6 empty] | 166.1341 | 134.825
|  |  | 166.2087 | 43.935 | [6 empty] | 166.2087 | 48.755 | [6 empty] | 166.2087 | 134.44
|  |  | 166.2834 | 42.105 | [6 empty] | 166.2834 | 48.505 | [6 empty] | 166.2834 | 134.36
|  |  | 166.3581 | 43.455 | [6 empty] | 166.3581 | 48.005 | [6 empty] | 166.3581 | 134.3
|  |  | 166.4327 | 43.18 | [6 empty] | 166.4327 | 47.66 | [6 empty] | 166.4327 | 134.11
|  |  | 166.5074 | 42.67 | [6 empty] | 166.5074 | 47.485 | [6 empty] | 166.5074 | 134.255
|  |  | 166.5821 | 42.455 | [6 empty] | 166.5821 | 47.7 | [6 empty] | 166.5821 | 134.43
|  |  | 166.6567 | 42.91 | [6 empty] | 166.6567 | 48.36 | [6 empty] | 166.6567 | 134.525
|  |  | 166.7314 | 43.7 | [6 empty] | 166.7314 | 49.05 | [6 empty] | 166.7314 | 134.625
|  |  | 166.8061 | 43.53 | [6 empty] | 166.8061 | 49.525 | [6 empty] | 166.8061 | 134.475
|  |  | 166.8807 | 42.58 | [6 empty] | 166.8807 | 49.57 | [6 empty] | 166.8807 | 134.18
|  |  | 166.9554 | 43.52 | [6 empty] | 166.9554 | 49.725 | [6 empty] | 166.9554 | 134.18
|  |  | 167.0301 | 43.805 | [6 empty] | 167.0301 | 49.87 | [6 empty] | 167.0301 | 134.485
|  |  | 167.1048 | 44.27 | [6 empty] | 167.1048 | 49.89 | [6 empty] | 167.1048 | 134.17
|  |  | 167.1794 | 44.335 | [6 empty] | 167.1794 | 49.81 | [6 empty] | 167.1794 | 133.47
|  |  | 167.2541 | 43.185 | [6 empty] | 167.2541 | 49.535 | [6 empty] | 167.2541 | 133.755
|  |  | 167.3288 | 43.48 | [6 empty] | 167.3288 | 49.3 | [6 empty] | 167.3288 | 134.52
|  |  | 167.4034 | 43.11 | [6 empty] | 167.4034 | 49.07 | [6 empty] | 167.4034 | 134.185
|  |  | 167.4781 | 42.975 | [6 empty] | 167.4781 | 48.68 | [6 empty] | 167.4781 | 133.715
|  |  | 167.5527 | 42.975 | [6 empty] | 167.5527 | 48.36 | [6 empty] | 167.5527 | 133.97
|  |  | 167.6274 | 42.235 | [6 empty] | 167.6274 | 47.93 | [6 empty] | 167.6274 | 134.35
|  |  | 167.7021 | 42.325 | [6 empty] | 167.7021 | 47.4 | [6 empty] | 167.7021 | 134.73
|  |  | 167.7767 | 41.945 | [6 empty] | 167.7767 | 47.04 | [6 empty] | 167.7767 | 134.415
|  |  | 167.8514 | 42.705 | [6 empty] | 167.8514 | 46.53 | [6 empty] | 167.8514 | 133.8
|  |  | 167.9261 | 43.185 | [6 empty] | 167.9261 | 45.835 | [6 empty] | 167.9261 | 133.81
|  |  | 168.0007 | 43.285 | [6 empty] | 168.0007 | 45.435 | [6 empty] | 168.0007 | 134.265
|  |  | 168.0754 | 43.45 | [6 empty] | 168.0754 | 45.245 | [6 empty] | 168.0754 | 134.105
|  |  | 168.1501 | 42.815 | [6 empty] | 168.1501 | 45.905 | [6 empty] | 168.1501 | 134.09
|  |  | 168.2247 | 42.955 | [6 empty] | 168.2247 | 46.43 | [6 empty] | 168.2247 | 134.205
|  |  | 168.2994 | 43.355 | [6 empty] | 168.2994 | 46.74 | [6 empty] | 168.2994 | 134.54
|  |  | 168.3741 | 42.61 | [6 empty] | 168.3741 | 46.76 | [6 empty] | 168.3741 | 134.78
|  |  | 168.4487 | 40.815 | [6 empty] | 168.4487 | 46.805 | [6 empty] | 168.4487 | 134.76
|  |  | 168.5234 | 42.58 | [6 empty] | 168.5234 | 46.75 | [6 empty] | 168.5234 | 134.58
|  |  | 168.5981 | 42.78 | [6 empty] | 168.5981 | 46.775 | [6 empty] | 168.5981 | 134.24
|  |  | 168.6727 | 42.675 | [6 empty] | 168.6727 | 46.79 | [6 empty] | 168.6727 | 134.31
|  |  | 168.7474 | 42.17 | [6 empty] | 168.7474 | 46.595 | [6 empty] | 168.7474 | 134.28
|  |  | 168.8221 | 42.255 | [6 empty] | 168.8221 | 46.265 | [6 empty] | 168.8221 | 134.17
|  |  | 168.8968 | 42.8 | [6 empty] | 168.8968 | 45.77 | [6 empty] | 168.8968 | 133.815
|  |  | 168.9714 | 42.615 | [6 empty] | 168.9714 | 45.41 | [6 empty] | 168.9714 | 133.39
|  |  | 169.0461 | 42.505 | [6 empty] | 169.0461 | 44.975 | [6 empty] | 169.0461 | 133.59
|  |  | 169.1208 | 43.04 | [6 empty] | 169.1208 | 44.655 | [6 empty] | 169.1208 | 134.135
|  |  | 169.1954 | 43.1 | [6 empty] | 169.1954 | 44.29 | [6 empty] | 169.1954 | 134.965
|  |  | 169.2701 | 42.945 | [6 empty] | 169.2701 | 43.985 | [6 empty] | 169.2701 | 134.93
|  |  | 169.3448 | 43.585 | [6 empty] | 169.3448 | 44.06 | [6 empty] | 169.3448 | 134.535
|  |  | 169.4194 | 43.145 | [6 empty] | 169.4194 | 44.31 | [6 empty] | 169.4194 | 134.515
|  |  | 169.4941 | 43.245 | [6 empty] | 169.4941 | 44.545 | [6 empty] | 169.4941 | 134.71
|  |  | 169.5688 | 43.785 | [6 empty] | 169.5688 | 45.04 | [6 empty] | 169.5688 | 135.065
|  |  | 169.6434 | 43.48 | [6 empty] | 169.6434 | 45.16 | [6 empty] | 169.6434 | 135.875
|  |  | 169.7181 | 40.985 | [6 empty] | 169.7181 | 45.25 | [6 empty] | 169.7181 | 136.045
|  |  | 169.7928 | 41.545 | [6 empty] | 169.7928 | 45.35 | [6 empty] | 169.7928 | 135.975
|  |  | 169.8674 | 42.135 | [6 empty] | 169.8674 | 45.38 | [6 empty] | 169.8674 | 135.69
|  |  | 169.9421 | 42.615 | [6 empty] | 169.9421 | 45.295 | [6 empty] | 169.9421 | 135.95
|  |  | 170.0168 | 43.635 | [6 empty] | 170.0168 | 45.29 | [6 empty] | 170.0168 | 136.39
|  |  | 170.0914 | 42.445 | [6 empty] | 170.0914 | 45.27 | [6 empty] | 170.0914 | 136.015
|  |  | 170.1661 | 42.785 | [6 empty] | 170.1661 | 45.055 | [6 empty] | 170.1661 | 136.2
|  |  | 170.2408 | 41.37 | [6 empty] | 170.2408 | 44.69 | [6 empty] | 170.2408 | 136.385
|  |  | 170.3154 | 41.445 | [6 empty] | 170.3154 | 44.155 | [6 empty] | 170.3154 | 136.345
|  |  | 170.3901 | 40.99 | [6 empty] | 170.3901 | 43.935 | [6 empty] | 170.3901 | 136.435
|  |  | 170.4648 | 40.355 | [6 empty] | 170.4648 | 43.585 | [6 empty] | 170.4648 | 136.725
|  |  | 170.5394 | 42.23 | [6 empty] | 170.5394 | 43.17 | [6 empty] | 170.5394 | 137.275
|  |  | 170.6141 | 41.405 | [6 empty] | 170.6141 | 42.685 | [6 empty] | 170.6141 | 137.825
|  |  | 170.6888 | 40.915 | [6 empty] | 170.6888 | 42.43 | [6 empty] | 170.6888 | 138.29
|  |  | 170.7634 | 41.43 | [6 empty] | 170.7634 | 42.695 | [6 empty] | 170.7634 | 138.47
|  |  | 170.8381 | 42.625 | [6 empty] | 170.8381 | 43.045 | [6 empty] | 170.8381 | 138.365
|  |  | 170.9128 | 42.04 | [6 empty] | 170.9128 | 43.44 | [6 empty] | 170.9128 | 138.015
|  |  | 170.9874 | 42.01 | [6 empty] | 170.9874 | 43.82 | [6 empty] | 170.9874 | 138.065
|  |  | 171.0621 | 41.465 | [6 empty] | 171.0621 | 44.13 | [6 empty] | 171.0621 | 138.66
|  |  | 171.1368 | 40.87 | [6 empty] | 171.1368 | 44.455 | [6 empty] | 171.1368 | 138.945
|  |  | 171.2114 | 40.745 | [6 empty] | 171.2114 | 44.805 | [6 empty] | 171.2114 | 138.5
|  |  | 171.2861 | 42.48 | [6 empty] | 171.2861 | 45.015 | [6 empty] | 171.2861 | 136.94
|  |  | 171.3608 | 42.065 | [6 empty] | 171.3608 | 45.035 | [6 empty] | 171.3608 | 136.745
|  |  | 171.4354 | 41.72 | [6 empty] | 171.4354 | 45.015 | [6 empty] | 171.4354 | 137.12
|  |  | 171.5101 | 40.79 | [6 empty] | 171.5101 | 44.92 | [6 empty] | 171.5101 | 137.395
|  |  | 171.5848 | 40.425 | [6 empty] | 171.5848 | 44.785 | [6 empty] | 171.5848 | 137.53
|  |  | 171.6594 | 41.82 | [6 empty] | 171.6594 | 44.91 | [6 empty] | 171.6594 | 137.43
|  |  | 171.7341 | 42.485 | [6 empty] | 171.7341 | 44.88 | [6 empty] | 171.7341 | 137.595
|  |  | 171.8088 | 41.615 | [6 empty] | 171.8088 | 44.635 | [6 empty] | 171.8088 | 137.885
|  |  | 171.8834 | 41.905 | [6 empty] | 171.8834 | 44.165 | [6 empty] | 171.8834 | 137.94
|  |  | 171.9581 | 41.86 | [6 empty] | 171.9581 | 43.89 | [6 empty] | 171.9581 | 137.535
|  |  | 172.0328 | 40.865 | [6 empty] | 172.0328 | 43.705 | [6 empty] | 172.0328 | 137.96
|  |  | 172.1074 | 41.05 | [6 empty] | 172.1074 | 43.945 | [6 empty] | 172.1074 | 138.425
|  |  | 172.1821 | 42.525 | [6 empty] | 172.1821 | 44.48 | [6 empty] | 172.1821 | 138.135
|  |  | 172.2568 | 41.905 | [6 empty] | 172.2568 | 45.25 | [6 empty] | 172.2568 | 138.145
|  |  | 172.3314 | 42.625 | [6 empty] | 172.3314 | 45.725 | [6 empty] | 172.3314 | 137.85
|  |  | 172.4061 | 42.005 | [6 empty] | 172.4061 | 46.375 | [6 empty] | 172.4061 | 137.325
|  |  | 172.4808 | 41.515 | [6 empty] | 172.4808 | 46.705 | [6 empty] | 172.4808 | 136.67
|  |  | 172.5554 | 41.055 | [6 empty] | 172.5554 | 46.95 | [6 empty] | 172.5554 | 136.48
|  |  | 172.6301 | 41.04 | [6 empty] | 172.6301 | 47.07 | [6 empty] | 172.6301 | 136.81
|  |  | 172.7048 | 40.79 | [6 empty] | 172.7048 | 47.275 | [6 empty] | 172.7048 | 136.62
|  |  | 172.7794 | 41.135 | [6 empty] | 172.7794 | 47.505 | [6 empty] | 172.7794 | 136.165
|  |  | 172.8541 | 41.91 | [6 empty] | 172.8541 | 47.58 | [6 empty] | 172.8541 | 135.51
|  |  | 172.9288 | 41.84 | [6 empty] | 172.9288 | 47.575 | [6 empty] | 172.9288 | 135.53
|  |  | 173.0034 | 41.74 | [6 empty] | 173.0034 | 47.55 | [6 empty] | 173.0034 | 136.27
|  |  | 173.0781 | 42.53 | [6 empty] | 173.0781 | 47.48 | [6 empty] | 173.0781 | 137.1
|  |  | 173.1528 | 41.435 | [6 empty] | 173.1528 | 47.4 | [6 empty] | 173.1528 | 136.875
|  |  | 173.2274 | 42.185 | [6 empty] | 173.2274 | 47.365 | [6 empty] | 173.2274 | 136.235
|  |  | 173.3021 | 42.025 | [6 empty] | 173.3021 | 47.24 | [6 empty] | 173.3021 | 135.875
|  |  | 173.3768 | 42.37 | [6 empty] | 173.3768 | 47.15 | [6 empty] | 173.3768 | 135.855
|  |  | 173.4514 | 42.175 | [6 empty] | 173.4514 | 47.03 | [6 empty] | 173.4514 | 136.11
|  |  | 173.5261 | 41.38 | [6 empty] | 173.5261 | 47.115 | [6 empty] | 173.5261 | 136.61
|  |  | 173.6008 | 42.065 | [6 empty] | 173.6008 | 47.82 | [6 empty] | 173.6008 | 136.845
|  |  | 173.6754 | 43.87 | [6 empty] | 173.6754 | 48.64 | [6 empty] | 173.6754 | 136.795
|  |  | 173.7501 | 43.635 | [6 empty] | 173.7501 | 49.415 | [6 empty] | 173.7501 | 136.275
|  |  | 173.8248 | 42.42 | [6 empty] | 173.8248 | 49.77 | [6 empty] | 173.8248 | 135.485
|  |  | 173.8994 | 42.5 | [6 empty] | 173.8994 | 50.105 | [6 empty] | 173.8994 | 134.46
|  |  | 173.9741 | 42.51 | [6 empty] | 173.9741 | 50.495 | [6 empty] | 173.9741 | 134.05
|  |  | 174.0488 | 43.085 | [6 empty] | 174.0488 | 50.7 | [6 empty] | 174.0488 | 134.375
|  |  | 174.1234 | 42.845 | [6 empty] | 174.1234 | 50.81 | [6 empty] | 174.1234 | 134.925
|  |  | 174.1981 | 42.905 | [6 empty] | 174.1981 | 50.9 | [6 empty] | 174.1981 | 134.795
|  |  | 174.2728 | 43.47 | [6 empty] | 174.2728 | 50.8 | [6 empty] | 174.2728 | 134.94
|  |  | 174.3474 | 43.205 | [6 empty] | 174.3474 | 50.72 | [6 empty] | 174.3474 | 134.875
|  |  | 174.4221 | 42.575 | [6 empty] | 174.4221 | 50.615 | [6 empty] | 174.4221 | 134.94
|  |  | 174.4968 | 42.485 | [6 empty] | 174.4968 | 50.455 | [6 empty] | 174.4968 | 134.945
|  |  | 174.5714 | 42.64 | [6 empty] | 174.5714 | 50.19 | [6 empty] | 174.5714 | 134.775
|  |  | 174.6461 | 43.39 | [6 empty] | 174.6461 | 49.835 | [6 empty] | 174.6461 | 134.545
|  |  | 174.7208 | 43.845 | [6 empty] | 174.7208 | 49.645 | [6 empty] | 174.7208 | 134.435
|  |  | 174.7954 | 43.83 | [6 empty] | 174.7954 | 49.6 | [6 empty] | 174.7954 | 134.105
|  |  | 174.8701 | 43.305 | [6 empty] | 174.8701 | 49.94 | [6 empty] | 174.8701 | 133.3
|  |  | 174.9448 | 44.235 | [6 empty] | 174.9448 | 50.225 | [6 empty] | 174.9448 | 133.15
|  |  | 175.0194 | 43.785 | [6 empty] | 175.0194 | 50.54 | [6 empty] | 175.0194 | 134.095
|  |  | 175.0941 | 43.545 | [6 empty] | 175.0941 | 50.895 | [6 empty] | 175.0941 | 133.49
|  |  | 175.1688 | 42.635 | [6 empty] | 175.1688 | 50.935 | [6 empty] | 175.1688 | 132.89
|  |  | 175.2435 | 43.505 | [6 empty] | 175.2435 | 50.895 | [6 empty] | 175.2435 | 132.66
|  |  | 175.3181 | 44.095 | [6 empty] | 175.3181 | 50.935 | [6 empty] | 175.3181 | 132.525
|  |  | 175.3928 | 42.79 | [6 empty] | 175.3928 | 51.065 | [6 empty] | 175.3928 | 132.475
|  |  | 175.4675 | 43.965 | [6 empty] | 175.4675 | 50.975 | [6 empty] | 175.4675 | 132.215
|  |  | 175.5421 | 43.92 | [6 empty] | 175.5421 | 50.74 | [6 empty] | 175.5421 | 132.29
|  |  | 175.6168 | 43.91 | [6 empty] | 175.6168 | 50.5 | [6 empty] | 175.6168 | 132.53
|  |  | 175.6915 | 43.625 | [6 empty] | 175.6915 | 50.385 | [6 empty] | 175.6915 | 132.875
|  |  | 175.7661 | 43.385 | [6 empty] | 175.7661 | 50.12 | [6 empty] | 175.7661 | 132.805
|  |  | 175.8408 | 43.425 | [6 empty] | 175.8408 | 49.48 | [6 empty] | 175.8408 | 132.915
|  |  | 175.9155 | 42.78 | [6 empty] | 175.9155 | 49.06 | [6 empty] | 175.9155 | 132.77
|  |  | 175.9901 | 44.15 | [6 empty] | 175.9901 | 48.675 | [6 empty] | 175.9901 | 132.705
|  |  | 176.0648 | 43.06 | [6 empty] | 176.0648 | 48.13 | [6 empty] | 176.0648 | 132.8
|  |  | 176.1395 | 42.325 | [6 empty] | 176.1395 | 47.665 | [6 empty] | 176.1395 | 133.37
|  |  | 176.2141 | 42.485 | [6 empty] | 176.2141 | 47.15 | [6 empty] | 176.2141 | 133.18
|  |  | 176.2888 | 43.875 | [6 empty] | 176.2888 | 47.04 | [6 empty] | 176.2888 | 132.96
|  |  | 176.3634 | 44.42 | [6 empty] | 176.3634 | 47.29 | [6 empty] | 176.3634 | 132.355
|  |  | 176.4381 | 42.895 | [6 empty] | 176.4381 | 47.475 | [6 empty] | 176.4381 | 132.16
|  |  | 176.5128 | 42.355 | [6 empty] | 176.5128 | 47.73 | [6 empty] | 176.5128 | 132.105
|  |  | 176.5874 | 41.98 | [6 empty] | 176.5874 | 47.84 | [6 empty] | 176.5874 | 132.015
|  |  | 176.6621 | 41.905 | [6 empty] | 176.6621 | 47.93 | [6 empty] | 176.6621 | 132.255
|  |  | 176.7368 | 42.11 | [6 empty] | 176.7368 | 48.015 | [6 empty] | 176.7368 | 132.5
|  |  | 176.8114 | 43.255 | [6 empty] | 176.8114 | 47.975 | [6 empty] | 176.8114 | 132.53
|  |  | 176.8861 | 43.715 | [6 empty] | 176.8861 | 47.835 | [6 empty] | 176.8861 | 132.415
|  |  | 176.9608 | 43.665 | [6 empty] | 176.9608 | 47.565 | [6 empty] | 176.9608 | 132.055
|  |  | 177.0355 | 44.64 | [6 empty] | 177.0355 | 47.435 | [6 empty] | 177.0355 | 132.39
|  |  | 177.1101 | 43.42 | [6 empty] | 177.1101 | 47.32 | [6 empty] | 177.1101 | 132.785
|  |  | 177.1848 | 43.125 | [6 empty] | 177.1848 | 47.065 | [6 empty] | 177.1848 | 133.035
|  |  | 177.2595 | 42.895 | [6 empty] | 177.2595 | 46.81 | [6 empty] | 177.2595 | 132.95
|  |  | 177.3341 | 42 | [6 empty] | 177.3341 | 46.515 | [6 empty] | 177.3341 | 132.985
|  |  | 177.4088 | 41.705 | [6 empty] | 177.4088 | 45.99 | [6 empty] | 177.4088 | 133.25
|  |  | 177.4835 | 42.765 | [6 empty] | 177.4835 | 45.65 | [6 empty] | 177.4835 | 133.495
|  |  | 177.5581 | 42.765 | [6 empty] | 177.5581 | 45.385 | [6 empty] | 177.5581 | 133.43
|  |  | 177.6328 | 42.5 | [6 empty] | 177.6328 | 44.99 | [6 empty] | 177.6328 | 133.815
|  |  | 177.7075 | 42.16 | [6 empty] | 177.7075 | 44.805 | [6 empty] | 177.7075 | 135.065
|  |  | 177.7821 | 42.355 | [6 empty] | 177.7821 | 44.975 | [6 empty] | 177.7821 | 135.235
|  |  | 177.8568 | 42.545 | [6 empty] | 177.8568 | 45.61 | [6 empty] | 177.8568 | 135.18
|  |  | 177.9315 | 43.23 | [6 empty] | 177.9315 | 45.965 | [6 empty] | 177.9315 | 135.245
|  |  | 178.0061 | 42.395 | [6 empty] | 178.0061 | 46.27 | [6 empty] | 178.0061 | 135.125
|  |  | 178.0808 | 42.33 | [6 empty] | 178.0808 | 46.375 | [6 empty] | 178.0808 | 135.42
|  |  | 178.1555 | 43.645 | [6 empty] | 178.1555 | 46.405 | [6 empty] | 178.1555 | 136.02
|  |  | 178.2301 | 43.2 | [6 empty] | 178.2301 | 46.335 | [6 empty] | 178.2301 | 136.705
|  |  | 178.3048 | 44 | [6 empty] | 178.3048 | 46.38 | [6 empty] | 178.3048 | 136.725
|  |  | 178.3795 | 42.795 | [6 empty] | 178.3795 | 46.42 | [6 empty] | 178.3795 | 136.815
|  |  | 178.4541 | 42.335 | [6 empty] | 178.4541 | 46.255 | [6 empty] | 178.4541 | 136.725
|  |  | 178.5288 | 41.345 | [6 empty] | 178.5288 | 45.94 | [6 empty] | 178.5288 | 136.42
|  |  | 178.6035 | 41.45 | [6 empty] | 178.6035 | 45.475 | [6 empty] | 178.6035 | 136.4
|  |  | 178.6781 | 41.805 | [6 empty] | 178.6781 | 45.12 | [6 empty] | 178.6781 | 137.02
|  |  | 178.7528 | 41.79 | [6 empty] | 178.7528 | 44.795 | [6 empty] | 178.7528 | 137.275
|  |  | 178.8275 | 41.005 | [6 empty] | 178.8275 | 44.405 | [6 empty] | 178.8275 | 137.29
|  |  | 178.9021 | 41.03 | [6 empty] | 178.9021 | 43.935 | [6 empty] | 178.9021 | 137.12
|  |  | 178.9768 | 40.855 | [6 empty] | 178.9768 | 43.35 | [6 empty] | 178.9768 | 137.155
|  |  | 179.0515 | 42.05 | [6 empty] | 179.0515 | 42.72 | [6 empty] | 179.0515 | 137.75
|  |  | 179.1261 | 42.06 | [6 empty] | 179.1261 | 42.725 | [6 empty] | 179.1261 | 138.54
|  |  | 179.2008 | 41.28 | [6 empty] | 179.2008 | 43.38 | [6 empty] | 179.2008 | 138.65
|  |  | 179.2755 | 41.005 | [6 empty] | 179.2755 | 44.645 | [6 empty] | 179.2755 | 138.87
|  |  | 179.3501 | 42.055 | [6 empty] | 179.3501 | 44.985 | [6 empty] | 179.3501 | 138.98
|  |  | 179.4248 | 40.86 | [6 empty] | 179.4248 | 45.07 | [6 empty] | 179.4248 | 138.865
|  |  | 179.4995 | 40.925 | [6 empty] | 179.4995 | 45.325 | [6 empty] | 179.4995 | 138.715
|  |  | 179.5741 | 41.225 | [6 empty] | 179.5741 | 45.545 | [6 empty] | 179.5741 | 138.585
|  |  | 179.6488 | 42.305 | [6 empty] | 179.6488 | 45.68 | [6 empty] | 179.6488 | 138.7
|  |  | 179.7235 | 42.96 | [6 empty] | 179.7235 | 45.58 | [6 empty] | 179.7235 | 138.765
|  |  | 179.7981 | 42.5 | [6 empty] | 179.7981 | 45.46 | [6 empty] | 179.7981 | 138.53
|  |  | 179.8728 | 42.19 | [6 empty] | 179.8728 | 45.48 | [6 empty] | 179.8728 | 138.2
|  |  | 179.9475 | 41.725 | [6 empty] | 179.9475 | 45.45 | [6 empty] | 179.9475 | 138
|  |  | 180.0221 | 42.04 | [6 empty] | 180.0221 | 45.315 | [6 empty] | 180.0221 | 138.075
|  |  | 180.0968 | 41.71 | [6 empty] | 180.0968 | 45.055 | [6 empty] | 180.0968 | 137.945
|  |  | 180.1715 | 41.85 | [6 empty] | 180.1715 | 44.805 | [6 empty] | 180.1715 | 138.14
|  |  | 180.2461 | 42.615 | [6 empty] | 180.2461 | 44.49 | [6 empty] | 180.2461 | 138.69
|  |  | 180.3208 | 42.015 | [6 empty] | 180.3208 | 44.425 | [6 empty] | 180.3208 | 138.46
|  |  | 180.3955 | 41.12 | [6 empty] | 180.3955 | 44.775 | [6 empty] | 180.3955 | 138.345
|  |  | 180.4701 | 43.245 | [6 empty] | 180.4701 | 45.575 | [6 empty] | 180.4701 | 138.52
|  |  | 180.5448 | 43.47 | [6 empty] | 180.5448 | 45.98 | [6 empty] | 180.5448 | 138.5
|  |  | 180.6195 | 43.38 | [6 empty] | 180.6195 | 46.325 | [6 empty] | 180.6195 | 137.595
|  |  | 180.6941 | 42.82 | [6 empty] | 180.6941 | 46.625 | [6 empty] | 180.6941 | 137.115
|  |  | 180.7688 | 41.62 | [6 empty] | 180.7688 | 46.645 | [6 empty] | 180.7688 | 137.18
|  |  | 180.8435 | 42.58 | [6 empty] | 180.8435 | 46.69 | [6 empty] | 180.8435 | 137.215
|  |  | 180.9181 | 42.255 | [6 empty] | 180.9181 | 46.88 | [6 empty] | 180.9181 | 137.51
|  |  | 180.9928 | 41.925 | [6 empty] | 180.9928 | 46.975 | [6 empty] | 180.9928 | 137.38
|  |  | 181.0675 | 40.92 | [6 empty] | 181.0675 | 47.045 | [6 empty] | 181.0675 | 136.585
|  |  | 181.1421 | 42.35 | [6 empty] | 181.1421 | 47.02 | [6 empty] | 181.1421 | 135.795
|  |  | 181.2168 | 42.615 | [6 empty] | 181.2168 | 46.905 | [6 empty] | 181.2168 | 136.07
|  |  | 181.2915 | 41.885 | [6 empty] | 181.2915 | 46.755 | [6 empty] | 181.2915 | 136.685
|  |  | 181.3661 | 42.105 | [6 empty] | 181.3661 | 46.64 | [6 empty] | 181.3661 | 137.14
|  |  | 181.4408 | 42.46 | [6 empty] | 181.4408 | 46.6 | [6 empty] | 181.4408 | 136.2
|  |  | 181.5155 | 42.73 | [6 empty] | 181.5155 | 46.5 | [6 empty] | 181.5155 | 136.15
|  |  | 181.5901 | 41.265 | [6 empty] | 181.5901 | 46.135 | [6 empty] | 181.5901 | 136.275
|  |  | 181.6648 | 42.09 | [6 empty] | 181.6648 | 45.97 | [6 empty] | 181.6648 | 136.185
|  |  | 181.7395 | 42.66 | [6 empty] | 181.7395 | 45.975 | [6 empty] | 181.7395 | 135.71
|  |  | 181.8141 | 42.38 | [6 empty] | 181.8141 | 46.425 | [6 empty] | 181.8141 | 135.37
|  |  | 181.8888 | 41.675 | [6 empty] | 181.8888 | 46.805 | [6 empty] | 181.8888 | 135.39
|  |  | 181.9635 | 43.255 | [6 empty] | 181.9635 | 47.34 | [6 empty] | 181.9635 | 135.745
|  |  | 182.0381 | 45.455 | [6 empty] | 182.0381 | 48.065 | [6 empty] | 182.0381 | 135.44
|  |  | 182.1128 | 44.03 | [6 empty] | 182.1128 | 48.39 | [6 empty] | 182.1128 | 135.265
|  |  | 182.1875 | 42.82 | [6 empty] | 182.1875 | 48.59 | [6 empty] | 182.1875 | 135.57
|  |  | 182.2621 | 42.65 | [6 empty] | 182.2621 | 48.67 | [6 empty] | 182.2621 | 135.9
|  |  | 182.3368 | 43.68 | [6 empty] | 182.3368 | 48.97 | [6 empty] | 182.3368 | 135.705
|  |  | 182.4115 | 43.405 | [6 empty] | 182.4115 | 49.165 | [6 empty] | 182.4115 | 135.19
|  |  | 182.4861 | 44.955 | [6 empty] | 182.4861 | 49.235 | [6 empty] | 182.4861 | 134.795
|  |  | 182.5608 | 45.145 | [6 empty] | 182.5608 | 49.22 | [6 empty] | 182.5608 | 134.6
|  |  | 182.6355 | 44.77 | [6 empty] | 182.6355 | 49.245 | [6 empty] | 182.6355 | 134.895
|  |  | 182.7101 | 43.77 | [6 empty] | 182.7101 | 49.22 | [6 empty] | 182.7101 | 134.69
|  |  | 182.7848 | 43.15 | [6 empty] | 182.7848 | 49.095 | [6 empty] | 182.7848 | 134.065
|  |  | 182.8595 | 42.675 | [6 empty] | 182.8595 | 48.91 | [6 empty] | 182.8595 | 133.94
|  |  | 182.9341 | 43.11 | [6 empty] | 182.9341 | 48.675 | [6 empty] | 182.9341 | 134.55
|  |  | 183.0088 | 42.55 | [6 empty] | 183.0088 | 48.35 | [6 empty] | 183.0088 | 134.505
|  |  | 183.0835 | 42.515 | [6 empty] | 183.0835 | 48.12 | [6 empty] | 183.0835 | 133.84
|  |  | 183.1581 | 44.32 | [6 empty] | 183.1581 | 47.89 | [6 empty] | 183.1581 | 134.015
|  |  | 183.2328 | 44.41 | [6 empty] | 183.2328 | 47.97 | [6 empty] | 183.2328 | 134.19
|  |  | 183.3075 | 43.385 | [6 empty] | 183.3075 | 48.52 | [6 empty] | 183.3075 | 134.165
|  |  | 183.3822 | 43.48 | [6 empty] | 183.3822 | 49.17 | [6 empty] | 183.3822 | 134.34
|  |  | 183.4568 | 43.46 | [6 empty] | 183.4568 | 49.72 | [6 empty] | 183.4568 | 134.475
|  |  | 183.5315 | 43.47 | [6 empty] | 183.5315 | 49.875 | [6 empty] | 183.5315 | 134.39
|  |  | 183.6062 | 42.035 | [6 empty] | 183.6062 | 49.995 | [6 empty] | 183.6062 | 134.09
|  |  | 183.6808 | 43.575 | [6 empty] | 183.6808 | 50.14 | [6 empty] | 183.6808 | 133.82
|  |  | 183.7555 | 44.02 | [6 empty] | 183.7555 | 50.065 | [6 empty] | 183.7555 | 133.8
|  |  | 183.8302 | 44.145 | [6 empty] | 183.8302 | 50.015 | [6 empty] | 183.8302 | 133.935
|  |  | 183.9048 | 42.415 | [6 empty] | 183.9048 | 49.87 | [6 empty] | 183.9048 | 133.98
|  |  | 183.9795 | 42.58 | [6 empty] | 183.9795 | 49.69 | [6 empty] | 183.9795 | 133.92
|  |  | 184.0542 | 43.415 | [6 empty] | 184.0542 | 49.42 | [6 empty] | 184.0542 | 134.065
|  |  | 184.1288 | 44.22 | [6 empty] | 184.1288 | 49.205 | [6 empty] | 184.1288 | 134.085
|  |  | 184.2035 | 42.45 | [6 empty] | 184.2035 | 48.82 | [6 empty] | 184.2035 | 133.7
|  |  | 184.2782 | 42.99 | [6 empty] | 184.2782 | 48.355 | [6 empty] | 184.2782 | 133.525
|  |  | 184.3528 | 43.12 | [6 empty] | 184.3528 | 47.865 | [6 empty] | 184.3528 | 133.905
|  |  | 184.4275 | 42.55 | [6 empty] | 184.4275 | 47.27 | [6 empty] | 184.4275 | 134.325
|  |  | 184.5022 | 42.52 | [6 empty] | 184.5022 | 46.86 | [6 empty] | 184.5022 | 134.35
|  |  | 184.5768 | 42.365 | [6 empty] | 184.5768 | 46.35 | [6 empty] | 184.5768 | 134.315
|  |  | 184.6515 | 42.64 | [6 empty] | 184.6515 | 45.68 | [6 empty] | 184.6515 | 134.41
|  |  | 184.7262 | 42.84 | [6 empty] | 184.7262 | 45.38 | [6 empty] | 184.7262 | 134.41
|  |  | 184.8008 | 43.38 | [6 empty] | 184.8008 | 45.71 | [6 empty] | 184.8008 | 134.685
|  |  | 184.8755 | 42.64 | [6 empty] | 184.8755 | 46.71 | [6 empty] | 184.8755 | 134.845
|  |  | 184.9502 | 42.3 | [6 empty] | 184.9502 | 47.155 | [6 empty] | 184.9502 | 135.1
|  |  | 185.0248 | 41.89 | [6 empty] | 185.0248 | 47.205 | [6 empty] | 185.0248 | 134.595
|  |  | 185.0995 | 42.41 | [6 empty] | 185.0995 | 47.095 | [6 empty] | 185.0995 | 134.13
|  |  | 185.1742 | 42.375 | [6 empty] | 185.1742 | 47.1 | [6 empty] | 185.1742 | 134.34
|  |  | 185.2488 | 43.215 | [6 empty] | 185.2488 | 46.98 | [6 empty] | 185.2488 | 134.375
|  |  | 185.3235 | 42.615 | [6 empty] | 185.3235 | 46.84 | [6 empty] | 185.3235 | 134.245
|  |  | 185.3982 | 43.355 | [6 empty] | 185.3982 | 46.735 | [6 empty] | 185.3982 | 134.085
|  |  | 185.4728 | 43.895 | [6 empty] | 185.4728 | 46.56 | [6 empty] | 185.4728 | 133.88
|  |  | 185.5475 | 43.63 | [6 empty] | 185.5475 | 46.165 | [6 empty] | 185.5475 | 134.035
|  |  | 185.6222 | 43.1 | [6 empty] | 185.6222 | 45.88 | [6 empty] | 185.6222 | 134.605
|  |  | 185.6968 | 42.995 | [6 empty] | 185.6968 | 45.63 | [6 empty] | 185.6968 | 134.82
|  |  | 185.7715 | 43.44 | [6 empty] | 185.7715 | 45.28 | [6 empty] | 185.7715 | 134.695
|  |  | 185.8462 | 42.205 | [6 empty] | 185.8462 | 44.905 | [6 empty] | 185.8462 | 134.18
|  |  | 185.9208 | 43.36 | [6 empty] | 185.9208 | 44.47 | [6 empty] | 185.9208 | 134.13
|  |  | 185.9955 | 43.585 | [6 empty] | 185.9955 | 44.205 | [6 empty] | 185.9955 | 134.7
|  |  | 186.0702 | 40.65 | [6 empty] | 186.0702 | 44.545 | [6 empty] | 186.0702 | 134.925
|  |  | 186.1448 | 42.205 | [6 empty] | 186.1448 | 44.93 | [6 empty] | 186.1448 | 134.655
|  |  | 186.2195 | 43.07 | [6 empty] | 186.2195 | 45.155 | [6 empty] | 186.2195 | 134.4
|  |  | 186.2942 | 43.005 | [6 empty] | 186.2942 | 45.265 | [6 empty] | 186.2942 | 134.86
|  |  | 186.3688 | 42.165 | [6 empty] | 186.3688 | 45.285 | [6 empty] | 186.3688 | 135.235
|  |  | 186.4435 | 42.995 | [6 empty] | 186.4435 | 45.405 | [6 empty] | 186.4435 | 135.175
|  |  | 186.5182 | 40.885 | [6 empty] | 186.5182 | 45.545 | [6 empty] | 186.5182 | 135.535
|  |  | 186.5928 | 40.895 | [6 empty] | 186.5928 | 45.555 | [6 empty] | 186.5928 | 136.355
|  |  | 186.6675 | 41.005 | [6 empty] | 186.6675 | 45.49 | [6 empty] | 186.6675 | 136.66
|  |  | 186.7422 | 42.955 | [6 empty] | 186.7422 | 45.285 | [6 empty] | 186.7422 | 136.345
|  |  | 186.8168 | 43.22 | [6 empty] | 186.8168 | 45.095 | [6 empty] | 186.8168 | 136.68
|  |  | 186.8915 | 42.53 | [6 empty] | 186.8915 | 44.88 | [6 empty] | 186.8915 | 137.28
|  |  | 186.9662 | 42.44 | [6 empty] | 186.9662 | 44.475 | [6 empty] | 186.9662 | 137.42
|  |  | 187.0408 | 42.52 | [6 empty] | 187.0408 | 44.11 | [6 empty] | 187.0408 | 136.65
|  |  | 187.1155 | 42.275 | [6 empty] | 187.1155 | 43.775 | [6 empty] | 187.1155 | 136.22
|  |  | 187.1902 | 41.705 | [6 empty] | 187.1902 | 43.295 | [6 empty] | 187.1902 | 136.295
|  |  | 187.2648 | 41.48 | [6 empty] | 187.2648 | 42.9 | [6 empty] | 187.2648 | 136.37
|  |  | 187.3395 | 42.425 | [6 empty] | 187.3395 | 42.525 | [6 empty] | 187.3395 | 136.38
|  |  | 187.4142 | 42.375 | [6 empty] | 187.4142 | 42.57 | [6 empty] | 187.4142 | 136.645
|  |  | 187.4888 | 42.125 | [6 empty] | 187.4888 | 43.035 | [6 empty] | 187.4888 | 137.585
|  |  | 187.5635 | 42.075 | [6 empty] | 187.5635 | 43.58 | [6 empty] | 187.5635 | 138.22
|  |  | 187.6382 | 42.96 | [6 empty] | 187.6382 | 44.035 | [6 empty] | 187.6382 | 138.285
|  |  | 187.7128 | 41.53 | [6 empty] | 187.7128 | 44.22 | [6 empty] | 187.7128 | 138.2
|  |  | 187.7875 | 41.745 | [6 empty] | 187.7875 | 44.405 | [6 empty] | 187.7875 | 138.055
|  |  | 187.8622 | 41.15 | [6 empty] | 187.8622 | 44.64 | [6 empty] | 187.8622 | 137.7
|  |  | 187.9368 | 41.27 | [6 empty] | 187.9368 | 44.835 | [6 empty] | 187.9368 | 137.87
|  |  | 188.0115 | 40.785 | [6 empty] | 188.0115 | 44.99 | [6 empty] | 188.0115 | 138.26
|  |  | 188.0862 | 41.4 | [6 empty] | 188.0862 | 45.02 | [6 empty] | 188.0862 | 138.54
|  |  | 188.1608 | 40.785 | [6 empty] | 188.1608 | 44.985 | [6 empty] | 188.1608 | 138.2
|  |  | 188.2355 | 41.435 | [6 empty] | 188.2355 | 44.86 | [6 empty] | 188.2355 | 137.86
|  |  | 188.3102 | 41.06 | [6 empty] | 188.3102 | 44.795 | [6 empty] | 188.3102 | 137.86
|  |  | 188.3848 | 41.345 | [6 empty] | 188.3848 | 44.835 | [6 empty] | 188.3848 | 138.265
|  |  | 188.4595 | 41.55 | [6 empty] | 188.4595 | 44.73 | [6 empty] | 188.4595 | 138.14
|  |  | 188.5342 | 41.635 | [6 empty] | 188.5342 | 44.51 | [6 empty] | 188.5342 | 137.96
|  |  | 188.6088 | 41.23 | [6 empty] | 188.6088 | 44.31 | [6 empty] | 188.6088 | 137.655
|  |  | 188.6835 | 41.345 | [6 empty] | 188.6835 | 44.175 | [6 empty] | 188.6835 | 137.96
|  |  | 188.7582 | 41.935 | [6 empty] | 188.7582 | 43.96 | [6 empty] | 188.7582 | 138.025
|  |  | 188.8328 | 42.76 | [6 empty] | 188.8328 | 44.27 | [6 empty] | 188.8328 | 137.87
|  |  | 188.9075 | 42.29 | [6 empty] | 188.9075 | 45.14 | [6 empty] | 188.9075 | 137.825
|  |  | 188.9822 | 41.725 | [6 empty] | 188.9822 | 45.945 | [6 empty] | 188.9822 | 138.27
|  |  | 189.0568 | 41.92 | [6 empty] | 189.0568 | 46.77 | [6 empty] | 189.0568 | 138.16
|  |  | 189.1315 | 41.56 | [6 empty] | 189.1315 | 47.05 | [6 empty] | 189.1315 | 137.62
|  |  | 189.2062 | 42.535 | [6 empty] | 189.2062 | 47.145 | [6 empty] | 189.2062 | 137.085
|  |  | 189.2808 | 41.22 | [6 empty] | 189.2808 | 47.215 | [6 empty] | 189.2808 | 136.905
|  |  | 189.3555 | 40.815 | [6 empty] | 189.3555 | 47.42 | [6 empty] | 189.3555 | 137.055
|  |  | 189.4302 | 40.935 | [6 empty] | 189.4302 | 47.525 | [6 empty] | 189.4302 | 137.445
|  |  | 189.5048 | 42.275 | [6 empty] | 189.5048 | 47.55 | [6 empty] | 189.5048 | 136.78
|  |  | 189.5795 | 42.36 | [6 empty] | 189.5795 | 47.545 | [6 empty] | 189.5795 | 136.295
|  |  | 189.6542 | 42.15 | [6 empty] | 189.6542 | 47.495 | [6 empty] | 189.6542 | 135.965
|  |  | 189.7289 | 41.43 | [6 empty] | 189.7289 | 47.395 | [6 empty] | 189.7289 | 135.92
|  |  | 189.8035 | 40.82 | [6 empty] | 189.8035 | 47.445 | [6 empty] | 189.8035 | 136.32
|  |  | 189.8782 | 42.005 | [6 empty] | 189.8782 | 47.57 | [6 empty] | 189.8782 | 136.71
|  |  | 189.9529 | 41.715 | [6 empty] | 189.9529 | 47.3 | [6 empty] | 189.9529 | 136.375
|  |  | 190.0275 | 43.235 | [6 empty] | 190.0275 | 47.05 | [6 empty] | 190.0275 | 135.495
|  |  | 190.1022 | 42.45 | [6 empty] | 190.1022 | 46.905 | [6 empty] | 190.1022 | 135.48
|  |  | 190.1768 | 42.335 | [6 empty] | 190.1768 | 46.985 | [6 empty] | 190.1768 | 135.905
|  |  | 190.2515 | 42.925 | [6 empty] | 190.2515 | 47.67 | [6 empty] | 190.2515 | 136.25
|  |  | 190.3262 | 43.725 | [6 empty] | 190.3262 | 48.545 | [6 empty] | 190.3262 | 136
|  |  | 190.4008 | 43.215 | [6 empty] | 190.4008 | 49.275 | [6 empty] | 190.4008 | 135.7
|  |  | 190.4755 | 42.39 | [6 empty] | 190.4755 | 50.02 | [6 empty] | 190.4755 | 135.48
|  |  | 190.5502 | 42.14 | [6 empty] | 190.5502 | 50.33 | [6 empty] | 190.5502 | 135.55
|  |  | 190.6248 | 43.39 | [6 empty] | 190.6248 | 50.55 | [6 empty] | 190.6248 | 135.495
|  |  | 190.6995 | 43.125 | [6 empty] | 190.6995 | 50.75 | [6 empty] | 190.6995 | 134.74
|  |  | 190.7742 | 42.07 | [6 empty] | 190.7742 | 50.805 | [6 empty] | 190.7742 | 134.1
|  |  | 190.8488 | 42.775 | [6 empty] | 190.8488 | 50.855 | [6 empty] | 190.8488 | 133.96
|  |  | 190.9235 | 43.13 | [6 empty] | 190.9235 | 50.815 | [6 empty] | 190.9235 | 134.725
|  |  | 190.9982 | 41.915 | [6 empty] | 190.9982 | 50.78 | [6 empty] | 190.9982 | 135.43
|  |  | 191.0728 | 41.955 | [6 empty] | 191.0728 | 50.76 | [6 empty] | 191.0728 | 135.65
|  |  | 191.1475 | 42.39 | [6 empty] | 191.1475 | 50.52 | [6 empty] | 191.1475 | 135
|  |  | 191.2222 | 42.83 | [6 empty] | 191.2222 | 50.275 | [6 empty] | 191.2222 | 134.455
|  |  | 191.2969 | 42.515 | [6 empty] | 191.2969 | 49.995 | [6 empty] | 191.2969 | 134.11
|  |  | 191.3715 | 43.78 | [6 empty] | 191.3715 | 49.58 | [6 empty] | 191.3715 | 133.435
|  |  | 191.4462 | 43.655 | [6 empty] | 191.4462 | 49.485 | [6 empty] | 191.4462 | 133.21
|  |  | 191.5209 | 42.78 | [6 empty] | 191.5209 | 49.655 | [6 empty] | 191.5209 | 132.97
|  |  | 191.5955 | 43.415 | [6 empty] | 191.5955 | 50.325 | [6 empty] | 191.5955 | 132.77
|  |  | 191.6702 | 43.52 | [6 empty] | 191.6702 | 50.665 | [6 empty] | 191.6702 | 132.97
|  |  | 191.7449 | 42.94 | [6 empty] | 191.7449 | 50.795 | [6 empty] | 191.7449 | 133.215
|  |  | 191.8195 | 43.63 | [6 empty] | 191.8195 | 51.005 | [6 empty] | 191.8195 | 133.33
|  |  | 191.8942 | 42.93 | [6 empty] | 191.8942 | 51.085 | [6 empty] | 191.8942 | 133.375
|  |  | 191.9689 | 43.905 | [6 empty] | 191.9689 | 51.13 | [6 empty] | 191.9689 | 133.06
|  |  | 192.0435 | 44.835 | [6 empty] | 192.0435 | 51.12 | [6 empty] | 192.0435 | 132.825
|  |  | 192.1182 | 45.09 | [6 empty] | 192.1182 | 50.97 | [6 empty] | 192.1182 | 132.145
|  |  | 192.1929 | 44.33 | [6 empty] | 192.1929 | 50.725 | [6 empty] | 192.1929 | 132
|  |  | 192.2675 | 43.16 | [6 empty] | 192.2675 | 50.41 | [6 empty] | 192.2675 | 132.17
|  |  | 192.3422 | 41.98 | [6 empty] | 192.3422 | 50.2 | [6 empty] | 192.3422 | 132.435
|  |  | 192.4169 | 43.925 | [6 empty] | 192.4169 | 49.94 | [6 empty] | 192.4169 | 132.035
|  |  | 192.4915 | 43.49 | [6 empty] | 192.4915 | 49.63 | [6 empty] | 192.4915 | 131.315
|  |  | 192.5662 | 43.525 | [6 empty] | 192.5662 | 49.365 | [6 empty] | 192.5662 | 131.105
|  |  | 192.6409 | 43.01 | [6 empty] | 192.6409 | 49.005 | [6 empty] | 192.6409 | 131.465
|  |  | 192.7155 | 43.39 | [6 empty] | 192.7155 | 48.275 | [6 empty] | 192.7155 | 132.235
|  |  | 192.7902 | 42.925 | [6 empty] | 192.7902 | 47.72 | [6 empty] | 192.7902 | 132.49
|  |  | 192.8649 | 43.38 | [6 empty] | 192.8649 | 47.17 | [6 empty] | 192.8649 | 132.385
|  |  | 192.9395 | 42.84 | [6 empty] | 192.9395 | 46.93 | [6 empty] | 192.9395 | 132.755
|  |  | 193.0142 | 43.67 | [6 empty] | 193.0142 | 47.095 | [6 empty] | 193.0142 | 133.005
|  |  | 193.0889 | 42.7 | [6 empty] | 193.0889 | 47.41 | [6 empty] | 193.0889 | 133.22
|  |  | 193.1635 | 41.795 | [6 empty] | 193.1635 | 47.595 | [6 empty] | 193.1635 | 132.865
|  |  | 193.2382 | 42.18 | [6 empty] | 193.2382 | 47.755 | [6 empty] | 193.2382 | 132.105
|  |  | 193.3129 | 43.01 | [6 empty] | 193.3129 | 47.855 | [6 empty] | 193.3129 | 131.965
|  |  | 193.3875 | 43.735 | [6 empty] | 193.3875 | 47.925 | [6 empty] | 193.3875 | 131.955
|  |  | 193.4622 | 43.91 | [6 empty] | 193.4622 | 47.875 | [6 empty] | 193.4622 | 132.675
|  |  | 193.5369 | 43.06 | [6 empty] | 193.5369 | 47.81 | [6 empty] | 193.5369 | 132.805
|  |  | 193.6115 | 42.745 | [6 empty] | 193.6115 | 47.725 | [6 empty] | 193.6115 | 132.89
|  |  | 193.6862 | 43.135 | [6 empty] | 193.6862 | 47.585 | [6 empty] | 193.6862 | 132.55
|  |  | 193.7609 | 42.74 | [6 empty] | 193.7609 | 47.35 | [6 empty] | 193.7609 | 132.51
|  |  | 193.8355 | 42.095 | [6 empty] | 193.8355 | 47.035 | [6 empty] | 193.8355 | 132.825
|  |  | 193.9102 | 42.405 | [6 empty] | 193.9102 | 46.86 | [6 empty] | 193.9102 | 133.28
|  |  | 193.9849 | 42.865 | [6 empty] | 193.9849 | 46.635 | [6 empty] | 193.9849 | 133.44
|  |  | 194.0595 | 42.72 | [6 empty] | 194.0595 | 46.145 | [6 empty] | 194.0595 | 133.52
|  |  | 194.1342 | 43.015 | [6 empty] | 194.1342 | 45.83 | [6 empty] | 194.1342 | 134.025
|  |  | 194.2089 | 42.875 | [6 empty] | 194.2089 | 45.49 | [6 empty] | 194.2089 | 134.255
|  |  | 194.2835 | 42.915 | [6 empty] | 194.2835 | 45.04 | [6 empty] | 194.2835 | 134.155
|  |  | 194.3582 | 42.5 | [6 empty] | 194.3582 | 44.41 | [6 empty] | 194.3582 | 134.48
|  |  | 194.4329 | 43.825 | [6 empty] | 194.4329 | 44.44 | [6 empty] | 194.4329 | 135.205
|  |  | 194.5075 | 43.475 | [6 empty] | 194.5075 | 44.925 | [6 empty] | 194.5075 | 135.515
|  |  | 194.5822 | 43.995 | [6 empty] | 194.5822 | 45.84 | [6 empty] | 194.5822 | 135.385
|  |  | 194.6569 | 42.685 | [6 empty] | 194.6569 | 46.135 | [6 empty] | 194.6569 | 135.265
|  |  | 194.7315 | 42.475 | [6 empty] | 194.7315 | 46.18 | [6 empty] | 194.7315 | 135.535
|  |  | 194.8062 | 42.3 | [6 empty] | 194.8062 | 46.325 | [6 empty] | 194.8062 | 135.66
|  |  | 194.8809 | 42.475 | [6 empty] | 194.8809 | 46.405 | [6 empty] | 194.8809 | 135.435
|  |  | 194.9555 | 42.36 | [6 empty] | 194.9555 | 46.33 | [6 empty] | 194.9555 | 135.42
|  |  | 195.0302 | 42.385 | [6 empty] | 195.0302 | 46.295 | [6 empty] | 195.0302 | 135.555
|  |  | 195.1049 | 42.575 | [6 empty] | 195.1049 | 46.145 | [6 empty] | 195.1049 | 135.785
|  |  | 195.1795 | 42.035 | [6 empty] | 195.1795 | 45.945 | [6 empty] | 195.1795 | 136.025
|  |  | 195.2542 | 42.655 | [6 empty] | 195.2542 | 45.3 | [6 empty] | 195.2542 | 136.39
|  |  | 195.3289 | 42.41 | [6 empty] | 195.3289 | 44.97 | [6 empty] | 195.3289 | 137.27
|  |  | 195.4035 | 41.05 | [6 empty] | 195.4035 | 44.73 | [6 empty] | 195.4035 | 137.965
|  |  | 195.4782 | 42.07 | [6 empty] | 195.4782 | 44.34 | [6 empty] | 195.4782 | 138.245
|  |  | 195.5529 | 42.16 | [6 empty] | 195.5529 | 43.965 | [6 empty] | 195.5529 | 138.455
|  |  | 195.6275 | 42.09 | [6 empty] | 195.6275 | 43.54 | [6 empty] | 195.6275 | 138.41
|  |  | 195.7022 | 42.05 | [6 empty] | 195.7022 | 43.04 | [6 empty] | 195.7022 | 138.205
|  |  | 195.7769 | 41.695 | [6 empty] | 195.7769 | 42.84 | [6 empty] | 195.7769 | 138.315
|  |  | 195.8515 | 42.25 | [6 empty] | 195.8515 | 43.265 | [6 empty] | 195.8515 | 138.93
|  |  | 195.9262 | 42.29 | [6 empty] | 195.9262 | 44.2 | [6 empty] | 195.9262 | 139.04
|  |  | 196.0009 | 41.43 | [6 empty] | 196.0009 | 44.73 | [6 empty] | 196.0009 | 139.11
|  |  | 196.0755 | 41.465 | [6 empty] | 196.0755 | 45.12 | [6 empty] | 196.0755 | 138.815
|  |  | 196.1502 | 42 | [6 empty] | 196.1502 | 45.4 | [6 empty] | 196.1502 | 138.7
|  |  | 196.2249 | 42.495 | [6 empty] | 196.2249 | 45.625 | [6 empty] | 196.2249 | 139.33
|  |  | 196.2995 | 42.81 | [6 empty] | 196.2995 | 45.6 | [6 empty] | 196.2995 | 139.235
|  |  | 196.3742 | 42.04 | [6 empty] | 196.3742 | 45.57 | [6 empty] | 196.3742 | 138.945
|  |  | 196.4489 | 41.67 | [6 empty] | 196.4489 | 45.52 | [6 empty] | 196.4489 | 138.085
|  |  | 196.5235 | 41.935 | [6 empty] | 196.5235 | 45.475 | [6 empty] | 196.5235 | 137.455
|  |  | 196.5982 | 43.4 | [6 empty] | 196.5982 | 45.385 | [6 empty] | 196.5982 | 137.865
|  |  | 196.6729 | 42.99 | [6 empty] | 196.6729 | 45.28 | [6 empty] | 196.6729 | 138.72
|  |  | 196.7475 | 42.58 | [6 empty] | 196.7475 | 45.1 | [6 empty] | 196.7475 | 138.26
|  |  | 196.8222 | 41.72 | [6 empty] | 196.8222 | 44.71 | [6 empty] | 196.8222 | 138.005
|  |  | 196.8969 | 41.5 | [6 empty] | 196.8969 | 44.315 | [6 empty] | 196.8969 | 137.765
|  |  | 196.9715 | 42.765 | [6 empty] | 196.9715 | 44.36 | [6 empty] | 196.9715 | 137.77
|  |  | 197.0462 | 43.17 | [6 empty] | 197.0462 | 44.64 | [6 empty] | 197.0462 | 137.91
|  |  | 197.1209 | 42.355 | [6 empty] | 197.1209 | 45.175 | [6 empty] | 197.1209 | 138.53
|  |  | 197.1955 | 42.47 | [6 empty] | 197.1955 | 46.015 | [6 empty] | 197.1955 | 138.935
|  |  | 197.2702 | 42.895 | [6 empty] | 197.2702 | 46.25 | [6 empty] | 197.2702 | 138.39
|  |  | 197.3449 | 43.59 | [6 empty] | 197.3449 | 46.575 | [6 empty] | 197.3449 | 137.115
|  |  | 197.4195 | 43.975 | [6 empty] | 197.4195 | 46.78 | [6 empty] | 197.4195 | 137.33
|  |  | 197.4942 | 43.015 | [6 empty] | 197.4942 | 46.965 | [6 empty] | 197.4942 | 137.99
|  |  | 197.5689 | 41.805 | [6 empty] | 197.5689 | 47.095 | [6 empty] | 197.5689 | 137.655
|  |  | 197.6436 | 41.085 | [6 empty] | 197.6436 | 47.185 | [6 empty] | 197.6436 | 136.73
|  |  | 197.7182 | 41.57 | [6 empty] | 197.7182 | 47.165 | [6 empty] | 197.7182 | 135.84
|  |  | 197.7929 | 41.085 | [6 empty] | 197.7929 | 47.105 | [6 empty] | 197.7929 | 135.775
|  |  | 197.8676 | 41.375 | [6 empty] | 197.8676 | 47.105 | [6 empty] | 197.8676 | 135.995
|  |  | 197.9422 | 41.125 | [6 empty] | 197.9422 | 47.165 | [6 empty] | 197.9422 | 136.24
|  |  | 198.0169 | 41.655 | [6 empty] | 198.0169 | 47.065 | [6 empty] | 198.0169 | 136.495
|  |  | 198.0916 | 42.07 | [6 empty] | 198.0916 | 46.86 | [6 empty] | 198.0916 | 136.41
|  |  | 198.1662 | 42.54 | [6 empty] | 198.1662 | 46.575 | [6 empty] | 198.1662 | 136.43
|  |  | 198.2409 | 43.445 | [6 empty] | 198.2409 | 46.28 | [6 empty] | 198.2409 | 136.545
|  |  | 198.3156 | 42.775 | [6 empty] | 198.3156 | 46.205 | [6 empty] | 198.3156 | 136.515
|  |  | 198.3902 | 41.775 | [6 empty] | 198.3902 | 46.12 | [6 empty] | 198.3902 | 136.685
|  |  | 198.4649 | 41.695 | [6 empty] | 198.4649 | 46.195 | [6 empty] | 198.4649 | 136.825
|  |  | 198.5396 | 41.99 | [6 empty] | 198.5396 | 46.905 | [6 empty] | 198.5396 | 136.875
|  |  | 198.6142 | 41.535 | [6 empty] | 198.6142 | 47.525 | [6 empty] | 198.6142 | 136.38
|  |  | 198.6889 | 41.575 | [6 empty] | 198.6889 | 47.915 | [6 empty] | 198.6889 | 135.74
|  |  | 198.7635 | 41.72 | [6 empty] | 198.7635 | 48.515 | [6 empty] | 198.7635 | 135.42
|  |  | 198.8382 | 42.395 | [6 empty] | 198.8382 | 48.785 | [6 empty] | 198.8382 | 135.665
|  |  | 198.9129 | 43.755 | [6 empty] | 198.9129 | 48.845 | [6 empty] | 198.9129 | 135.875
|  |  | 198.9875 | 42.42 | [6 empty] | 198.9875 | 48.68 | [6 empty] | 198.9875 | 135.515
|  |  | 199.0622 | 41.815 | [6 empty] | 199.0622 | 48.51 | [6 empty] | 199.0622 | 134.92
|  |  | 199.1369 | 40.63 | [6 empty] | 199.1369 | 48.35 | [6 empty] | 199.1369 | 134.5
|  |  | 199.2115 | 40.405 | [6 empty] | 199.2115 | 48.2 | [6 empty] | 199.2115 | 134.595
|  |  | 199.2862 | 39.835 | [6 empty] | 199.2862 | 47.845 | [6 empty] | 199.2862 | 135.04
|  |  | 199.3609 | 39.21 | [6 empty] | 199.3609 | 47.525 | [6 empty] | 199.3609 | 135.135
|  |  | 199.4356 | 41.24 | [6 empty] | 199.4356 | 47.265 | [6 empty] | 199.4356 | 134.63
|  |  | 199.5102 | 41.27 | [6 empty] | 199.5102 | 46.92 | [6 empty] | 199.5102 | 134.46
|  |  | 199.5849 | 40.035 | [6 empty] | 199.5849 | 46.59 | [6 empty] | 199.5849 | 134.165
|  |  | 199.6596 | 39.2 | [6 empty] | 199.6596 | 46.225 | [6 empty] | 199.6596 | 133.975
|  |  | 199.7342 | 39.2 | [6 empty] | 199.7342 | 45.69 | [6 empty] | 199.7342 | 134.06
|  |  | 199.8089 | 0 | [6 empty] | 199.8089 | 45.34 | [6 empty] | 199.8089 | 134.145
|  |  | 199.8836 | 0 | [6 empty] | 199.8836 | 44.855 | [6 empty] | 199.8836 | 134.495
|  |  | 199.9582 | 0 | [6 empty] | 199.9582 | 43.995 | [6 empty] | 199.9582 | 134.98
|  |  | 200.0329 | 0 | [6 empty] | 200.0329 | 43.215 | [6 empty] | 200.0329 | 134.6
[11 empty] | 200.1076 | 42.4 | [6 empty] | 200.1076 | 134.105
[11 empty] | 200.1822 | 40.865 | [6 empty] | 200.1822 | 133.96
[11 empty] | 200.2569 | 39.525 | [6 empty] | 200.2569 | 134.185
[11 empty] | 200.3316 | 37.76 | [6 empty] | 200.3316 | 134.195
[11 empty] | 200.4062 | 37.76 | [6 empty] | 200.4062 | 134.255
[11 empty] | 200.4809 | 37.76 | [6 empty] | 200.4809 | 134.495
[11 empty] | 200.5556 | 37.76 | [6 empty] | 200.5556 | 134.375
[11 empty] | 200.6302 | 37.76 | [6 empty] | 200.6302 | 134.01
[11 empty] | 200.7049 | 0 | [6 empty] | 200.7049 | 133.67
[11 empty] | 200.7796 | 0 | [6 empty] | 200.7796 | 133.315
[11 empty] | 200.8542 | 0 | [6 empty] | 200.8542 | 133.445
[11 empty] | 200.9289 | 0 | [6 empty] | 200.9289 | 133.66
[19 empty] | 201.0036 | 134.305
[19 empty] | 201.0782 | 134.935
[19 empty] | 201.1529 | 135.105
[19 empty] | 201.2276 | 135.135
[19 empty] | 201.3022 | 135.075
[19 empty] | 201.3769 | 135.35
[19 empty] | 201.4516 | 135.54
[19 empty] | 201.5262 | 135.075
[19 empty] | 201.6009 | 134.785
[19 empty] | 201.6756 | 134.885
[19 empty] | 201.7502 | 134.82
[19 empty] | 201.8249 | 134.56
[19 empty] | 201.8996 | 134.135
[19 empty] | 201.9742 | 133.725
[19 empty] | 202.0489 | 133.935
[19 empty] | 202.1236 | 134.37
[19 empty] | 202.1982 | 134.965
[19 empty] | 202.2729 | 134.55
[19 empty] | 202.3476 | 133.9
[19 empty] | 202.4222 | 134.005
[19 empty] | 202.4969 | 134.55
[19 empty] | 202.5716 | 134.7
[19 empty] | 202.6462 | 134.14
[19 empty] | 202.7209 | 133.71
[19 empty] | 202.7956 | 133.655
[19 empty] | 202.8702 | 134.155
[19 empty] | 202.9449 | 134.56
[19 empty] | 203.0196 | 133.92
[19 empty] | 203.0942 | 132.71
[19 empty] | 203.1689 | 131.1
[19 empty] | 203.2436 | 130.88
[19 empty] | 203.3182 | 131.045
[19 empty] | 203.3929 | 131.295
[19 empty] | 203.4676 | 130.89
[19 empty] | 203.5422 | 129.84
[19 empty] | 203.6169 | 128.415
[19 empty] | 203.6916 | 127.85
[19 empty] | 203.7662 | 127.445
[19 empty] | 203.8409 | 127.405
[19 empty] | 203.9156 | 127.425
[19 empty] | 203.9902 | 127.005
[19 empty] | 204.0649 | 126.84
[19 empty] | 204.1396 | 126.96
[19 empty] | 204.2142 | 126.48
[19 empty] | 204.2889 | 125.79
[19 empty] | 204.3636 | 125.425
[19 empty] | 204.4382 | 124.83
[19 empty] | 204.5129 | 124.275
[19 empty] | 204.5876 | 124.55
[19 empty] | 204.6622 | 124.365
[19 empty] | 204.7369 | 123.665
[19 empty] | 204.8116 | 123.05
[19 empty] | 204.8862 | 122.45
[19 empty] | 204.9609 | 122.5
[19 empty] | 205.0356 | 122.375
[19 empty] | 205.1102 | 121.685
[19 empty] | 205.1849 | 121.585
[19 empty] | 205.2596 | 121.375
[19 empty] | 205.3342 | 120.455
[19 empty] | 205.4089 | 119.93
[19 empty] | 205.4836 | 119.78
[19 empty] | 205.5582 | 119.73
[19 empty] | 205.6329 | 119.61
[19 empty] | 205.7076 | 119.46
[19 empty] | 205.7823 | 118.785
[19 empty] | 205.8569 | 118.35
[19 empty] | 205.9316 | 118.14
[19 empty] | 206.0063 | 117.46
[19 empty] | 206.0809 | 117.625
[19 empty] | 206.1556 | 117.99
[19 empty] | 206.2303 | 117.03
[19 empty] | 206.3049 | 116.225
[19 empty] | 206.3796 | 115.305
[19 empty] | 206.4543 | 114.59
[19 empty] | 206.5289 | 114.85
[19 empty] | 206.6036 | 115.07
[19 empty] | 206.6783 | 114.475
[19 empty] | 206.7529 | 113.815
[19 empty] | 206.8276 | 113.555
[19 empty] | 206.9023 | 112.915
[19 empty] | 206.9769 | 112.17
[19 empty] | 207.0516 | 111.97
[19 empty] | 207.1263 | 112.16
[19 empty] | 207.2009 | 111.97
[19 empty] | 207.2756 | 111.295
[19 empty] | 207.3503 | 111.19
[19 empty] | 207.4249 | 111.355
[19 empty] | 207.4996 | 110.385
[19 empty] | 207.5743 | 109.395
[19 empty] | 207.6489 | 109.34
[19 empty] | 207.7236 | 109.55
[19 empty] | 207.7983 | 109.455
[19 empty] | 207.8729 | 108.965
[19 empty] | 207.9476 | 108.23
[19 empty] | 208.0223 | 107.815
[19 empty] | 208.0969 | 108.305
[19 empty] | 208.1716 | 107.91
[19 empty] | 208.2463 | 106.605
[19 empty] | 208.3209 | 106.085
[19 empty] | 208.3956 | 105.815
[19 empty] | 208.4703 | 106.045
[19 empty] | 208.5449 | 106.26
[19 empty] | 208.6196 | 105.995
[19 empty] | 208.6943 | 104.72
[19 empty] | 208.7689 | 103.335
[19 empty] | 208.8436 | 103.165
[19 empty] | 208.9183 | 103.12
[19 empty] | 208.9929 | 103.165
[19 empty] | 209.0676 | 102.76
[19 empty] | 209.1423 | 102.41
[19 empty] | 209.2169 | 102.04
[19 empty] | 209.2916 | 101.305
[19 empty] | 209.3663 | 100.615
[19 empty] | 209.4409 | 99.74
[19 empty] | 209.5156 | 99.54
[19 empty] | 209.5903 | 99.775
[19 empty] | 209.6649 | 99.705
[19 empty] | 209.7396 | 99.44
[19 empty] | 209.8143 | 98.475
[19 empty] | 209.8889 | 97.855
[19 empty] | 209.9636 | 97.735
[19 empty] | 210.0383 | 97.405
[19 empty] | 210.1129 | 96.765
[19 empty] | 210.1876 | 95.945
[19 empty] | 210.2623 | 96.015
[19 empty] | 210.3369 | 95.97
[19 empty] | 210.4116 | 95.585
[19 empty] | 210.4863 | 95.43
[19 empty] | 210.5609 | 95.3
[19 empty] | 210.6356 | 94.565
[19 empty] | 210.7103 | 94.1
[19 empty] | 210.7849 | 93.87
[19 empty] | 210.8596 | 93.8
[19 empty] | 210.9343 | 93.05
[19 empty] | 211.0089 | 93.05
[19 empty] | 211.0836 | 92.19
[19 empty] | 211.1583 | 91.245
[19 empty] | 211.2329 | 91.245
[19 empty] | 211.3076 | 90.68
[19 empty] | 211.3823 | 90.255
[19 empty] | 211.4569 | 89.575
[19 empty] | 211.5316 | 89.285
[19 empty] | 211.6063 | 88.77
[19 empty] | 211.6809 | 88.585
[19 empty] | 211.7556 | 89.03
[19 empty] | 211.8303 | 88.94
[19 empty] | 211.9049 | 88.395
[19 empty] | 211.9796 | 87.81
[19 empty] | 212.0543 | 87.25
[19 empty] | 212.129 | 86.575
[19 empty] | 212.2036 | 86.71
[19 empty] | 212.2783 | 86.16
[19 empty] | 212.353 | 85.595
[19 empty] | 212.4276 | 84.215
[19 empty] | 212.5023 | 83.38
[19 empty] | 212.577 | 82.745
[19 empty] | 212.6516 | 83.15
[19 empty] | 212.7263 | 83.025
[19 empty] | 212.8009 | 82.89
[19 empty] | 212.8756 | 83.035
[19 empty] | 212.9503 | 82.35
[19 empty] | 213.0249 | 81.85
[19 empty] | 213.0996 | 81.45
[19 empty] | 213.1743 | 81.505
[19 empty] | 213.2489 | 81.345
[19 empty] | 213.3236 | 80.89
[19 empty] | 213.3983 | 80.09
[19 empty] | 213.4729 | 79.785
[19 empty] | 213.5476 | 79.355
[19 empty] | 213.6223 | 78.96
[19 empty] | 213.6969 | 79.535
[19 empty] | 213.7716 | 79.445
[19 empty] | 213.8463 | 78.78
[19 empty] | 213.921 | 78.495
[19 empty] | 213.9956 | 77.88
[19 empty] | 214.0703 | 76.755
[19 empty] | 214.145 | 76.605
[19 empty] | 214.2196 | 76.37
[19 empty] | 214.2943 | 75.42
[19 empty] | 214.369 | 74.97
[19 empty] | 214.4436 | 75.895
[19 empty] | 214.5183 | 74.495
[19 empty] | 214.593 | 73.695
[19 empty] | 214.6676 | 73.385
[19 empty] | 214.7423 | 73.67
[19 empty] | 214.817 | 73.79
[19 empty] | 214.8916 | 74.06
[19 empty] | 214.9663 | 73.26
[19 empty] | 215.041 | 72.14
[19 empty] | 215.1156 | 72.49
[19 empty] | 215.1903 | 72.68
[19 empty] | 215.265 | 71.07
[19 empty] | 215.3396 | 69.73
[19 empty] | 215.4143 | 69.74
[19 empty] | 215.489 | 69.345
[19 empty] | 215.5636 | 68.54
[19 empty] | 215.6383 | 68.425
[19 empty] | 215.713 | 67.485
[19 empty] | 215.7876 | 66.925
[19 empty] | 215.8623 | 66.9
[19 empty] | 215.937 | 65.945
[19 empty] | 216.0116 | 64.535
[19 empty] | 216.0863 | 63.435
[19 empty] | 216.161 | 62.89
[19 empty] | 216.2356 | 62.89
[19 empty] | 216.3103 | 66.86
[19 empty] | 216.385 | 66.86
[19 empty] | 216.4596 | 62.275
[19 empty] | 216.5343 | 63.895
[19 empty] | 216.609 | 65.735
[19 empty] | 216.6836 | 65.305
[19 empty] | 216.7583 | 63.595
[19 empty] | 216.833 | 63.265
[19 empty] | 216.9076 | 62.97
[19 empty] | 216.9823 | 62.55
[19 empty] | 217.057 | 61.7
[19 empty] | 217.1316 | 60.775
[19 empty] | 217.2063 | 60.21
[19 empty] | 217.281 | 60.735
[19 empty] | 217.3556 | 60.915
[19 empty] | 217.4303 | 60.34
[19 empty] | 217.505 | 59.355
[19 empty] | 217.5796 | 58.78
[19 empty] | 217.6543 | 58.08
[19 empty] | 217.729 | 57.355
[19 empty] | 217.8036 | 57.305
[19 empty] | 217.8783 | 56.6
[19 empty] | 217.953 | 55.625
[19 empty] | 218.0276 | 55.625
[19 empty] | 218.1023 | 0
[19 empty] | 218.177 | 0
[19 empty] | 218.2516 | 0
[19 empty] | 218.3263 | 0
Note on E16: Place curser in this cell. 1. Press Ctrl-Shift-C 2. Expose to x-rays

## Settings
XLPRO ™ Version 4.06 |  |  | BW-2 | DEFAULT
Operation Settings Area |  |  | 3500 | 3500 | Max. Freq. (Hz)
|  |  | 20 | 20 | Initial Delay (ms)
Units |  |  | 7 | 7 | End Delay (ms)
Radiation: |  | R | 100 | 100 | Width (ms)
/Time: |  | /min | 7 | 7 | Pulses in window
Scale: |  |  | 3 | 3 | Initial Interval (ms)
Pulse Rate: |  | 15 | 20 | 20 | Noise threshold (kV)
Mode |  |  | 0.9 | 0.9 | Upper threshold
Units Display: |  | Rescale | 0.7 | 0.7 | Lower threshold
Cell Update: |  | Incremental | 3 | 3 | FT-resolution control
Auto Zero: |  | Disabled |  | INSTRUMENT PARAMETERS
Capture Mode: |  | Auto | 0.09838 | 0.09838 | 9000 Sample Interval (ms)
Screen Tracking: |  | Yes | 38.08 | 38.08 | 9000 Input Filter (us) |  | 4081-4082/4083
kVp |  |  | 37.3 | 37.3 | 33.6 | 33.6 | Sample Interval (us)
Config: | kVS | Time+kV | 2.9764 | 2.9764 | 2.9764 | 2.9764 | Input Filter -3dB (kHz)
Scale: |  | 4083 | 16.2 | 16.2 | 16.2 | 16.2 | Output Filter (us)
Start delay (ms): |  | 0 | 4 | 4 | 2.725 | 2.725 | Averaging Intervals
Width (ms): |  | 5000 | 0.75 | 0.75 | Width fraction of kVp
Filtration (mm-Al): |  | 2.5 | 0.17 | 0.17 | Duplicate width (ms)
Dental threshold: |  | 0
Device Index |  | 0
Monitor Configuration |  |  |  |  | COEFFICIENTS FOR PRACTICAL kV
Serial No: |  | N/A | Power | Values
ROM Ver: |  | N/A | Min. kV | 20 | 36 | 150 | 161
Detector: |  | N/A | 0 | -18.6452 | -0.017472 | -0.017472 | 0
Chamb Kq: |  | N/A | 1 | 0.85068 | 1.0308e-05 | 1.0308e-05 | 0
Atm Corr: |  | Automatic | 2 | -0.019681 | 2.3082e-05 | 2.3082e-05 | 0
Temp (°C): |  | 22 | 3 | 0.00025663 | -1.662e-07 | -1.662e-07 | 0
Press (kPa): |  | 101.3 | 4 | -2.1424e-06 | 4.3106e-10 | 4.3106e-10 | 0
|  |  | 5 | 0 | 0 | 0 | 0
|  |  | EXP? | 1 | 0 | 0 | 0
NON-LINEAR CORRECTION FACTORS
|  |  | 1030 | 2700 | 0.12
|  |  | 0.343 | -0.569 | 0.2809 | -0.0331
|  |  | 2.9804 | -5.4582 | 4.6956 | -1.4557