## Readme
General Notes:
These worksheets in this workbook were derived from the old "Examp9010.xls" workbook. They have
been updated for use with 9000XL Version 3 and now include automatic kVp measurment capture and
analysis. In some cases where the kVp data might be erroneous (e.g. a hvl data point with added filtration)
the kVp values are hidden by coloring the font white.
Font colors have been used to help identify different types of data. In general, all user entry data is in a medium-blue color.
All text that should not need to be changed in is black. Dose data is dark-blue, time data is medium-green, and kVp data is dark-red
| EXAMPLE |  | Measured
|  | SET | Dose | Time | HV
|  | mA | mR | ms | kVp
|  | 25 | 6.68 | 44.3 | 58.5
To protect formatting and foumulas, all cells that should not need to be changed are locked and protected.
Those cells that contain formulas have also been hidden. You may of course unlock them and edit as you see fit.
Before unlocking or making changes, be sure you have retained a copy of the original workbook. To remove the protection,
Select "Unprotect Sheet" from the Tools\Protection menu.
The worksheets have been formatted to print a report and exclude comments and extraneous data. To view how a
sheet will look when printed, select "File" and then "Print Preview".
For additional information see the Radcal "XLPROManual.pdf" file either on the install CD or
in your 9000PXL directory.
Note on E13: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.

## Worksheet Information
Sample Worksheets in "9000 Sample Template.xls"
The three sample worksheets in the workbook 9000 Sample Template.xls delivered with XLPRO have been
designed to demonstrate the ease of use of XLPRO with Excel and a formatted worksheet. These
worksheets capture dose and rate from a 9010 or 9015 with an Accu-kV divider attached. These worksheets
implement typical QA measurements and data-collection requirements for a radiology department.
The degree of macro automation used in these worksheets is minimal, which makes it very easy
for users to modify the sheets to fit their individual data-collection needs.
Most of the usefulness of these worksheets comes from the formulas and graphs that are in the
worksheets themselves. The user places the cursor at the cell into which a measurement is to be
inserted and then initiates a measurement through a control-key sequence. This control-key
sequence initiates communication with the radiation monitor through the XLPRO control
interface. Once XLPRO inserts the measurement into the cell, control is returned to the user. The
user then ensures that the cursor is in the correct cell for the next measurement and initiates a new
measurement. This sequence continues until all the required measurements are completed.
For each section of these worksheets, there is a setup area to input machine parameters, a data collection
area for either manual or automatic data collection, and a comment area to outline the
results of the test and the action to be taken. All three of these worksheets come with the
protection property of Excel turned on so that only the data-entry areas of the worksheet are
unprotected and modifiable by the user. You may turn the protection property off so that you can
modify other areas easily.
You should keep in mind when modifying these worksheets that there are items such as the
Settings Area and intermediate calculation areas which are located outside of the main worksheet.
These areas must not be broken by inserting rows into the middle of them.
Macros
The workbook contains the macros listed below. The Auto_Open subroutine initializes
combination CONTROL-SHIFT keys to run the subroutines that they are associated with. These
subroutines first determine the cell at the top-left corner of the Settings Area and use these coordinates to
record the initial settings for Scale and Cell Update. Then they modify the settings area cells as needed to
execute the command, and execute the appropriate 9000XLMacro.xls macro (e.g.MacroRunMaxRate),
which does not exit if no argument is passed to the macro. The subroutines end after restoring the
Settings Area to its initial configuration and advancing the selection by one row. The Auto_Close
subroutine allows you to delete the Auto_Activate commands when the workbook closes.
FOR EXAMPLE
Sub Auto_Open()
Application.OnKey "+^o", "RateOverwrite"
Application.OnKey "+^f", "FluorokVpOverwrite"
End Sub
Sub RateOverwrite()
On Error GoTo Quit 'Trap no Settings Area or its write protection.
'Find the coordinates of the top-left cell of the Settings Area.
r1 = ActiveSheet.Names("_RCTL").RefersToRange.Row
c1 = ActiveSheet.Names("_RCTL").RefersToRange.Column
'Store the current values from the Settings Area.
incr = Cells(r1 + 10, c1 + 2)
units = Cells(r1 + 6, c1 + 2)
'Modify the Settings Area.
Cells(r1 + 10, c1 + 2) = "Overwrite"
Cells(r1 + 6, c1 + 2) = ""
'Run the MaxRate macro.
Application.Run "9000XLMacro.xls!MacroRunMaxRate"
'Restore the Settings Area to its old values and advance the selection by one row.
Cells(r1 + 10, c1 + 2) = incr
Cells(r1 + 6, c1 + 2) = units
ActiveCell.Offset(1, 0).Select
Quit:
If (Err) Then Call MsgBox("No recorded Settings Area or it is protected.", _
vbCritical, "Local Error")
On Error GoTo 0
End Sub
Sub Auto_Close()
If (Right(ActiveWorkbook.Name, 3) = "xlt") Then Exit Sub
i = 0
On Error Resume Next
For j = 1 To ActiveWorkbook.Sheets.Count 'Look for Auto_Activates.
ws = ActiveWorkbook.Sheets(j).Name
If (ws <> "Settings") Then i = Sheets(j).Names("Auto_Activate").Index
Next j
On Error GoTo 0
If (i = 0) Then Exit Sub 'Avoid query if no Auto_Activates are present.
If (MsgBox("Do you want to keep auto activation?", vbYesNo) <> vbNo) Then Exit Sub
On Error Resume Next
For j = 1 To ActiveWorkbook.Sheets.Count 'Delete Auto_Activates.
ws = ActiveWorkbook.Sheets(j).Name
If (ws <> "Settings") Then Sheets(j).Names("Auto_Activate").Delete
Next j
On Error GoTo 0
End Sub
Sheet named Radiographic
The worksheet Radiographic is set up for a strictly radiographic room or station. The red comment indicators,
in the yellow cells indicate comments that describe the steps required for executing the Measure Pulse
| macro (CONTROL-SHIFT P). |  |  | Example | Note: these comments will not be visible when you print.
A. Facility Information
Facility information contains data areas for equipment location, x-ray equipment type and
equipment operational ratings. There are no formulas associated with this section of the
worksheet.
B. kVp Accuracy and Exposure Linearity
kVp Accuracy and Exposure Linearity provides automatic input of the Dose Time & kVp
data from the radiation monitor. Formulas for this section are for kV error
calculation & Pass/Fail criteria ($24:G37), linear-regression calculation of the best-fit
straight line to the measured dose data, and the error of the measured dose data
relative to the best-fit dose data (H24:I37). Linear-regression calculations are performed just to
the right of the graph in an area of the worksheet (T21:AA37) which is not in the print area.
Formulas in this area must be modified by the user if the full data area is not used.
C. mA and Exposure Linearity
mA and Exposure Linearity allows for the automatic input of exposure data from the radiation
monitor. The Set mA data area is unprotected, and users may change these numbers to reflect the
settings on their x-ray machines. Formulas for this section are for the linear-regression calculation
of the best-fit straight line to the measured dose data (F57:F63), the error of the measured dose
data relative to the best-fit dose data (G57:G63), and average dose/mAs (B65). Again, the linefitting
calculations are performed just to the right of the graph in an area of the worksheet
(X56:AB64) that is not in the print area. Formulas in this area must be modified by the user if the
full data area is not used.
D. Timer and Exposure Linearity
Timer and Exposure Linearity allows for the automatic input of exposure data from the radiation
monitor. The Set Time data area is unprotected, and users may change these numbers
to reflect the settings on their x-ray machines. Formulas for this section are for the calculation
of timer error (F79:F94), the linear regression calculation of the best-fit straight
line to the measured dose data (G79:G94), and the error of the measured dose data relative to
the best-fit dose data (H79:H94). Again, the linefitting calculations are performed just to the right
of the graph in an area of the worksheet (X79:AB87) that is not in the print area.
Formulas in this area must be modified by the user if the full data area is not used.
E. Exposure Reproducibility
Exposure Reproducibility is a summary area where the user may enter information about the x-ray
techniques used, a summary of the results, and comments about the action required.
F. Half Value Layer
Formulas for this section are for the calculation of mR/s (E134:F139), the normalization of the mR/s values
(G134:G139), and the calculation of the HVL (D141) done in cells T132:X139. The HVL is calculated
using the mR/s values. The formulas for mR/s calculation use the time in the cell
just to the left of the mR/s cell. Again, the two-point log interpolation for
determining the calculated HVL is done just to the right of this section in a non-printing area of
the worksheet. If the user wishes to do fewer measurements in determining HVL, then the
interpolation formula must be changed also.
G. Phototimer
Phototimer allows capture of data for three different cell areas of the phototimer area.
The last four sections (H., I., J. and K.) are text-entry areas for users to insert their observations
and comments about various non-radiation related aspects of the x-ray machine performance.
Sheet named Special
The "Special" worksheet is set up for a fluoroscopic or radiographic room. The red indicators
in the yellow cells indicate comments that describe the steps required for executing the MeasureFluorokVp
| macro (CONTROL-SHIFT F). |  |  | Example | Note: these comments will not be visible when you print.
A. Facility Information
Facility information contains data areas for equipment location, x-ray equipment type and
equipment operational ratings. There are no formulas associated with this section of the
worksheet.
B. Fluoroscopy
The fluoroscopy section has data-entry areas for machine-setup parameters, for dose-rate or
exposure information, and result and action comments. There are no formulas associated with this
section of the worksheet.
1) Maximum Table Top Exposure Rate
2) Automatic Dose Control
3) Automatic Brightness Control
4) Resolution
5) Timer
6) kV Accuracy
7) Half Value Layer
8) Spot Films
9) Exposure Reproducibility
10) Mechanical
C. Radiographic Overhead X-Ray Tube
See Radiographic Sheet discussion above
Sheet named Special
The "Mobile" worksheet is set up for a mobile x-ray machine. The red comment indicators,
in the yellow cells indicate comments that describe the steps required for executing the Measure kVp
| macro (CONTROL-SHIFT K). |  |  | Example | Note: these comments will not be visible when you print.
This sheet follows the format and measurement principles found in the Radiographic sheet
discussed above.
Sheet named Waveforms
This sheet contains a number of waveforms that have been captured from different types of generators.
In all cases, the data was collected by selecting Measure kVp from the 9000XL menu (CONTROL-SHIFT K)
Sheet named Test
This sheet is provided as a place to make simple "scratch measurments" as you learn how
to use 9000XL.
Sheet named Settings
This sheet is found in all 9000XL workbooks and contains the basic settings information needed to start
making measurments. For further information regarding the Settings sheet see the Radcal "XLPROManual.pdf"
file either on the install CD or in your 9000PXL directory.
Note on F89: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on F147: 1. Place cursor in this cell and press "Control-Shift f" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make a fluoro exposure. 3. Wait for kV and Rate values to settle to stable numbers. 4. Press Escape to end.
Note on F175: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.

## Radiographic
2026 ANNUAL CALIBRATION REPORT | [25 empty] | XLPRO ™ Version 4.06
RADIOLOGY DEPARTMENT | [25 empty] | Operation Settings Area
[7 empty] | DATE: | 2026-07-29 21:20
Units
A. | FACILITY |  |  | Almost anywhere USA | [6 empty] | Room No: | B243 |  | Other: | [11 empty] | Radiation: | R
|  | State Reg. No: |  | CA423-619 |  |  |  | UNIT: | Radiographic | [16 empty] | /Time: | /min
[26 empty] | Scale: | milli
| EQUIPMENT |  |  | Manufacturer | [5 empty] | Model |  |  |  | Serial No. | [11 empty] | Pulse Rate: | 0.25
|  |  | CONSOLE: | General Electric |  |  |  |  | DXR-750 |  |  |  | A417176 | [12 empty] | Mode
|  |  | TUBE: | Machlett |  |  |  |  | Dynamax 69 B |  |  |  | 99828288 | [12 empty] | Units Display: | Omit
|  |  | RATINGS: |  |  | Maximum tube kV: | 150 | kVp. |  |  | Maximum tube mA: | 700 | mA. | Phase: | 3 | [10 empty] | Cell Update: | Incremental
[6 empty] | Inherent filtration: | 1.2 | mm Al |  |  | Total Filtration | 3.1 | mm Al | [12 empty] | Auto Zero: | Disabled
[26 empty] | Capture Mode: | Auto
[26 empty] | Screen Tracking: | Yes
B. | kVp ACCURACY AND EXPOSURE LINEARITY | [24 empty] | kVp | 9000
|  | Focal Spot: | lg |  | Field Size: | Max | [19 empty] | Config: | Dose+Time+kV
|  | 60 | mA | 100 | ms | 6 | mAs | [18 empty] | Scale: | DAccu-kV
| Radcal Accu-kV diagnostic sensor placed on table . | [24 empty] | Gain: | 32.77
|  | @ | 20.5 | inches from target, | [21 empty] | Offset: | 0.00
| Radcal Chamber at: |  |  | 15.25 | inches from target | [20 empty] | Comm
|  | Measured |  |  | kV Accuracy |  | Linearity | [10 empty] | Exposure time: | 0.1 | sec | [5 empty] | Index: | 0
| Set | Dose | Time | HV | Error | Pass/ | mR | % | [8 empty] | Ln Values |  | Reg. |  |  | Regression Output: |  |  |  | Baud: | 9600
| kVp | mR | ms | kVp | kV | Fail | regrs. | Error | [8 empty] | kVp | Dose | Dose |  |  |  | b | a |  | Monitor Configuration
| 50 | 198 | 99.9 | 49.1 | -0.9 |  | 192.6169 | -2.795% | [8 empty] | 3.8939 | 5.2883 | 192.6169 |  |  | Coeff Value | 2.079 | -2.8348 |  | Serial No: | A 1
| 55 | 225 | 99.5 | 52 | -3 |  | 217.0242 | -3.675% | [8 empty] | 3.9512 | 5.4161 | 217.0242 |  |  | Std Error | 0.029373 | 0.12877 |  | ROM Ver: | 3.24
| 60 | 260 | 99.7 | 58 | -2 |  | 272.3367 | 4.53% | [8 empty] | 4.0604 | 5.5607 | 272.3367 |  |  | R squared | 0.99761 | 0.032073 |  | Detector: | 90x6
| 65 | 291 | 100 | 61 | -4 | FAIL | 302.4414 | 3.783% | [8 empty] | 4.1109 | 5.6733 | 302.4414 | [6 empty] | Chamb Kq: | 1.00
| 70 | 381 | 101 | 69 | -1 |  | 390.7602 | 2.498% | [8 empty] | 4.2341 | 5.9428 | 390.7602 |  | y = a + (b * x) |  |  |  |  | Atm Corr: | Uncorrected
| 75 | 421 | 101 | 72 | -3 |  | 426.9119 | 1.385% | [8 empty] | 4.2767 | 6.0426 | 426.9119 |  | a= | -2.8348 |  |  |  | Temp (°C): | 21.7
| 80 | 525 | 101 | 79 | -1 |  | 517.7412 | -1.402% | [8 empty] | 4.3694 | 6.2634 | 517.7412 |  | b= | 2.079 |  |  |  | Press (kPa): | 99.6
| 85 | 560 | 100 | 83 | -2 |  | 573.7337 | 2.394% | [8 empty] | 4.4188 | 6.3279 | 573.7337
| 90 | 677 | 101 | 88 | -2 |  | 647.9295 | -4.487% | [8 empty] | 4.4773 | 6.5177 | 647.9295
| 95 | 731 | 102 | 91 | -4 | FAIL | 694.698 | -5.226% | [8 empty] | 4.5109 | 6.5944 | 694.698
| 100 | 832 | 100 | 99 | -1 |  | 827.7064 | -0.5187% | [8 empty] | 4.5951 | 6.7238 | 827.7064
| 110 | 1030 | 101 | 110 | 0 |  | 1030.406 | 0.0394% | [8 empty] | 4.7005 | 6.9373 | 1030.406
| 120 | 1210 | 102 | 120 | 0 |  | 1234.7318 | 2.003% | [8 empty] | 4.7875 | 7.0984 | 1234.7318
| 130 | 1400 | 103 | 128 | -2 |  | 1412.0359 | 0.8524% | [8 empty] | 4.852 | 7.2442 | 1412.0359
|  |  |  | Max.: | -4 | kV | Max: | -5.226%
|  |  |  | Limit ± | 3 | kV
Comments:
|  |  |  | kV Accuracy: | Service recommended for the 65 & 95 kV stations.
GE's specifications for kV accuracy need to be checked.
|  |  |  | Exposure Linearity: | Exposures linear to within 6% at all kV stations tested
Acceptable results
C. | mA AND EXPOSURE LINEARITY
|  | Technique: | 60 | kV | 50 | ms
| 21 | in. target to chamber distance
Measured
| SET | Dose | Time | HV | Regression
| mA | mR | ms | kVp | mR | Error | [16 empty] | Regression Output:
| 25 | 6.68 | 44.3 | 58.5 | 8.4324 | -20.78% | [14 empty] | 6.7459 |  |  | b | a
| 50 | 13.2 | 45 | 59 | 17.3399 | -23.88% | [14 empty] | 6.936 |  | Coeff Value | 0.3563 | -0.47518
| 100 | 33.7 | 47.5 | 59.1 | 35.155 | -4.139% | [14 empty] | 7.031 |  | Std Error | 0.015556 | 4.7879
| 200 | 70.9 | 49.2 | 59.4 | 70.7853 | 0.1621% | [14 empty] | 7.0785 |  | R squared | 0.99056 | 7.9673
| 300 | 111 | 48.9 | 59.4 | 106.4155 | 4.308% | [14 empty] | 7.0944
| 400 | 155 | 48.6 | 59.5 | 142.0457 | 9.12% | [14 empty] | 7.1023 | y = a + (b * x)
| 600 | 203 | 48.3 | 59.4 | 213.3062 | -4.832% | [14 empty] | 7.1102 | a= | -0.47518
| Regression Avg.: |  |  |  | Max. Error: | -23.88% | [15 empty] | b= | 0.3563
| 7.014 | mR/mAs @ 21 inches
Comments:
|  |  | Results: | Exposures should be linear to within +/-10%
At 25 and 50 mA the exposures are too low
|  |  | Action: | SERVICE REQUIRED
D. | TIMER AND EXPOSURE LINEARITY
|  | Technique: | 75 | kV | 300 | mA
| 21 | in. target to chamber distance
Measured
| SET | Dose | Time | HV | Timer | Regression
| Time | mR | ms | kVp | Error | mR | Error
| 17 | 78 | 18 | 74.5 | 5.882% | 79.2031 | -1.519% | [14 empty] | Regression Output:
| 25 | 89.8 | 24 | 74 | -4% | 106.4557 | -15.65% | [15 empty] | b | a
| 33 | 118 | 35 | 73.7 | 6.061% | 133.7083 | -11.75% | [14 empty] | Coeff Value | 3.4066 | 21.2914
| 50 | 178 | 52 | 73.6 | 4% | 191.62 | -7.108% | [14 empty] | Std Error | 0.019324 | 7.8785
| 67 | 233 | 65 | 74.2 | -2.985% | 249.5317 | -6.625% | [14 empty] | R squared | 0.99955 | 22.2629
| 100 | 352 | 102 | 74.17 | 2% | 361.9485 | -2.749%
| 125 | 448 | 128 | 74.14 | 2.4% | 447.1128 | 0.1984% | [13 empty] | y = a + (b * x)
| 150 | 535 | 155 | 74.11 | 3.333% | 532.277 | 0.5116% | [13 empty] | a= | 21.2914
| 200 | 704 | 205 | 74.08 | 2.5% | 702.6056 | 0.1985% | [13 empty] | b= | 3.4066
| 250 | 913 | 255 | 74.05 | 2% | 872.9341 | 4.59%
| 300 | 1060 | 306 | 74.02 | 2% | 1043.2627 | 1.604%
| 400 | 1400 | 408 | 73.99 | 2% | 1383.9197 | 1.162%
| 500 | 1760 | 508 | 73.96 | 1.6% | 1724.5768 | 2.054%
| 600 | 2080 | 610 | 73.93 | 1.667% | 2065.2339 | 0.715%
| 800 | 2700 | 815 | 73.9 | 1.875% | 2746.548 | -1.695%
| 1000 | 3420 | 1008 | 73.87 | 0.8% | 3427.8622 | -0.2294%
Comments:
|  |  | Results: | Exposures should be linear to within +/-10%
At .0167 and .025 seconds the exposures are too low
|  |  | Action: | SERVICE REQUIRED
E. | EXPOSURE REPRODUCIBILITY
|  |  |  | Dose | Time | kVp | mR/mAs
|  |  | PASS/FAIL Criteria: | 5% | 5% | 5% | 5%
|  |  | mA: | 200 | ms: | 100 |  | mA: | 200 | ms: | 100 |  | mA: | 200 | ms: | 100
|  |  | kVp: | 60 | mAs: | 20 |  | kVp: | 60 | mAs: | 20 |  | kVp: | 60 | mAs: | 20
|  |  | Measured |  |  |  |  | Measured |  |  |  |  | Measured
|  |  | Dose | Time | HV | mR/ |  | Dose | Time | HV | mR/ |  | Dose | Time | HV | mR/
|  |  | mR | ms | kVp | mAs |  | mR | ms | kVp | mAs |  | mR | ms | kVp | mAs
|  | X1 | 100 | 100 | 59.4 | 5 | X1 | 100 | 100 | 59.4 | 5 | X1 | 100 | 100 | 59.4 | 5
|  | X2 | 99 | 98 | 60 | 5.051 | X2 | 99 | 98 | 60 | 5.051 | X2 | 99 | 98 | 60 | 5.051
|  | X3 | 98 | 98 | 60.1 | 5 | X3 | 98 | 98 | 60.1 | 5 | X3 | 98 | 98 | 60.1 | 5
|  | X4 | 99 | 98.5 | 59.8 | 5.0254 | X4 | 99 | 98.5 | 59.8 | 5.0254 | X4 | 99 | 98.5 | 59.8 | 5.0254
|  | X5 | 94 | 99.6 | 60.3 | 4.7189 | X5 | 94 | 99.6 | 60.3 | 4.7189 | X5 | 94 | 99.6 | 60.3 | 4.7189
|  | Avg | 98 | 98.82 | 59.92 | 4.9591 | Avg | 98 | 98.82 | 59.92 | 4.9591 | Avg | 98 | 98.82 | 59.92 | 4.9591
|  | CV | 2.393% | 0.9395% | 0.5708% | 2.741% | CV | 2.393% | 0.9395% | 0.5708% | 2.741% | CV | 2.393% | 0.9395% | 0.5708% | 2.741%
|  | Pass/Fail | Pass | Pass | Pass | Pass |  | Pass | Pass | Pass | Pass |  | Pass | Pass | Pass | Pass
|  |  | RESULTS: | Reproducibility was within the required +/- 5%.
|  |  | COMMENT: | Acceptable results
F. | HALF VALUE LAYER
Technique
| kV | mA | ms | added filtration |  |  | 74.5 kV (75 kV setting) 200 mA, 2.0 mmAl Added
| 75 | 200 | 100 | 2 | mmAl |  | Calculated HVL @75.0 kV setting is 2.5 mmAl
| 15.25 | in. target to chamber distance
Measured
| mm | Dose | Time | HV | Calculated | [12 empty] | HVL (2 point Log interpolation)
| Al | mR | ms | kVp | mR/s | Atten | [13 empty] | (a)Just > | (b)Just < |  | 0.5 line data
| 0 | 218 | 100.1 | 74.5 | 2177.8222 | 1 | [10 empty] | 1 |  | Target Value | 0.5 |  |  | 0 | 0.5
| 1 | 158 | 99.8 | 74.8 | 1583.1663 | 0.72695 | [10 empty] | 0.72695 |  | Row | 3 | 4 |  | 5 | 0.5
| 2 | 122 | 99.9 | 75 | 1221.2212 | 0.56075 | [10 empty] | 0.56075 |  | Exposure Rate | 1221.2212 | 970
| 3 | 97 | 100 | 74.1 | 970 | 0.4454 | [10 empty] | 0.4454 |  | Thickness (mm) | 2 | 3
| 4 | 78.1 | 99.9 | 84.9 | 781.7818 | 0.35897 | [10 empty] | 0.35897
| 5 | 65.1 | 99.9 | 75 | 651.6517 | 0.29922 | [10 empty] | 0.29922 |  | Calc. HVL | 2.4979 | mmAl
|  | Calculated HVL: | 2.4979 | mmAl
Comments:
|  |  | Results: | Acceptable
|  |  | Action: | None Required
G. | PHOTOTIMER
Technique
| kV | mA | ms | cassette |  |  | 72 | inch SFD |  | 67.5 | inch SCD, Lead Apron
| 121 | 200 | 500 | 14 x 17 | in. |  |  | Results: | Chest Stand + Lead Apron
|  | Center Cell |  |  |  |  | Right Cell |  |  |  |  | Left Cell
|  | Measured |  |  |  |  | Measured |  |  |  |  | Measured
|  | Dose | Time | HV | Dose |  | Dose | Time | HV | Dose |  | Dose | Time | HV | Dose
|  | mR | ms | kVp | Ratio |  | mR | ms | kVp | Ratio |  | mR | ms | kVp | Ratio
| NORM | 44 | 77 | 121.1 | N/A |  | 31 | 52 | 121 | N/A |  | 48 | 85 | 120.8 | N/A
| 1.25 | 52 | 94 | 120.8 | 1.1818 |  | 39 | 68 | 120.8667 | 1.2581 |  | 56 | 101 | 121 | 1.1667
| 1.5 | 65 | 118 | 121 | 1.4773 |  | 47 | 77 | 120.8167 | 1.5161 |  | 72 | 134 | 120.8667 | 1.5
| 2 | 73 | 134 | 120.8667 | 1.6591 |  | 56 | 101 | 120.7667 | 1.8065 |  | 85 | 159 | 121 | 1.7708
| 0.75 | 30 | 51 | 120.8167 | 0.68182 |  | 22 | 35 | 120.7167 | 0.70968 |  | 39 | 68 | 120.8667 | 0.8125
| 0.5 | 22 | 35 | 120.7667 | 0.5 |  | 17 | 27 | 120.8667 | 0.54839 |  | 26 | 43 | 120.8167 | 0.54167
| 0.25 | 13 | 18 | 120.7167 | 0.29545 |  | 9 | 10 | 120.8167 | 0.29032 |  | 13 | 18 | 120.7667 | 0.27083
Comments:
|  |  | Results: | On both normal and 3/4 density settings, Exposures are not balanced
for the three cells. There should be better agreement in % variation
among the three cells.
|  |  | Action: | SERVICE RECOMMENDED
H. | LIGHT FIELD/X-RAY FIELD ALIGNMENT
|  | TABLE: | At 37 inch SFD - Alignment was not within 2% of SFD. The x-ray field is 1/2 inch smaller
and 1/4 inch larger than the light field (cross table), and shifted ~ 1/4 inch longitudinally.
|  | CHEST STAND: | At 72 inch SFD, the alignment is poor: 3/4 inch of the x-ray field is cut off (or shifted)
on the left side (facing the chest unit) and ~3/4 inch too large on the right side.
|  | Comments: | SERVICE REQUIRED
I. | CENTERING OF X-RAY FIELD TO CASSETTE IN BUCKY TRAY
|  | TABLE: | At 40 inch SFD, Center of x-ray field and cassette are within 2% of SFD.
|  | Comments: | Acceptable Results
|  | WALL: | Center of x-ray field and cassette were not within 2% of SFD
|  | Comments: | SERVICE REQUIRED
J. | AUTOMATIC COLLIMATION
SERVICE REQUIRED, this system was not operational.
K. | MECHANICAL
| 1 | Ready light on control panel did not light
| 2 | Chest stand cassette size indicator light did not light.
| 3 | When using low technique (~ 50 - 55 kV, 5 mAs) the RESET indicator came on.
| 4 | Several times, during testing no exposure occurred when the exposure switch was engaged.
L. | IN AIR EXPOSURE FROM CLINICAL PROTOCOLS
Measured
|  |  |  | Generator Settings |  |  | Dose | Time | HV
|  | STUDY |  | KV | mA | ms | mR | ms | kVp
|  | ABDOMEN |  | 75 | 200 | 300 | 704 | 305 | 74.8
|  |  |  | 105 | 200 | 50 | 242 | 48 | 105.1
|  |  |  | 68 | 300 | 100 | 278 | 97 | 68.7
CHEST
|  |  | a) Erect | 75 | 300 | 25 | 89 | 22 | 74.3
|  |  |  | 85 | 400 | 100 | 667 | 98 | 84.2
|  |  |  | 88 | 400 | 100 | 712 | 7 | 87.2
|  |  | b) Supine | 63 | 300 | 25 | 64 | 22 | 62.8
|  |  |  | 65 | 300 | 17 | 42 | 13 | 64.9
|  |  |  | 76 | 300 | 33 | 124 | 31 | 75.8
|  |  | LAT | 95 | 600 | 25 | 3050 | 256 | 93.7
|  |  |  | 110 | 300 | 33 | 278 | 32 | 109.6
c) PHOTOTIMED - on 3/4 Density
|  |  |  | 80 | 300 | 100 | 433 | 97 | 79.7
|  |  |  | 100 | 200 | 500 | 2240 | 506 | 100.1
HIP & PELVIS
|  |  |  | 69 | 100 | 200 | 183 | 203 | 69.2
|  |  |  | 70 | 300 | 200 | 635 | 197 | 69.9
|  |  |  | 74 | 200 | 250 | 608 | 255 | 74.2
|  |  |  | 90 | 300 | 250 | 1400 | 256 | 89.9
|  | L-SPINE |  | 75 | 200 | 250 | 585 | 255 | 75
|  |  |  | 80 | 200 | 250 | 719 | 255 | 80.1
|  |  |  | 85 | 200 | 800 | 2330 | 797 | 85
|  |  |  | 90 | 200 | 800 | 2810 | 797 | 89.8
OBSTRUCTION SERIES
|  |  |  | 78 | 300 | 150 | 604 | 147 | 77.3
|  |  |  | 80 | 300 | 400 | 1680 | 397 | 79.2
|  |  |  | 85 | 300 | 400 | 1770 | 397 | 84.4
CALIBRATION PERFORMED BY:
CERTIFIED RADIOLOGICAL PHYSICIST
Note on A3: NOTE: This worksheet requires Radcal 9000XL ver 2.1 or higher
Note on C24: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on C57: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on C79: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on C134: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on C157: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on D111: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on H157: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on M157: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.

## Special
2026 ANNUAL CALIBRATION REPORT | [22 empty] | XLPRO ™ Version 4.06
RADIOLOGY DEPARTMENT | [22 empty] | Operation Settings Area
[6 empty] | DATE: | 2026-07-29 21:20
Units
A. | FACILITY |  |  | Almost anywhere USA | [6 empty] | Room No: | 453B |  | Other: | [8 empty] | Radiation: |  | R
|  | State Reg. No: |  | No. 705544 (Fluro) |  |  |  | UNIT: | Special Procedures | [13 empty] | /Time: |  | /min
|  |  |  | No 703190 (Radiographic) | [18 empty] | Scale:
| EQUIPMENT |  |  | Manufacturer | [5 empty] | Model |  |  |  | Serial No. | [8 empty] | Pulse Rate: |  | 15
|  |  | CONSOLE: | General Electric |  |  |  |  | DXR-750 |  |  |  | A417176 | [9 empty] | Mode
|  |  | TUBE: | Machlett |  |  |  |  | Dynamax 69 B |  |  |  | 99828288 | [9 empty] | Units Display: |  | Omit
|  |  | RATINGS: |  |  | Maximum tube kV: | 165 | kVp. |  |  | Maximum tube mA: | 100 | mA. | Phase: | 1 | [7 empty] | Cell Update: |  | Overwrite
[6 empty] | Inherent filtration: |  | mm Al |  |  | Total Filtration |  | mm Al | [9 empty] | Auto Zero: |  | Disabled
Capture Mode:
B. | FLUOROSCOPY | [21 empty] | Screen Tracking: |  | Yes
| 1) | Maximum Table Top Exposure Rate | [20 empty] | kVp |  | 9000F
|  | R/min | kV: | [19 empty] | Config: |  | Dose+kV
|  | 8.6 | 120.1 | [19 empty] | Scale: |  | DAccu-kV
|  | mA: | None seen | [19 empty] | Gain: |  | 32.77
[23 empty] | Offset: |  | 0.00
|  | COMMENT: |  | Acceptable. This complies with federal and state regulations | [18 empty] | Comm
[23 empty] | Index: |  | 0
[23 empty] | Baud: |  | 9600
| 2) | Automatic Dose Control | [20 empty] | Monitor Configuration
|  | Technique: | [20 empty] | Serial No: |  | A 1
|  | kV | mA |  | Auto mode will override manual mode if phantom's thickness and density | [17 empty] | ROM Ver: |  | 3.25
|  | 70 | MAX |  | drive system beyond the selected kV (via dial on image intensifier) | [17 empty] | Detector: |  | 90x6
[23 empty] | Chamb Kq: |  | 1.00
[5 empty] | Table Top Exposures | [17 empty] | Atm Corr: |  | Automatic
|  |  | Phantom |  | kV |  | R/min | kVp | [14 empty] | Temp (°C): |  | 2
|  |  | Copper |  | 75 |  | 1.65 | 74.3 | [14 empty] | Press (kPa): |  | 100.5
[5 empty] | 80 |  | 0.73 | 80.1
[5 empty] | 90 |  | 0.46 | 91
|  |  | 0.75" Al + Cu |  | 72 |  | 7.11 | 70.4
[5 empty] | 80 |  | 2.97 | 80.4
[5 empty] | 120 |  | 0.63 | 121.3
|  |  | 1.5" Al + Cu |  | 86.5 |  | 6.955 | 80.8
[5 empty] | 121 |  | 1.6 | 120.2
|  |  | Lead + Cu + Al |  | 121 |  | 8 | 119.7
|  | COMMENT: |  | This unit is performing as per GE's Design. It will operate on the kV selected
(ie 70 kV) unless the patient/phantom's thickness or density drives the system
to maximum mA first. Then it will increase the kV. The above data illustrates
how patient dose is lowered at higher kV's. Conversely, as kV increases
contrast decreases.
| 3) | Automatic Brightness Control
|  | Technique: |  | Different thicknesses of different materials were placed on the table in the beam
The ion chamber was placed as close as possible to the outside entrance to image
intensifier assembly
[5 empty] | Phantom Material |  |  | R/min | kV
[5 empty] | 0.75 " Aluminum |  |  | 0.0065 | 70
[5 empty] | 1.5 " Aluminum |  |  | 0.006 | 70
[5 empty] | Copper |  |  | 0.005 | 70
[5 empty] | Cu + 0.75"Al |  |  | 0.005 | 70
[5 empty] | 1.5" Al + Cu |  |  | 0.006 | 85
|  | COMMENT: |  | Acceptable Results
| 4) | Resolution
|  |  | Technique: | Line pair phantom was taped to the outside entrance of the image intensifier.
|  |  | Results: | 1.1 line pairs per millimeter
|  | COMMENT: |  | Acceptable Results. This is not as good as in Room 1
| 5) | Timer
|  |  | Results: | Has audible indication at the completion of nominal 5 minutes of Fluoro time.
|  |  | COMMENT: | Acceptable Results.
| 6) | kV Accuracy
|  | Technique: |  | Radcal AccukV sensor was placed on the table at 90 degree angle with the
anode/cathode of the x-ray tube.
[6 empty] | Radcal |  | Apparent
|  |  | SET | kV (TV) |  | R/min | kVp | kV Error
|  |  | 75 | 75 |  |  | 76.1 | 1.1
|  |  | 80 | 80 |  |  | 80 | 0
|  |  | 85 | 85,85 |  |  | 82.5 | -2.5
|  |  | 90 | 89 |  |  | 93 | 3
|  |  | 100 | 99,100 |  |  | 95 | -5
|  |  | 110 |  |  |  | 103 | -7
|  |  | 120 | 121 |  |  | 115 | -5
|  |  | COMMENT: | Service Required for the high kV's
NOTE: The kV meter used to measure the kV drove the automatic system to a
minimum kV of 75.
| 7) | Half Value Layer
|  |  | Result: | 2.9 mm of Al at 75 kV
|  |  | COMMENT: | Acceptable Results
| 8) | Spot Films
|  | Technique: |  | Routine clinical settings of 110 kV, 1.5 seconds back-up for phototimer, phantom.
a) Spacing
No everlap noted on the 4/1 and 2/1 films.
b) Film Density
Variation in density corresponded with exposure variations and phantoms.
See patient exposures below.
c) Patient Exposures
|  |  |  | Technique: | Automatic Mode, Copper, Copper plus Aluminum phantoms.
70 kV on Fluoro, 95 kV, 400 mA, 0.5 sec backup.
|  |  |  | Results: | [5 empty] | Dose | Time
|  |  |  | Phantom |  |  | Spacing |  | Grid | mR | ms
[6 empty] | Copper + 1.5" Al |  | 1/1 | No | 1330 | 77
[9 empty] | Yes | 1670 | 93
[6 empty] | Copper + 3/4" Al |  | 2/1 | No | 917 | 56
[9 empty] | Yes | 874 | 56
[6 empty] | Copper + 1.5" Al |  | 2/1 | No | 2010 | 124
[9 empty] | Yes | 1920 | 119
[6 empty] | Copper + 1.5" Al |  | 4/1 | No | 2670 | 1610
[9 empty] | Yes | 2310 | 1450
[6 empty] | Copper +3/4" Al |  | 4/1 | No | 954 | 64
[6 empty] | Copper Alone |  | 4/1 | Yes | 309 | 26
| 9) | Exposure Reproducibility
|  |  | Result: | Exposures were not within the required 5% for the Fluoro mode using Copper
+ 1.5 in Aluminum phantom. The fluoro kV was 70 kV selected via the kV dial on
the image intensifier assembly. Table top exposures varried as follows:
7.51, 6.7, 6.92, 6.7, 6.38, 6.95, 6.91 R/min
|  |  | COMMENT: | Service Needed
| 10) | Mechanical
|  |  | NOTE: | This system appears to have only the 9 inch fluoro field of view
in spite of 6 and 4.5 inch FOV indicators on the equipment.
C. | RADIOGRAPHIC OVERHEAD X-RAY TUBE
| 1). | kVp Accuracy and Exposure Linearity
|  | Focal Spot: | lg |  | Field Size: | Max
Radcal Accu-kV diagnostic sensor placed on table .
| @ | 20.5 | inches from target, |  |  | 6 | mAs
| Radcal Chamber at: |  |  | 15.25 | inches from target
|  | Measured |  |  | kVp Accuracy |  | Linearity | [9 empty] | Exposure time: | 0.999 | sec
| Set | Dose | Time | HV | Error | Pass/ | mR | % | [8 empty] | Ln Values |  | Reg. |  |  | Regression Output:
| kVp | mR | ms | kVp | kV | Fail | regrs. | Error | [8 empty] | kVp | Dose | Dose |  |  |  | b | a
| 50 | 185 | 99.9 | 49 | -1 |  | 181.9213 | -1.692% | [8 empty] | 3.8918 | 5.2204 | 181.9213 |  |  | Coeff Value | 2.1598 | -3.2019
| 55 | 195 | 99.5 | 52 | -3 |  | 206.8338 | 5.721% | [8 empty] | 3.9512 | 5.273 | 206.8338 |  |  | Std Error | 0.029362 | 0.12873
| 60 | 260 | 99.7 | 58 | -2 |  | 261.8475 | 0.7056% | [8 empty] | 4.0604 | 5.5607 | 261.8475 |  |  | R squared | 0.99779 | 0.032187
| 65 | 291 | 100 | 61 | -4 | FAIL | 291.979 | 0.3353% | [8 empty] | 4.1109 | 5.6733 | 291.979
| 70 | 381 | 101 | 69 | -1 |  | 381.0146 | 0.00384% | [8 empty] | 4.2341 | 5.9428 | 381.0146 |  | y = a + (b * x)
| 75 | 421 | 101 | 72 | -3 |  | 417.6974 | -0.7907% | [8 empty] | 4.2767 | 6.0426 | 417.6974 |  | a= | -3.2019
| 80 | 525 | 101 | 79 | -1 |  | 510.375 | -2.866% | [8 empty] | 4.3694 | 6.2634 | 510.375 |  | b= | 2.1598
| 85 | 560 | 100 | 83 | -2 |  | 567.8307 | 1.379% | [8 empty] | 4.4188 | 6.3279 | 567.8307
| 90 | 677 | 101 | 88 | -2 |  | 644.2987 | -5.075% | [8 empty] | 4.4773 | 6.5177 | 644.2987
| 95 | 731 | 102 | 91 | -4 | FAIL | 692.6772 | -5.533% | [8 empty] | 4.5109 | 6.5944 | 692.6772
| 100 | 832 | 100 | 99 | -1 |  | 830.9322 | -0.1285% | [8 empty] | 4.5951 | 6.7238 | 830.9322
| 110 | 1030 | 101 | 110 | 0 |  | 1043.2582 | 1.271% | [8 empty] | 4.7005 | 6.9373 | 1043.2582
| 120 | 1210 | 102 | 120 | 0 |  | 1258.9454 | 3.888% | [8 empty] | 4.7875 | 7.0984 | 1258.9454
| 130 | 1440 | 103 | 129 | -1 |  | 1471.7781 | 2.159% | [8 empty] | 4.8598 | 7.2724 | 1471.7781
|  |  |  | Max.: | -4 | kV | Max: | 5.721%
Comments:
|  |  |  | kV Accuracy: | Service recommended for the 65 & 95 kV stations.
GE's specifications for kV accuracy need to be checked.
|  |  |  | Exposure Linearity: | Exposures linear to within 6% at all kV stations tested
Acceptable results
| 2). | mA and Exposure Linearity
|  | Technique: | 60 | kV | 50 | ms
| 21 | in. target to chamber distance
Measured
| SET | Dose | Time | HV | Regression
| mA | mR | ms | kVp | mR | Error | [15 empty] | Regression Output:
| 25 | 6.68 | 55 | 58.5 | 8.4324 | -20.78% | [13 empty] | 6.7459 |  |  | b | a
| 50 | 13.2 | 52 | 59 | 17.3399 | -23.88% | [13 empty] | 6.936 |  | Coeff Value | 0.3563 | -0.47518
| 100 | 33.7 | 54 | 59.1 | 35.155 | -4.139% | [13 empty] | 7.031 |  | Std Error | 0.015556 | 4.7879
| 200 | 70.9 | 53 | 59.4 | 70.7853 | 0.1621% | [13 empty] | 7.0785 |  | R squared | 0.99056 | 7.9673
| 300 | 111 | 52 | 59.4 | 106.4155 | 4.308% | [13 empty] | 7.0944
| 400 | 155 | 53 | 59.5 | 142.0457 | 9.12% | [13 empty] | 7.1023 | y = a + (b * x)
| 600 | 203 | 55 | 59.4 | 213.3062 | -4.832% | [13 empty] | 7.1102 | a= | -0.47518
[5 empty] | Max. Error: | -23.88% | [14 empty] | b= | 0.3563
Regression Avg.:
| 7.014 | mR/mAs @ 21 inches
Comments:
|  |  | Results: | Exposures should be linear to within +/-10%
At 25 and 50 mA the exposures are too low
|  |  | Action: | SERVICE REQUIRED
| 3). | Timer and Exposure Linearity
|  | Technique: | 75 | kV | 300 | mA
| 21 | in. target to chamber distance
Measured
| SET | Dose | Time | HV | Timer | Regression
| Time | mR | ms | kVp | Error | mR | Error
| 17 | 59.5 | 18 | 74.5 | 5.882% | 77.019 | -22.75% | [12 empty] | Regression Output:
| 25 | 89.8 | 22 | 74 | -12% | 104.3019 | -13.9% | [13 empty] | b | a
| 33 | 118 | 35 | 73.7 | 6.061% | 131.5847 | -10.32% | [12 empty] | Coeff Value | 3.4104 | 19.043
| 50 | 178 | 52 | 73.6 | 4% | 189.5608 | -6.099% | [12 empty] | Std Error | 0.0198 | 8.0727
| 67 | 233 | 65 | 74.2 | -2.985% | 247.5368 | -5.873% | [12 empty] | R squared | 0.99953 | 22.8117
| 100 | 352 | 102 | 74.17 | 2% | 360.0785 | -2.244%
| 125 | 448 | 128 | 74.14 | 2.4% | 445.3374 | 0.5979% | [11 empty] | y = a + (b * x)
| 150 | 535 | 155 | 74.11 | 3.333% | 530.5963 | 0.8299% | [11 empty] | a= | 19.043
| 200 | 704 | 205 | 74.08 | 2.5% | 701.1141 | 0.4116% | [11 empty] | b= | 3.4104
| 250 | 913 | 255 | 74.05 | 2% | 871.6319 | 4.746%
| 300 | 1060 | 306 | 74.02 | 2% | 1042.1497 | 1.713%
| 400 | 1400 | 408 | 73.99 | 2% | 1383.1853 | 1.216%
| 500 | 1760 | 508 | 73.96 | 1.6% | 1724.2208 | 2.075%
| 600 | 2080 | 610 | 73.93 | 1.667% | 2065.2564 | 0.7139%
| 800 | 2700 | 815 | 73.9 | 1.875% | 2747.3276 | -1.723%
| 1000 | 3420 | 1008 | 73.87 | 0.8% | 3429.3987 | -0.2741%
Comments:
|  |  | Results: | Exposures should be linear to within +/-10%
At .0167 and .025 seconds the exposures are too low
|  |  | Action: | SERVICE REQUIRED
| 4). | Exposure Reproducibility
|  |  | TECHNIQUES USED = | 55, 65, 75, 85, 95 kV, 200 Ma, 0.1s
75 kV, 300 Ma, 4/5 s
|  |  | RESULTS= | Reproducibility was within the required +/- 5%.
Exposures varied by 8%
|  |  | COMMENT: | Acceptable results
| 5). | Half Value Layer
Technique
| kV | mA | ms | added filtration |  |  | 74.5 kV (75 kV setting) 200 mA, 2.0 mmAl Added
| 75 | 200 | 100 | 2 | mmAl |  | Calculated HVL @75.0 kV setting is 2.5 mmAl
| 15.25 | in. target to chamber distance
Measured
| mm | Dose | Time | HV | Calculated | [11 empty] | HVL (2 point Log interpolation)
| Al | mR | ms | kVp | mR/s | Atten | [12 empty] | (a)Just > | (b)Just < |  | 0.5 line data
| 0 | 218 | 100.1 | 74.5 | 2177.8222 | 1 | [9 empty] | 1 |  | Target Value | 0.5 |  |  | 0 | 0.5
| 1 | 158 | 99.8 | 74.8 | 1583.1663 | 0.72695 | [9 empty] | 0.72695 |  | Row | 3 | 4 |  | 5 | 0.5
| 2 | 122 | 99.9 | 75 | 1221.2212 | 0.56075 | [9 empty] | 0.56075 |  | Exposure Rate | 1221.2212 | 970
| 3 | 97 | 100 | 74.1 | 970 | 0.4454 | [9 empty] | 0.4454 |  | Thickness (mm) | 2 | 3
| 4 | 78.1 | 99.9 | 84.9 | 781.7818 | 0.35897 | [9 empty] | 0.35897
| 5 | 65.1 | 99.9 | 75 | 651.6517 | 0.29922 | [9 empty] | 0.29922 |  | Calc. HVL | 2.4979 | mmAl
|  | Calculated HVL: | 2.4979 | mmAl
Comments:
|  |  | Results: | Acceptable
|  |  | Action: | None Required
| 6). | Phototimer
Technique
| kV | mA | ms | cassette |  |  | 72 | inch SFD |  | 67.5 | inch SCD, Lead Apron
| 121 | 200 | 500 | 14 x 17 | in. |  |  | Results: | Chest Stand + Lead Apron
|  | Center Cell |  |  |  |  | Right Cell |  |  |  |  | Left Cell
|  | Measured |  |  |  |  | Measured |  |  |  |  | Measured
|  | Dose | Time | HV | Dose |  | Dose | Time | HV | Dose |  | Dose | Time | HV | Dose
|  | mR | ms | kVp | Ratio |  | mR | ms | kVp | Ratio |  | mR | ms | kVp | Ratio
| NORM | 44 | 77 | 121.1 | N/A |  | 31 | 52 | 121 | N/A |  | 48 | 85 | 120.8 | N/A
| 1.25 | 52 | 94 | 120.8 | 1.1818 |  | 39 | 68 | 120.8667 | 1.2581 |  | 56 | 101 | 121 | 1.1667
| 1.5 | 65 | 118 | 121 | 1.4773 |  | 47 | 77 | 120.8167 | 1.5161 |  | 72 | 134 | 120.8667 | 1.5
| 2 | 73 | 134 | 120.8667 | 1.6591 |  | 56 | 101 | 120.7667 | 1.8065 |  | 85 | 159 | 121 | 1.7708
| 0.75 | 30 | 51 | 120.8167 | 0.68182 |  | 22 | 35 | 120.7167 | 0.70968 |  | 39 | 68 | 120.8667 | 0.8125
| 0.5 | 22 | 35 | 120.7667 | 0.5 |  | 17 | 27 | 120.8667 | 0.54839 |  | 26 | 43 | 120.8167 | 0.54167
| 0.25 | 13 | 18 | 120.7167 | 0.29545 |  | 9 | 10 | 120.8167 | 0.29032 |  | 13 | 18 | 120.7667 | 0.27083
|  |  | Results: | On both normal and 3/4 density settings, Exposures are not balanced
for the three cells. There should be better agreement in % variation
among the three cells.
|  |  | Action: | SERVICE RECOMMENDED
| 7). | Light Field/X-ray Field Alignment
|  | TABLE: | At 37 inch SFD - Alignment was not within 2% of SFD. The x-ray field is 1/2 inch smaller
and 1/4 inch larger than the light field (cross table), and shifted ~ 1/4 inch longitudinally.
|  | CHEST STAND: | At 72 inch SFD, the alignment is poor: 3/4 inch of the x-ray field is cut off (or shifted)
on the left side (facing the chest unit) and ~ 3/4 inch too large on the right side.
|  | Comments: | SERVICE REQUIRED
| 8). | Centering of X-ray field to Cassette in Bucky Tray
|  | TABLE: | At 40 inch SFD, Center of x-ray field and cassette are within 2% of SFD.
|  | Comments: | Acceptable Results
|  | WALL: | Center of x-ray field and cassette were not within 2% of SFD
|  | Comments: | SERVICE REQUIRED
| 9). | Automatic Collimation
SERVICE REQUIRED, this system was not operational.
| 10). | Mechanical
| 1 | Ready light on control panel did not light
| 2 | Chest stand cassette size indicator light did not light.
| 3 | When using low technique (~ 50 - 55 kV, 5 mAs) the RESET indicator came on.
| 4 | Several times, during testing no exposure occurred when the exposure switch was engaged.
| 11). | In air exposure from clinical protocols
Measured
|  |  |  | Generator Settings |  |  | Dose | Time | HV
|  | STUDY |  | KV | mA | ms | mR | ms | kVp
|  | ABDOMEN |  | 75 | 200 | 300 | 704 | 305 | 74.8
|  |  |  | 105 | 200 | 50 | 242 | 48 | 105.1
|  |  |  | 68 | 300 | 100 | 278 | 97 | 68.7
CHEST
|  |  | a) Erect | 75 | 300 | 25 | 89 | 22 | 74.3
|  |  |  | 85 | 400 | 100 | 667 | 98 | 84.2
|  |  |  | 88 | 400 | 100 | 712 | 7 | 87.2
|  |  | b) Supine | 63 | 300 | 25 | 64 | 22 | 62.8
|  |  |  | 65 | 300 | 17 | 42 | 13 | 64.9
|  |  |  | 76 | 300 | 33 | 124 | 31 | 75.8
|  |  | LAT | 95 | 600 | 25 | 3050 | 256 | 93.7
|  |  |  | 110 | 300 | 33 | 278 | 32 | 109.6
c) PHOTOTIMED - on 3/4 Density
|  |  |  | 80 | 300 | 100 | 433 | 97 | 79.7
|  |  |  | 100 | 200 | 500 | 2240 | 506 | 100.1
HIP & PELVIS
|  |  |  | 69 | 100 | 200 | 183 | 203 | 69.2
|  |  |  | 70 | 300 | 200 | 635 | 197 | 69.9
|  |  |  | 74 | 200 | 250 | 608 | 255 | 74.2
|  |  |  | 90 | 300 | 250 | 1400 | 256 | 89.9
|  | L-SPINE |  | 75 | 200 | 250 | 585 | 255 | 75
|  |  |  | 80 | 200 | 250 | 719 | 255 | 80.1
|  |  |  | 85 | 200 | 800 | 2330 | 797 | 85
|  |  |  | 90 | 200 | 800 | 2810 | 797 | 89.8
OBSTRUCTION SERIES
|  |  |  | 78 | 300 | 150 | 604 | 147 | 77.3
|  |  |  | 80 | 300 | 400 | 1680 | 397 | 79.2
|  |  |  | 85 | 300 | 400 | 1770 | 397 | 84.4
CALIBRATION PERFORMED BY:
CERTIFIED RADIOLOGICAL PHYSICIST
Note on A3: NOTE: This worksheet requires Radcal 9000XL ver 2.1 or higher
Note on C17: 1. Place cursor in this cell and press "Control-Shift f" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make a fluoro exposure. 3. Wait for kV and Rate values to settle to stable numbers. 4. Press Escape to end.
Note on C159: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on C187: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on C209: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on C249: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on C271: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on G87: To Make a Fluoro Rate & kVp Measurement 1. Place cursor in this cell and press "Control-Shift-f" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make a fluoro exposure. 3. Wait for kV and Rate values to settle to stable numbers. 4. Press Escape to end. 5. Move cursor to next cell
Note on H30: To Make a Fluoro Rate & kVp Measurement 1. Place cursor in this cell and press "Control-Shift-f" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make a fluoro exposure. 3. Wait for kV and Rate values to settle to stable numbers. 4. Press Escape to end. 5. Move cursor to next cell
Note on H271: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on H325: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on I55: To Make a Fluoro Rate & kVp Measurement 1. Place cursor in this cell and press "Control-Shift-f" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make a fluoro exposure. 3. Wait for kV and Rate values to settle to stable numbers. 4. Press Escape to end. 5. Move cursor to next cell
Note on K121: To make a dose & time measurement (BE SURE THE kVp SENSOR IS IN THE BEAM): 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Press Escape to end.
Note on M271: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.

## Mobile
2026 ANNUAL CALIBRATION REPORT | [24 empty] | XLPRO ™ Version 4.06
RADIOLOGY DEPARTMENT | [24 empty] | Operation Settings Area
[7 empty] | DATE: | 2026-07-29 21:20
Units
A. | FACILITY |  |  | Almost anywhere USA | [6 empty] | Room No: | 223 |  | Other: | [10 empty] | Radiation: | R
|  | State Reg. No: |  | CA 324-916 |  |  |  | UNIT: | Radiographic Mobile | [15 empty] | /Time: | /min
[25 empty] | Scale: | milli
| EQUIPMENT |  |  | Manufacturer | [5 empty] | Model |  |  |  | Serial No. | [10 empty] | Pulse Rate: | 15
|  |  | CONSOLE: | General Electric |  |  |  |  | AMX-4 |  |  |  | 371294 | [11 empty] | Mode
|  |  | TUBE: | General Electric |  |  |  |  | GE MX75 |  |  |  | MX9184 | [11 empty] | Units Display: | Omit
|  |  | RATINGS: |  |  | Maximum tube kV: | 75 | kVp. Maximum tube mA: |  |  |  | 40 | mA. | Phase: | 1 | [9 empty] | Cell Update: | Incremental
[25 empty] | Auto Zero: | Disabled
[25 empty] | Capture Mode: | Manual
B. | kVp ACCURACY AND EXPOSURE LINEARITY | [23 empty] | Screen Tracking: | Yes
|  | Focal Spot: | lg |  | Field Size: | Max | [18 empty] | kVp | 9000
| Radcal Accu-kV diagnostic sensor placed on table . | [23 empty] | Config: | Dose+Time
| @ | 20.5 | inches from target, |  |  | 6 | mAs | [17 empty] | Scale: | DAccu-kV
| Radcal Chamber at: |  |  | 15 | inches from target | [19 empty] | Gain: | 32.77
| Radcal Chamber at: |  |  | 15.25 | inches from target | [19 empty] | Offset: | 0.00
|  | Measured |  |  | kV Accuracy |  | Linearity | [9 empty] | Exposure time: | 0.0999 | sec | [5 empty] | Comm
| Set | Dose | Time | HV | Error | Pass/ | mR | % | [7 empty] | Ln Values |  | Reg. |  |  | Regression Output: |  |  |  | Index: | 0
| kVp | mR | ms | kVp | kV | Fail | regrs. | Error | [7 empty] | kVp | Dose | Dose |  |  |  | b | a |  | Baud: | 9600
| 50 | 81 | 99.9 | 50.5 | 0.5 |  | 83.8508 | 3.4% | [7 empty] | 3.922 | 4.3944 | 83.8508 |  |  | Coeff Value | 2.1689 | -4.0772 |  | Monitor Configuration
| 55 | 100 | 99.5 | 55 | 0 |  | 100.9043 | 0.8962% | [7 empty] | 4.0073 | 4.6052 | 100.9043 |  |  | Std Error | 0.037879 | 0.16778 |  | Serial No: | a1801
| 60 | 125 | 99.7 | 61 | 1 |  | 126.3098 | 1.037% | [7 empty] | 4.1109 | 4.8283 | 126.3098 |  |  | R squared | 0.99635 | 0.039822 |  | ROM Ver: | 3.24
| 65 | 157 | 100 | 68 | 3 |  | 159.868 | 1.794% | [7 empty] | 4.2195 | 5.0562 | 159.868 | [6 empty] | Detector: | 90x6
| 70 | 181 | 101 | 73 | 3 |  | 186.4629 | 2.93% | [7 empty] | 4.2905 | 5.1985 | 186.4629 |  | y = a + (b * x) |  |  |  |  | Chamb Kq: | 1.00
| 75 | 218 | 101 | 77 | 2 |  | 209.3342 | -4.14% | [7 empty] | 4.3438 | 5.3845 | 209.3342 |  | a= | -4.0772 |  |  |  | Atm Corr: | Uncorrected
| 80 | 248 | 101 | 83 | 3 |  | 246.3301 | -0.6779% | [7 empty] | 4.4188 | 5.5134 | 246.3301 |  | b= | 2.1689 |  |  |  | Temp (°C): | 24.3
| 85 | 284 | 100 | 86.5 | 1.5 |  | 269.4155 | -5.413% | [7 empty] | 4.4601 | 5.649 | 269.4155 | [6 empty] | Press (kPa): | 99.9
| 90 | 325 | 101 | 91 | 1 |  | 300.7407 | -8.067% | [7 empty] | 4.5109 | 5.7838 | 300.7407
| 95 | 330 | 102 | 95 | 0 |  | 330.1499 | 0.04541% | [7 empty] | 4.5539 | 5.7991 | 330.1499
| 100 | 404 | 100 | 103 | 3 |  | 393.4301 | -2.687% | [7 empty] | 4.6347 | 6.0014 | 393.4301
| 110 | 481 | 101 | 113 | 3 |  | 480.9999 | -2.893e-05% | [7 empty] | 4.7274 | 6.1759 | 480.9999
| 120 | 572 | 102 | 125 | 5 | FAIL | 598.7 | 4.46% | [7 empty] | 4.8283 | 6.3491 | 598.7
| 124 | 606 | 103 | 129 | 5 | FAIL | 641.0303 | 5.465% | [7 empty] | 4.8598 | 6.4069 | 641.0303
|  |  |  | Max.: | 5 | kV | Max: | -8.067%
|  |  |  | Limit ± | 3 | kV
Comments:
|  |  |  | kV Accuracy: | Results acceptable below 120kV.
SERVICE REQUIRED FOR kV STATIONS ABOVE 110kV
|  |  |  | Exposure Linearity: | Exposures linear to within 6% at all kV stations tested
Very Good Results, which are better than the other Capmobile
C. | mAs AND EXPOSURE LINEARITY
|  | Technique: | 70 | kV | 50 | ms
| 20 | in. target to chamber distance
Measured
| SET | Data | Time | HV | Regression |  | mR/
| mAs | mR | ms | kVp | mR | Error | mAs | [13 empty] | Regression Output:
| 1 | 33.4 | 55 | 70.3 | 52.6483 | -36.56% | 52.6483 | [14 empty] | b | a
| 2 | 34.6 | 52 | 70.2 | 63.9551 | -45.9% | 31.9776 | [13 empty] | Coeff Value | 11.3069 | 41.3414
| 3 | 51.2 | 54 | 70.1 | 75.262 | -31.97% | 25.0873 | [13 empty] | Std Error | 0.7334 | 10.9405
| 4 | 65.7 | 53 | 70.2 | 86.5688 | -24.11% | 21.6422 | [13 empty] | R squared | 0.94814 | 27.1034
| 5 | 82 | 52 | 70.4 | 97.8757 | -16.22% | 19.5751
| 6 | 112 | 53 | 79.2 | 109.1825 | 2.581% | 18.1971 | [12 empty] | y = a + (b * x)
| 7 | 128 | 55 | 70 | 120.4894 | 6.233% | 17.2128 | [12 empty] | a= | 41.3414
| 8 | 157 | 55 | 70.3 | 131.7962 | 19.12% | 16.4745 | [12 empty] | b= | 11.3069
| 10 | 170 | 55 | 70 | 154.4099 | 10.1% | 15.441
| 12 | 202 | 55 | 70.1 | 177.0237 | 14.11% | 14.752
| 14 | 228 | 55 | 70 | 199.6374 | 14.21% | 14.2598
| 16 | 260 | 55 | 69.9 | 222.2511 | 16.98% | 13.8907
| 22 | 311 | 55 | 70 | 290.0922 | 7.207% | 13.186
| 27 | 346 | 55 | 69.9 | 346.6265 | -0.1807% | 12.838
| 35 | 384 | 55 | 70 | 437.0813 | -12.14% | 12.488
[5 empty] | Max. Error: | -45.9%
Regression Avg.:
| 19.978 | mR/mAs @ 20 inches
Comments:
|  |  | Results: | These results are not as good as a the other Capmobile.
Even though this is a capacitor discharge unit.
Exposures should be linear to within +/- 10% from 1 to 10 or 14 mAs
|  |  | Action: | SERVICE REQUIRED
D. | EXPOSURE REPRODUCIBILITY
|  |  | TECHNIQUES USED= | 55, 70,85,100 kV, 6 mAs
75 kV, 8 mAs
70 kV, 3, 5, 10 mAs
|  |  | RESULTS= | Reproducibility was within the required +/- 5%.
Exposures varied by 8%
|  |  | Comment: | Acceptable results
E. | HALF VALUE LAYER
Technique
| kV | mAs | ms | added filtration |  |  | 75.7 kV (75 kV setting) 8 mA, 2.0 mmAl Added
| 75 | 8 | 100 | 2 | mmAl |  | Calculated HVL @75.0 kV setting is 2.7 mmAl
| 15.25 | in. target to chamber distance
HVL (2 point Log interpolation)
|  | Measured | [17 empty] | (a)Just > | (b)Just < |  | 0.5 line data
| mm | Dose | Time | HV | Calculated | [11 empty] | 1 |  | Target Value | 0.5 |  |  | 0 | 0.5
| Al | mR | ms | kVp | mR/s | Atten | [10 empty] | 0.74854 |  | Row | 3 | 4 |  | 5 | 0.5
| 0 | 171 | 100 | 75.7 | 171 | 1 | [10 empty] | 0.5848 |  | Exposure Rate | 100 | 81
| 1 | 128 | 100 | 75.9 | 128 | 0.74854 | [10 empty] | 0.47368 |  | Thickness (mm) | 2 | 3
| 2 | 100 | 100 | 76.3 | 100 | 0.5848 | [10 empty] | 0.39708
| 3 | 81 | 100 | 76.5 | 81 | 0.47368 | [10 empty] | 0.3345 |  | Calc. HVL | 2.7434 | mmAl
| 4 | 67.9 | 100 | 76.8 | 67.9 | 0.39708
| 5 | 57.2 | 100 | 76.9 | 57.2 | 0.3345
|  | Calculated HVL: | 2.7434 | mmAl
Comments:
|  |  | Results: | Acceptable
|  |  | Action: | None Required
F. | LIGHT FIELD/X-RAY FIELD ALIGNMENT
|  | Findings: | At 27 inch SFD - Alignment was not within 2% of SFD. The x-ray field is 1/2 inch smaller
than the light field, as well as shifted/
|  | Comments: | While this is in compliance, it is at the limit. Service recommended to prevent
it from being out of compliance.
G. | MECHANICAL
| 1 | Locks and light fields were working porperly.
| 2 | Audible beam on indicator needs to be louder.
H. | IN AIR EXPOSURE FROM CLINICAL PROTOCOLS
|  | Measurement Distance from Source: |  |  |  |  | 15 | inches
Measured
|  |  |  | Generator Settings |  |  | Dose | Time | HV
|  | STUDY |  | KV | mAs |  | mR | ms | kVp
|  | ABDOMEN |  | 80 | 22 |  | 756 | 274 | 78.6
|  | & C-Spine |  | 80 | 22 |  | 756 | 274 | 79
CHEST
|  |  |  | 70 | 5 |  | 142 | 14 | 69.8
|  |  |  | 75 | 8 |  | 172 | 980 | 74 | at 20 inches
|  | COLANG. |  | 90 | 35 |  | 1520 | 500 | 89.5
|  | HIP |  | 108 | 27 |  | 1720 | 461 | 107.5
|  |  |  | 85 | 27 |  | 999 | 360 | 84
|  | L-SPINE |  | 120 | 35 |  | 2590 | 618 | 119
|  |  |  | 120 | 50 |  | 3230 | 900 | 118.5
|  | PELVIS |  | 75 | 22 |  | 629 | 242 | 74.5
CALIBRATION PERFORMED BY:
CERTIFIED RADIOLOGICAL PHYSICIST
Note on A3: NOTE: This worksheet requires Radcal 9000XL ver 2.1 or higher
Note on C23: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on C55: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on C102: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on H135: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.

## Waveforms
Waveform Examples
|  |  |  | Single Phase 60 Hz Dental |  |  |  |  | Pulse Fluro |  |  |  |  | High frequency Radiographic | [5 empty] | 3-Phase 12-Pulse Radiographic | [5 empty] | AMX-4 Mobile
XLPRO ™ Version 4.06 |  | 0.255 | R | 220 | ms-9000 | 60.5 | kVp-9000 | 0.113 | R | 0.5 | ms-9000 | 80.1 | kVp-9000 | 0.152 | R | 220 | ms-9000 | 55.6 | kVp-9000 | 0.743 | R | 254 | ms-9000 | 113.5 | kVp-9000 | 0.1681 | R | 201 | ms-9000 | 83.5 | kVp-9000
Operation Settings Area |  | Peak= | 62.2 | kVp-FT1 | PPV= | 56.6 | kV-Prac. | Peak= | 82.4418 | kVp-FT2 | PPV= | 77.3458 | kV-Prac. | Peak= | 56.8 | kVp-FT1 | PPV= | 51.9 | kV-Prac. | Peak= | 113 | kVp-FT1 | PPV= | 109.8 | kV-Prac. | Peak= | 90.7 | kVp-FT1 | PPV= | 81.5 | kV-Prac.
|  | Avg= | 60.186 | kV | Ripple= | 4.0638 | kV p-p. | Avg= | 80.3854 | kV | Ripple= | 4.1127 | kV | Avg= | 54.9727 | kV | Ripple= | 3.7425 | kV p-p. | Avg= | 110.9141 | kV | Ripple= | 4.2152 | kV p-p. | Avg= | 79.7452 | kV | Ripple= | 21.9083 | kV p-p.
Units |  | Pk Freq= | 59.6637 | Hz | Rate= | 60.2 | Hz | Pk Freq= | 486.3952 | Hz | Rate= | 35.9 | Hz | Pk Freq= | 59.1816 | Hz | Rate= | 60.2 | Hz | Pk Freq= | 95.2938 | Hz | Rate= | 1.2 | Hz | Pk Freq= | 1987.7683 | Hz | Rate= | 1.5 | Hz
Radiation: | R | Width= | 237.4893 | ms-Mem. | Pulses= | 15 |  | Width= | 370.5975 | ms-Mem. | Pulses= | 14 |  | Width= | 237.4893 | ms-Mem. | Pulses= | 15 |  | Width= | 255.5912 | ms-Mem. | Pulses= | 1 |  | Width= | 201.0887 | ms-Mem. | Pulses= | 1
/Time: | /min | ms | kV |  | Hz | kV | LINES | ms | kV |  | Hz | kV | LINES |  | Picture 112 | Picture 113 | Picture 114 |  |  |  | Picture 106 | Picture 107 | Picture 108 |  |  |  | Picture 118 | Picture 119 | Picture 120
Scale: |  | 0 | 0 |  | 0 | 1.8442e-15 | 0 | 0 | 0 |  | 0 | 5.2604e-15 | 0
Pulse Rate: | 0.25 | 0.09838 | 0 |  | 1.5298 | 0.090122 | 238.3747 | 0.09838 | 0 |  | 0.0099264 | 0.44162 | 437.791
Mode |  | 0.19676 | 0 |  | 3.0597 | 0.17264 | 62.2179 | 0.19676 | 0 |  | 19.8529 | 0.80152 | 82.4418
Units Display: | Rescale | 0.29514 | 0 |  | 4.5895 | 0.24014 | 62.2179 | 0.29514 | 0 |  | 29.7793 | 1.0184 | 82.4418
Cell Update: | Incremental | 0.39352 | 0 |  | 6.1194 | 0.28583 | 60.186 | 0.39352 | 0 |  | 39.7057 | 1.0673 | 80.3854
Auto Zero: | Disabled | 0.4919 | 40.6524 |  | 7.6492 | 0.30432 | 60.186 | 0.4919 | 26.248 |  | 49.6322 | 0.96785 | 80.3854
Capture Mode: | Auto | 0.59028 | 41.3726 |  | 9.179 | 0.2927 |  | 0.59028 | 36.8913 |  | 59.5586 | 0.78765 | 69.1831
Screen Tracking: | Yes | 0.68866 | 42.1729 |  | 10.7089 | 0.25209 |  | 0.68866 | 45.2138 |  | 69.485 | 0.64129 | 69.1831
kVp | Full | 0.78704 | 42.573 |  | 12.2387 | 0.19041 | 2430 | 0.78704 | 48.2547 |  | 79.4115 | 0.6342 | 4463
Config: | Dose+Time+kV | 0.88542 | 42.8131 |  | 13.7685 | 0.13224 |  | 0.88542 | 46.1741 |  | 89.3379 | 0.72739
Scale: | DAccu-kV | 0.9838 | 42.8531 |  | 15.2984 | 0.13541 |  | 0.9838 | 45.4539 |  | 99.2643 | 0.81262
Gain: | 32.77 | 1.0822 | 43.1332 |  | 16.8282 | 0.20317 |  | 1.0822 | 45.6539 |  | 109.1908 | 0.83764
Offset: | 0.00 | 1.1806 | 43.6133 |  | 18.3581 | 0.279 |  | 1.1806 | 45.6539 |  | 119.1172 | 0.79574
Comm |  | 1.2789 | 44.0534 |  | 19.8879 | 0.33439 |  | 1.2789 | 45.6539 |  | 129.0436 | 0.69574
Index: | 0 | 1.3773 | 44.0134 |  | 21.4177 | 0.3563 |  | 1.3773 | 0 |  | 138.9701 | 0.54718
Baud: | 9600 | 1.4757 | 43.3732 |  | 22.9476 | 0.33957 |  | 1.4757 | 0 |  | 148.8965 | 0.36466
Monitor Configuration |  | 1.5741 | 43.0531 |  | 24.4774 | 0.28571 |  | 1.5741 | 0 |  | 158.8229 | 0.2069
Serial No: | A 1 | 1.6725 | 43.3732 |  | 26.0073 | 0.20255 |  | 1.6725 | 0 |  | 168.7494 | 0.24829
ROM Ver: | 3.25 | 1.7708 | 44.3335 |  | 27.5371 | 0.1059 |  | 1.7708 | 0 |  | 178.6758 | 0.39797
Detector: | 90x6 | 1.8692 | 44.7737 |  | 29.0669 | 0.061212 |  | 1.8692 | 0 |  | 188.6022 | 0.49186
Chamb Kq: | 1.00 | 1.9676 | 44.3735 |  | 30.5968 | 0.14009 |  | 1.9676 | 0 |  | 198.5287 | 0.47793
Atm Corr: | Automatic | 2.066 | 44.0935 |  | 32.1266 | 0.22788 |  | 2.066 | 0 |  | 208.4551 | 0.34757
Temp (°C): | 2 | 2.1644 | 44.1735 |  | 33.6564 | 0.30568 |  | 2.1644 | 0 |  | 218.3815 | 0.13942
Press (kPa): | 100.5 | 2.2627 | 43.8134 |  | 35.1863 | 0.37863 |  | 2.2627 | 0 |  | 228.308 | 0.19599
|  | 2.3611 | 42.8531 |  | 36.7161 | 0.45529 |  | 2.3611 | 0 |  | 238.2344 | 0.44409
|  | 2.4595 | 42.613 |  | 38.246 | 0.53738 |  | 2.4595 | 0 |  | 248.1608 | 0.62556
|  | 2.5579 | 43.5733 |  | 39.7758 | 0.61442 |  | 2.5579 | 0 |  | 258.0873 | 0.69024
|  | 2.6562 | 43.8534 |  | 41.3056 | 0.6661 |  | 2.6562 | 0 |  | 268.0137 | 0.62627
|  | 2.7546 | 42.653 |  | 42.8355 | 0.66875 |  | 2.7546 | 0 |  | 277.9401 | 0.46062
|  | 2.853 | 40.8925 |  | 44.3653 | 0.60249 |  | 2.853 | 0 |  | 287.8666 | 0.27536
|  | 2.9514 | 40.8125 |  | 45.8952 | 0.46078 |  | 2.9514 | 0 |  | 297.793 | 0.26534
|  | 3.0498 | 25.4478 |  | 47.425 | 0.28869 |  | 3.0498 | 0 |  | 307.7194 | 0.41238
|  | 3.1481 | 25.4478 |  | 48.9548 | 0.35035 |  | 3.1481 | 0 |  | 317.6459 | 0.52659
|  | 3.2465 | 25.4478 |  | 50.4847 | 0.70171 |  | 3.2465 | 0 |  | 327.5723 | 0.55931
|  | 3.3449 | 0 |  | 52.0145 | 1.1408 |  | 3.3449 | 0 |  | 337.4987 | 0.51627
|  | 3.4433 | 0 |  | 53.5443 | 1.5936 |  | 3.4433 | 0 |  | 347.4252 | 0.42179
|  | 3.5417 | 0 |  | 55.0742 | 2.0111 |  | 3.5417 | 0 |  | 357.3516 | 0.29861
|  | 3.64 | 0 |  | 56.604 | 2.3503 |  | 3.64 | 0 |  | 367.278 | 0.16393
|  | 3.7384 | 0 |  | 58.1339 | 2.5755 |  | 3.7384 | 0 |  | 377.2045 | 0.12248
|  | 3.8368 | 0 |  | 59.6637 | 2.6621 |  | 3.8368 | 0 |  | 387.1309 | 0.27619
|  | 3.9352 | 0 |  | 61.1935 | 2.599 |  | 3.9352 | 0 |  | 397.0573 | 0.44659
|  | 4.0336 | 0 |  | 62.7234 | 2.3905 |  | 4.0336 | 0 |  | 406.9838 | 0.55843
|  | 4.1319 | 0 |  | 64.2532 | 2.0557 |  | 4.1319 | 0 |  | 416.9102 | 0.56236
|  | 4.2303 | 0 |  | 65.7831 | 1.627 |  | 4.2303 | 0 |  | 426.8366 | 0.43359
|  | 4.3287 | 0 |  | 67.3129 | 1.146 |  | 4.3287 | 0 |  | 436.7631 | 0.20754
|  | 4.4271 | 0 |  | 68.8427 | 0.66073 |  | 4.4271 | 0 |  | 446.6895 | 0.28858
|  | 4.5255 | 0 |  | 70.3726 | 0.23747 |  | 4.5255 | 0 |  | 456.6159 | 0.65593
|  | 4.6238 | 0 |  | 71.9024 | 0.24115 |  | 4.6238 | 0 |  | 466.5424 | 0.9914
|  | 4.7222 | 0 |  | 73.4322 | 0.49465 |  | 4.7222 | 0 |  | 476.4688 | 1.2011
|  | 4.8206 | 0 |  | 74.9621 | 0.65224 |  | 4.8206 | 0 |  | 486.3952 | 1.2318
|  | 4.919 | 0 |  | 76.4919 | 0.69645 |  | 4.919 | 0 |  | 496.3217 | 1.0709
|  | 5.0174 | 0 |  | 78.0218 | 0.63789 |  | 5.0174 | 0 |  | 506.2481 | 0.74932
|  | 5.1157 | 0 |  | 79.5516 | 0.50001 |  | 5.1157 | 0 |  | 516.1745 | 0.33616
|  | 5.2141 | 0 |  | 81.0814 | 0.31434 |  | 5.2141 | 0 |  | 526.101 | 0.13429
|  | 5.3125 | 0 |  | 82.6113 | 0.11896 |  | 5.3125 | 0 |  | 536.0274 | 0.46975
|  | 5.4109 | 0 |  | 84.1411 | 0.089891 |  | 5.4109 | 0 |  | 545.9538 | 0.68945
|  | 5.5093 | 0 |  | 85.671 | 0.22277 |  | 5.5093 | 0 |  | 555.8803 | 0.74838
|  | 5.6076 | 0 |  | 87.2008 | 0.30317 |  | 5.6076 | 0 |  | 565.8067 | 0.65857
|  | 5.706 | 0 |  | 88.7306 | 0.31934 |  | 5.706 | 0 |  | 575.7331 | 0.47167
|  | 5.8044 | 0 |  | 90.2605 | 0.27676 |  | 5.8044 | 0 |  | 585.6596 | 0.2839
|  | 5.9028 | 0 |  | 91.7903 | 0.19437 |  | 5.9028 | 0 |  | 595.586 | 0.25503
|  | 6.0012 | 0 |  | 93.3202 | 0.11911 |  | 6.0012 | 0 |  | 605.5124 | 0.34953
|  | 6.0995 | 0 |  | 94.85 | 0.14635 |  | 6.0995 | 0 |  | 615.4389 | 0.40807
|  | 6.1979 | 0 |  | 96.3798 | 0.23055 |  | 6.1979 | 0 |  | 625.3653 | 0.3978
|  | 6.2963 | 0 |  | 97.9097 | 0.2951 |  | 6.2963 | 0 |  | 635.2917 | 0.33804
|  | 6.3947 | 0 |  | 99.4395 | 0.31683 |  | 6.3947 | 0 |  | 645.2182 | 0.26086
|  | 6.4931 | 0 |  | 100.9693 | 0.29032 |  | 6.4931 | 0 |  | 655.1446 | 0.18801
|  | 6.5914 | 0 |  | 102.4992 | 0.22153 |  | 6.5914 | 0 |  | 665.071 | 0.12717
|  | 6.6898 | 0 |  | 104.029 | 0.13168 |  | 6.6898 | 0 |  | 674.9975 | 0.11071
|  | 6.7882 | 0 |  | 105.5589 | 0.10409 |  | 6.7882 | 0 |  | 684.9239 | 0.16631
|  | 6.8866 | 0 |  | 107.0887 | 0.1988 |  | 6.8866 | 0 |  | 694.8503 | 0.23005
|  | 6.985 | 0 |  | 108.6185 | 0.31507 |  | 6.985 | 0 |  | 704.7768 | 0.24979
|  | 7.0833 | 0 |  | 110.1484 | 0.42254 |  | 7.0833 | 0 |  | 714.7032 | 0.19921
|  | 7.1817 | 0 |  | 111.6782 | 0.51801 |  | 7.1817 | 0 |  | 724.6296 | 0.075698
|  | 7.2801 | 0 |  | 113.2081 | 0.60655 |  | 7.2801 | 0 |  | 734.5561 | 0.10153
|  | 7.3785 | 0 |  | 114.7379 | 0.6941 |  | 7.3785 | 0 |  | 744.4825 | 0.29023
|  | 7.4769 | 0 |  | 116.2677 | 0.78099 |  | 7.4769 | 0 |  | 754.4089 | 0.44473
|  | 7.5752 | 0 |  | 117.7976 | 0.8592 |  | 7.5752 | 0 |  | 764.3354 | 0.52342
|  | 7.6736 | 0 |  | 119.3274 | 0.91465 |  | 7.6736 | 0 |  | 774.2618 | 0.50274
|  | 7.772 | 0 |  | 120.8572 | 0.93205 |  | 7.772 | 0 |  | 784.1882 | 0.38405
|  | 7.8704 | 0 |  | 122.3871 | 0.89961 |  | 7.8704 | 0 |  | 794.1147 | 0.19349
|  | 7.9688 | 0 |  | 123.9169 | 0.8123 |  | 7.9688 | 0 |  | 804.0411 | 0.025485
|  | 8.0671 | 0 |  | 125.4468 | 0.67342 |  | 8.0671 | 0 |  | 813.9675 | 0.22372
|  | 8.1655 | 0 |  | 126.9766 | 0.49501 |  | 8.1655 | 0 |  | 823.894 | 0.36088
|  | 8.2639 | 0 |  | 128.5064 | 0.29868 |  | 8.2639 | 0 |  | 833.8204 | 0.41494
|  | 8.3623 | 0 |  | 130.0363 | 0.13471 |  | 8.3623 | 0 |  | 843.7468 | 0.38683
|  | 8.4606 | 0 |  | 131.5661 | 0.16999 |  | 8.4606 | 0 |  | 853.6733 | 0.2968
|  | 8.559 | 0 |  | 133.096 | 0.29544 |  | 8.559 | 0 |  | 863.5997 | 0.17395
|  | 8.6574 | 0 |  | 134.6258 | 0.38219 |  | 8.6574 | 0 |  | 873.5261 | 0.043019
|  | 8.7558 | 0 |  | 136.1556 | 0.40975 |  | 8.7558 | 0 |  | 883.4526 | 0.086904
|  | 8.8542 | 0 |  | 137.6855 | 0.37625 |  | 8.8542 | 0 |  | 893.379 | 0.21105
|  | 8.9525 | 0 |  | 139.2153 | 0.28967 |  | 8.9525 | 0 |  | 903.3054 | 0.32194
|  | 9.0509 | 0 |  | 140.7451 | 0.16587 |  | 9.0509 | 0 |  | 913.2319 | 0.39518
|  | 9.1493 | 0 |  | 142.275 | 0.041407 |  | 9.1493 | 0 |  | 923.1583 | 0.39743
|  | 9.2477 | 0 |  | 143.8048 | 0.13183 |  | 9.2477 | 0 |  | 933.0847 | 0.30141
|  | 9.3461 | 0 |  | 145.3347 | 0.25116 |  | 9.3461 | 0 |  | 943.0112 | 0.10149
|  | 9.4444 | 0 |  | 146.8645 | 0.33709 |  | 9.4444 | 0 |  | 952.9376 | 0.18986
|  | 9.5428 | 0 |  | 148.3943 | 0.38096 |  | 9.5428 | 0 |  | 962.864 | 0.51327
|  | 9.6412 | 0 |  | 149.9242 | 0.38431 |  | 9.6412 | 0 |  | 972.7905 | 0.81073
|  | 9.7396 | 0 |  | 151.454 | 0.35917 |  | 9.7396 | 0 |  | 982.7169 | 1.016
|  | 9.838 | 0 |  | 152.9839 | 0.32782 |  | 9.838 | 0 |  | 992.6433 | 1.0793
|  | 9.9363 | 0 |  | 154.5137 | 0.31612 |  | 9.9363 | 0 |  | 1002.5698 | 0.98047
|  | 10.0347 | 0 |  | 156.0435 | 0.33371 |  | 10.0347 | 0 |  | 1012.4962 | 0.73594
|  | 10.1331 | 0 |  | 157.5734 | 0.36397 |  | 10.1331 | 0 |  | 1022.4226 | 0.3957
|  | 10.2315 | 0 |  | 159.1032 | 0.38233 |  | 10.2315 | 0 |  | 1032.3491 | 0.036253
|  | 10.3299 | 0 |  | 160.633 | 0.3722 |  | 10.3299 | 0 |  | 1042.2755 | 0.28115
|  | 10.4282 | 0 |  | 162.1629 | 0.32794 |  | 10.4282 | 0 |  | 1052.2019 | 0.47683
|  | 10.5266 | 0 |  | 163.6927 | 0.25438 |  | 10.5266 | 0 |  | 1062.1284 | 0.52393
|  | 10.625 | 0 |  | 165.2226 | 0.16935 |  | 10.625 | 0 |  | 1072.0548 | 0.42746
|  | 10.7234 | 0 |  | 166.7524 | 0.12264 |  | 10.7234 | 0 |  | 1081.9812 | 0.24204
|  | 10.8218 | 0 |  | 168.2822 | 0.16803 |  | 10.8218 | 0 |  | 1091.9077 | 0.19261
|  | 10.9201 | 0 |  | 169.8121 | 0.24512 |  | 10.9201 | 0 |  | 1101.8341 | 0.39617
|  | 11.0185 | 0 |  | 171.3419 | 0.3088 |  | 11.0185 | 0 |  | 1111.7605 | 0.58018
|  | 11.1169 | 0 |  | 172.8718 | 0.34734 |  | 11.1169 | 0 |  | 1121.687 | 0.67577
|  | 11.2153 | 0 |  | 174.4016 | 0.36215 |  | 11.2153 | 0 |  | 1131.6134 | 0.67718
|  | 11.3137 | 0 |  | 175.9314 | 0.36284 |  | 11.3137 | 0 |  | 1141.5398 | 0.60956
|  | 11.412 | 0 |  | 177.4613 | 0.36246 |  | 11.412 | 0 |  | 1151.4663 | 0.51157
|  | 11.5104 | 0 |  | 178.9911 | 0.36957 |  | 11.5104 | 0 |  | 1161.3927 | 0.41159
|  | 11.6088 | 0 |  | 180.5209 | 0.38152 |  | 11.6088 | 0 |  | 1171.3191 | 0.31007
|  | 11.7072 | 0 |  | 182.0508 | 0.38645 |  | 11.7072 | 0 |  | 1181.2456 | 0.19104
|  | 11.8056 | 0 |  | 183.5806 | 0.37103 |  | 11.8056 | 0 |  | 1191.172 | 0.06357
|  | 11.9039 | 0 |  | 185.1105 | 0.32643 |  | 11.9039 | 0 |  | 1201.0984 | 0.13139
|  | 12.0023 | 0 |  | 186.6403 | 0.25038 |  | 12.0023 | 0 |  | 1211.0249 | 0.25736
|  | 12.1007 | 0 |  | 188.1701 | 0.14719 |  | 12.1007 | 0 |  | 1220.9513 | 0.32696
|  | 12.1991 | 0 |  | 189.7 | 0.027369 |  | 12.1991 | 0 |  | 1230.8777 | 0.31353
|  | 12.2975 | 0 |  | 191.2298 | 0.10037 |  | 12.2975 | 0 |  | 1240.8042 | 0.21431
|  | 12.3958 | 0 |  | 192.7597 | 0.2169 |  | 12.3958 | 0 |  | 1250.7306 | 0.051029
|  | 12.4942 | 0 |  | 194.2895 | 0.3116 |  | 12.4942 | 0 |  | 1260.657 | 0.14223
|  | 12.5926 | 0 |  | 195.8193 | 0.37464 |  | 12.5926 | 0 |  | 1270.5835 | 0.31347
|  | 12.691 | 0 |  | 197.3492 | 0.40085 |  | 12.691 | 0 |  | 1280.5099 | 0.42448
|  | 12.7894 | 0 |  | 198.879 | 0.39032 |  | 12.7894 | 0 |  | 1290.4363 | 0.44994
|  | 12.8877 | 0 |  | 200.4088 | 0.34822 |  | 12.8877 | 0 |  | 1300.3628 | 0.38676
|  | 12.9861 | 0 |  | 201.9387 | 0.28379 |  | 12.9861 | 0 |  | 1310.2892 | 0.25518
|  | 13.0845 | 0 |  | 203.4685 | 0.2087 |  | 13.0845 | 0 |  | 1320.2156 | 0.11125
|  | 13.1829 | 0 |  | 204.9984 | 0.13481 |  | 13.1829 | 0 |  | 1330.1421 | 0.15652
|  | 13.2812 | 0 |  | 206.5282 | 0.071584 |  | 13.2812 | 0 |  | 1340.0685 | 0.29443
|  | 13.3796 | 0 |  | 208.058 | 0.024252 |  | 13.3796 | 0 |  | 1349.9949 | 0.39747
|  | 13.478 | 0 |  | 209.5879 | 0.025124 |  | 13.478 | 0 |  | 1359.9214 | 0.45243
|  | 13.5764 | 0 |  | 211.1177 | 0.066115 |  | 13.5764 | 0 |  | 1369.8478 | 0.46501
|  | 13.6748 | 0 |  | 212.6476 | 0.11777 |  | 13.6748 | 0 |  | 1379.7742 | 0.44604
|  | 13.7731 | 0 |  | 214.1774 | 0.17636 |  | 13.7731 | 0 |  | 1389.7007 | 0.40451
|  | 13.8715 | 0 |  | 215.7072 | 0.23315 |  | 13.8715 | 0 |  | 1399.6271 | 0.34634
|  | 13.9699 | 0 |  | 217.2371 | 0.27724 |  | 13.9699 | 0 |  | 1409.5535 | 0.27635
|  | 14.0683 | 0 |  | 218.7669 | 0.29843 |  | 14.0683 | 0 |  | 1419.48 | 0.1976
|  | 14.1667 | 0 |  | 220.2967 | 0.28922 |  | 14.1667 | 0 |  | 1429.4064 | 0.10838
|  | 14.265 | 0 |  | 221.8266 | 0.24617 |  | 14.265 | 0 |  | 1439.3328 | 0.049935
|  | 14.3634 | 0 |  | 223.3564 | 0.17057 |  | 14.3634 | 0 |  | 1449.2592 | 0.18049
|  | 14.4618 | 0 |  | 224.8863 | 0.069744 |  | 14.4618 | 0 |  | 1459.1857 | 0.36212
|  | 14.5602 | 0 |  | 226.4161 | 0.062111 |  | 14.5602 | 0 |  | 1469.1121 | 0.55386
|  | 14.6586 | 0 |  | 227.9459 | 0.18744 |  | 14.6586 | 0 |  | 1479.0385 | 0.71828
|  | 14.7569 | 0 |  | 229.4758 | 0.3105 |  | 14.7569 | 0 |  | 1488.965 | 0.81578
|  | 14.8553 | 0 |  | 231.0056 | 0.41761 |  | 14.8553 | 0 |  | 1498.8914 | 0.81608
|  | 14.9537 | 0 |  | 232.5355 | 0.49963 |  | 14.9537 | 0 |  | 1508.8178 | 0.70854
|  | 15.0521 | 0 |  | 234.0653 | 0.55092 |  | 15.0521 | 0 |  | 1518.7443 | 0.50777
|  | 15.1505 | 0 |  | 235.5951 | 0.56965 |  | 15.1505 | 0 |  | 1528.6707 | 0.25307
|  | 15.2488 | 0 |  | 237.125 | 0.55753 |  | 15.2488 | 0 |  | 1538.5971 | 0.04596
|  | 15.3472 | 0 |  | 238.6548 | 0.51916 |  | 15.3472 | 0 |  | 1548.5236 | 0.21681
|  | 15.4456 | 0 |  | 240.1846 | 0.46089 |  | 15.4456 | 0 |  | 1558.45 | 0.31748
|  | 15.544 | 0 |  | 241.7145 | 0.38969 |  | 15.544 | 0 |  | 1568.3764 | 0.29684
|  | 15.6424 | 0 |  | 243.2443 | 0.31216 |  | 15.6424 | 0 |  | 1578.3029 | 0.16172
|  | 15.7407 | 0 |  | 244.7742 | 0.2342 |  | 15.7407 | 0 |  | 1588.2293 | 0.065734
|  | 15.8391 | 0 |  | 246.304 | 0.1619 |  | 15.8391 | 0 |  | 1598.1557 | 0.30595
|  | 15.9375 | 0 |  | 247.8338 | 0.10546 |  | 15.9375 | 0 |  | 1608.0822 | 0.52148
|  | 16.0359 | 0 |  | 249.3637 | 0.086501 |  | 16.0359 | 0 |  | 1618.0086 | 0.65996
|  | 16.1343 | 0 |  | 250.8935 | 0.1092 |  | 16.1343 | 0 |  | 1627.935 | 0.6959
|  | 16.2326 | 0 |  | 252.4234 | 0.14206 |  | 16.2326 | 0 |  | 1637.8615 | 0.6331
|  | 16.331 | 0 |  | 253.9532 | 0.16676 |  | 16.331 | 0 |  | 1647.7879 | 0.50417
|  | 16.4294 | 0 |  | 255.483 | 0.17688 |  | 16.4294 | 0 |  | 1657.7143 | 0.36409
|  | 16.5278 | 0 |  | 257.0129 | 0.17061 |  | 16.5278 | 0 |  | 1667.6408 | 0.27163
|  | 16.6262 | 0 |  | 258.5427 | 0.14897 |  | 16.6262 | 0 |  | 1677.5672 | 0.23412
|  | 16.7245 | 0 |  | 260.0726 | 0.11551 |  | 16.7245 | 0 |  | 1687.4936 | 0.19235
|  | 16.8229 | 0 |  | 261.6024 | 0.076573 |  | 16.8229 | 0 |  | 1697.4201 | 0.10838
|  | 16.9213 | 0 |  | 263.1322 | 0.04609 |  | 16.9213 | 0 |  | 1707.3465 | 0.027583
|  | 17.0197 | 0 |  | 264.6621 | 0.052329 |  | 17.0197 | 0 |  | 1717.2729 | 0.14117
|  | 17.1181 | 0 |  | 266.1919 | 0.082202 |  | 17.1181 | 0 |  | 1727.1994 | 0.23319
|  | 17.2164 | 45.934 |  | 267.7217 | 0.11223 |  | 17.2164 | 0 |  | 1737.1258 | 0.27402
|  | 17.3148 | 46.1341 |  | 269.2516 | 0.13788 |  | 17.3148 | 0 |  | 1747.0522 | 0.2779
|  | 17.4132 | 46.6542 |  | 270.7814 | 0.15969 |  | 17.4132 | 0 |  | 1756.9787 | 0.30652
|  | 17.5116 | 48.5748 |  | 272.3113 | 0.1788 |  | 17.5116 | 0 |  | 1766.9051 | 0.40671
|  | 17.61 | 49.4551 |  | 273.8411 | 0.195 |  | 17.61 | 0 |  | 1776.8315 | 0.54042
|  | 17.7083 | 49.055 |  | 275.3709 | 0.20647 |  | 17.7083 | 0 |  | 1786.758 | 0.64863
|  | 17.8067 | 49.3351 |  | 276.9008 | 0.21086 |  | 17.8067 | 0 |  | 1796.6844 | 0.6934
|  | 17.9051 | 49.5751 |  | 278.4306 | 0.20718 |  | 17.9051 | 0 |  | 1806.6108 | 0.66253
|  | 18.0035 | 49.6151 |  | 279.9605 | 0.19763 |  | 18.0035 | 0 |  | 1816.5373 | 0.56904
|  | 18.1019 | 49.6952 |  | 281.4903 | 0.1888 |  | 18.1019 | 0 |  | 1826.4637 | 0.45019
|  | 18.2002 | 49.3351 |  | 283.0201 | 0.19005 |  | 18.2002 | 0 |  | 1836.3901 | 0.36787
|  | 18.2986 | 48.8949 |  | 284.55 | 0.20676 |  | 18.2986 | 0 |  | 1846.3166 | 0.38014
|  | 18.397 | 48.5348 |  | 286.0798 | 0.23464 |  | 18.397 | 0 |  | 1856.243 | 0.467
|  | 18.4954 | 48.6548 |  | 287.6096 | 0.26354 |  | 18.4954 | 0 |  | 1866.1694 | 0.57143
|  | 18.5938 | 49.295 |  | 289.1395 | 0.28376 |  | 18.5938 | 0 |  | 1876.0959 | 0.65852
|  | 18.6921 | 49.015 |  | 290.6693 | 0.28868 |  | 18.6921 | 0 |  | 1886.0223 | 0.7063
|  | 18.7905 | 48.3748 |  | 292.1992 | 0.27527 |  | 18.7905 | 0 |  | 1895.9487 | 0.69704
|  | 18.8889 | 48.6548 |  | 293.729 | 0.24383 |  | 18.8889 | 0 |  | 1905.8752 | 0.61982
|  | 18.9873 | 48.4948 |  | 295.2588 | 0.19754 |  | 18.9873 | 0 |  | 1915.8016 | 0.47615
|  | 19.0856 | 47.5345 |  | 296.7887 | 0.14173 |  | 19.0856 | 0 |  | 1925.728 | 0.28302
|  | 19.184 | 47.2944 |  | 298.3185 | 0.083519 |  | 19.184 | 0 |  | 1935.6545 | 0.075714
|  | 19.2824 | 48.2547 |  | 299.8484 | 0.038778 |  | 19.2824 | 0 |  | 1945.5809 | 0.14263
|  | 19.3808 | 48.3347 |  | 301.3782 | 0.05631 |  | 19.3808 | 0 |  | 1955.5073 | 0.29909
|  | 19.4792 | 45.774 |  | 302.908 | 0.10189 |  | 19.4792 | 0 |  | 1965.4338 | 0.39908
|  | 19.5775 | 42.733 |  | 304.4379 | 0.14728 |  | 19.5775 | 0 |  | 1975.3602 | 0.44504
|  | 19.6759 | 41.8528 |  | 305.9677 | 0.19028 |  | 19.6759 | 0 |  | 1985.2866 | 0.45983
|  | 19.7743 | 42.9731 |  | 307.4975 | 0.23004 |  | 19.7743 | 0 |  | 1995.2131 | 0.47173
|  | 19.8727 | 42.533 |  | 309.0274 | 0.26397 |  | 19.8727 | 0 |  | 2005.1395 | 0.49083
|  | 19.9711 | 41.1726 |  | 310.5572 | 0.28776 |  | 19.9711 | 0 |  | 2015.0659 | 0.50111
|  | 20.0694 | 41.1726 |  | 312.0871 | 0.29658 |  | 20.0694 | 0 |  | 2024.9924 | 0.47827
|  | 20.1678 | 41.1726 |  | 313.6169 | 0.28673 |  | 20.1678 | 0 |  | 2034.9188 | 0.40754
|  | 20.2662 | 0 |  | 315.1467 | 0.25697 |  | 20.2662 | 0 |  | 2044.8452 | 0.28887
|  | 20.3646 | 0 |  | 316.6766 | 0.20971 |  | 20.3646 | 0 |  | 2054.7717 | 0.13825
|  | 20.463 | 0 |  | 318.2064 | 0.15272 |  | 20.463 | 0 |  | 2064.6981 | 0.09443
|  | 20.5613 | 0 |  | 319.7363 | 0.10492 |  | 20.5613 | 0 |  | 2074.6245 | 0.26263
|  | 20.6597 | 0 |  | 321.2661 | 0.10193 |  | 20.6597 | 0 |  | 2084.551 | 0.43274
|  | 20.7581 | 0 |  | 322.7959 | 0.13989 |  | 20.7581 | 0 |  | 2094.4774 | 0.56944
|  | 20.8565 | 0 |  | 324.3258 | 0.17974 |  | 20.8565 | 0 |  | 2104.4038 | 0.65171
|  | 20.9549 | 0 |  | 325.8556 | 0.20286 |  | 20.9549 | 0 |  | 2114.3303 | 0.66494
|  | 21.0532 | 0 |  | 327.3854 | 0.20305 |  | 21.0532 | 0 |  | 2124.2567 | 0.60529
|  | 21.1516 | 0 |  | 328.9153 | 0.18051 |  | 21.1516 | 0 |  | 2134.1831 | 0.48373
|  | 21.25 | 0 |  | 330.4451 | 0.14119 |  | 21.25 | 0 |  | 2144.1096 | 0.327
|  | 21.3484 | 0 |  | 331.975 | 0.09934 |  | 21.3484 | 0 |  | 2154.036 | 0.17631
|  | 21.4468 | 0 |  | 333.5048 | 0.08462 |  | 21.4468 | 0 |  | 2163.9624 | 0.093288
|  | 21.5451 | 0 |  | 335.0346 | 0.11053 |  | 21.5451 | 0 |  | 2173.8889 | 0.10032
|  | 21.6435 | 0 |  | 336.5645 | 0.1451 |  | 21.6435 | 0 |  | 2183.8153 | 0.11519
|  | 21.7419 | 0 |  | 338.0943 | 0.1666 |  | 21.7419 | 0 |  | 2193.7417 | 0.17525
|  | 21.8403 | 0 |  | 339.6242 | 0.16712 |  | 21.8403 | 0 |  | 2203.6682 | 0.30324
|  | 21.9387 | 0 |  | 341.154 | 0.14552 |  | 21.9387 | 0 |  | 2213.5946 | 0.44574
|  | 22.037 | 0 |  | 342.6838 | 0.10568 |  | 22.037 | 0 |  | 2223.521 | 0.54998
|  | 22.1354 | 0 |  | 344.2137 | 0.058081 |  | 22.1354 | 0 |  | 2233.4475 | 0.5823
|  | 22.2338 | 0 |  | 345.7435 | 0.042888 |  | 22.2338 | 0 |  | 2243.3739 | 0.543
|  | 22.3322 | 0 |  | 347.2733 | 0.083216 |  | 22.3322 | 0 |  | 2253.3003 | 0.4895
|  | 22.4306 | 0 |  | 348.8032 | 0.12365 |  | 22.4306 | 0 |  | 2263.2268 | 0.53099
|  | 22.5289 | 0 |  | 350.333 | 0.14847 |  | 22.5289 | 0 |  | 2273.1532 | 0.68938
|  | 22.6273 | 0 |  | 351.8629 | 0.15397 |  | 22.6273 | 0 |  | 2283.0796 | 0.8694
|  | 22.7257 | 0 |  | 353.3927 | 0.14197 |  | 22.7257 | 0 |  | 2293.0061 | 0.98706
|  | 22.8241 | 0 |  | 354.9225 | 0.12006 |  | 22.8241 | 0 |  | 2302.9325 | 1.0001
|  | 22.9225 | 0 |  | 356.4524 | 0.1025 |  | 22.9225 | 0 |  | 2312.8589 | 0.9034
|  | 23.0208 | 0 |  | 357.9822 | 0.10381 |  | 23.0208 | 0 |  | 2322.7854 | 0.72432
|  | 23.1192 | 0 |  | 359.5121 | 0.11991 |  | 23.1192 | 0 |  | 2332.7118 | 0.519
|  | 23.2176 | 0 |  | 361.0419 | 0.1339 |  | 23.2176 | 0 |  | 2342.6382 | 0.37571
|  | 23.316 | 0 |  | 362.5717 | 0.13378 |  | 23.316 | 0 |  | 2352.5647 | 0.3757
|  | 23.4144 | 0 |  | 364.1016 | 0.11488 |  | 23.4144 | 0 |  | 2362.4911 | 0.4612
|  | 23.5127 | 0 |  | 365.6314 | 0.079479 |  | 23.5127 | 0 |  | 2372.4175 | 0.5436
|  | 23.6111 | 0 |  | 367.1612 | 0.045651 |  | 23.6111 | 0 |  | 2382.344 | 0.59339
|  | 23.7095 | 0 |  | 368.6911 | 0.069191 |  | 23.7095 | 0 |  | 2392.2704 | 0.60126
|  | 23.8079 | 0 |  | 370.2209 | 0.12376 |  | 23.8079 | 0 |  | 2402.1968 | 0.55745
|  | 23.9062 | 0 |  | 371.7508 | 0.17365 |  | 23.9062 | 0 |  | 2412.1233 | 0.45366
|  | 24.0046 | 0 |  | 373.2806 | 0.20731 |  | 24.0046 | 0 |  | 2422.0497 | 0.29306
|  | 24.103 | 0 |  | 374.8104 | 0.2192 |  | 24.103 | 0 |  | 2431.9761 | 0.097231
|  | 24.2014 | 0 |  | 376.3403 | 0.20856 |  | 24.2014 | 0 |  | 2441.9026 | 0.10247
|  | 24.2998 | 0 |  | 377.8701 | 0.18023 |  | 24.2998 | 0 |  | 2451.829 | 0.2538
|  | 24.3981 | 0 |  | 379.4 | 0.14673 |  | 24.3981 | 0 |  | 2461.7554 | 0.32398
|  | 24.4965 | 0 |  | 380.9298 | 0.12987 |  | 24.4965 | 0 |  | 2471.6819 | 0.29266
|  | 24.5949 | 0 |  | 382.4596 | 0.14447 |  | 24.5949 | 0 |  | 2481.6083 | 0.16906
|  | 24.6933 | 0 |  | 383.9895 | 0.17533 |  | 24.6933 | 0 |  | 2491.5347 | 0.089588
|  | 24.7917 | 0 |  | 385.5193 | 0.20045 |  | 24.7917 | 0 |  | 2501.4612 | 0.27794
|  | 24.89 | 0 |  | 387.0491 | 0.20786 |  | 24.89 | 0 |  | 2511.3876 | 0.4551
|  | 24.9884 | 0 |  | 388.579 | 0.19447 |  | 24.9884 | 0 |  | 2521.314 | 0.54834
|  | 25.0868 | 0 |  | 390.1088 | 0.16544 |  | 25.0868 | 0 |  | 2531.2405 | 0.52754
|  | 25.1852 | 0 |  | 391.6387 | 0.13685 |  | 25.1852 | 0 |  | 2541.1669 | 0.39813
|  | 25.2836 | 0 |  | 393.1685 | 0.13478 |  | 25.2836 | 0 |  | 2551.0933 | 0.24052
|  | 25.3819 | 0 |  | 394.6983 | 0.16556 |  | 25.3819 | 0 |  | 2561.0198 | 0.32202
|  | 25.4803 | 0 |  | 396.2282 | 0.20532 |  | 25.4803 | 0 |  | 2570.9462 | 0.57529
|  | 25.5787 | 0 |  | 397.758 | 0.23257 |  | 25.5787 | 0 |  | 2580.8726 | 0.80488
|  | 25.6771 | 0 |  | 399.2879 | 0.23553 |  | 25.6771 | 0 |  | 2590.7991 | 0.94686
|  | 25.7755 | 0 |  | 400.8177 | 0.20922 |  | 25.7755 | 0 |  | 2600.7255 | 0.97879
|  | 25.8738 | 0 |  | 402.3475 | 0.15411 |  | 25.8738 | 0 |  | 2610.6519 | 0.90626
|  | 25.9722 | 0 |  | 403.8774 | 0.075173 |  | 25.9722 | 0 |  | 2620.5784 | 0.75796
|  | 26.0706 | 0 |  | 405.4072 | 0.019164 |  | 26.0706 | 0 |  | 2630.5048 | 0.58001
|  | 26.169 | 0 |  | 406.937 | 0.11802 |  | 26.169 | 0 |  | 2640.4312 | 0.42914
|  | 26.2674 | 0 |  | 408.4669 | 0.21034 |  | 26.2674 | 0 |  | 2650.3577 | 0.34941
|  | 26.3657 | 0 |  | 409.9967 | 0.28584 |  | 26.3657 | 0 |  | 2660.2841 | 0.31975
|  | 26.4641 | 0 |  | 411.5266 | 0.33674 |  | 26.4641 | 0 |  | 2670.2105 | 0.27284
|  | 26.5625 | 0 |  | 413.0564 | 0.35874 |  | 26.5625 | 0 |  | 2680.137 | 0.17664
|  | 26.6609 | 0 |  | 414.5862 | 0.35152 |  | 26.6609 | 0 |  | 2690.0634 | 0.13295
|  | 26.7593 | 0 |  | 416.1161 | 0.31865 |  | 26.7593 | 0 |  | 2699.9898 | 0.31108
|  | 26.8576 | 0 |  | 417.6459 | 0.26703 |  | 26.8576 | 0 |  | 2709.9163 | 0.53789
|  | 26.956 | 0 |  | 419.1758 | 0.20579 |  | 26.956 | 0 |  | 2719.8427 | 0.73114
|  | 27.0544 | 0 |  | 420.7056 | 0.14519 |  | 27.0544 | 0 |  | 2729.7691 | 0.84397
|  | 27.1528 | 0 |  | 422.2354 | 0.095238 |  | 27.1528 | 0 |  | 2739.6956 | 0.85169
|  | 27.2512 | 0 |  | 423.7653 | 0.063376 |  | 27.2512 | 0 |  | 2749.622 | 0.76485
|  | 27.3495 | 0 |  | 425.2951 | 0.047373 |  | 27.3495 | 0 |  | 2759.5484 | 0.64993
|  | 27.4479 | 0 |  | 426.825 | 0.034095 |  | 27.4479 | 0 |  | 2769.4749 | 0.6332
|  | 27.5463 | 0 |  | 428.3548 | 0.017208 |  | 27.5463 | 0 |  | 2779.4013 | 0.76581
|  | 27.6447 | 0 |  | 429.8846 | 0.025093 |  | 27.6447 | 0 |  | 2789.3277 | 0.93895
|  | 27.7431 | 0 |  | 431.4145 | 0.055254 |  | 27.7431 | 0 |  | 2799.2542 | 1.0424
|  | 27.8414 | 0 |  | 432.9443 | 0.08449 |  | 27.8414 | 0 |  | 2809.1806 | 1.0226
|  | 27.9398 | 0 |  | 434.4741 | 0.10531 |  | 27.9398 | 0 |  | 2819.107 | 0.87289
|  | 28.0382 | 0 |  | 436.004 | 0.11324 |  | 28.0382 | 0 |  | 2829.0335 | 0.62528
|  | 28.1366 | 0 |  | 437.5338 | 0.10653 |  | 28.1366 | 0 |  | 2838.9599 | 0.34713
|  | 28.235 | 0 |  | 439.0637 | 0.086371 |  | 28.235 | 0 |  | 2848.8863 | 0.20156
|  | 28.3333 | 0 |  | 440.5935 | 0.056944 |  | 28.3333 | 0 |  | 2858.8128 | 0.32331
|  | 28.4317 | 0 |  | 442.1233 | 0.026891 |  | 28.4317 | 0 |  | 2868.7392 | 0.44105
|  | 28.5301 | 0 |  | 443.6532 | 0.024376 |  | 28.5301 | 0 |  | 2878.6656 | 0.47084
|  | 28.6285 | 0 |  | 445.183 | 0.043561 |  | 28.6285 | 0 |  | 2888.5921 | 0.41844
|  | 28.7269 | 0 |  | 446.7129 | 0.053882 |  | 28.7269 | 0 |  | 2898.5185 | 0.31379
|  | 28.8252 | 0 |  | 448.2427 | 0.051007 |  | 28.8252 | 0 |  | 2908.4449 | 0.19399
|  | 28.9236 | 0 |  | 449.7725 | 0.042007 |  | 28.9236 | 0 |  | 2918.3714 | 0.10646
|  | 29.022 | 0 |  | 451.3024 | 0.052479 |  | 29.022 | 0 |  | 2928.2978 | 0.13019
|  | 29.1204 | 0 |  | 452.8322 | 0.089729 |  | 29.1204 | 0 |  | 2938.2242 | 0.21436
|  | 29.2188 | 0 |  | 454.362 | 0.13545 |  | 29.2188 | 0 |  | 2948.1507 | 0.29796
|  | 29.3171 | 0 |  | 455.8919 | 0.17724 |  | 29.3171 | 0 |  | 2958.0771 | 0.36906
|  | 29.4155 | 0 |  | 457.4217 | 0.20652 |  | 29.4155 | 0 |  | 2968.0035 | 0.43659
|  | 29.5139 | 0 |  | 458.9516 | 0.21718 |  | 29.5139 | 0 |  | 2977.93 | 0.52573
|  | 29.6123 | 0 |  | 460.4814 | 0.20598 |  | 29.6123 | 0 |  | 2987.8564 | 0.64957
|  | 29.7106 | 0 |  | 462.0112 | 0.17292 |  | 29.7106 | 0 |  | 2997.7828 | 0.78244
|  | 29.809 | 0 |  | 463.5411 | 0.12157 |  | 29.809 | 0 |  | 3007.7093 | 0.87158
|  | 29.9074 | 0 |  | 465.0709 | 0.06086 |  | 29.9074 | 0 |  | 3017.6357 | 0.8657
|  | 30.0058 | 0 |  | 466.6008 | 0.04039 |  | 30.0058 | 0 |  | 3027.5621 | 0.74154
|  | 30.1042 | 0 |  | 468.1306 | 0.097158 |  | 30.1042 | 0 |  | 3037.4886 | 0.54253
|  | 30.2025 | 0 |  | 469.6604 | 0.15336 |  | 30.2025 | 0 |  | 3047.415 | 0.48101
|  | 30.3009 | 0 |  | 471.1903 | 0.19377 |  | 30.3009 | 0 |  | 3057.3414 | 0.73679
|  | 30.3993 | 0 |  | 472.7201 | 0.21329 |  | 30.3993 | 0 |  | 3067.2679 | 1.0792
|  | 30.4977 | 0 |  | 474.2499 | 0.21056 |  | 30.4977 | 0 |  | 3077.1943 | 1.3328
|  | 30.5961 | 0 |  | 475.7798 | 0.1875 |  | 30.5961 | 0 |  | 3087.1207 | 1.4158
|  | 30.6944 | 0 |  | 477.3096 | 0.14891 |  | 30.6944 | 0 |  | 3097.0472 | 1.3014
|  | 30.7928 | 0 |  | 478.8395 | 0.10227 |  | 30.7928 | 0 |  | 3106.9736 | 1.0143
|  | 30.8912 | 0 |  | 480.3693 | 0.05933 |  | 30.8912 | 0 |  | 3116.9 | 0.63259
|  | 30.9896 | 0 |  | 481.8991 | 0.045255 |  | 30.9896 | 0 |  | 3126.8265 | 0.34015
|  | 31.088 | 0 |  | 483.429 | 0.064893 |  | 31.088 | 0 |  | 3136.7529 | 0.44282
|  | 31.1863 | 0 |  | 484.9588 | 0.084401 |  | 31.1863 | 0 |  | 3146.6793 | 0.63904
|  | 31.2847 | 0 |  | 486.4887 | 0.09224 |  | 31.2847 | 0 |  | 3156.6058 | 0.70141
|  | 31.3831 | 0 |  | 488.0185 | 0.087457 |  | 31.3831 | 0 |  | 3166.5322 | 0.61288
|  | 31.4815 | 0 |  | 489.5483 | 0.072658 |  | 31.4815 | 0 |  | 3176.4586 | 0.45949
|  | 31.5799 | 0 |  | 491.0782 | 0.052155 |  | 31.5799 | 0 |  | 3186.3851 | 0.46808
|  | 31.6782 | 0 |  | 492.608 | 0.03075 |  | 31.6782 | 0 |  | 3196.3115 | 0.70305
|  | 31.7766 | 0 |  | 494.1378 | 0.01264 |  | 31.7766 | 0 |  | 3206.2379 | 0.96411
|  | 31.875 | 0 |  | 495.6677 | 0.0043786 |  | 31.875 | 0 |  | 3216.1644 | 1.1396
|  | 31.9734 | 0 |  | 497.1975 | 0.013141 |  | 31.9734 | 0 |  | 3226.0908 | 1.1885
|  | 32.0718 | 0 |  | 498.7274 | 0.025042 |  | 32.0718 | 0 |  | 3236.0172 | 1.1071
|  | 32.1701 | 0 |  | 500.2572 | 0.040645 |  | 32.1701 | 0 |  | 3245.9437 | 0.91742
|  | 32.2685 | 0 |  | 501.787 | 0.057829 |  | 32.2685 | 0 |  | 3255.8701 | 0.66216
|  | 32.3669 | 0 |  | 503.3169 | 0.07256 |  | 32.3669 | 0 |  | 3265.7965 | 0.42094
|  | 32.4653 | 0 |  | 504.8467 | 0.080831 |  | 32.4653 | 0 |  | 3275.723 | 0.36931
|  | 32.5637 | 0 |  | 506.3766 | 0.080056 |  | 32.5637 | 0 |  | 3285.6494 | 0.53743
|  | 32.662 | 0 |  | 507.9064 | 0.070594 |  | 32.662 | 0 |  | 3295.5758 | 0.72365
|  | 32.7604 | 0 |  | 509.4362 | 0.058655 |  | 32.7604 | 0 |  | 3305.5023 | 0.84052
|  | 32.8588 | 0 |  | 510.9661 | 0.059568 |  | 32.8588 | 34.4505 |  | 3315.4287 | 0.86621
|  | 32.9572 | 0 |  | 512.4959 | 0.081055 |  | 32.9572 | 48.5348 |  | 3325.3551 | 0.80942
|  | 33.0556 | 0 |  | 514.0257 | 0.11056 |  | 33.0556 | 50.3754 |  | 3335.2816 | 0.69906
|  | 33.1539 | 0 |  | 515.5556 | 0.13574 |  | 33.1539 | 52.6161 |  | 3345.208 | 0.57161
|  | 33.2523 | 0 |  | 517.0854 | 0.14887 |  | 33.2523 | 53.9765 |  | 3355.1344 | 0.45422
|  | 33.3507 | 0 |  | 518.6153 | 0.1454 |  | 33.3507 | 54.2966 |  | 3365.0609 | 0.35779
|  | 33.4491 | 0 |  | 520.1451 | 0.12384 |  | 33.4491 | 57.4975 |  | 3374.9873 | 0.30338
|  | 33.5475 | 0 |  | 521.6749 | 0.086586 |  | 33.5475 | 59.5782 |  | 3384.9137 | 0.34193
|  | 33.6458 | 0 |  | 523.2048 | 0.04582 |  | 33.6458 | 59.9783 |  | 3394.8402 | 0.46205
|  | 33.7442 | 0 |  | 524.7346 | 0.056649 |  | 33.7442 | 60.0983 |  | 3404.7666 | 0.58748
|  | 33.8426 | 0 |  | 526.2645 | 0.1108 |  | 33.8426 | 60.6185 |  | 3414.693 | 0.65634
|  | 33.941 | 49.6552 |  | 527.7943 | 0.16456 |  | 33.941 | 60.5385 |  | 3424.6195 | 0.63793
|  | 34.0394 | 50.1353 |  | 529.3241 | 0.20613 |  | 34.0394 | 59.058 |  | 3434.5459 | 0.54571
|  | 34.1377 | 50.0953 |  | 530.854 | 0.22958 |  | 34.1377 | 59.018 |  | 3444.4723 | 0.47157
|  | 34.2361 | 50.1353 |  | 532.3838 | 0.23218 |  | 34.2361 | 59.2181 |  | 3454.3988 | 0.5667
|  | 34.3345 | 50.5354 |  | 533.9136 | 0.2144 |  | 34.3345 | 59.7782 |  | 3464.3252 | 0.81059
|  | 34.4329 | 51.2556 |  | 535.4435 | 0.17979 |  | 34.4329 | 60.6585 |  | 3474.2516 | 1.0799
|  | 34.5312 | 51.3357 |  | 536.9733 | 0.13502 |  | 34.5312 | 61.2987 |  | 3484.1781 | 1.298
|  | 34.6296 | 51.3357 |  | 538.5032 | 0.090796 |  | 34.6296 | 60.2584 |  | 3494.1045 | 1.4214
|  | 34.728 | 51.3357 |  | 540.033 | 0.065704 |  | 34.728 | 58.8179
|  | 34.8264 | 51.3357 |  | 541.5628 | 0.073433 |  | 34.8264 | 58.8179
|  | 34.9248 | 54.0965 |  | 543.0927 | 0.092301 |  | 34.9248 | 58.8179
|  | 35.0231 | 53.7764 |  | 544.6225 | 0.10359 |  | 35.0231 | 0
|  | 35.1215 | 54.2966 |  | 546.1524 | 0.1028 |  | 35.1215 | 0
|  | 35.2199 | 55.0968 |  | 547.6822 | 0.092056 |  | 35.2199 | 0
|  | 35.3183 | 54.8567 |  | 549.212 | 0.077534 |  | 35.3183 | 0
|  | 35.4167 | 54.5366 |  | 550.7419 | 0.067518 |  | 35.4167 | 0
|  | 35.515 | 53.4963 |  | 552.2717 | 0.066189 |  | 35.515 | 0
|  | 35.6134 | 52.376 |  | 553.8015 | 0.068074 |  | 35.6134 | 0
|  | 35.7118 | 52.456 |  | 555.3314 | 0.064842 |  | 35.7118 | 0
|  | 35.8102 | 53.0962 |  | 556.8612 | 0.05222 |  | 35.8102 | 0
|  | 35.9086 | 53.2563 |  | 558.3911 | 0.03199 |  | 35.9086 | 0
|  | 36.0069 | 52.2159 |  | 559.9209 | 0.024852 |  | 36.0069 | 0
|  | 36.1053 | 50.4554 |  | 561.4507 | 0.051242 |  | 36.1053 | 0
|  | 36.2037 | 48.5348 |  | 562.9806 | 0.082351 |  | 36.2037 | 0
|  | 36.3021 | 47.4145 |  | 564.5104 | 0.10636 |  | 36.3021 | 0
|  | 36.4005 | 46.0541 |  | 566.0403 | 0.11772 |  | 36.4005 | 0
|  | 36.4988 | 45.814 |  | 567.5701 | 0.11367 |  | 36.4988 | 0
|  | 36.5972 | 45.5739 |  | 569.0999 | 0.094681 |  | 36.5972 | 0
|  | 36.6956 | 45.1738 |  | 570.6298 | 0.066251 |  | 36.6956 | 0
|  | 36.794 | 44.8137 |  | 572.1596 | 0.047789 |  | 36.794 | 0
|  | 36.8924 | 41.0925 |  | 573.6894 | 0.067974 |  | 36.8924 | 0
|  | 36.9907 | 41.0925 |  | 575.2193 | 0.10463 |  | 36.9907 | 0
|  | 37.0891 | 41.0925 |  | 576.7491 | 0.1357 |  | 37.0891 | 0
|  | 37.1875 | 0 |  | 578.279 | 0.15243 |  | 37.1875 | 0
|  | 37.2859 | 0 |  | 579.8088 | 0.15102 |  | 37.2859 | 0
|  | 37.3843 | 0 |  | 581.3386 | 0.13154 |  | 37.3843 | 0
|  | 37.4826 | 0 |  | 582.8685 | 0.098916 |  | 37.4826 | 0
|  | 37.581 | 0 |  | 584.3983 | 0.068349 |  | 37.581 | 0
|  | 37.6794 | 0 |  | 585.9282 | 0.073441 |  | 37.6794 | 0
|  | 37.7778 | 0 |  | 587.458 | 0.11165 |  | 37.7778 | 0
|  | 37.8762 | 0 |  | 588.9878 | 0.15196 |  | 37.8762 | 0
|  | 37.9745 | 0 |  | 590.5177 | 0.1803 |  | 37.9745 | 0
|  | 38.0729 | 0 |  | 592.0475 | 0.19075 |  | 38.0729 | 0
|  | 38.1713 | 0 |  | 593.5774 | 0.18159 |  | 38.1713 | 0
|  | 38.2697 | 0 |  | 595.1072 | 0.15451 |  | 38.2697 | 0
|  | 38.3681 | 0 |  | 596.637 | 0.11431 |  | 38.3681 | 0
|  | 38.4664 | 0 |  | 598.1669 | 0.069552 |  | 38.4664 | 0
|  | 38.5648 | 0 |  | 599.6967 | 0.040861 |  | 38.5648 | 0
|  | 38.6632 | 0 |  | 601.2265 | 0.057422 |  | 38.6632 | 0
|  | 38.7616 | 0 |  | 602.7564 | 0.085161 |  | 38.7616 | 0
|  | 38.86 | 0 |  | 604.2862 | 0.10239 |  | 38.86 | 0
|  | 38.9583 | 0 |  | 605.8161 | 0.10508 |  | 38.9583 | 0
|  | 39.0567 | 0 |  | 607.3459 | 0.094385 |  | 39.0567 | 0
|  | 39.1551 | 0 |  | 608.8757 | 0.074901 |  | 39.1551 | 0
|  | 39.2535 | 0 |  | 610.4056 | 0.054853 |  | 39.2535 | 0
|  | 39.3519 | 0 |  | 611.9354 | 0.045875 |  | 39.3519 | 0
|  | 39.4502 | 0 |  | 613.4653 | 0.050715 |  | 39.4502 | 0
|  | 39.5486 | 0 |  | 614.9951 | 0.056885 |  | 39.5486 | 0
|  | 39.647 | 0 |  | 616.5249 | 0.055373 |  | 39.647 | 0
|  | 39.7454 | 0 |  | 618.0548 | 0.044412 |  | 39.7454 | 0
|  | 39.8438 | 0 |  | 619.5846 | 0.02921 |  | 39.8438 | 0
|  | 39.9421 | 0 |  | 621.1144 | 0.029783 |  | 39.9421 | 0
|  | 40.0405 | 0 |  | 622.6443 | 0.050771 |  | 40.0405 | 0
|  | 40.1389 | 0 |  | 624.1741 | 0.072403 |  | 40.1389 | 0
|  | 40.2373 | 0 |  | 625.704 | 0.085683 |  | 40.2373 | 0
|  | 40.3356 | 0 |  | 627.2338 | 0.086524 |  | 40.3356 | 0
|  | 40.434 | 0 |  | 628.7636 | 0.073974 |  | 40.434 | 0
|  | 40.5324 | 0 |  | 630.2935 | 0.051258 |  | 40.5324 | 0
|  | 40.6308 | 0 |  | 631.8233 | 0.034136 |  | 40.6308 | 0
|  | 40.7292 | 0 |  | 633.3532 | 0.053433 |  | 40.7292 | 0
|  | 40.8275 | 0 |  | 634.883 | 0.088337 |  | 40.8275 | 0
|  | 40.9259 | 0 |  | 636.4128 | 0.11887 |  | 40.9259 | 0
|  | 41.0243 | 0 |  | 637.9427 | 0.13724 |  | 41.0243 | 0
|  | 41.1227 | 0 |  | 639.4725 | 0.13932 |  | 41.1227 | 0
|  | 41.2211 | 0 |  | 641.0023 | 0.12382 |  | 41.2211 | 0
|  | 41.3194 | 0 |  | 642.5322 | 0.092186 |  | 41.3194 | 0
|  | 41.4178 | 0 |  | 644.062 | 0.048411 |  | 41.4178 | 0
|  | 41.5162 | 0 |  | 645.5919 | 0.0061784 |  | 41.5162 | 0
|  | 41.6146 | 0 |  | 647.1217 | 0.052813 |  | 41.6146 | 0
|  | 41.713 | 0 |  | 648.6515 | 0.095827 |  | 41.713 | 0
|  | 41.8113 | 0 |  | 650.1814 | 0.1263 |  | 41.8113 | 0
|  | 41.9097 | 0 |  | 651.7112 | 0.14052 |  | 41.9097 | 0
|  | 42.0081 | 0 |  | 653.2411 | 0.13754 |  | 42.0081 | 0
|  | 42.1065 | 0 |  | 654.7709 | 0.11977 |  | 42.1065 | 0
|  | 42.2049 | 0 |  | 656.3007 | 0.094074 |  | 42.2049 | 0
|  | 42.3032 | 0 |  | 657.8306 | 0.074651 |  | 42.3032 | 0
|  | 42.4016 | 0 |  | 659.3604 | 0.079107 |  | 42.4016 | 0
|  | 42.5 | 0 |  | 660.8902 | 0.10183 |  | 42.5 | 0
|  | 42.5984 | 0 |  | 662.4201 | 0.12453 |  | 42.5984 | 0
|  | 42.6968 | 0 |  | 663.9499 | 0.13711 |  | 42.6968 | 0
|  | 42.7951 | 0 |  | 665.4798 | 0.1359 |  | 42.7951 | 0
|  | 42.8935 | 0 |  | 667.0096 | 0.12104 |  | 42.8935 | 0
|  | 42.9919 | 0 |  | 668.5394 | 0.095556 |  | 42.9919 | 0
|  | 43.0903 | 0 |  | 670.0693 | 0.064735 |  | 43.0903 | 0
|  | 43.1887 | 0 |  | 671.5991 | 0.03672 |  | 43.1887 | 0
|  | 43.287 | 0 |  | 673.129 | 0.026534 |  | 43.287 | 0
|  | 43.3854 | 0 |  | 674.6588 | 0.034724 |  | 43.3854 | 0
|  | 43.4838 | 0 |  | 676.1886 | 0.039531 |  | 43.4838 | 0
|  | 43.5822 | 0 |  | 677.7185 | 0.034095 |  | 43.5822 | 0
|  | 43.6806 | 0 |  | 679.2483 | 0.020012 |  | 43.6806 | 0
|  | 43.7789 | 0 |  | 680.7781 | 0.015942 |  | 43.7789 | 0
|  | 43.8773 | 0 |  | 682.308 | 0.03795 |  | 43.8773 | 0
|  | 43.9757 | 0 |  | 683.8378 | 0.061329 |  | 43.9757 | 0
|  | 44.0741 | 0 |  | 685.3677 | 0.07816 |  | 44.0741 | 0
|  | 44.1725 | 0 |  | 686.8975 | 0.084522 |  | 44.1725 | 0
|  | 44.2708 | 0 |  | 688.4273 | 0.078814 |  | 44.2708 | 0
|  | 44.3692 | 0 |  | 689.9572 | 0.062265 |  | 44.3692 | 0
|  | 44.4676 | 0 |  | 691.487 | 0.041109 |  | 44.4676 | 0
|  | 44.566 | 0 |  | 693.0169 | 0.035784 |  | 44.566 | 0
|  | 44.6644 | 0 |  | 694.5467 | 0.057679 |  | 44.6644 | 0
|  | 44.7627 | 0 |  | 696.0765 | 0.083953 |  | 44.7627 | 0
|  | 44.8611 | 0 |  | 697.6064 | 0.10281 |  | 44.8611 | 0
|  | 44.9595 | 0 |  | 699.1362 | 0.10944 |  | 44.9595 | 0
|  | 45.0579 | 0 |  | 700.666 | 0.10222 |  | 45.0579 | 0
|  | 45.1562 | 0 |  | 702.1959 | 0.082246 |  | 45.1562 | 0
|  | 45.2546 | 0 |  | 703.7257 | 0.053768 |  | 45.2546 | 0
|  | 45.353 | 0 |  | 705.2556 | 0.029363 |  | 45.353 | 0
|  | 45.4514 | 0 |  | 706.7854 | 0.040286 |  | 45.4514 | 0
|  | 45.5498 | 0 |  | 708.3152 | 0.067303 |  | 45.5498 | 0
|  | 45.6481 | 0 |  | 709.8451 | 0.08832 |  | 45.6481 | 0
|  | 45.7465 | 0 |  | 711.3749 | 0.098168 |  | 45.7465 | 0
|  | 45.8449 | 0 |  | 712.9048 | 0.09658 |  | 45.8449 | 0
|  | 45.9433 | 0 |  | 714.4346 | 0.087517 |  | 45.9433 | 0
|  | 46.0417 | 0 |  | 715.9644 | 0.079876 |  | 46.0417 | 0
|  | 46.14 | 0 |  | 717.4943 | 0.083976 |  | 46.14 | 0
|  | 46.2384 | 0 |  | 719.0241 | 0.099982 |  | 46.2384 | 0
|  | 46.3368 | 0 |  | 720.5539 | 0.11831 |  | 46.3368 | 0
|  | 46.4352 | 0 |  | 722.0838 | 0.1302 |  | 46.4352 | 0
|  | 46.5336 | 0 |  | 723.6136 | 0.13076 |  | 46.5336 | 0
|  | 46.6319 | 0 |  | 725.1435 | 0.1186 |  | 46.6319 | 0
|  | 46.7303 | 0 |  | 726.6733 | 0.095515 |  | 46.7303 | 0
|  | 46.8287 | 0 |  | 728.2031 | 0.06685 |  | 46.8287 | 0
|  | 46.9271 | 0 |  | 729.733 | 0.044546 |  | 46.9271 | 0
|  | 47.0255 | 0 |  | 731.2628 | 0.047004 |  | 47.0255 | 0
|  | 47.1238 | 0 |  | 732.7927 | 0.064369 |  | 47.1238 | 0
|  | 47.2222 | 0 |  | 734.3225 | 0.077306 |  | 47.2222 | 0
|  | 47.3206 | 0 |  | 735.8523 | 0.079241 |  | 47.3206 | 0
|  | 47.419 | 0 |  | 737.3822 | 0.070071 |  | 47.419 | 0
|  | 47.5174 | 0 |  | 738.912 | 0.055388 |  | 47.5174 | 0
|  | 47.6157 | 0 |  | 740.4418 | 0.049704 |  | 47.6157 | 0
|  | 47.7141 | 0 |  | 741.9717 | 0.064616 |  | 47.7141 | 0
|  | 47.8125 | 0 |  | 743.5015 | 0.088309 |  | 47.8125 | 0
|  | 47.9109 | 0 |  | 745.0314 | 0.10792 |  | 47.9109 | 0
|  | 48.0093 | 0 |  | 746.5612 | 0.11707 |  | 48.0093 | 0
|  | 48.1076 | 0 |  | 748.091 | 0.11351 |  | 48.1076 | 0
|  | 48.206 | 0 |  | 749.6209 | 0.098797 |  | 48.206 | 0
|  | 48.3044 | 0 |  | 751.1507 | 0.079379 |  | 48.3044 | 0
|  | 48.4028 | 0 |  | 752.6806 | 0.06881 |  | 48.4028 | 0
|  | 48.5012 | 0 |  | 754.2104 | 0.078822 |  | 48.5012 | 0
|  | 48.5995 | 0 |  | 755.7402 | 0.10038 |  | 48.5995 | 0
|  | 48.6979 | 0 |  | 757.2701 | 0.11913 |  | 48.6979 | 0
|  | 48.7963 | 0 |  | 758.7999 | 0.12722 |  | 48.7963 | 0
|  | 48.8947 | 0 |  | 760.3298 | 0.12175 |  | 48.8947 | 0
|  | 48.9931 | 0 |  | 761.8596 | 0.1037 |  | 48.9931 | 0
|  | 49.0914 | 0 |  | 763.3894 | 0.078152 |  | 49.0914 | 0
|  | 49.1898 | 0 |  | 764.9193 | 0.056888 |  | 49.1898 | 0
|  | 49.2882 | 0 |  | 766.4491 | 0.057624 |  | 49.2882 | 0
|  | 49.3866 | 0 |  | 767.9789 | 0.07553 |  | 49.3866 | 0
|  | 49.485 | 0 |  | 769.5088 | 0.091459 |  | 49.485 | 0
|  | 49.5833 | 0 |  | 771.0386 | 0.09608 |  | 49.5833 | 0
|  | 49.6817 | 0 |  | 772.5685 | 0.08696 |  | 49.6817 | 0
|  | 49.7801 | 0 |  | 774.0983 | 0.067115 |  | 49.7801 | 0
|  | 49.8785 | 0 |  | 775.6281 | 0.049753 |  | 49.8785 | 0
|  | 49.9768 | 0 |  | 777.158 | 0.060779 |  | 49.9769 | 0
|  | 50.0752 | 0 |  | 778.6878 | 0.093456 |  | 50.0752 | 0
|  | 50.1736 | 0 |  | 780.2177 | 0.1264 |  | 50.1736 | 0
|  | 50.272 | 0 |  | 781.7475 | 0.14944 |  | 50.272 | 0
|  | 50.3704 | 0 |  |  |  |  | 50.3704 | 0
|  | 50.4688 | 0 |  |  |  |  | 50.4688 | 0
|  | 50.5671 | 0 |  |  |  |  | 50.5671 | 0
|  | 50.6655 | 51.0156 |  |  |  |  | 50.6655 | 0
|  | 50.7639 | 52.2159 |  |  |  |  | 50.7639 | 0
|  | 50.8623 | 53.8164 |  |  |  |  | 50.8623 | 0
|  | 50.9606 | 54.7767 |  |  |  |  | 50.9606 | 0
|  | 51.059 | 54.9768 |  |  |  |  | 51.059 | 0
|  | 51.1574 | 54.3366 |  |  |  |  | 51.1574 | 0
|  | 51.2558 | 54.1365 |  |  |  |  | 51.2558 | 0
|  | 51.3542 | 54.6967 |  |  |  |  | 51.3542 | 0
|  | 51.4525 | 55.737 |  |  |  |  | 51.4525 | 0
|  | 51.5509 | 56.4572 |  |  |  |  | 51.5509 | 0
|  | 51.6493 | 56.2972 |  |  |  |  | 51.6493 | 0
|  | 51.7477 | 56.1371 |  |  |  |  | 51.7477 | 0
|  | 51.8461 | 56.1371 |  |  |  |  | 51.8461 | 0
|  | 51.9444 | 56.1371 |  |  |  |  | 51.9444 | 0
|  | 52.0428 | 56.1371 |  |  |  |  | 52.0428 | 0
|  | 52.1412 | 56.1371 |  |  |  |  | 52.1412 | 0
|  | 52.2396 | 56.1371 |  |  |  |  | 52.2396 | 0
|  | 52.338 | 56.1371 |  |  |  |  | 52.338 | 0
|  | 52.4363 | 56.1371 |  |  |  |  | 52.4363 | 0
|  | 52.5347 | 56.1371 |  |  |  |  | 52.5347 | 0
|  | 52.6331 | 53.6964 |  |  |  |  | 52.6331 | 0
|  | 52.7315 | 53.2563 |  |  |  |  | 52.7315 | 0
|  | 52.8299 | 52.8961 |  |  |  |  | 52.8299 | 0
|  | 52.9282 | 51.9358 |  |  |  |  | 52.9282 | 0
|  | 53.0266 | 51.1356 |  |  |  |  | 53.0266 | 0
|  | 53.125 | 49.8552 |  |  |  |  | 53.125 | 0
|  | 53.2234 | 48.3347 |  |  |  |  | 53.2234 | 0
|  | 53.3218 | 47.6145 |  |  |  |  | 53.3218 | 0
|  | 53.4201 | 46.5742 |  |  |  |  | 53.4201 | 0
|  | 53.5185 | 44.3335 |  |  |  |  | 53.5185 | 0
|  | 53.6169 | 41.7327 |  |  |  |  | 53.6169 | 0
|  | 53.7153 | 40.9325 |  |  |  |  | 53.7153 | 0
|  | 53.8137 | 40.9325 |  |  |  |  | 53.8137 | 0
|  | 53.912 | 40.9325 |  |  |  |  | 53.912 | 0
|  | 54.0104 | 0 |  |  |  |  | 54.0104 | 0
|  | 54.1088 | 0 |  |  |  |  | 54.1088 | 0
|  | 54.2072 | 0 |  |  |  |  | 54.2072 | 0
|  | 54.3056 | 0 |  |  |  |  | 54.3056 | 0
|  | 54.4039 | 0 |  |  |  |  | 54.4039 | 0
|  | 54.5023 | 0 |  |  |  |  | 54.5023 | 0
|  | 54.6007 | 0 |  |  |  |  | 54.6007 | 0
|  | 54.6991 | 0 |  |  |  |  | 54.6991 | 0
|  | 54.7975 | 0 |  |  |  |  | 54.7975 | 0
|  | 54.8958 | 0 |  |  |  |  | 54.8958 | 0
|  | 54.9942 | 0 |  |  |  |  | 54.9942 | 0
|  | 55.0926 | 0 |  |  |  |  | 55.0926 | 0
|  | 55.191 | 0 |  |  |  |  | 55.191 | 0
|  | 55.2893 | 0 |  |  |  |  | 55.2894 | 0
|  | 55.3877 | 0 |  |  |  |  | 55.3877 | 0
|  | 55.4861 | 0 |  |  |  |  | 55.4861 | 0
|  | 55.5845 | 0 |  |  |  |  | 55.5845 | 0
|  | 55.6829 | 0 |  |  |  |  | 55.6829 | 0
|  | 55.7812 | 0 |  |  |  |  | 55.7812 | 0
|  | 55.8796 | 0 |  |  |  |  | 55.8796 | 0
|  | 55.978 | 0 |  |  |  |  | 55.978 | 0
|  | 56.0764 | 0 |  |  |  |  | 56.0764 | 0
|  | 56.1748 | 0 |  |  |  |  | 56.1748 | 0
|  | 56.2731 | 0 |  |  |  |  | 56.2731 | 0
|  | 56.3715 | 0 |  |  |  |  | 56.3715 | 0
|  | 56.4699 | 0 |  |  |  |  | 56.4699 | 0
|  | 56.5683 | 0 |  |  |  |  | 56.5683 | 0
|  | 56.6667 | 0 |  |  |  |  | 56.6667 | 0
|  | 56.765 | 0 |  |  |  |  | 56.765 | 0
|  | 56.8634 | 0 |  |  |  |  | 56.8634 | 0
|  | 56.9618 | 0 |  |  |  |  | 56.9618 | 0
|  | 57.0602 | 0 |  |  |  |  | 57.0602 | 0
|  | 57.1586 | 0 |  |  |  |  | 57.1586 | 0
|  | 57.2569 | 0 |  |  |  |  | 57.2569 | 0
|  | 57.3553 | 0 |  |  |  |  | 57.3553 | 0
|  | 57.4537 | 0 |  |  |  |  | 57.4537 | 0
|  | 57.5521 | 0 |  |  |  |  | 57.5521 | 0
|  | 57.6505 | 0 |  |  |  |  | 57.6505 | 0
|  | 57.7488 | 0 |  |  |  |  | 57.7488 | 0
|  | 57.8472 | 0 |  |  |  |  | 57.8472 | 0
|  | 57.9456 | 0 |  |  |  |  | 57.9456 | 0
|  | 58.044 | 0 |  |  |  |  | 58.044 | 0
|  | 58.1424 | 0 |  |  |  |  | 58.1424 | 0
|  | 58.2407 | 0 |  |  |  |  | 58.2407 | 0
|  | 58.3391 | 0 |  |  |  |  | 58.3391 | 0
|  | 58.4375 | 0 |  |  |  |  | 58.4375 | 0
|  | 58.5359 | 0 |  |  |  |  | 58.5359 | 0
|  | 58.6343 | 0 |  |  |  |  | 58.6343 | 0
|  | 58.7326 | 0 |  |  |  |  | 58.7326 | 0
|  | 58.831 | 0 |  |  |  |  | 58.831 | 0
|  | 58.9294 | 0 |  |  |  |  | 58.9294 | 0
|  | 59.0278 | 0 |  |  |  |  | 59.0278 | 0
|  | 59.1262 | 0 |  |  |  |  | 59.1262 | 0
|  | 59.2245 | 0 |  |  |  |  | 59.2245 | 0
|  | 59.3229 | 0 |  |  |  |  | 59.3229 | 0
|  | 59.4213 | 0 |  |  |  |  | 59.4213 | 0
|  | 59.5197 | 0 |  |  |  |  | 59.5197 | 0
|  | 59.6181 | 0 |  |  |  |  | 59.6181 | 0
|  | 59.7164 | 0 |  |  |  |  | 59.7164 | 0
|  | 59.8148 | 0 |  |  |  |  | 59.8148 | 0
|  | 59.9132 | 0 |  |  |  |  | 59.9132 | 0
|  | 60.0116 | 0 |  |  |  |  | 60.0116 | 0
|  | 60.11 | 0 |  |  |  |  | 60.11 | 0
|  | 60.2083 | 0 |  |  |  |  | 60.2083 | 0
|  | 60.3067 | 0 |  |  |  |  | 60.3067 | 0
|  | 60.4051 | 0 |  |  |  |  | 60.4051 | 0
|  | 60.5035 | 0 |  |  |  |  | 60.5035 | 0
|  | 60.6018 | 0 |  |  |  |  | 60.6019 | 0
|  | 60.7002 | 0 |  |  |  |  | 60.7002 | 0
|  | 60.7986 | 0 |  |  |  |  | 60.7986 | 0
|  | 60.897 | 0 |  |  |  |  | 60.897 | 0
|  | 60.9954 | 0 |  |  |  |  | 60.9954 | 0
|  | 61.0938 | 0 |  |  |  |  | 61.0938 | 0
|  | 61.1921 | 0 |  |  |  |  | 61.1921 | 0
|  | 61.2905 | 0 |  |  |  |  | 61.2905 | 0
|  | 61.3889 | 0 |  |  |  |  | 61.3889 | 0
|  | 61.4873 | 0 |  |  |  |  | 61.4873 | 0
|  | 61.5856 | 0 |  |  |  |  | 61.5856 | 0
|  | 61.684 | 0 |  |  |  |  | 61.684 | 0
|  | 61.7824 | 0 |  |  |  |  | 61.7824 | 0
|  | 61.8808 | 0 |  |  |  |  | 61.8808 | 0
|  | 61.9792 | 0 |  |  |  |  | 61.9792 | 0
|  | 62.0775 | 0 |  |  |  |  | 62.0775 | 0
|  | 62.1759 | 0 |  |  |  |  | 62.1759 | 0
|  | 62.2743 | 0 |  |  |  |  | 62.2743 | 0
|  | 62.3727 | 0 |  |  |  |  | 62.3727 | 0
|  | 62.4711 | 0 |  |  |  |  | 62.4711 | 0
|  | 62.5694 | 0 |  |  |  |  | 62.5694 | 0
|  | 62.6678 | 0 |  |  |  |  | 62.6678 | 0
|  | 62.7662 | 0 |  |  |  |  | 62.7662 | 0
|  | 62.8646 | 0 |  |  |  |  | 62.8646 | 0
|  | 62.963 | 0 |  |  |  |  | 62.963 | 0
|  | 63.0613 | 0 |  |  |  |  | 63.0613 | 0
|  | 63.1597 | 0 |  |  |  |  | 63.1597 | 0
|  | 63.2581 | 0 |  |  |  |  | 63.2581 | 0
|  | 63.3565 | 0 |  |  |  |  | 63.3565 | 0
|  | 63.4549 | 0 |  |  |  |  | 63.4549 | 0
|  | 63.5532 | 0 |  |  |  |  | 63.5532 | 0
|  | 63.6516 | 0 |  |  |  |  | 63.6516 | 0
|  | 63.75 | 0 |  |  |  |  | 63.75 | 0
|  | 63.8484 | 0 |  |  |  |  | 63.8484 | 0
|  | 63.9468 | 0 |  |  |  |  | 63.9468 | 0
|  | 64.0451 | 0 |  |  |  |  | 64.0451 | 0
|  | 64.1435 | 0 |  |  |  |  | 64.1435 | 0
|  | 64.2419 | 0 |  |  |  |  | 64.2419 | 0
|  | 64.3403 | 0 |  |  |  |  | 64.3403 | 0
|  | 64.4387 | 0 |  |  |  |  | 64.4387 | 0
|  | 64.537 | 0 |  |  |  |  | 64.537 | 0
|  | 64.6354 | 0 |  |  |  |  | 64.6354 | 0
|  | 64.7338 | 0 |  |  |  |  | 64.7338 | 0
|  | 64.8322 | 0 |  |  |  |  | 64.8322 | 0
|  | 64.9306 | 0 |  |  |  |  | 64.9306 | 0
|  | 65.0289 | 0 |  |  |  |  | 65.0289 | 0
|  | 65.1273 | 0 |  |  |  |  | 65.1273 | 0
|  | 65.2257 | 0 |  |  |  |  | 65.2257 | 0
|  | 65.3241 | 0 |  |  |  |  | 65.3241 | 0
|  | 65.4225 | 0 |  |  |  |  | 65.4225 | 0
|  | 65.5208 | 0 |  |  |  |  | 65.5208 | 0
|  | 65.6192 | 0 |  |  |  |  | 65.6192 | 0
|  | 65.7176 | 0 |  |  |  |  | 65.7176 | 0
|  | 65.816 | 0 |  |  |  |  | 65.816 | 0
|  | 65.9144 | 0 |  |  |  |  | 65.9144 | 0
|  | 66.0127 | 0 |  |  |  |  | 66.0127 | 0
|  | 66.1111 | 0 |  |  |  |  | 66.1111 | 0
|  | 66.2095 | 0 |  |  |  |  | 66.2095 | 48.5348
|  | 66.3079 | 0 |  |  |  |  | 66.3079 | 60.0583
|  | 66.4062 | 0 |  |  |  |  | 66.4062 | 59.8183
|  | 66.5046 | 0 |  |  |  |  | 66.5046 | 60.0183
|  | 66.603 | 0 |  |  |  |  | 66.603 | 60.4985
|  | 66.7014 | 0 |  |  |  |  | 66.7014 | 61.7789
|  | 66.7998 | 0 |  |  |  |  | 66.7998 | 63.3793
|  | 66.8981 | 0 |  |  |  |  | 66.8981 | 64.9798
|  | 66.9965 | 0 |  |  |  |  | 66.9965 | 66.1002
|  | 67.0949 | 0 |  |  |  |  | 67.0949 | 66.9004
|  | 67.1933 | 0 |  |  |  |  | 67.1933 | 67.3806
|  | 67.2917 | 0 |  |  |  |  | 67.2917 | 67.1005
|  | 67.39 | 53.8164 |  |  |  |  | 67.39 | 69.0211
|  | 67.4884 | 54.8567 |  |  |  |  | 67.4884 | 68.4609
|  | 67.5868 | 55.9771 |  |  |  |  | 67.5868 | 68.5409
|  | 67.6852 | 56.7773 |  |  |  |  | 67.6852 | 68.3809
|  | 67.7836 | 56.9774 |  |  |  |  | 67.7836 | 71.7419
|  | 67.8819 | 56.4172 |  |  |  |  | 67.8819 | 71.2617
|  | 67.9803 | 56.1371 |  |  |  |  | 67.9803 | 67.9007
|  | 68.0787 | 57.0974 |  |  |  |  | 68.0787 | 67.4606
|  | 68.1771 | 58.2978 |  |  |  |  | 68.1771 | 66.2602
|  | 68.2755 | 58.8179 |  |  |  |  | 68.2755 | 63.6994
|  | 68.3738 | 58.7379 |  |  |  |  | 68.3738 | 63.6994
|  | 68.4722 | 59.1781 |  |  |  |  | 68.4722 | 63.6994
|  | 68.5706 | 58.6179 |  |  |  |  | 68.5706 | 0
|  | 68.669 | 58.5779 |  |  |  |  | 68.669 | 0
|  | 68.7674 | 59.018 |  |  |  |  | 68.7674 | 0
|  | 68.8657 | 58.6179 |  |  |  |  | 68.8657 | 0
|  | 68.9641 | 58.7379 |  |  |  |  | 68.9641 | 0
|  | 69.0625 | 58.6179 |  |  |  |  | 69.0625 | 0
|  | 69.1609 | 58.1778 |  |  |  |  | 69.1609 | 0
|  | 69.2593 | 57.1374 |  |  |  |  | 69.2593 | 0
|  | 69.3576 | 55.8971 |  |  |  |  | 69.3576 | 0
|  | 69.456 | 56.3772 |  |  |  |  | 69.456 | 0
|  | 69.5544 | 56.5373 |  |  |  |  | 69.5544 | 0
|  | 69.6528 | 54.8567 |  |  |  |  | 69.6528 | 0
|  | 69.7512 | 52.456 |  |  |  |  | 69.7512 | 0
|  | 69.8495 | 51.0956 |  |  |  |  | 69.8495 | 0
|  | 69.9479 | 50.1353 |  |  |  |  | 69.9479 | 0
|  | 70.0463 | 48.4548 |  |  |  |  | 70.0463 | 0
|  | 70.1447 | 47.2544 |  |  |  |  | 70.1447 | 0
|  | 70.2431 | 46.9743 |  |  |  |  | 70.2431 | 0
|  | 70.3414 | 46.4142 |  |  |  |  | 70.3414 | 0
|  | 70.4398 | 44.4936 |  |  |  |  | 70.4398 | 0
|  | 70.5382 | 41.2926 |  |  |  |  | 70.5382 | 0
|  | 70.6366 | 41.2926 |  |  |  |  | 70.6366 | 0
|  | 70.735 | 41.2926 |  |  |  |  | 70.735 | 0
|  | 70.8333 | 0 |  |  |  |  | 70.8333 | 0
|  | 70.9317 | 0 |  |  |  |  | 70.9317 | 0
|  | 71.0301 | 0 |  |  |  |  | 71.0301 | 0
|  | 71.1285 | 0 |  |  |  |  | 71.1285 | 0
|  | 71.2269 | 0 |  |  |  |  | 71.2269 | 0
|  | 71.3252 | 0 |  |  |  |  | 71.3252 | 0
|  | 71.4236 | 0 |  |  |  |  | 71.4236 | 0
|  | 71.522 | 0 |  |  |  |  | 71.522 | 0
|  | 71.6204 | 0 |  |  |  |  | 71.6204 | 0
|  | 71.7188 | 0 |  |  |  |  | 71.7188 | 0
|  | 71.8171 | 0 |  |  |  |  | 71.8171 | 0
|  | 71.9155 | 0 |  |  |  |  | 71.9155 | 0
|  | 72.0139 | 0 |  |  |  |  | 72.0139 | 0
|  | 72.1123 | 0 |  |  |  |  | 72.1123 | 0
|  | 72.2106 | 0 |  |  |  |  | 72.2106 | 0
|  | 72.309 | 0 |  |  |  |  | 72.309 | 0
|  | 72.4074 | 0 |  |  |  |  | 72.4074 | 0
|  | 72.5058 | 0 |  |  |  |  | 72.5058 | 0
|  | 72.6042 | 0 |  |  |  |  | 72.6042 | 0
|  | 72.7025 | 0 |  |  |  |  | 72.7025 | 0
|  | 72.8009 | 0 |  |  |  |  | 72.8009 | 0
|  | 72.8993 | 0 |  |  |  |  | 72.8993 | 0
|  | 72.9977 | 0 |  |  |  |  | 72.9977 | 0
|  | 73.0961 | 0 |  |  |  |  | 73.0961 | 0
|  | 73.1944 | 0 |  |  |  |  | 73.1944 | 0
|  | 73.2928 | 0 |  |  |  |  | 73.2928 | 0
|  | 73.3912 | 0 |  |  |  |  | 73.3912 | 0
|  | 73.4896 | 0 |  |  |  |  | 73.4896 | 0
|  | 73.588 | 0 |  |  |  |  | 73.588 | 0
|  | 73.6863 | 0 |  |  |  |  | 73.6863 | 0
|  | 73.7847 | 0 |  |  |  |  | 73.7847 | 0
|  | 73.8831 | 0 |  |  |  |  | 73.8831 | 0
|  | 73.9815 | 0 |  |  |  |  | 73.9815 | 0
|  | 74.0799 | 0 |  |  |  |  | 74.0799 | 0
|  | 74.1782 | 0 |  |  |  |  | 74.1782 | 0
|  | 74.2766 | 0 |  |  |  |  | 74.2766 | 0
|  | 74.375 | 0 |  |  |  |  | 74.375 | 0
|  | 74.4734 | 0 |  |  |  |  | 74.4734 | 0
|  | 74.5718 | 0 |  |  |  |  | 74.5718 | 0
|  | 74.6701 | 0 |  |  |  |  | 74.6701 | 0
|  | 74.7685 | 0 |  |  |  |  | 74.7685 | 0
|  | 74.8669 | 0 |  |  |  |  | 74.8669 | 0
|  | 74.9653 | 0 |  |  |  |  | 74.9653 | 0
|  | 75.0637 | 0 |  |  |  |  | 75.0637 | 0
|  | 75.162 | 0 |  |  |  |  | 75.162 | 0
|  | 75.2604 | 0 |  |  |  |  | 75.2604 | 0
|  | 75.3588 | 0 |  |  |  |  | 75.3588 | 0
|  | 75.4572 | 0 |  |  |  |  | 75.4572 | 0
|  | 75.5556 | 0 |  |  |  |  | 75.5556 | 0
|  | 75.6539 | 0 |  |  |  |  | 75.6539 | 0
|  | 75.7523 | 0 |  |  |  |  | 75.7523 | 0
|  | 75.8507 | 0 |  |  |  |  | 75.8507 | 0
|  | 75.9491 | 0 |  |  |  |  | 75.9491 | 0
|  | 76.0475 | 0 |  |  |  |  | 76.0475 | 0
|  | 76.1458 | 0 |  |  |  |  | 76.1458 | 0
|  | 76.2442 | 0 |  |  |  |  | 76.2442 | 0
|  | 76.3426 | 0 |  |  |  |  | 76.3426 | 0
|  | 76.441 | 0 |  |  |  |  | 76.441 | 0
|  | 76.5394 | 0 |  |  |  |  | 76.5394 | 0
|  | 76.6377 | 0 |  |  |  |  | 76.6377 | 0
|  | 76.7361 | 0 |  |  |  |  | 76.7361 | 0
|  | 76.8345 | 0 |  |  |  |  | 76.8345 | 0
|  | 76.9329 | 0 |  |  |  |  | 76.9329 | 0
|  | 77.0312 | 0 |  |  |  |  | 77.0312 | 0
|  | 77.1296 | 0 |  |  |  |  | 77.1296 | 0
|  | 77.228 | 0 |  |  |  |  | 77.228 | 0
|  | 77.3264 | 0 |  |  |  |  | 77.3264 | 0
|  | 77.4248 | 0 |  |  |  |  | 77.4248 | 0
|  | 77.5231 | 0 |  |  |  |  | 77.5231 | 0
|  | 77.6215 | 0 |  |  |  |  | 77.6215 | 0
|  | 77.7199 | 0 |  |  |  |  | 77.7199 | 0
|  | 77.8183 | 0 |  |  |  |  | 77.8183 | 0
|  | 77.9167 | 0 |  |  |  |  | 77.9167 | 0
|  | 78.015 | 0 |  |  |  |  | 78.015 | 0
|  | 78.1134 | 0 |  |  |  |  | 78.1134 | 0
|  | 78.2118 | 0 |  |  |  |  | 78.2118 | 0
|  | 78.3102 | 0 |  |  |  |  | 78.3102 | 0
|  | 78.4086 | 0 |  |  |  |  | 78.4086 | 0
|  | 78.5069 | 0 |  |  |  |  | 78.5069 | 0
|  | 78.6053 | 0 |  |  |  |  | 78.6053 | 0
|  | 78.7037 | 0 |  |  |  |  | 78.7037 | 0
|  | 78.8021 | 0 |  |  |  |  | 78.8021 | 0
|  | 78.9005 | 0 |  |  |  |  | 78.9005 | 0
|  | 78.9988 | 0 |  |  |  |  | 78.9988 | 0
|  | 79.0972 | 0 |  |  |  |  | 79.0972 | 0
|  | 79.1956 | 0 |  |  |  |  | 79.1956 | 0
|  | 79.294 | 0 |  |  |  |  | 79.294 | 0
|  | 79.3924 | 0 |  |  |  |  | 79.3924 | 0
|  | 79.4907 | 0 |  |  |  |  | 79.4907 | 0
|  | 79.5891 | 0 |  |  |  |  | 79.5891 | 0
|  | 79.6875 | 0 |  |  |  |  | 79.6875 | 0
|  | 79.7859 | 0 |  |  |  |  | 79.7859 | 0
|  | 79.8843 | 0 |  |  |  |  | 79.8843 | 0
|  | 79.9826 | 0 |  |  |  |  | 79.9826 | 0
|  | 80.081 | 0 |  |  |  |  | 80.081 | 0
|  | 80.1794 | 0 |  |  |  |  | 80.1794 | 0
|  | 80.2778 | 0 |  |  |  |  | 80.2778 | 0
|  | 80.3762 | 0 |  |  |  |  | 80.3762 | 0
|  | 80.4745 | 0 |  |  |  |  | 80.4745 | 0
|  | 80.5729 | 0 |  |  |  |  | 80.5729 | 0
|  | 80.6713 | 0 |  |  |  |  | 80.6713 | 0
|  | 80.7697 | 0 |  |  |  |  | 80.7697 | 0
|  | 80.868 | 0 |  |  |  |  | 80.8681 | 0
|  | 80.9664 | 0 |  |  |  |  | 80.9664 | 0
|  | 81.0648 | 0 |  |  |  |  | 81.0648 | 0
|  | 81.1632 | 0 |  |  |  |  | 81.1632 | 0
|  | 81.2616 | 0 |  |  |  |  | 81.2616 | 0
|  | 81.36 | 0 |  |  |  |  | 81.36 | 0
|  | 81.4583 | 0 |  |  |  |  | 81.4583 | 0
|  | 81.5567 | 0 |  |  |  |  | 81.5567 | 0
|  | 81.6551 | 0 |  |  |  |  | 81.6551 | 0
|  | 81.7535 | 0 |  |  |  |  | 81.7535 | 0
|  | 81.8519 | 0 |  |  |  |  | 81.8519 | 0
|  | 81.9502 | 0 |  |  |  |  | 81.9502 | 0
|  | 82.0486 | 0 |  |  |  |  | 82.0486 | 0
|  | 82.147 | 0 |  |  |  |  | 82.147 | 0
|  | 82.2454 | 0 |  |  |  |  | 82.2454 | 0
|  | 82.3438 | 0 |  |  |  |  | 82.3438 | 0
|  | 82.4421 | 0 |  |  |  |  | 82.4421 | 0
|  | 82.5405 | 0 |  |  |  |  | 82.5405 | 0
|  | 82.6389 | 0 |  |  |  |  | 82.6389 | 0
|  | 82.7373 | 0 |  |  |  |  | 82.7373 | 0
|  | 82.8356 | 0 |  |  |  |  | 82.8356 | 0
|  | 82.934 | 0 |  |  |  |  | 82.934 | 0
|  | 83.0324 | 0 |  |  |  |  | 83.0324 | 0
|  | 83.1308 | 0 |  |  |  |  | 83.1308 | 0
|  | 83.2292 | 0 |  |  |  |  | 83.2292 | 0
|  | 83.3275 | 0 |  |  |  |  | 83.3275 | 0
|  | 83.4259 | 0 |  |  |  |  | 83.4259 | 0
|  | 83.5243 | 0 |  |  |  |  | 83.5243 | 0
|  | 83.6227 | 0 |  |  |  |  | 83.6227 | 0
|  | 83.7211 | 0 |  |  |  |  | 83.7211 | 0
|  | 83.8194 | 0 |  |  |  |  | 83.8194 | 0
|  | 83.9178 | 0 |  |  |  |  | 83.9178 | 0
|  | 84.0162 | 0 |  |  |  |  | 84.0162 | 0
|  | 84.1146 | 56.6173 |  |  |  |  | 84.1146 | 0
|  | 84.213 | 56.8574 |  |  |  |  | 84.213 | 0
|  | 84.3113 | 57.0574 |  |  |  |  | 84.3113 | 0
|  | 84.4097 | 57.0974 |  |  |  |  | 84.4097 | 0
|  | 84.5081 | 57.2175 |  |  |  |  | 84.5081 | 0
|  | 84.6065 | 58.4178 |  |  |  |  | 84.6065 | 0
|  | 84.7049 | 59.6982 |  |  |  |  | 84.7049 | 0
|  | 84.8032 | 60.5785 |  |  |  |  | 84.8032 | 0
|  | 84.9016 | 61.1387 |  |  |  |  | 84.9016 | 0
|  | 85 | 61.2987 |  |  |  |  | 85 | 0
|  | 85.0984 | 61.1387 |  |  |  |  | 85.0984 | 0
|  | 85.1968 | 60.9786 |  |  |  |  | 85.1968 | 0
|  | 85.2951 | 60.6185 |  |  |  |  | 85.2951 | 0
|  | 85.3935 | 60.1784 |  |  |  |  | 85.3935 | 0
|  | 85.4919 | 60.0183 |  |  |  |  | 85.4919 | 0
|  | 85.5903 | 59.7782 |  |  |  |  | 85.5903 | 0
|  | 85.6887 | 59.2981 |  |  |  |  | 85.6887 | 0
|  | 85.787 | 58.938 |  |  |  |  | 85.787 | 0
|  | 85.8854 | 58.5779 |  |  |  |  | 85.8854 | 0
|  | 85.9838 | 58.0977 |  |  |  |  | 85.9838 | 0
|  | 86.0822 | 57.4575 |  |  |  |  | 86.0822 | 0
|  | 86.1805 | 56.6573 |  |  |  |  | 86.1806 | 0
|  | 86.2789 | 55.4569 |  |  |  |  | 86.2789 | 0
|  | 86.3773 | 54.1765 |  |  |  |  | 86.3773 | 0
|  | 86.4757 | 52.9762 |  |  |  |  | 86.4757 | 0
|  | 86.5741 | 52.0159 |  |  |  |  | 86.5741 | 0
|  | 86.6725 | 51.4557 |  |  |  |  | 86.6725 | 0
|  | 86.7708 | 50.7355 |  |  |  |  | 86.7708 | 0
|  | 86.8692 | 49.4551 |  |  |  |  | 86.8692 | 0
|  | 86.9676 | 46.7343 |  |  |  |  | 86.9676 | 0
|  | 87.066 | 43.9734 |  |  |  |  | 87.066 | 0
|  | 87.1644 | 42.8531 |  |  |  |  | 87.1644 | 0
|  | 87.2627 | 42.3729 |  |  |  |  | 87.2627 | 0
|  | 87.3611 | 24.4475 |  |  |  |  | 87.3611 | 0
|  | 87.4595 | 24.4475 |  |  |  |  | 87.4595 | 0
|  | 87.5579 | 24.4475 |  |  |  |  | 87.5579 | 0
|  | 87.6562 | 0 |  |  |  |  | 87.6562 | 0
|  | 87.7546 | 0 |  |  |  |  | 87.7546 | 0
|  | 87.853 | 0 |  |  |  |  | 87.853 | 0
|  | 87.9514 | 0 |  |  |  |  | 87.9514 | 0
|  | 88.0498 | 0 |  |  |  |  | 88.0498 | 0
|  | 88.1481 | 0 |  |  |  |  | 88.1481 | 0
|  | 88.2465 | 0 |  |  |  |  | 88.2465 | 0
|  | 88.3449 | 0 |  |  |  |  | 88.3449 | 0
|  | 88.4433 | 0 |  |  |  |  | 88.4433 | 0
|  | 88.5417 | 0 |  |  |  |  | 88.5417 | 0
|  | 88.64 | 0 |  |  |  |  | 88.64 | 0
|  | 88.7384 | 0 |  |  |  |  | 88.7384 | 0
|  | 88.8368 | 0 |  |  |  |  | 88.8368 | 0
|  | 88.9352 | 0 |  |  |  |  | 88.9352 | 0
|  | 89.0336 | 0 |  |  |  |  | 89.0336 | 0
|  | 89.1319 | 0 |  |  |  |  | 89.1319 | 0
|  | 89.2303 | 0 |  |  |  |  | 89.2303 | 0
|  | 89.3287 | 0 |  |  |  |  | 89.3287 | 0
|  | 89.4271 | 0 |  |  |  |  | 89.4271 | 0
|  | 89.5255 | 0 |  |  |  |  | 89.5255 | 0
|  | 89.6238 | 0 |  |  |  |  | 89.6238 | 0
|  | 89.7222 | 0 |  |  |  |  | 89.7222 | 0
|  | 89.8206 | 0 |  |  |  |  | 89.8206 | 0
|  | 89.919 | 0 |  |  |  |  | 89.919 | 0
|  | 90.0174 | 0 |  |  |  |  | 90.0174 | 0
|  | 90.1157 | 0 |  |  |  |  | 90.1157 | 0
|  | 90.2141 | 0 |  |  |  |  | 90.2141 | 0
|  | 90.3125 | 0 |  |  |  |  | 90.3125 | 0
|  | 90.4109 | 0 |  |  |  |  | 90.4109 | 0
|  | 90.5093 | 0 |  |  |  |  | 90.5093 | 0
|  | 90.6076 | 0 |  |  |  |  | 90.6076 | 0
|  | 90.706 | 0 |  |  |  |  | 90.706 | 0
|  | 90.8044 | 0 |  |  |  |  | 90.8044 | 0
|  | 90.9028 | 0 |  |  |  |  | 90.9028 | 0
|  | 91.0012 | 0 |  |  |  |  | 91.0012 | 0
|  | 91.0995 | 0 |  |  |  |  | 91.0995 | 0
|  | 91.1979 | 0 |  |  |  |  | 91.1979 | 0
|  | 91.2963 | 0 |  |  |  |  | 91.2963 | 0
|  | 91.3947 | 0 |  |  |  |  | 91.3947 | 0
|  | 91.493 | 0 |  |  |  |  | 91.4931 | 0
|  | 91.5914 | 0 |  |  |  |  | 91.5914 | 0
|  | 91.6898 | 0 |  |  |  |  | 91.6898 | 0
|  | 91.7882 | 0 |  |  |  |  | 91.7882 | 0
|  | 91.8866 | 0 |  |  |  |  | 91.8866 | 0
|  | 91.985 | 0 |  |  |  |  | 91.985 | 0
|  | 92.0833 | 0 |  |  |  |  | 92.0833 | 0
|  | 92.1817 | 0 |  |  |  |  | 92.1817 | 0
|  | 92.2801 | 0 |  |  |  |  | 92.2801 | 0
|  | 92.3785 | 0 |  |  |  |  | 92.3785 | 0
|  | 92.4769 | 0 |  |  |  |  | 92.4769 | 0
|  | 92.5752 | 0 |  |  |  |  | 92.5752 | 0
|  | 92.6736 | 0 |  |  |  |  | 92.6736 | 0
|  | 92.772 | 0 |  |  |  |  | 92.772 | 0
|  | 92.8704 | 0 |  |  |  |  | 92.8704 | 0
|  | 92.9688 | 0 |  |  |  |  | 92.9688 | 0
|  | 93.0671 | 0 |  |  |  |  | 93.0671 | 0
|  | 93.1655 | 0 |  |  |  |  | 93.1655 | 0
|  | 93.2639 | 0 |  |  |  |  | 93.2639 | 0
|  | 93.3623 | 0 |  |  |  |  | 93.3623 | 0
|  | 93.4606 | 0 |  |  |  |  | 93.4606 | 0
|  | 93.559 | 0 |  |  |  |  | 93.559 | 0
|  | 93.6574 | 0 |  |  |  |  | 93.6574 | 0
|  | 93.7558 | 0 |  |  |  |  | 93.7558 | 0
|  | 93.8542 | 0 |  |  |  |  | 93.8542 | 0
|  | 93.9525 | 0 |  |  |  |  | 93.9525 | 0
|  | 94.0509 | 0 |  |  |  |  | 94.0509 | 0
|  | 94.1493 | 0 |  |  |  |  | 94.1493 | 0
|  | 94.2477 | 0 |  |  |  |  | 94.2477 | 0
|  | 94.3461 | 0 |  |  |  |  | 94.3461 | 0
|  | 94.4444 | 0 |  |  |  |  | 94.4444 | 0
|  | 94.5428 | 0 |  |  |  |  | 94.5428 | 0
|  | 94.6412 | 0 |  |  |  |  | 94.6412 | 0
|  | 94.7396 | 0 |  |  |  |  | 94.7396 | 0
|  | 94.838 | 0 |  |  |  |  | 94.838 | 0
|  | 94.9363 | 0 |  |  |  |  | 94.9363 | 0
|  | 95.0347 | 0 |  |  |  |  | 95.0347 | 0
|  | 95.1331 | 0 |  |  |  |  | 95.1331 | 0
|  | 95.2315 | 0 |  |  |  |  | 95.2315 | 0
|  | 95.3299 | 0 |  |  |  |  | 95.3299 | 0
|  | 95.4282 | 0 |  |  |  |  | 95.4282 | 0
|  | 95.5266 | 0 |  |  |  |  | 95.5266 | 0
|  | 95.625 | 0 |  |  |  |  | 95.625 | 0
|  | 95.7234 | 0 |  |  |  |  | 95.7234 | 0
|  | 95.8218 | 0 |  |  |  |  | 95.8218 | 0
|  | 95.9201 | 0 |  |  |  |  | 95.9201 | 0
|  | 96.0185 | 0 |  |  |  |  | 96.0185 | 0
|  | 96.1169 | 0 |  |  |  |  | 96.1169 | 0
|  | 96.2153 | 0 |  |  |  |  | 96.2153 | 0
|  | 96.3137 | 0 |  |  |  |  | 96.3137 | 0
|  | 96.412 | 0 |  |  |  |  | 96.412 | 0
|  | 96.5104 | 0 |  |  |  |  | 96.5104 | 0
|  | 96.6088 | 0 |  |  |  |  | 96.6088 | 0
|  | 96.7072 | 0 |  |  |  |  | 96.7072 | 0
|  | 96.8055 | 0 |  |  |  |  | 96.8056 | 0
|  | 96.9039 | 0 |  |  |  |  | 96.9039 | 0
|  | 97.0023 | 0 |  |  |  |  | 97.0023 | 0
|  | 97.1007 | 0 |  |  |  |  | 97.1007 | 0
|  | 97.1991 | 0 |  |  |  |  | 97.1991 | 0
|  | 97.2974 | 0 |  |  |  |  | 97.2975 | 0
|  | 97.3958 | 0 |  |  |  |  | 97.3958 | 0
|  | 97.4942 | 0 |  |  |  |  | 97.4942 | 0
|  | 97.5926 | 0 |  |  |  |  | 97.5926 | 0
|  | 97.691 | 0 |  |  |  |  | 97.691 | 0
|  | 97.7894 | 0 |  |  |  |  | 97.7894 | 0
|  | 97.8877 | 0 |  |  |  |  | 97.8877 | 0
|  | 97.9861 | 0 |  |  |  |  | 97.9861 | 0
|  | 98.0845 | 0 |  |  |  |  | 98.0845 | 0
|  | 98.1829 | 0 |  |  |  |  | 98.1829 | 0
|  | 98.2812 | 0 |  |  |  |  | 98.2812 | 0
|  | 98.3796 | 0 |  |  |  |  | 98.3796 | 0
|  | 98.478 | 0 |  |  |  |  | 98.478 | 0
|  | 98.5764 | 0 |  |  |  |  | 98.5764 | 0
|  | 98.6748 | 0 |  |  |  |  | 98.6748 | 0
|  | 98.7731 | 0 |  |  |  |  | 98.7731 | 0
|  | 98.8715 | 0 |  |  |  |  | 98.8715 | 0
|  | 98.9699 | 0 |  |  |  |  | 98.9699 | 0
|  | 99.0683 | 0 |  |  |  |  | 99.0683 | 0
|  | 99.1667 | 0 |  |  |  |  | 99.1667 | 0
|  | 99.265 | 0 |  |  |  |  | 99.265 | 0
|  | 99.3634 | 0 |  |  |  |  | 99.3634 | 0
|  | 99.4618 | 0 |  |  |  |  | 99.4618 | 0
|  | 99.5602 | 0 |  |  |  |  | 99.5602 | 0
|  | 99.6586 | 0 |  |  |  |  | 99.6586 | 34.4505
|  | 99.7569 | 0 |  |  |  |  | 99.7569 | 64.5397
|  | 99.8553 | 0 |  |  |  |  | 99.8553 | 68.861
|  | 99.9537 | 0 |  |  |  |  | 99.9537 | 67.6606
|  | 100.0521 | 0 |  |  |  |  | 100.0521 | 67.8207
|  | 100.1505 | 0 |  |  |  |  | 100.1505 | 69.2611
|  | 100.2488 | 0 |  |  |  |  | 100.2488 | 72.9423
|  | 100.3472 | 0 |  |  |  |  | 100.3472 | 72.7022
|  | 100.4456 | 0 |  |  |  |  | 100.4456 | 75.263
|  | 100.544 | 0 |  |  |  |  | 100.544 | 77.1836
|  | 100.6424 | 0 |  |  |  |  | 100.6424 | 77.9838
|  | 100.7407 | 0 |  |  |  |  | 100.7407 | 78.9441
|  | 100.8391 | 56.1771 |  |  |  |  | 100.8391 | 77.7437
|  | 100.9375 | 57.9377 |  |  |  |  | 100.9375 | 78.3839
|  | 101.0359 | 58.2978 |  |  |  |  | 101.0359 | 78.4239
|  | 101.1343 | 57.4175 |  |  |  |  | 101.1343 | 78.664
|  | 101.2326 | 57.7776 |  |  |  |  | 101.2326 | 78.9841
|  | 101.331 | 58.7379 |  |  |  |  | 101.331 | 78.8241
|  | 101.4294 | 59.7782 |  |  |  |  | 101.4294 | 78.2639
|  | 101.5278 | 60.8186 |  |  |  |  | 101.5278 | 77.3036
|  | 101.6262 | 60.9786 |  |  |  |  | 101.6262 | 84.3457
|  | 101.7245 | 60.8986 |  |  |  |  | 101.7245 | 91.3879
|  | 101.8229 | 60.5385 |  |  |  |  | 101.8229 | 30.5693
|  | 101.9213 | 59.7382 |  |  |  |  | 101.9213 | 30.5693
|  | 102.0197 | 59.2581 |  |  |  |  | 102.0197 | 30.5693
|  | 102.118 | 59.4581 |  |  |  |  | 102.1181 | 0
|  | 102.2164 | 60.8986 |  |  |  |  | 102.2164 | 0
|  | 102.3148 | 61.2187 |  |  |  |  | 102.3148 | 0
|  | 102.4132 | 60.1784 |  |  |  |  | 102.4132 | 0
|  | 102.5116 | 59.2981 |  |  |  |  | 102.5116 | 0
|  | 102.6099 | 59.7382 |  |  |  |  | 102.61 | 0
|  | 102.7083 | 60.4584 |  |  |  |  | 102.7083 | 0
|  | 102.8067 | 60.7385 |  |  |  |  | 102.8067 | 0
|  | 102.9051 | 60.6585 |  |  |  |  | 102.9051 | 0
|  | 103.0035 | 58.8179 |  |  |  |  | 103.0035 | 0
|  | 103.1019 | 55.9371 |  |  |  |  | 103.1019 | 0
|  | 103.2002 | 53.6964 |  |  |  |  | 103.2002 | 0
|  | 103.2986 | 53.0562 |  |  |  |  | 103.2986 | 0
|  | 103.397 | 52.0959 |  |  |  |  | 103.397 | 0
|  | 103.4954 | 51.0556 |  |  |  |  | 103.4954 | 0
|  | 103.5938 | 50.4954 |  |  |  |  | 103.5938 | 0
|  | 103.6921 | 48.6548 |  |  |  |  | 103.6921 | 0
|  | 103.7905 | 45.6139 |  |  |  |  | 103.7905 | 0
|  | 103.8889 | 43.6933 |  |  |  |  | 103.8889 | 0
|  | 103.9873 | 41.6127 |  |  |  |  | 103.9873 | 0
|  | 104.0856 | 41.6127 |  |  |  |  | 104.0856 | 0
|  | 104.184 | 41.6127 |  |  |  |  | 104.184 | 0
|  | 104.2824 | 0 |  |  |  |  | 104.2824 | 0
|  | 104.3808 | 0 |  |  |  |  | 104.3808 | 0
|  | 104.4792 | 0 |  |  |  |  | 104.4792 | 0
|  | 104.5775 | 0 |  |  |  |  | 104.5775 | 0
|  | 104.6759 | 0 |  |  |  |  | 104.6759 | 0
|  | 104.7743 | 0 |  |  |  |  | 104.7743 | 0
|  | 104.8727 | 0 |  |  |  |  | 104.8727 | 0
|  | 104.9711 | 0 |  |  |  |  | 104.9711 | 0
|  | 105.0694 | 0 |  |  |  |  | 105.0694 | 0
|  | 105.1678 | 0 |  |  |  |  | 105.1678 | 0
|  | 105.2662 | 0 |  |  |  |  | 105.2662 | 0
|  | 105.3646 | 0 |  |  |  |  | 105.3646 | 0
|  | 105.463 | 0 |  |  |  |  | 105.463 | 0
|  | 105.5613 | 0 |  |  |  |  | 105.5613 | 0
|  | 105.6597 | 0 |  |  |  |  | 105.6597 | 0
|  | 105.7581 | 0 |  |  |  |  | 105.7581 | 0
|  | 105.8565 | 0 |  |  |  |  | 105.8565 | 0
|  | 105.9549 | 0 |  |  |  |  | 105.9549 | 0
|  | 106.0532 | 0 |  |  |  |  | 106.0532 | 0
|  | 106.1516 | 0 |  |  |  |  | 106.1516 | 0
|  | 106.25 | 0 |  |  |  |  | 106.25 | 0
|  | 106.3484 | 0 |  |  |  |  | 106.3484 | 0
|  | 106.4468 | 0 |  |  |  |  | 106.4468 | 0
|  | 106.5451 | 0 |  |  |  |  | 106.5451 | 0
|  | 106.6435 | 0 |  |  |  |  | 106.6435 | 0
|  | 106.7419 | 0 |  |  |  |  | 106.7419 | 0
|  | 106.8403 | 0 |  |  |  |  | 106.8403 | 0
|  | 106.9387 | 0 |  |  |  |  | 106.9387 | 0
|  | 107.037 | 0 |  |  |  |  | 107.037 | 0
|  | 107.1354 | 0 |  |  |  |  | 107.1354 | 0
|  | 107.2338 | 0 |  |  |  |  | 107.2338 | 0
|  | 107.3322 | 0 |  |  |  |  | 107.3322 | 0
|  | 107.4305 | 0 |  |  |  |  | 107.4306 | 0
|  | 107.5289 | 0 |  |  |  |  | 107.5289 | 0
|  | 107.6273 | 0 |  |  |  |  | 107.6273 | 0
|  | 107.7257 | 0 |  |  |  |  | 107.7257 | 0
|  | 107.8241 | 0 |  |  |  |  | 107.8241 | 0
|  | 107.9224 | 0 |  |  |  |  | 107.9225 | 0
|  | 108.0208 | 0 |  |  |  |  | 108.0208 | 0
|  | 108.1192 | 0 |  |  |  |  | 108.1192 | 0
|  | 108.2176 | 0 |  |  |  |  | 108.2176 | 0
|  | 108.316 | 0 |  |  |  |  | 108.316 | 0
|  | 108.4144 | 0 |  |  |  |  | 108.4144 | 0
|  | 108.5127 | 0 |  |  |  |  | 108.5127 | 0
|  | 108.6111 | 0 |  |  |  |  | 108.6111 | 0
|  | 108.7095 | 0 |  |  |  |  | 108.7095 | 0
|  | 108.8079 | 0 |  |  |  |  | 108.8079 | 0
|  | 108.9062 | 0 |  |  |  |  | 108.9062 | 0
|  | 109.0046 | 0 |  |  |  |  | 109.0046 | 0
|  | 109.103 | 0 |  |  |  |  | 109.103 | 0
|  | 109.2014 | 0 |  |  |  |  | 109.2014 | 0
|  | 109.2998 | 0 |  |  |  |  | 109.2998 | 0
|  | 109.3981 | 0 |  |  |  |  | 109.3981 | 0
|  | 109.4965 | 0 |  |  |  |  | 109.4965 | 0
|  | 109.5949 | 0 |  |  |  |  | 109.5949 | 0
|  | 109.6933 | 0 |  |  |  |  | 109.6933 | 0
|  | 109.7917 | 0 |  |  |  |  | 109.7917 | 0
|  | 109.89 | 0 |  |  |  |  | 109.89 | 0
|  | 109.9884 | 0 |  |  |  |  | 109.9884 | 0
|  | 110.0868 | 0 |  |  |  |  | 110.0868 | 0
|  | 110.1852 | 0 |  |  |  |  | 110.1852 | 0
|  | 110.2836 | 0 |  |  |  |  | 110.2836 | 0
|  | 110.3819 | 0 |  |  |  |  | 110.3819 | 0
|  | 110.4803 | 0 |  |  |  |  | 110.4803 | 0
|  | 110.5787 | 0 |  |  |  |  | 110.5787 | 0
|  | 110.6771 | 0 |  |  |  |  | 110.6771 | 0
|  | 110.7755 | 0 |  |  |  |  | 110.7755 | 0
|  | 110.8738 | 0 |  |  |  |  | 110.8738 | 0
|  | 110.9722 | 0 |  |  |  |  | 110.9722 | 0
|  | 111.0706 | 0 |  |  |  |  | 111.0706 | 0
|  | 111.169 | 0 |  |  |  |  | 111.169 | 0
|  | 111.2674 | 0 |  |  |  |  | 111.2674 | 0
|  | 111.3657 | 0 |  |  |  |  | 111.3657 | 0
|  | 111.4641 | 0 |  |  |  |  | 111.4641 | 0
|  | 111.5625 | 0 |  |  |  |  | 111.5625 | 0
|  | 111.6609 | 0 |  |  |  |  | 111.6609 | 0
|  | 111.7593 | 0 |  |  |  |  | 111.7593 | 0
|  | 111.8576 | 0 |  |  |  |  | 111.8576 | 0
|  | 111.956 | 0 |  |  |  |  | 111.956 | 0
|  | 112.0544 | 0 |  |  |  |  | 112.0544 | 0
|  | 112.1528 | 0 |  |  |  |  | 112.1528 | 0
|  | 112.2512 | 0 |  |  |  |  | 112.2512 | 0
|  | 112.3495 | 0 |  |  |  |  | 112.3495 | 0
|  | 112.4479 | 0 |  |  |  |  | 112.4479 | 0
|  | 112.5463 | 0 |  |  |  |  | 112.5463 | 0
|  | 112.6447 | 0 |  |  |  |  | 112.6447 | 0
|  | 112.743 | 0 |  |  |  |  | 112.7431 | 0
|  | 112.8414 | 0 |  |  |  |  | 112.8414 | 0
|  | 112.9398 | 0 |  |  |  |  | 112.9398 | 0
|  | 113.0382 | 0 |  |  |  |  | 113.0382 | 0
|  | 113.1366 | 0 |  |  |  |  | 113.1366 | 0
|  | 113.2349 | 0 |  |  |  |  | 113.235 | 0
|  | 113.3333 | 0 |  |  |  |  | 113.3333 | 0
|  | 113.4317 | 0 |  |  |  |  | 113.4317 | 0
|  | 113.5301 | 0 |  |  |  |  | 113.5301 | 0
|  | 113.6285 | 0 |  |  |  |  | 113.6285 | 0
|  | 113.7269 | 0 |  |  |  |  | 113.7269 | 0
|  | 113.8252 | 0 |  |  |  |  | 113.8252 | 0
|  | 113.9236 | 0 |  |  |  |  | 113.9236 | 0
|  | 114.022 | 0 |  |  |  |  | 114.022 | 0
|  | 114.1204 | 0 |  |  |  |  | 114.1204 | 0
|  | 114.2188 | 0 |  |  |  |  | 114.2188 | 0
|  | 114.3171 | 0 |  |  |  |  | 114.3171 | 0
|  | 114.4155 | 0 |  |  |  |  | 114.4155 | 0
|  | 114.5139 | 0 |  |  |  |  | 114.5139 | 0
|  | 114.6123 | 0 |  |  |  |  | 114.6123 | 0
|  | 114.7106 | 0 |  |  |  |  | 114.7106 | 0
|  | 114.809 | 0 |  |  |  |  | 114.809 | 0
|  | 114.9074 | 0 |  |  |  |  | 114.9074 | 0
|  | 115.0058 | 0 |  |  |  |  | 115.0058 | 0
|  | 115.1042 | 0 |  |  |  |  | 115.1042 | 0
|  | 115.2025 | 0 |  |  |  |  | 115.2025 | 0
|  | 115.3009 | 0 |  |  |  |  | 115.3009 | 0
|  | 115.3993 | 0 |  |  |  |  | 115.3993 | 0
|  | 115.4977 | 0 |  |  |  |  | 115.4977 | 0
|  | 115.5961 | 0 |  |  |  |  | 115.5961 | 0
|  | 115.6944 | 0 |  |  |  |  | 115.6944 | 0
|  | 115.7928 | 0 |  |  |  |  | 115.7928 | 0
|  | 115.8912 | 0 |  |  |  |  | 115.8912 | 0
|  | 115.9896 | 0 |  |  |  |  | 115.9896 | 0
|  | 116.088 | 0 |  |  |  |  | 116.088 | 0
|  | 116.1863 | 0 |  |  |  |  | 116.1863 | 0
|  | 116.2847 | 0 |  |  |  |  | 116.2847 | 0
|  | 116.3831 | 0 |  |  |  |  | 116.3831 | 0
|  | 116.4815 | 0 |  |  |  |  | 116.4815 | 0
|  | 116.5799 | 0 |  |  |  |  | 116.5799 | 0
|  | 116.6782 | 0 |  |  |  |  | 116.6782 | 0
|  | 116.7766 | 0 |  |  |  |  | 116.7766 | 0
|  | 116.875 | 0 |  |  |  |  | 116.875 | 0
|  | 116.9734 | 0 |  |  |  |  | 116.9734 | 0
|  | 117.0718 | 0 |  |  |  |  | 117.0718 | 0
|  | 117.1701 | 0 |  |  |  |  | 117.1701 | 0
|  | 117.2685 | 0 |  |  |  |  | 117.2685 | 0
|  | 117.3669 | 0 |  |  |  |  | 117.3669 | 0
|  | 117.4653 | 0 |  |  |  |  | 117.4653 | 0
|  | 117.5637 | 56.0171 |  |  |  |  | 117.5637 | 0
|  | 117.662 | 58.0977 |  |  |  |  | 117.662 | 0
|  | 117.7604 | 58.898 |  |  |  |  | 117.7604 | 0
|  | 117.8588 | 59.2581 |  |  |  |  | 117.8588 | 0
|  | 117.9572 | 59.5382 |  |  |  |  | 117.9572 | 0
|  | 118.0555 | 60.7385 |  |  |  |  | 118.0556 | 0
|  | 118.1539 | 62.4991 |  |  |  |  | 118.1539 | 0
|  | 118.2523 | 61.9789 |  |  |  |  | 118.2523 | 0
|  | 118.3507 | 60.8586 |  |  |  |  | 118.3507 | 0
|  | 118.4491 | 61.2987 |  |  |  |  | 118.4491 | 0
|  | 118.5474 | 62.139 |  |  |  |  | 118.5475 | 0
|  | 118.6458 | 61.8189 |  |  |  |  | 118.6458 | 0
|  | 118.7442 | 61.2987 |  |  |  |  | 118.7442 | 0
|  | 118.8426 | 62.0189 |  |  |  |  | 118.8426 | 0
|  | 118.941 | 62.339 |  |  |  |  | 118.941 | 0
|  | 119.0393 | 61.8189 |  |  |  |  | 119.0394 | 0
|  | 119.1377 | 60.8586 |  |  |  |  | 119.1377 | 0
|  | 119.2361 | 60.4184 |  |  |  |  | 119.2361 | 0
|  | 119.3345 | 61.0586 |  |  |  |  | 119.3345 | 0
|  | 119.4329 | 60.5785 |  |  |  |  | 119.4329 | 0
|  | 119.5312 | 59.1781 |  |  |  |  | 119.5312 | 0
|  | 119.6296 | 58.5379 |  |  |  |  | 119.6296 | 0
|  | 119.728 | 57.7376 |  |  |  |  | 119.728 | 0
|  | 119.8264 | 56.3772 |  |  |  |  | 119.8264 | 0
|  | 119.9248 | 54.7767 |  |  |  |  | 119.9248 | 0
|  | 120.0231 | 53.4163 |  |  |  |  | 120.0231 | 0
|  | 120.1215 | 51.9759 |  |  |  |  | 120.1215 | 0
|  | 120.2199 | 51.0156 |  |  |  |  | 120.2199 | 0
|  | 120.3183 | 50.6955 |  |  |  |  | 120.3183 | 0
|  | 120.4167 | 50.5754 |  |  |  |  | 120.4167 | 0
|  | 120.515 | 49.5751 |  |  |  |  | 120.515 | 0
|  | 120.6134 | 47.4945 |  |  |  |  | 120.6134 | 0
|  | 120.7118 | 40.7724 |  |  |  |  | 120.7118 | 0
|  | 120.8102 | 40.7724 |  |  |  |  | 120.8102 | 0
|  | 120.9086 | 40.7724 |  |  |  |  | 120.9086 | 0
|  | 121.0069 | 0 |  |  |  |  | 121.0069 | 0
|  | 121.1053 | 0 |  |  |  |  | 121.1053 | 0
|  | 121.2037 | 0 |  |  |  |  | 121.2037 | 0
|  | 121.3021 | 0 |  |  |  |  | 121.3021 | 0
|  | 121.4005 | 0 |  |  |  |  | 121.4005 | 0
|  | 121.4988 | 0 |  |  |  |  | 121.4988 | 0
|  | 121.5972 | 0 |  |  |  |  | 121.5972 | 0
|  | 121.6956 | 0 |  |  |  |  | 121.6956 | 0
|  | 121.794 | 0 |  |  |  |  | 121.794 | 0
|  | 121.8924 | 0 |  |  |  |  | 121.8924 | 0
|  | 121.9907 | 0 |  |  |  |  | 121.9907 | 0
|  | 122.0891 | 0 |  |  |  |  | 122.0891 | 0
|  | 122.1875 | 0 |  |  |  |  | 122.1875 | 0
|  | 122.2859 | 0 |  |  |  |  | 122.2859 | 0
|  | 122.3843 | 0 |  |  |  |  | 122.3843 | 0
|  | 122.4826 | 0 |  |  |  |  | 122.4826 | 0
|  | 122.581 | 0 |  |  |  |  | 122.581 | 0
|  | 122.6794 | 0 |  |  |  |  | 122.6794 | 0
|  | 122.7778 | 0 |  |  |  |  | 122.7778 | 0
|  | 122.8762 | 0 |  |  |  |  | 122.8762 | 0
|  | 122.9745 | 0 |  |  |  |  | 122.9745 | 0
|  | 123.0729 | 0 |  |  |  |  | 123.0729 | 0
|  | 123.1713 | 0 |  |  |  |  | 123.1713 | 0
|  | 123.2697 | 0 |  |  |  |  | 123.2697 | 0
|  | 123.368 | 0 |  |  |  |  | 123.3681 | 0
|  | 123.4664 | 0 |  |  |  |  | 123.4664 | 0
|  | 123.5648 | 0 |  |  |  |  | 123.5648 | 0
|  | 123.6632 | 0 |  |  |  |  | 123.6632 | 0
|  | 123.7616 | 0 |  |  |  |  | 123.7616 | 0
|  | 123.8599 | 0 |  |  |  |  | 123.86 | 0
|  | 123.9583 | 0 |  |  |  |  | 123.9583 | 0
|  | 124.0567 | 0 |  |  |  |  | 124.0567 | 0
|  | 124.1551 | 0 |  |  |  |  | 124.1551 | 0
|  | 124.2535 | 0 |  |  |  |  | 124.2535 | 0
|  | 124.3518 | 0 |  |  |  |  | 124.3519 | 0
|  | 124.4502 | 0 |  |  |  |  | 124.4502 | 0
|  | 124.5486 | 0 |  |  |  |  | 124.5486 | 0
|  | 124.647 | 0 |  |  |  |  | 124.647 | 0
|  | 124.7454 | 0 |  |  |  |  | 124.7454 | 0
|  | 124.8438 | 0 |  |  |  |  | 124.8438 | 0
|  | 124.9421 | 0 |  |  |  |  | 124.9421 | 0
|  | 125.0405 | 0 |  |  |  |  | 125.0405 | 0
|  | 125.1389 | 0 |  |  |  |  | 125.1389 | 0
|  | 125.2373 | 0 |  |  |  |  | 125.2373 | 0
|  | 125.3356 | 0 |  |  |  |  | 125.3356 | 0
|  | 125.434 | 0 |  |  |  |  | 125.434 | 0
|  | 125.5324 | 0 |  |  |  |  | 125.5324 | 0
|  | 125.6308 | 0 |  |  |  |  | 125.6308 | 0
|  | 125.7292 | 0 |  |  |  |  | 125.7292 | 0
|  | 125.8275 | 0 |  |  |  |  | 125.8275 | 0
|  | 125.9259 | 0 |  |  |  |  | 125.9259 | 0
|  | 126.0243 | 0 |  |  |  |  | 126.0243 | 0
|  | 126.1227 | 0 |  |  |  |  | 126.1227 | 0
|  | 126.2211 | 0 |  |  |  |  | 126.2211 | 0
|  | 126.3194 | 0 |  |  |  |  | 126.3194 | 0
|  | 126.4178 | 0 |  |  |  |  | 126.4178 | 0
|  | 126.5162 | 0 |  |  |  |  | 126.5162 | 0
|  | 126.6146 | 0 |  |  |  |  | 126.6146 | 0
|  | 126.713 | 0 |  |  |  |  | 126.713 | 0
|  | 126.8113 | 0 |  |  |  |  | 126.8113 | 0
|  | 126.9097 | 0 |  |  |  |  | 126.9097 | 0
|  | 127.0081 | 0 |  |  |  |  | 127.0081 | 0
|  | 127.1065 | 0 |  |  |  |  | 127.1065 | 0
|  | 127.2049 | 0 |  |  |  |  | 127.2049 | 0
|  | 127.3032 | 0 |  |  |  |  | 127.3032 | 0
|  | 127.4016 | 0 |  |  |  |  | 127.4016 | 0
|  | 127.5 | 0 |  |  |  |  | 127.5 | 0
|  | 127.5984 | 0 |  |  |  |  | 127.5984 | 0
|  | 127.6968 | 0 |  |  |  |  | 127.6968 | 0
|  | 127.7951 | 0 |  |  |  |  | 127.7951 | 0
|  | 127.8935 | 0 |  |  |  |  | 127.8935 | 0
|  | 127.9919 | 0 |  |  |  |  | 127.9919 | 0
|  | 128.0903 | 0 |  |  |  |  | 128.0903 | 0
|  | 128.1887 | 0 |  |  |  |  | 128.1887 | 0
|  | 128.287 | 0 |  |  |  |  | 128.287 | 0
|  | 128.3854 | 0 |  |  |  |  | 128.3854 | 0
|  | 128.4838 | 0 |  |  |  |  | 128.4838 | 0
|  | 128.5822 | 0 |  |  |  |  | 128.5822 | 0
|  | 128.6806 | 0 |  |  |  |  | 128.6806 | 0
|  | 128.7789 | 0 |  |  |  |  | 128.7789 | 0
|  | 128.8773 | 0 |  |  |  |  | 128.8773 | 0
|  | 128.9757 | 0 |  |  |  |  | 128.9757 | 0
|  | 129.0741 | 0 |  |  |  |  | 129.0741 | 0
|  | 129.1725 | 0 |  |  |  |  | 129.1725 | 0
|  | 129.2708 | 0 |  |  |  |  | 129.2708 | 0
|  | 129.3692 | 0 |  |  |  |  | 129.3692 | 0
|  | 129.4676 | 0 |  |  |  |  | 129.4676 | 0
|  | 129.566 | 0 |  |  |  |  | 129.566 | 0
|  | 129.6644 | 0 |  |  |  |  | 129.6644 | 0
|  | 129.7627 | 0 |  |  |  |  | 129.7627 | 0
|  | 129.8611 | 0 |  |  |  |  | 129.8611 | 0
|  | 129.9595 | 0 |  |  |  |  | 129.9595 | 0
|  | 130.0579 | 0 |  |  |  |  | 130.0579 | 0
|  | 130.1562 | 0 |  |  |  |  | 130.1562 | 0
|  | 130.2546 | 0 |  |  |  |  | 130.2546 | 0
|  | 130.353 | 0 |  |  |  |  | 130.353 | 0
|  | 130.4514 | 0 |  |  |  |  | 130.4514 | 0
|  | 130.5498 | 0 |  |  |  |  | 130.5498 | 0
|  | 130.6481 | 0 |  |  |  |  | 130.6481 | 0
|  | 130.7465 | 0 |  |  |  |  | 130.7465 | 0
|  | 130.8449 | 0 |  |  |  |  | 130.8449 | 0
|  | 130.9433 | 0 |  |  |  |  | 130.9433 | 0
|  | 131.0417 | 0 |  |  |  |  | 131.0417 | 0
|  | 131.14 | 0 |  |  |  |  | 131.14 | 0
|  | 131.2384 | 0 |  |  |  |  | 131.2384 | 0
|  | 131.3368 | 0 |  |  |  |  | 131.3368 | 0
|  | 131.4352 | 0 |  |  |  |  | 131.4352 | 0
|  | 131.5336 | 0 |  |  |  |  | 131.5336 | 0
|  | 131.6319 | 0 |  |  |  |  | 131.6319 | 0
|  | 131.7303 | 0 |  |  |  |  | 131.7303 | 0
|  | 131.8287 | 0 |  |  |  |  | 131.8287 | 0
|  | 131.9271 | 0 |  |  |  |  | 131.9271 | 0
|  | 132.0255 | 0 |  |  |  |  | 132.0255 | 0
|  | 132.1238 | 0 |  |  |  |  | 132.1238 | 0
|  | 132.2222 | 0 |  |  |  |  | 132.2222 | 0
|  | 132.3206 | 0 |  |  |  |  | 132.3206 | 0
|  | 132.419 | 0 |  |  |  |  | 132.419 | 0
|  | 132.5174 | 0 |  |  |  |  | 132.5174 | 0
|  | 132.6157 | 0 |  |  |  |  | 132.6157 | 0
|  | 132.7141 | 0 |  |  |  |  | 132.7141 | 0
|  | 132.8125 | 0 |  |  |  |  | 132.8125 | 0
|  | 132.9109 | 0 |  |  |  |  | 132.9109 | 0
|  | 133.0093 | 0 |  |  |  |  | 133.0093 | 0
|  | 133.1076 | 0 |  |  |  |  | 133.1076 | 35.6909
|  | 133.206 | 0 |  |  |  |  | 133.206 | 73.4624
|  | 133.3044 | 0 |  |  |  |  | 133.3044 | 78.584
|  | 133.4028 | 0 |  |  |  |  | 133.4028 | 78.4639
|  | 133.5012 | 0 |  |  |  |  | 133.5012 | 79.4242
|  | 133.5995 | 0 |  |  |  |  | 133.5995 | 78.704
|  | 133.6979 | 0 |  |  |  |  | 133.6979 | 77.5037
|  | 133.7963 | 0 |  |  |  |  | 133.7963 | 78.624
|  | 133.8947 | 0 |  |  |  |  | 133.8947 | 79.1442
|  | 133.9931 | 0 |  |  |  |  | 133.9931 | 79.0241
|  | 134.0914 | 0 |  |  |  |  | 134.0914 | 78.504
|  | 134.1898 | 0 |  |  |  |  | 134.1898 | 79.2242
|  | 134.2882 | 58.7779 |  |  |  |  | 134.2882 | 79.9444
|  | 134.3866 | 59.8183 |  |  |  |  | 134.3866 | 80.1445
|  | 134.485 | 60.1384 |  |  |  |  | 134.485 | 80.3045
|  | 134.5833 | 60.6585 |  |  |  |  | 134.5833 | 80.4245
|  | 134.6817 | 60.8186 |  |  |  |  | 134.6817 | 80.5846
|  | 134.7801 | 61.0186 |  |  |  |  | 134.7801 | 80.9847
|  | 134.8785 | 61.7388 |  |  |  |  | 134.8785 | 80.8647
|  | 134.9769 | 62.179 |  |  |  |  | 134.9769 | 81.0647
|  | 135.0752 | 62.299 |  |  |  |  | 135.0752 | 81.3848
|  | 135.1736 | 62.6591 |  |  |  |  | 135.1736 | 81.2248
|  | 135.272 | 62.7792 |  |  |  |  | 135.272 | 81.2248
|  | 135.3704 | 62.7391 |  |  |  |  | 135.3704 | 81.2248
|  | 135.4688 | 63.0993 |  |  |  |  | 135.4688 | 0
|  | 135.5671 | 63.4594 |  |  |  |  | 135.5671 | 0
|  | 135.6655 | 63.7395 |  |  |  |  | 135.6655 | 0
|  | 135.7639 | 63.1793 |  |  |  |  | 135.7639 | 0
|  | 135.8623 | 61.8189 |  |  |  |  | 135.8623 | 0
|  | 135.9606 | 60.8586 |  |  |  |  | 135.9606 | 0
|  | 136.059 | 60.4184 |  |  |  |  | 136.059 | 0
|  | 136.1574 | 60.1384 |  |  |  |  | 136.1574 | 0
|  | 136.2558 | 60.1784 |  |  |  |  | 136.2558 | 0
|  | 136.3542 | 59.6582 |  |  |  |  | 136.3542 | 0
|  | 136.4525 | 58.0577 |  |  |  |  | 136.4525 | 0
|  | 136.5509 | 56.5373 |  |  |  |  | 136.5509 | 0
|  | 136.6493 | 55.817 |  |  |  |  | 136.6493 | 0
|  | 136.7477 | 54.5366 |  |  |  |  | 136.7477 | 0
|  | 136.8461 | 53.0962 |  |  |  |  | 136.8461 | 0
|  | 136.9444 | 52.1359 |  |  |  |  | 136.9444 | 0
|  | 137.0428 | 51.2556 |  |  |  |  | 137.0428 | 0
|  | 137.1412 | 49.3351 |  |  |  |  | 137.1412 | 0
|  | 137.2396 | 46.4542 |  |  |  |  | 137.2396 | 0
|  | 137.338 | 43.5733 |  |  |  |  | 137.338 | 0
|  | 137.4363 | 41.6527 |  |  |  |  | 137.4363 | 0
|  | 137.5347 | 41.6527 |  |  |  |  | 137.5347 | 0
|  | 137.6331 | 41.6527 |  |  |  |  | 137.6331 | 0
|  | 137.7315 | 0 |  |  |  |  | 137.7315 | 0
|  | 137.8299 | 0 |  |  |  |  | 137.8299 | 0
|  | 137.9282 | 0 |  |  |  |  | 137.9282 | 0
|  | 138.0266 | 0 |  |  |  |  | 138.0266 | 0
|  | 138.125 | 0 |  |  |  |  | 138.125 | 0
|  | 138.2234 | 0 |  |  |  |  | 138.2234 | 0
|  | 138.3218 | 0 |  |  |  |  | 138.3218 | 0
|  | 138.4201 | 0 |  |  |  |  | 138.4201 | 0
|  | 138.5185 | 0 |  |  |  |  | 138.5185 | 0
|  | 138.6169 | 0 |  |  |  |  | 138.6169 | 0
|  | 138.7153 | 0 |  |  |  |  | 138.7153 | 0
|  | 138.8137 | 0 |  |  |  |  | 138.8137 | 0
|  | 138.912 | 0 |  |  |  |  | 138.912 | 0
|  | 139.0104 | 0 |  |  |  |  | 139.0104 | 0
|  | 139.1088 | 0 |  |  |  |  | 139.1088 | 0
|  | 139.2072 | 0 |  |  |  |  | 139.2072 | 0
|  | 139.3056 | 0 |  |  |  |  | 139.3056 | 0
|  | 139.4039 | 0 |  |  |  |  | 139.4039 | 0
|  | 139.5023 | 0 |  |  |  |  | 139.5023 | 0
|  | 139.6007 | 0 |  |  |  |  | 139.6007 | 0
|  | 139.6991 | 0 |  |  |  |  | 139.6991 | 0
|  | 139.7975 | 0 |  |  |  |  | 139.7975 | 0
|  | 139.8958 | 0 |  |  |  |  | 139.8958 | 0
|  | 139.9942 | 0 |  |  |  |  | 139.9942 | 0
|  | 140.0926 | 0 |  |  |  |  | 140.0926 | 0
|  | 140.191 | 0 |  |  |  |  | 140.191 | 0
|  | 140.2894 | 0 |  |  |  |  | 140.2894 | 0
|  | 140.3877 | 0 |  |  |  |  | 140.3877 | 0
|  | 140.4861 | 0 |  |  |  |  | 140.4861 | 0
|  | 140.5845 | 0 |  |  |  |  | 140.5845 | 0
|  | 140.6829 | 0 |  |  |  |  | 140.6829 | 0
|  | 140.7812 | 0 |  |  |  |  | 140.7812 | 0
|  | 140.8796 | 0 |  |  |  |  | 140.8796 | 0
|  | 140.978 | 0 |  |  |  |  | 140.978 | 0
|  | 141.0764 | 0 |  |  |  |  | 141.0764 | 0
|  | 141.1748 | 0 |  |  |  |  | 141.1748 | 0
|  | 141.2731 | 0 |  |  |  |  | 141.2731 | 0
|  | 141.3715 | 0 |  |  |  |  | 141.3715 | 0
|  | 141.4699 | 0 |  |  |  |  | 141.4699 | 0
|  | 141.5683 | 0 |  |  |  |  | 141.5683 | 0
|  | 141.6667 | 0 |  |  |  |  | 141.6667 | 0
|  | 141.765 | 0 |  |  |  |  | 141.765 | 0
|  | 141.8634 | 0 |  |  |  |  | 141.8634 | 0
|  | 141.9618 | 0 |  |  |  |  | 141.9618 | 0
|  | 142.0602 | 0 |  |  |  |  | 142.0602 | 0
|  | 142.1586 | 0 |  |  |  |  | 142.1586 | 0
|  | 142.2569 | 0 |  |  |  |  | 142.2569 | 0
|  | 142.3553 | 0 |  |  |  |  | 142.3553 | 0
|  | 142.4537 | 0 |  |  |  |  | 142.4537 | 0
|  | 142.5521 | 0 |  |  |  |  | 142.5521 | 0
|  | 142.6505 | 0 |  |  |  |  | 142.6505 | 0
|  | 142.7488 | 0 |  |  |  |  | 142.7488 | 0
|  | 142.8472 | 0 |  |  |  |  | 142.8472 | 0
|  | 142.9456 | 0 |  |  |  |  | 142.9456 | 0
|  | 143.044 | 0 |  |  |  |  | 143.044 | 0
|  | 143.1424 | 0 |  |  |  |  | 143.1424 | 0
|  | 143.2407 | 0 |  |  |  |  | 143.2407 | 0
|  | 143.3391 | 0 |  |  |  |  | 143.3391 | 0
|  | 143.4375 | 0 |  |  |  |  | 143.4375 | 0
|  | 143.5359 | 0 |  |  |  |  | 143.5359 | 0
|  | 143.6343 | 0 |  |  |  |  | 143.6343 | 0
|  | 143.7326 | 0 |  |  |  |  | 143.7326 | 0
|  | 143.831 | 0 |  |  |  |  | 143.831 | 0
|  | 143.9294 | 0 |  |  |  |  | 143.9294 | 0
|  | 144.0278 | 0 |  |  |  |  | 144.0278 | 0
|  | 144.1262 | 0 |  |  |  |  | 144.1262 | 0
|  | 144.2245 | 0 |  |  |  |  | 144.2245 | 0
|  | 144.3229 | 0 |  |  |  |  | 144.3229 | 0
|  | 144.4213 | 0 |  |  |  |  | 144.4213 | 0
|  | 144.5197 | 0 |  |  |  |  | 144.5197 | 0
|  | 144.6181 | 0 |  |  |  |  | 144.6181 | 0
|  | 144.7164 | 0 |  |  |  |  | 144.7164 | 0
|  | 144.8148 | 0 |  |  |  |  | 144.8148 | 0
|  | 144.9132 | 0 |  |  |  |  | 144.9132 | 0
|  | 145.0116 | 0 |  |  |  |  | 145.0116 | 0
|  | 145.11 | 0 |  |  |  |  | 145.11 | 0
|  | 145.2083 | 0 |  |  |  |  | 145.2083 | 0
|  | 145.3067 | 0 |  |  |  |  | 145.3067 | 0
|  | 145.4051 | 0 |  |  |  |  | 145.4051 | 0
|  | 145.5035 | 0 |  |  |  |  | 145.5035 | 0
|  | 145.6019 | 0 |  |  |  |  | 145.6019 | 0
|  | 145.7002 | 0 |  |  |  |  | 145.7002 | 0
|  | 145.7986 | 0 |  |  |  |  | 145.7986 | 0
|  | 145.897 | 0 |  |  |  |  | 145.897 | 0
|  | 145.9954 | 0 |  |  |  |  | 145.9954 | 0
|  | 146.0938 | 0 |  |  |  |  | 146.0938 | 0
|  | 146.1921 | 0 |  |  |  |  | 146.1921 | 0
|  | 146.2905 | 0 |  |  |  |  | 146.2905 | 0
|  | 146.3889 | 0 |  |  |  |  | 146.3889 | 0
|  | 146.4873 | 0 |  |  |  |  | 146.4873 | 0
|  | 146.5856 | 0 |  |  |  |  | 146.5856 | 0
|  | 146.684 | 0 |  |  |  |  | 146.684 | 0
|  | 146.7824 | 0 |  |  |  |  | 146.7824 | 0
|  | 146.8808 | 0 |  |  |  |  | 146.8808 | 0
|  | 146.9792 | 0 |  |  |  |  | 146.9792 | 0
|  | 147.0775 | 0 |  |  |  |  | 147.0775 | 0
|  | 147.1759 | 0 |  |  |  |  | 147.1759 | 0
|  | 147.2743 | 0 |  |  |  |  | 147.2743 | 0
|  | 147.3727 | 0 |  |  |  |  | 147.3727 | 0
|  | 147.4711 | 0 |  |  |  |  | 147.4711 | 0
|  | 147.5694 | 0 |  |  |  |  | 147.5694 | 0
|  | 147.6678 | 0 |  |  |  |  | 147.6678 | 0
|  | 147.7662 | 0 |  |  |  |  | 147.7662 | 0
|  | 147.8646 | 0 |  |  |  |  | 147.8646 | 0
|  | 147.963 | 0 |  |  |  |  | 147.963 | 0
|  | 148.0613 | 0 |  |  |  |  | 148.0613 | 0
|  | 148.1597 | 0 |  |  |  |  | 148.1597 | 0
|  | 148.2581 | 0 |  |  |  |  | 148.2581 | 0
|  | 148.3565 | 0 |  |  |  |  | 148.3565 | 0
|  | 148.4549 | 0 |  |  |  |  | 148.4549 | 0
|  | 148.5532 | 0 |  |  |  |  | 148.5532 | 0
|  | 148.6516 | 0 |  |  |  |  | 148.6516 | 0
|  | 148.75 | 0 |  |  |  |  | 148.75 | 0
|  | 148.8484 | 0 |  |  |  |  | 148.8484 | 0
|  | 148.9468 | 0 |  |  |  |  | 148.9468 | 0
|  | 149.0451 | 0 |  |  |  |  | 149.0451 | 0
|  | 149.1435 | 0 |  |  |  |  | 149.1435 | 0
|  | 149.2419 | 0 |  |  |  |  | 149.2419 | 0
|  | 149.3403 | 0 |  |  |  |  | 149.3403 | 0
|  | 149.4387 | 0 |  |  |  |  | 149.4387 | 0
|  | 149.537 | 0 |  |  |  |  | 149.537 | 0
|  | 149.6354 | 0 |  |  |  |  | 149.6354 | 0
|  | 149.7338 | 0 |  |  |  |  | 149.7338 | 0
|  | 149.8322 | 0 |  |  |  |  | 149.8322 | 0
|  | 149.9306 | 0 |  |  |  |  | 149.9306 | 0
|  | 150.0289 | 0 |  |  |  |  | 150.0289 | 0
|  | 150.1273 | 0 |  |  |  |  | 150.1273 | 0
|  | 150.2257 | 0 |  |  |  |  | 150.2257 | 0
|  | 150.3241 | 0 |  |  |  |  | 150.3241 | 0
|  | 150.4225 | 0 |  |  |  |  | 150.4225 | 0
|  | 150.5208 | 0 |  |  |  |  | 150.5208 | 0
|  | 150.6192 | 0 |  |  |  |  | 150.6192 | 0
|  | 150.7176 | 0 |  |  |  |  | 150.7176 | 0
|  | 150.816 | 0 |  |  |  |  | 150.816 | 0
|  | 150.9144 | 0 |  |  |  |  | 150.9144 | 0
|  | 151.0127 | 57.8577 |  |  |  |  | 151.0127 | 0
|  | 151.1111 | 58.898 |  |  |  |  | 151.1111 | 0
|  | 151.2095 | 58.978 |  |  |  |  | 151.2095 | 0
|  | 151.3079 | 59.3781 |  |  |  |  | 151.3079 | 0
|  | 151.4062 | 60.2584 |  |  |  |  | 151.4062 | 0
|  | 151.5046 | 61.0586 |  |  |  |  | 151.5046 | 0
|  | 151.603 | 61.6988 |  |  |  |  | 151.603 | 0
|  | 151.7014 | 61.2587 |  |  |  |  | 151.7014 | 0
|  | 151.7998 | 61.4187 |  |  |  |  | 151.7998 | 0
|  | 151.8981 | 61.6588 |  |  |  |  | 151.8981 | 0
|  | 151.9965 | 61.0986 |  |  |  |  | 151.9965 | 0
|  | 152.0949 | 61.0186 |  |  |  |  | 152.0949 | 0
|  | 152.1933 | 60.8586 |  |  |  |  | 152.1933 | 0
|  | 152.2917 | 60.4584 |  |  |  |  | 152.2917 | 0
|  | 152.39 | 60.9786 |  |  |  |  | 152.39 | 0
|  | 152.4884 | 62.4991 |  |  |  |  | 152.4884 | 0
|  | 152.5868 | 63.0592 |  |  |  |  | 152.5868 | 0
|  | 152.6852 | 61.6188 |  |  |  |  | 152.6852 | 0
|  | 152.7836 | 60.7385 |  |  |  |  | 152.7836 | 0
|  | 152.8819 | 60.4985 |  |  |  |  | 152.8819 | 0
|  | 152.9803 | 60.0183 |  |  |  |  | 152.9803 | 0
|  | 153.0787 | 58.978 |  |  |  |  | 153.0787 | 0
|  | 153.1771 | 57.7776 |  |  |  |  | 153.1771 | 0
|  | 153.2755 | 56.9774 |  |  |  |  | 153.2755 | 0
|  | 153.3738 | 56.0971 |  |  |  |  | 153.3738 | 0
|  | 153.4722 | 55.2969 |  |  |  |  | 153.4722 | 0
|  | 153.5706 | 53.0562 |  |  |  |  | 153.5706 | 0
|  | 153.669 | 51.5357 |  |  |  |  | 153.669 | 0
|  | 153.7673 | 50.8955 |  |  |  |  | 153.7674 | 0
|  | 153.8657 | 49.5351 |  |  |  |  | 153.8657 | 0
|  | 153.9641 | 47.0143 |  |  |  |  | 153.9641 | 0
|  | 154.0625 | 45.4139 |  |  |  |  | 154.0625 | 0
|  | 154.1609 | 44.3735 |  |  |  |  | 154.1609 | 0
|  | 154.2593 | 44.3735 |  |  |  |  | 154.2593 | 0
|  | 154.3576 | 44.3735 |  |  |  |  | 154.3576 | 0
|  | 154.456 | 0 |  |  |  |  | 154.456 | 0
|  | 154.5544 | 0 |  |  |  |  | 154.5544 | 0
|  | 154.6528 | 0 |  |  |  |  | 154.6528 | 0
|  | 154.7512 | 0 |  |  |  |  | 154.7512 | 0
|  | 154.8495 | 0 |  |  |  |  | 154.8495 | 0
|  | 154.9479 | 0 |  |  |  |  | 154.9479 | 0
|  | 155.0463 | 0 |  |  |  |  | 155.0463 | 0
|  | 155.1447 | 0 |  |  |  |  | 155.1447 | 0
|  | 155.2431 | 0 |  |  |  |  | 155.2431 | 0
|  | 155.3414 | 0 |  |  |  |  | 155.3414 | 0
|  | 155.4398 | 0 |  |  |  |  | 155.4398 | 0
|  | 155.5382 | 0 |  |  |  |  | 155.5382 | 0
|  | 155.6366 | 0 |  |  |  |  | 155.6366 | 0
|  | 155.735 | 0 |  |  |  |  | 155.735 | 0
|  | 155.8333 | 0 |  |  |  |  | 155.8333 | 0
|  | 155.9317 | 0 |  |  |  |  | 155.9317 | 0
|  | 156.0301 | 0 |  |  |  |  | 156.0301 | 0
|  | 156.1285 | 0 |  |  |  |  | 156.1285 | 0
|  | 156.2269 | 0 |  |  |  |  | 156.2269 | 0
|  | 156.3252 | 0 |  |  |  |  | 156.3252 | 0
|  | 156.4236 | 0 |  |  |  |  | 156.4236 | 0
|  | 156.522 | 0 |  |  |  |  | 156.522 | 0
|  | 156.6204 | 0 |  |  |  |  | 156.6204 | 0
|  | 156.7188 | 0 |  |  |  |  | 156.7188 | 0
|  | 156.8171 | 0 |  |  |  |  | 156.8171 | 0
|  | 156.9155 | 0 |  |  |  |  | 156.9155 | 0
|  | 157.0139 | 0 |  |  |  |  | 157.0139 | 0
|  | 157.1123 | 0 |  |  |  |  | 157.1123 | 0
|  | 157.2106 | 0 |  |  |  |  | 157.2106 | 0
|  | 157.309 | 0 |  |  |  |  | 157.309 | 0
|  | 157.4074 | 0 |  |  |  |  | 157.4074 | 0
|  | 157.5058 | 0 |  |  |  |  | 157.5058 | 0
|  | 157.6042 | 0 |  |  |  |  | 157.6042 | 0
|  | 157.7025 | 0 |  |  |  |  | 157.7025 | 0
|  | 157.8009 | 0 |  |  |  |  | 157.8009 | 0
|  | 157.8993 | 0 |  |  |  |  | 157.8993 | 0
|  | 157.9977 | 0 |  |  |  |  | 157.9977 | 0
|  | 158.0961 | 0 |  |  |  |  | 158.0961 | 0
|  | 158.1944 | 0 |  |  |  |  | 158.1944 | 0
|  | 158.2928 | 0 |  |  |  |  | 158.2928 | 0
|  | 158.3912 | 0 |  |  |  |  | 158.3912 | 0
|  | 158.4896 | 0 |  |  |  |  | 158.4896 | 0
|  | 158.588 | 0 |  |  |  |  | 158.588 | 0
|  | 158.6863 | 0 |  |  |  |  | 158.6863 | 0
|  | 158.7847 | 0 |  |  |  |  | 158.7847 | 0
|  | 158.8831 | 0 |  |  |  |  | 158.8831 | 0
|  | 158.9815 | 0 |  |  |  |  | 158.9815 | 0
|  | 159.0798 | 0 |  |  |  |  | 159.0799 | 0
|  | 159.1782 | 0 |  |  |  |  | 159.1782 | 0
|  | 159.2766 | 0 |  |  |  |  | 159.2766 | 0
|  | 159.375 | 0 |  |  |  |  | 159.375 | 0
|  | 159.4734 | 0 |  |  |  |  | 159.4734 | 0
|  | 159.5718 | 0 |  |  |  |  | 159.5718 | 0
|  | 159.6701 | 0 |  |  |  |  | 159.6701 | 0
|  | 159.7685 | 0 |  |  |  |  | 159.7685 | 0
|  | 159.8669 | 0 |  |  |  |  | 159.8669 | 0
|  | 159.9653 | 0 |  |  |  |  | 159.9653 | 0
|  | 160.0637 | 0 |  |  |  |  | 160.0637 | 0
|  | 160.162 | 0 |  |  |  |  | 160.162 | 0
|  | 160.2604 | 0 |  |  |  |  | 160.2604 | 0
|  | 160.3588 | 0 |  |  |  |  | 160.3588 | 0
|  | 160.4572 | 0 |  |  |  |  | 160.4572 | 0
|  | 160.5556 | 0 |  |  |  |  | 160.5556 | 0
|  | 160.6539 | 0 |  |  |  |  | 160.6539 | 0
|  | 160.7523 | 0 |  |  |  |  | 160.7523 | 0
|  | 160.8507 | 0 |  |  |  |  | 160.8507 | 0
|  | 160.9491 | 0 |  |  |  |  | 160.9491 | 0
|  | 161.0475 | 0 |  |  |  |  | 161.0475 | 0
|  | 161.1458 | 0 |  |  |  |  | 161.1458 | 0
|  | 161.2442 | 0 |  |  |  |  | 161.2442 | 0
|  | 161.3426 | 0 |  |  |  |  | 161.3426 | 0
|  | 161.441 | 0 |  |  |  |  | 161.441 | 0
|  | 161.5394 | 0 |  |  |  |  | 161.5394 | 0
|  | 161.6377 | 0 |  |  |  |  | 161.6377 | 0
|  | 161.7361 | 0 |  |  |  |  | 161.7361 | 0
|  | 161.8345 | 0 |  |  |  |  | 161.8345 | 0
|  | 161.9329 | 0 |  |  |  |  | 161.9329 | 0
|  | 162.0312 | 0 |  |  |  |  | 162.0312 | 0
|  | 162.1296 | 0 |  |  |  |  | 162.1296 | 0
|  | 162.228 | 0 |  |  |  |  | 162.228 | 0
|  | 162.3264 | 0 |  |  |  |  | 162.3264 | 0
|  | 162.4248 | 0 |  |  |  |  | 162.4248 | 0
|  | 162.5231 | 0 |  |  |  |  | 162.5231 | 0
|  | 162.6215 | 0 |  |  |  |  | 162.6215 | 0
|  | 162.7199 | 0 |  |  |  |  | 162.7199 | 0
|  | 162.8183 | 0 |  |  |  |  | 162.8183 | 0
|  | 162.9167 | 0 |  |  |  |  | 162.9167 | 0
|  | 163.015 | 0 |  |  |  |  | 163.015 | 0
|  | 163.1134 | 0 |  |  |  |  | 163.1134 | 0
|  | 163.2118 | 0 |  |  |  |  | 163.2118 | 0
|  | 163.3102 | 0 |  |  |  |  | 163.3102 | 0
|  | 163.4086 | 0 |  |  |  |  | 163.4086 | 0
|  | 163.5069 | 0 |  |  |  |  | 163.5069 | 0
|  | 163.6053 | 0 |  |  |  |  | 163.6053 | 0
|  | 163.7037 | 0 |  |  |  |  | 163.7037 | 0
|  | 163.8021 | 0 |  |  |  |  | 163.8021 | 0
|  | 163.9005 | 0 |  |  |  |  | 163.9005 | 0
|  | 163.9988 | 0 |  |  |  |  | 163.9988 | 0
|  | 164.0972 | 0 |  |  |  |  | 164.0972 | 0
|  | 164.1956 | 0 |  |  |  |  | 164.1956 | 0
|  | 164.294 | 0 |  |  |  |  | 164.294 | 0
|  | 164.3923 | 0 |  |  |  |  | 164.3924 | 0
|  | 164.4907 | 0 |  |  |  |  | 164.4907 | 0
|  | 164.5891 | 0 |  |  |  |  | 164.5891 | 0
|  | 164.6875 | 0 |  |  |  |  | 164.6875 | 0
|  | 164.7859 | 0 |  |  |  |  | 164.7859 | 0
|  | 164.8843 | 0 |  |  |  |  | 164.8843 | 0
|  | 164.9826 | 0 |  |  |  |  | 164.9826 | 0
|  | 165.081 | 0 |  |  |  |  | 165.081 | 0
|  | 165.1794 | 0 |  |  |  |  | 165.1794 | 0
|  | 165.2778 | 0 |  |  |  |  | 165.2778 | 0
|  | 165.3762 | 0 |  |  |  |  | 165.3762 | 0
|  | 165.4745 | 0 |  |  |  |  | 165.4745 | 0
|  | 165.5729 | 0 |  |  |  |  | 165.5729 | 0
|  | 165.6713 | 0 |  |  |  |  | 165.6713 | 0
|  | 165.7697 | 0 |  |  |  |  | 165.7697 | 0
|  | 165.8681 | 0 |  |  |  |  | 165.8681 | 0
|  | 165.9664 | 0 |  |  |  |  | 165.9664 | 0
|  | 166.0648 | 0 |  |  |  |  | 166.0648 | 0
|  | 166.1632 | 0 |  |  |  |  | 166.1632 | 0
|  | 166.2616 | 0 |  |  |  |  | 166.2616 | 0
|  | 166.36 | 0 |  |  |  |  | 166.36 | 0
|  | 166.4583 | 0 |  |  |  |  | 166.4583 | 0
|  | 166.5567 | 0 |  |  |  |  | 166.5567 | 36.3711
|  | 166.6551 | 0 |  |  |  |  | 166.6551 | 70.3015
|  | 166.7535 | 0 |  |  |  |  | 166.7535 | 80.5446
|  | 166.8519 | 0 |  |  |  |  | 166.8519 | 81.0647
|  | 166.9502 | 0 |  |  |  |  | 166.9502 | 80.8247
|  | 167.0486 | 0 |  |  |  |  | 167.0486 | 80.3045
|  | 167.147 | 0 |  |  |  |  | 167.147 | 80.0244
|  | 167.2454 | 0 |  |  |  |  | 167.2454 | 79.9844
|  | 167.3438 | 0 |  |  |  |  | 167.3438 | 80.4646
|  | 167.4421 | 0 |  |  |  |  | 167.4421 | 81.7449
|  | 167.5405 | 0 |  |  |  |  | 167.5405 | 81.6249
|  | 167.6389 | 0 |  |  |  |  | 167.6389 | 81.0647
|  | 167.7373 | 58.3378 |  |  |  |  | 167.7373 | 81.4649
|  | 167.8356 | 59.4581 |  |  |  |  | 167.8356 | 81.5049
|  | 167.934 | 59.4181 |  |  |  |  | 167.934 | 80.7847
|  | 168.0324 | 59.138 |  |  |  |  | 168.0324 | 80.7446
|  | 168.1308 | 59.5782 |  |  |  |  | 168.1308 | 81.5449
|  | 168.2292 | 60.5385 |  |  |  |  | 168.2292 | 81.0247
|  | 168.3275 | 61.4988 |  |  |  |  | 168.3275 | 80.9047
|  | 168.4259 | 62.299 |  |  |  |  | 168.4259 | 80.8247
|  | 168.5243 | 62.5391 |  |  |  |  | 168.5243 | 84.6658
|  | 168.6227 | 62.139 |  |  |  |  | 168.6227 | 88.0269
|  | 168.7211 | 61.6988 |  |  |  |  | 168.7211 | 22.0867
|  | 168.8194 | 62.0989 |  |  |  |  | 168.8194 | 22.0867
|  | 168.9178 | 62.5791 |  |  |  |  | 168.9178 | 22.0867
|  | 169.0162 | 62.299 |  |  |  |  | 169.0162 | 0
|  | 169.1146 | 61.9389 |  |  |  |  | 169.1146 | 0
|  | 169.213 | 61.8989 |  |  |  |  | 169.213 | 0
|  | 169.3113 | 61.7789 |  |  |  |  | 169.3113 | 0
|  | 169.4097 | 60.8186 |  |  |  |  | 169.4097 | 0
|  | 169.5081 | 59.7782 |  |  |  |  | 169.5081 | 0
|  | 169.6065 | 59.8183 |  |  |  |  | 169.6065 | 0
|  | 169.7048 | 60.3784 |  |  |  |  | 169.7049 | 0
|  | 169.8032 | 60.1384 |  |  |  |  | 169.8032 | 0
|  | 169.9016 | 58.7779 |  |  |  |  | 169.9016 | 0
|  | 170 | 57.4975 |  |  |  |  | 170 | 0
|  | 170.0984 | 56.4572 |  |  |  |  | 170.0984 | 0
|  | 170.1968 | 55.3369 |  |  |  |  | 170.1968 | 0
|  | 170.2951 | 53.7364 |  |  |  |  | 170.2951 | 0
|  | 170.3935 | 52.0559 |  |  |  |  | 170.3935 | 0
|  | 170.4919 | 50.6154 |  |  |  |  | 170.4919 | 0
|  | 170.5903 | 49.6552 |  |  |  |  | 170.5903 | 0
|  | 170.6887 | 48.5348 |  |  |  |  | 170.6887 | 0
|  | 170.787 | 46.6942 |  |  |  |  | 170.787 | 0
|  | 170.8854 | 43.3332 |  |  |  |  | 170.8854 | 0
|  | 170.9838 | 23.4872 |  |  |  |  | 170.9838 | 0
|  | 171.0822 | 23.4872 |  |  |  |  | 171.0822 | 0
|  | 171.1806 | 23.4872 |  |  |  |  | 171.1806 | 0
|  | 171.2789 | 0 |  |  |  |  | 171.2789 | 0
|  | 171.3773 | 0 |  |  |  |  | 171.3773 | 0
|  | 171.4757 | 0 |  |  |  |  | 171.4757 | 0
|  | 171.5741 | 0 |  |  |  |  | 171.5741 | 0
|  | 171.6725 | 0 |  |  |  |  | 171.6725 | 0
|  | 171.7708 | 0 |  |  |  |  | 171.7708 | 0
|  | 171.8692 | 0 |  |  |  |  | 171.8692 | 0
|  | 171.9676 | 0 |  |  |  |  | 171.9676 | 0
|  | 172.066 | 0 |  |  |  |  | 172.066 | 0
|  | 172.1644 | 0 |  |  |  |  | 172.1644 | 0
|  | 172.2627 | 0 |  |  |  |  | 172.2627 | 0
|  | 172.3611 | 0 |  |  |  |  | 172.3611 | 0
|  | 172.4595 | 0 |  |  |  |  | 172.4595 | 0
|  | 172.5579 | 0 |  |  |  |  | 172.5579 | 0
|  | 172.6562 | 0 |  |  |  |  | 172.6562 | 0
|  | 172.7546 | 0 |  |  |  |  | 172.7546 | 0
|  | 172.853 | 0 |  |  |  |  | 172.853 | 0
|  | 172.9514 | 0 |  |  |  |  | 172.9514 | 0
|  | 173.0498 | 0 |  |  |  |  | 173.0498 | 0
|  | 173.1481 | 0 |  |  |  |  | 173.1481 | 0
|  | 173.2465 | 0 |  |  |  |  | 173.2465 | 0
|  | 173.3449 | 0 |  |  |  |  | 173.3449 | 0
|  | 173.4433 | 0 |  |  |  |  | 173.4433 | 0
|  | 173.5417 | 0 |  |  |  |  | 173.5417 | 0
|  | 173.64 | 0 |  |  |  |  | 173.64 | 0
|  | 173.7384 | 0 |  |  |  |  | 173.7384 | 0
|  | 173.8368 | 0 |  |  |  |  | 173.8368 | 0
|  | 173.9352 | 0 |  |  |  |  | 173.9352 | 0
|  | 174.0336 | 0 |  |  |  |  | 174.0336 | 0
|  | 174.1319 | 0 |  |  |  |  | 174.1319 | 0
|  | 174.2303 | 0 |  |  |  |  | 174.2303 | 0
|  | 174.3287 | 0 |  |  |  |  | 174.3287 | 0
|  | 174.4271 | 0 |  |  |  |  | 174.4271 | 0
|  | 174.5255 | 0 |  |  |  |  | 174.5255 | 0
|  | 174.6238 | 0 |  |  |  |  | 174.6238 | 0
|  | 174.7222 | 0 |  |  |  |  | 174.7222 | 0
|  | 174.8206 | 0 |  |  |  |  | 174.8206 | 0
|  | 174.919 | 0 |  |  |  |  | 174.919 | 0
|  | 175.0173 | 0 |  |  |  |  | 175.0174 | 0
|  | 175.1157 | 0 |  |  |  |  | 175.1157 | 0
|  | 175.2141 | 0 |  |  |  |  | 175.2141 | 0
|  | 175.3125 | 0 |  |  |  |  | 175.3125 | 0
|  | 175.4109 | 0 |  |  |  |  | 175.4109 | 0
|  | 175.5092 | 0 |  |  |  |  | 175.5093 | 0
|  | 175.6076 | 0 |  |  |  |  | 175.6076 | 0
|  | 175.706 | 0 |  |  |  |  | 175.706 | 0
|  | 175.8044 | 0 |  |  |  |  | 175.8044 | 0
|  | 175.9028 | 0 |  |  |  |  | 175.9028 | 0
|  | 176.0012 | 0 |  |  |  |  | 176.0012 | 0
|  | 176.0995 | 0 |  |  |  |  | 176.0995 | 0
|  | 176.1979 | 0 |  |  |  |  | 176.1979 | 0
|  | 176.2963 | 0 |  |  |  |  | 176.2963 | 0
|  | 176.3947 | 0 |  |  |  |  | 176.3947 | 0
|  | 176.4931 | 0 |  |  |  |  | 176.4931 | 0
|  | 176.5914 | 0 |  |  |  |  | 176.5914 | 0
|  | 176.6898 | 0 |  |  |  |  | 176.6898 | 0
|  | 176.7882 | 0 |  |  |  |  | 176.7882 | 0
|  | 176.8866 | 0 |  |  |  |  | 176.8866 | 0
|  | 176.985 | 0 |  |  |  |  | 176.985 | 0
|  | 177.0833 | 0 |  |  |  |  | 177.0833 | 0
|  | 177.1817 | 0 |  |  |  |  | 177.1817 | 0
|  | 177.2801 | 0 |  |  |  |  | 177.2801 | 0
|  | 177.3785 | 0 |  |  |  |  | 177.3785 | 0
|  | 177.4769 | 0 |  |  |  |  | 177.4769 | 0
|  | 177.5752 | 0 |  |  |  |  | 177.5752 | 0
|  | 177.6736 | 0 |  |  |  |  | 177.6736 | 0
|  | 177.772 | 0 |  |  |  |  | 177.772 | 0
|  | 177.8704 | 0 |  |  |  |  | 177.8704 | 0
|  | 177.9688 | 0 |  |  |  |  | 177.9688 | 0
|  | 178.0671 | 0 |  |  |  |  | 178.0671 | 0
|  | 178.1655 | 0 |  |  |  |  | 178.1655 | 0
|  | 178.2639 | 0 |  |  |  |  | 178.2639 | 0
|  | 178.3623 | 0 |  |  |  |  | 178.3623 | 0
|  | 178.4606 | 0 |  |  |  |  | 178.4606 | 0
|  | 178.559 | 0 |  |  |  |  | 178.559 | 0
|  | 178.6574 | 0 |  |  |  |  | 178.6574 | 0
|  | 178.7558 | 0 |  |  |  |  | 178.7558 | 0
|  | 178.8542 | 0 |  |  |  |  | 178.8542 | 0
|  | 178.9525 | 0 |  |  |  |  | 178.9525 | 0
|  | 179.0509 | 0 |  |  |  |  | 179.0509 | 0
|  | 179.1493 | 0 |  |  |  |  | 179.1493 | 0
|  | 179.2477 | 0 |  |  |  |  | 179.2477 | 0
|  | 179.3461 | 0 |  |  |  |  | 179.3461 | 0
|  | 179.4444 | 0 |  |  |  |  | 179.4444 | 0
|  | 179.5428 | 0 |  |  |  |  | 179.5428 | 0
|  | 179.6412 | 0 |  |  |  |  | 179.6412 | 0
|  | 179.7396 | 0 |  |  |  |  | 179.7396 | 0
|  | 179.838 | 0 |  |  |  |  | 179.838 | 0
|  | 179.9363 | 0 |  |  |  |  | 179.9363 | 0
|  | 180.0347 | 0 |  |  |  |  | 180.0347 | 0
|  | 180.1331 | 0 |  |  |  |  | 180.1331 | 0
|  | 180.2315 | 0 |  |  |  |  | 180.2315 | 0
|  | 180.3298 | 0 |  |  |  |  | 180.3299 | 0
|  | 180.4282 | 0 |  |  |  |  | 180.4282 | 0
|  | 180.5266 | 0 |  |  |  |  | 180.5266 | 0
|  | 180.625 | 0 |  |  |  |  | 180.625 | 0
|  | 180.7234 | 0 |  |  |  |  | 180.7234 | 0
|  | 180.8217 | 0 |  |  |  |  | 180.8218 | 0
|  | 180.9201 | 0 |  |  |  |  | 180.9201 | 0
|  | 181.0185 | 0 |  |  |  |  | 181.0185 | 0
|  | 181.1169 | 0 |  |  |  |  | 181.1169 | 0
|  | 181.2153 | 0 |  |  |  |  | 181.2153 | 0
|  | 181.3137 | 0 |  |  |  |  | 181.3137 | 0
|  | 181.412 | 0 |  |  |  |  | 181.412 | 0
|  | 181.5104 | 0 |  |  |  |  | 181.5104 | 0
|  | 181.6088 | 0 |  |  |  |  | 181.6088 | 0
|  | 181.7072 | 0 |  |  |  |  | 181.7072 | 0
|  | 181.8056 | 0 |  |  |  |  | 181.8056 | 0
|  | 181.9039 | 0 |  |  |  |  | 181.9039 | 0
|  | 182.0023 | 0 |  |  |  |  | 182.0023 | 0
|  | 182.1007 | 0 |  |  |  |  | 182.1007 | 0
|  | 182.1991 | 0 |  |  |  |  | 182.1991 | 0
|  | 182.2975 | 0 |  |  |  |  | 182.2975 | 0
|  | 182.3958 | 0 |  |  |  |  | 182.3958 | 0
|  | 182.4942 | 0 |  |  |  |  | 182.4942 | 0
|  | 182.5926 | 0 |  |  |  |  | 182.5926 | 0
|  | 182.691 | 0 |  |  |  |  | 182.691 | 0
|  | 182.7894 | 0 |  |  |  |  | 182.7894 | 0
|  | 182.8877 | 0 |  |  |  |  | 182.8877 | 0
|  | 182.9861 | 0 |  |  |  |  | 182.9861 | 0
|  | 183.0845 | 0 |  |  |  |  | 183.0845 | 0
|  | 183.1829 | 0 |  |  |  |  | 183.1829 | 0
|  | 183.2812 | 0 |  |  |  |  | 183.2812 | 0
|  | 183.3796 | 0 |  |  |  |  | 183.3796 | 0
|  | 183.478 | 0 |  |  |  |  | 183.478 | 0
|  | 183.5764 | 0 |  |  |  |  | 183.5764 | 0
|  | 183.6748 | 0 |  |  |  |  | 183.6748 | 0
|  | 183.7731 | 0 |  |  |  |  | 183.7731 | 0
|  | 183.8715 | 0 |  |  |  |  | 183.8715 | 0
|  | 183.9699 | 0 |  |  |  |  | 183.9699 | 0
|  | 184.0683 | 0 |  |  |  |  | 184.0683 | 0
|  | 184.1667 | 0 |  |  |  |  | 184.1667 | 0
|  | 184.265 | 0 |  |  |  |  | 184.265 | 0
|  | 184.3634 | 0 |  |  |  |  | 184.3634 | 0
|  | 184.4618 | 57.2575 |  |  |  |  | 184.4618 | 0
|  | 184.5602 | 58.2978 |  |  |  |  | 184.5602 | 0
|  | 184.6586 | 59.4982 |  |  |  |  | 184.6586 | 0
|  | 184.7569 | 60.9386 |  |  |  |  | 184.7569 | 0
|  | 184.8553 | 60.7785 |  |  |  |  | 184.8553 | 0
|  | 184.9537 | 59.7382 |  |  |  |  | 184.9537 | 0
|  | 185.0521 | 60.3784 |  |  |  |  | 185.0521 | 0
|  | 185.1505 | 61.5788 |  |  |  |  | 185.1505 | 0
|  | 185.2488 | 62.6991 |  |  |  |  | 185.2488 | 0
|  | 185.3472 | 62.6191 |  |  |  |  | 185.3472 | 0
|  | 185.4456 | 61.9789 |  |  |  |  | 185.4456 | 0
|  | 185.544 | 61.7789 |  |  |  |  | 185.544 | 0
|  | 185.6423 | 62.219 |  |  |  |  | 185.6424 | 0
|  | 185.7407 | 63.1793 |  |  |  |  | 185.7407 | 0
|  | 185.8391 | 63.0192 |  |  |  |  | 185.8391 | 0
|  | 185.9375 | 62.0589 |  |  |  |  | 185.9375 | 0
|  | 186.0359 | 61.3387 |  |  |  |  | 186.0359 | 0
|  | 186.1342 | 61.5788 |  |  |  |  | 186.1343 | 0
|  | 186.2326 | 61.5788 |  |  |  |  | 186.2326 | 0
|  | 186.331 | 61.5788 |  |  |  |  | 186.331 | 0
|  | 186.4294 | 59.8583 |  |  |  |  | 186.4294 | 0
|  | 186.5278 | 58.978 |  |  |  |  | 186.5278 | 0
|  | 186.6262 | 57.8577 |  |  |  |  | 186.6262 | 0
|  | 186.7245 | 57.0574 |  |  |  |  | 186.7245 | 0
|  | 186.8229 | 56.7373 |  |  |  |  | 186.8229 | 0
|  | 186.9213 | 55.697 |  |  |  |  | 186.9213 | 0
|  | 187.0197 | 54.4166 |  |  |  |  | 187.0197 | 0
|  | 187.1181 | 53.2963 |  |  |  |  | 187.1181 | 0
|  | 187.2164 | 52.8961 |  |  |  |  | 187.2164 | 0
|  | 187.3148 | 52.7361 |  |  |  |  | 187.3148 | 0
|  | 187.4132 | 49.6552 |  |  |  |  | 187.4132 | 0
|  | 187.5116 | 44.8537 |  |  |  |  | 187.5116 | 0
|  | 187.61 | 42.1329 |  |  |  |  | 187.61 | 0
|  | 187.7083 | 41.0125 |  |  |  |  | 187.7083 | 0
|  | 187.8067 | 41.0125 |  |  |  |  | 187.8067 | 0
|  | 187.9051 | 41.0125 |  |  |  |  | 187.9051 | 0
|  | 188.0035 | 0 |  |  |  |  | 188.0035 | 0
|  | 188.1019 | 0 |  |  |  |  | 188.1019 | 0
|  | 188.2002 | 0 |  |  |  |  | 188.2002 | 0
|  | 188.2986 | 0 |  |  |  |  | 188.2986 | 0
|  | 188.397 | 0 |  |  |  |  | 188.397 | 0
|  | 188.4954 | 0 |  |  |  |  | 188.4954 | 0
|  | 188.5938 | 0 |  |  |  |  | 188.5938 | 0
|  | 188.6921 | 0 |  |  |  |  | 188.6921 | 0
|  | 188.7905 | 0 |  |  |  |  | 188.7905 | 0
|  | 188.8889 | 0 |  |  |  |  | 188.8889 | 0
|  | 188.9873 | 0 |  |  |  |  | 188.9873 | 0
|  | 189.0856 | 0 |  |  |  |  | 189.0856 | 0
|  | 189.184 | 0 |  |  |  |  | 189.184 | 0
|  | 189.2824 | 0 |  |  |  |  | 189.2824 | 0
|  | 189.3808 | 0 |  |  |  |  | 189.3808 | 0
|  | 189.4792 | 0 |  |  |  |  | 189.4792 | 0
|  | 189.5775 | 0 |  |  |  |  | 189.5775 | 0
|  | 189.6759 | 0 |  |  |  |  | 189.6759 | 0
|  | 189.7743 | 0 |  |  |  |  | 189.7743 | 0
|  | 189.8727 | 0 |  |  |  |  | 189.8727 | 0
|  | 189.9711 | 0 |  |  |  |  | 189.9711 | 0
|  | 190.0694 | 0 |  |  |  |  | 190.0694 | 0
|  | 190.1678 | 0 |  |  |  |  | 190.1678 | 0
|  | 190.2662 | 0 |  |  |  |  | 190.2662 | 0
|  | 190.3646 | 0 |  |  |  |  | 190.3646 | 0
|  | 190.463 | 0 |  |  |  |  | 190.463 | 0
|  | 190.5613 | 0 |  |  |  |  | 190.5613 | 0
|  | 190.6597 | 0 |  |  |  |  | 190.6597 | 0
|  | 190.7581 | 0 |  |  |  |  | 190.7581 | 0
|  | 190.8565 | 0 |  |  |  |  | 190.8565 | 0
|  | 190.9548 | 0 |  |  |  |  | 190.9549 | 0
|  | 191.0532 | 0 |  |  |  |  | 191.0532 | 0
|  | 191.1516 | 0 |  |  |  |  | 191.1516 | 0
|  | 191.25 | 0 |  |  |  |  | 191.25 | 0
|  | 191.3484 | 0 |  |  |  |  | 191.3484 | 0
|  | 191.4467 | 0 |  |  |  |  | 191.4468 | 0
|  | 191.5451 | 0 |  |  |  |  | 191.5451 | 0
|  | 191.6435 | 0 |  |  |  |  | 191.6435 | 0
|  | 191.7419 | 0 |  |  |  |  | 191.7419 | 0
|  | 191.8403 | 0 |  |  |  |  | 191.8403 | 0
|  | 191.9386 | 0 |  |  |  |  | 191.9387 | 0
|  | 192.037 | 0 |  |  |  |  | 192.037 | 0
|  | 192.1354 | 0 |  |  |  |  | 192.1354 | 0
|  | 192.2338 | 0 |  |  |  |  | 192.2338 | 0
|  | 192.3322 | 0 |  |  |  |  | 192.3322 | 0
|  | 192.4306 | 0 |  |  |  |  | 192.4306 | 0
|  | 192.5289 | 0 |  |  |  |  | 192.5289 | 0
|  | 192.6273 | 0 |  |  |  |  | 192.6273 | 0
|  | 192.7257 | 0 |  |  |  |  | 192.7257 | 0
|  | 192.8241 | 0 |  |  |  |  | 192.8241 | 0
|  | 192.9225 | 0 |  |  |  |  | 192.9225 | 0
|  | 193.0208 | 0 |  |  |  |  | 193.0208 | 0
|  | 193.1192 | 0 |  |  |  |  | 193.1192 | 0
|  | 193.2176 | 0 |  |  |  |  | 193.2176 | 0
|  | 193.316 | 0 |  |  |  |  | 193.316 | 0
|  | 193.4144 | 0 |  |  |  |  | 193.4144 | 0
|  | 193.5127 | 0 |  |  |  |  | 193.5127 | 0
|  | 193.6111 | 0 |  |  |  |  | 193.6111 | 0
|  | 193.7095 | 0 |  |  |  |  | 193.7095 | 0
|  | 193.8079 | 0 |  |  |  |  | 193.8079 | 0
|  | 193.9062 | 0 |  |  |  |  | 193.9062 | 0
|  | 194.0046 | 0 |  |  |  |  | 194.0046 | 0
|  | 194.103 | 0 |  |  |  |  | 194.103 | 0
|  | 194.2014 | 0 |  |  |  |  | 194.2014 | 0
|  | 194.2998 | 0 |  |  |  |  | 194.2998 | 0
|  | 194.3981 | 0 |  |  |  |  | 194.3981 | 0
|  | 194.4965 | 0 |  |  |  |  | 194.4965 | 0
|  | 194.5949 | 0 |  |  |  |  | 194.5949 | 0
|  | 194.6933 | 0 |  |  |  |  | 194.6933 | 0
|  | 194.7917 | 0 |  |  |  |  | 194.7917 | 0
|  | 194.89 | 0 |  |  |  |  | 194.89 | 0
|  | 194.9884 | 0 |  |  |  |  | 194.9884 | 0
|  | 195.0868 | 0 |  |  |  |  | 195.0868 | 0
|  | 195.1852 | 0 |  |  |  |  | 195.1852 | 0
|  | 195.2836 | 0 |  |  |  |  | 195.2836 | 0
|  | 195.3819 | 0 |  |  |  |  | 195.3819 | 0
|  | 195.4803 | 0 |  |  |  |  | 195.4803 | 0
|  | 195.5787 | 0 |  |  |  |  | 195.5787 | 0
|  | 195.6771 | 0 |  |  |  |  | 195.6771 | 0
|  | 195.7755 | 0 |  |  |  |  | 195.7755 | 0
|  | 195.8738 | 0 |  |  |  |  | 195.8738 | 0
|  | 195.9722 | 0 |  |  |  |  | 195.9722 | 0
|  | 196.0706 | 0 |  |  |  |  | 196.0706 | 0
|  | 196.169 | 0 |  |  |  |  | 196.169 | 0
|  | 196.2673 | 0 |  |  |  |  | 196.2674 | 0
|  | 196.3657 | 0 |  |  |  |  | 196.3657 | 0
|  | 196.4641 | 0 |  |  |  |  | 196.4641 | 0
|  | 196.5625 | 0 |  |  |  |  | 196.5625 | 0
|  | 196.6609 | 0 |  |  |  |  | 196.6609 | 0
|  | 196.7592 | 0 |  |  |  |  | 196.7593 | 0
|  | 196.8576 | 0 |  |  |  |  | 196.8576 | 0
|  | 196.956 | 0 |  |  |  |  | 196.956 | 0
|  | 197.0544 | 0 |  |  |  |  | 197.0544 | 0
|  | 197.1528 | 0 |  |  |  |  | 197.1528 | 0
|  | 197.2511 | 0 |  |  |  |  | 197.2512 | 0
|  | 197.3495 | 0 |  |  |  |  | 197.3495 | 0
|  | 197.4479 | 0 |  |  |  |  | 197.4479 | 0
|  | 197.5463 | 0 |  |  |  |  | 197.5463 | 0
|  | 197.6447 | 0 |  |  |  |  | 197.6447 | 0
|  | 197.7431 | 0 |  |  |  |  | 197.7431 | 0
|  | 197.8414 | 0 |  |  |  |  | 197.8414 | 0
|  | 197.9398 | 0 |  |  |  |  | 197.9398 | 0
|  | 198.0382 | 0 |  |  |  |  | 198.0382 | 0
|  | 198.1366 | 0 |  |  |  |  | 198.1366 | 0
|  | 198.235 | 0 |  |  |  |  | 198.235 | 0
|  | 198.3333 | 0 |  |  |  |  | 198.3333 | 0
|  | 198.4317 | 0 |  |  |  |  | 198.4317 | 0
|  | 198.5301 | 0 |  |  |  |  | 198.5301 | 0
|  | 198.6285 | 0 |  |  |  |  | 198.6285 | 0
|  | 198.7269 | 0 |  |  |  |  | 198.7269 | 0
|  | 198.8252 | 0 |  |  |  |  | 198.8252 | 0
|  | 198.9236 | 0 |  |  |  |  | 198.9236 | 0
|  | 199.022 | 0 |  |  |  |  | 199.022 | 0
|  | 199.1204 | 0 |  |  |  |  | 199.1204 | 0
|  | 199.2188 | 0 |  |  |  |  | 199.2188 | 0
|  | 199.3171 | 0 |  |  |  |  | 199.3171 | 0
|  | 199.4155 | 0 |  |  |  |  | 199.4155 | 0
|  | 199.5139 | 0 |  |  |  |  | 199.5139 | 0
|  | 199.6123 | 0 |  |  |  |  | 199.6123 | 0
|  | 199.7106 | 0 |  |  |  |  | 199.7106 | 0
|  | 199.809 | 0 |  |  |  |  | 199.809 | 0
|  | 199.9074 | 0 |  |  |  |  | 199.9074 | 31.8897
|  | 200.0058 | 0 |  |  |  |  | 200.0058 | 65.8201
|  | 200.1042 | 0 |  |  |  |  | 200.1042 | 63.0993
|  | 200.2025 | 0 |  |  |  |  | 200.2025 | 62.8992
|  | 200.3009 | 0 |  |  |  |  | 200.3009 | 75.6231
|  | 200.3993 | 0 |  |  |  |  | 200.3993 | 80.1044
|  | 200.4977 | 0 |  |  |  |  | 200.4977 | 80.6246
|  | 200.5961 | 0 |  |  |  |  | 200.5961 | 81.3448
|  | 200.6944 | 0 |  |  |  |  | 200.6944 | 81.3048
|  | 200.7928 | 0 |  |  |  |  | 200.7928 | 81.1047
|  | 200.8912 | 0 |  |  |  |  | 200.8912 | 80.5446
|  | 200.9896 | 0 |  |  |  |  | 200.9896 | 80.1845
|  | 201.088 | 0 |  |  |  |  | 201.088 | 80.4646
|  | 201.1863 | 56.5373 |  |  |  |  | 201.1863 | 80.3445
|  | 201.2847 | 58.3778 |  |  |  |  | 201.2847 | 80.3445
|  | 201.3831 | 59.018 |  |  |  |  | 201.3831 | 80.3445
|  | 201.4815 | 59.8183 |  |  |  |  | 201.4815 | 80.2245
|  | 201.5798 | 60.1784 |  |  |  |  | 201.5799 | 80.7046
|  | 201.6782 | 60.4184 |  |  |  |  | 201.6782 | 80.5846
|  | 201.7766 | 61.1387 |  |  |  |  | 201.7766 | 80.1845
|  | 201.875 | 61.4187 |  |  |  |  | 201.875 | 80.4245
|  | 201.9734 | 61.0986 |  |  |  |  | 201.9734 | 83.7856
|  | 202.0717 | 60.6985 |  |  |  |  | 202.0718 | 88.587
|  | 202.1701 | 61.0986 |  |  |  |  | 202.1701 | 32.89
|  | 202.2685 | 62.299 |  |  |  |  | 202.2685 | 32.89
|  | 202.3669 | 62.9392 |  |  |  |  | 202.3669 | 32.89
|  | 202.4653 | 62.6591 |  |  |  |  | 202.4653 | 0
|  | 202.5636 | 61.9389 |  |  |  |  | 202.5637 | 0
|  | 202.662 | 61.6188 |  |  |  |  | 202.662 | 0
|  | 202.7604 | 61.5388 |  |  |  |  | 202.7604 | 0
|  | 202.8588 | 61.2187 |  |  |  |  | 202.8588 | 0
|  | 202.9572 | 60.7785 |  |  |  |  | 202.9572 | 0
|  | 203.0556 | 60.4985 |  |  |  |  | 203.0556 | 0
|  | 203.1539 | 60.2584 |  |  |  |  | 203.1539 | 0
|  | 203.2523 | 59.6182 |  |  |  |  | 203.2523 | 0
|  | 203.3507 | 59.058 |  |  |  |  | 203.3507 | 0
|  | 203.4491 | 58.0177 |  |  |  |  | 203.4491 | 0
|  | 203.5475 | 56.3372 |  |  |  |  | 203.5475 | 0
|  | 203.6458 | 54.6567 |  |  |  |  | 203.6458 | 0
|  | 203.7442 | 53.5363 |  |  |  |  | 203.7442 | 0
|  | 203.8426 | 52.1759 |  |  |  |  | 203.8426 | 0
|  | 203.941 | 51.3757 |  |  |  |  | 203.941 | 0
|  | 204.0394 | 51.0156 |  |  |  |  | 204.0394 | 0
|  | 204.1377 | 49.9753 |  |  |  |  | 204.1377 | 0
|  | 204.2361 | 48.0547 |  |  |  |  | 204.2361 | 0
|  | 204.3345 | 47.7346 |  |  |  |  | 204.3345 | 0
|  | 204.4329 | 38.7718 |  |  |  |  | 204.4329 | 0
|  | 204.5312 | 38.7718 |  |  |  |  | 204.5312 | 0
|  | 204.6296 | 38.7718 |  |  |  |  | 204.6296 | 0
|  | 204.728 | 0 |  |  |  |  | 204.728 | 0
|  | 204.8264 | 0 |  |  |  |  | 204.8264 | 0
|  | 204.9248 | 0 |  |  |  |  | 204.9248 | 0
|  | 205.0231 | 0 |  |  |  |  | 205.0231 | 0
|  | 205.1215 | 0 |  |  |  |  | 205.1215 | 0
|  | 205.2199 | 0 |  |  |  |  | 205.2199 | 0
|  | 205.3183 | 0 |  |  |  |  | 205.3183 | 0
|  | 205.4167 | 0 |  |  |  |  | 205.4167 | 0
|  | 205.515 | 0 |  |  |  |  | 205.515 | 0
|  | 205.6134 | 0 |  |  |  |  | 205.6134 | 0
|  | 205.7118 | 0 |  |  |  |  | 205.7118 | 0
|  | 205.8102 | 0 |  |  |  |  | 205.8102 | 0
|  | 205.9086 | 0 |  |  |  |  | 205.9086 | 0
|  | 206.0069 | 0 |  |  |  |  | 206.0069 | 0
|  | 206.1053 | 0 |  |  |  |  | 206.1053 | 0
|  | 206.2037 | 0 |  |  |  |  | 206.2037 | 0
|  | 206.3021 | 0 |  |  |  |  | 206.3021 | 0
|  | 206.4005 | 0 |  |  |  |  | 206.4005 | 0
|  | 206.4988 | 0 |  |  |  |  | 206.4988 | 0
|  | 206.5972 | 0 |  |  |  |  | 206.5972 | 0
|  | 206.6956 | 0 |  |  |  |  | 206.6956 | 0
|  | 206.794 | 0 |  |  |  |  | 206.794 | 0
|  | 206.8923 | 0 |  |  |  |  | 206.8924 | 0
|  | 206.9907 | 0 |  |  |  |  | 206.9907 | 0
|  | 207.0891 | 0 |  |  |  |  | 207.0891 | 0
|  | 207.1875 | 0 |  |  |  |  | 207.1875 | 0
|  | 207.2859 | 0 |  |  |  |  | 207.2859 | 0
|  | 207.3842 | 0 |  |  |  |  | 207.3843 | 0
|  | 207.4826 | 0 |  |  |  |  | 207.4826 | 0
|  | 207.581 | 0 |  |  |  |  | 207.581 | 0
|  | 207.6794 | 0 |  |  |  |  | 207.6794 | 0
|  | 207.7778 | 0 |  |  |  |  | 207.7778 | 0
|  | 207.8761 | 0 |  |  |  |  | 207.8762 | 0
|  | 207.9745 | 0 |  |  |  |  | 207.9745 | 0
|  | 208.0729 | 0 |  |  |  |  | 208.0729 | 0
|  | 208.1713 | 0 |  |  |  |  | 208.1713 | 0
|  | 208.2697 | 0 |  |  |  |  | 208.2697 | 0
|  | 208.3681 | 0 |  |  |  |  | 208.3681 | 0
|  | 208.4664 | 0 |  |  |  |  | 208.4664 | 0
|  | 208.5648 | 0 |  |  |  |  | 208.5648 | 0
|  | 208.6632 | 0 |  |  |  |  | 208.6632 | 0
|  | 208.7616 | 0 |  |  |  |  | 208.7616 | 0
|  | 208.86 | 0 |  |  |  |  | 208.86 | 0
|  | 208.9583 | 0 |  |  |  |  | 208.9583 | 0
|  | 209.0567 | 0 |  |  |  |  | 209.0567 | 0
|  | 209.1551 | 0 |  |  |  |  | 209.1551 | 0
|  | 209.2535 | 0 |  |  |  |  | 209.2535 | 0
|  | 209.3519 | 0 |  |  |  |  | 209.3519 | 0
|  | 209.4502 | 0 |  |  |  |  | 209.4502 | 0
|  | 209.5486 | 0 |  |  |  |  | 209.5486 | 0
|  | 209.647 | 0 |  |  |  |  | 209.647 | 0
|  | 209.7454 | 0 |  |  |  |  | 209.7454 | 0
|  | 209.8438 | 0 |  |  |  |  | 209.8438 | 0
|  | 209.9421 | 0 |  |  |  |  | 209.9421 | 0
|  | 210.0405 | 0 |  |  |  |  | 210.0405 | 0
|  | 210.1389 | 0 |  |  |  |  | 210.1389 | 0
|  | 210.2373 | 0 |  |  |  |  | 210.2373 | 0
|  | 210.3356 | 0 |  |  |  |  | 210.3356 | 0
|  | 210.434 | 0 |  |  |  |  | 210.434 | 0
|  | 210.5324 | 0 |  |  |  |  | 210.5324 | 0
|  | 210.6308 | 0 |  |  |  |  | 210.6308 | 0
|  | 210.7292 | 0 |  |  |  |  | 210.7292 | 0
|  | 210.8275 | 0 |  |  |  |  | 210.8275 | 0
|  | 210.9259 | 0 |  |  |  |  | 210.9259 | 0
|  | 211.0243 | 0 |  |  |  |  | 211.0243 | 0
|  | 211.1227 | 0 |  |  |  |  | 211.1227 | 0
|  | 211.2211 | 0 |  |  |  |  | 211.2211 | 0
|  | 211.3194 | 0 |  |  |  |  | 211.3194 | 0
|  | 211.4178 | 0 |  |  |  |  | 211.4178 | 0
|  | 211.5162 | 0 |  |  |  |  | 211.5162 | 0
|  | 211.6146 | 0 |  |  |  |  | 211.6146 | 0
|  | 211.713 | 0 |  |  |  |  | 211.713 | 0
|  | 211.8113 | 0 |  |  |  |  | 211.8113 | 0
|  | 211.9097 | 0 |  |  |  |  | 211.9097 | 0
|  | 212.0081 | 0 |  |  |  |  | 212.0081 | 0
|  | 212.1065 | 0 |  |  |  |  | 212.1065 | 0
|  | 212.2048 | 0 |  |  |  |  | 212.2049 | 0
|  | 212.3032 | 0 |  |  |  |  | 212.3032 | 0
|  | 212.4016 | 0 |  |  |  |  | 212.4016 | 0
|  | 212.5 | 0 |  |  |  |  | 212.5 | 0
|  | 212.5984 | 0 |  |  |  |  | 212.5984 | 0
|  | 212.6967 | 0 |  |  |  |  | 212.6968 | 0
|  | 212.7951 | 0 |  |  |  |  | 212.7951 | 0
|  | 212.8935 | 0 |  |  |  |  | 212.8935 | 0
|  | 212.9919 | 0 |  |  |  |  | 212.9919 | 0
|  | 213.0903 | 0 |  |  |  |  | 213.0903 | 0
|  | 213.1886 | 0 |  |  |  |  | 213.1887 | 0
|  | 213.287 | 0 |  |  |  |  | 213.287 | 0
|  | 213.3854 | 0 |  |  |  |  | 213.3854 | 0
|  | 213.4838 | 0 |  |  |  |  | 213.4838 | 0
|  | 213.5822 | 0 |  |  |  |  | 213.5822 | 0
|  | 213.6805 | 0 |  |  |  |  | 213.6806 | 0
|  | 213.7789 | 0 |  |  |  |  | 213.7789 | 0
|  | 213.8773 | 0 |  |  |  |  | 213.8773 | 0
|  | 213.9757 | 0 |  |  |  |  | 213.9757 | 0
|  | 214.0741 | 0 |  |  |  |  | 214.0741 | 0
|  | 214.1725 | 0 |  |  |  |  | 214.1725 | 0
|  | 214.2708 | 0 |  |  |  |  | 214.2708 | 0
|  | 214.3692 | 0 |  |  |  |  | 214.3692 | 0
|  | 214.4676 | 0 |  |  |  |  | 214.4676 | 0
|  | 214.566 | 0 |  |  |  |  | 214.566 | 0
|  | 214.6644 | 0 |  |  |  |  | 214.6644 | 0
|  | 214.7627 | 0 |  |  |  |  | 214.7627 | 0
|  | 214.8611 | 0 |  |  |  |  | 214.8611 | 0
|  | 214.9595 | 0 |  |  |  |  | 214.9595 | 0
|  | 215.0579 | 0 |  |  |  |  | 215.0579 | 0
|  | 215.1562 | 0 |  |  |  |  | 215.1562 | 0
|  | 215.2546 | 0 |  |  |  |  | 215.2546 | 0
|  | 215.353 | 0 |  |  |  |  | 215.353 | 0
|  | 215.4514 | 0 |  |  |  |  | 215.4514 | 0
|  | 215.5498 | 0 |  |  |  |  | 215.5498 | 0
|  | 215.6481 | 0 |  |  |  |  | 215.6481 | 0
|  | 215.7465 | 0 |  |  |  |  | 215.7465 | 0
|  | 215.8449 | 0 |  |  |  |  | 215.8449 | 0
|  | 215.9433 | 0 |  |  |  |  | 215.9433 | 0
|  | 216.0417 | 0 |  |  |  |  | 216.0417 | 0
|  | 216.14 | 0 |  |  |  |  | 216.14 | 0
|  | 216.2384 | 0 |  |  |  |  | 216.2384 | 0
|  | 216.3368 | 0 |  |  |  |  | 216.3368 | 0
|  | 216.4352 | 0 |  |  |  |  | 216.4352 | 0
|  | 216.5336 | 0 |  |  |  |  | 216.5336 | 0
|  | 216.6319 | 0 |  |  |  |  | 216.6319 | 0
|  | 216.7303 | 0 |  |  |  |  | 216.7303 | 0
|  | 216.8287 | 0 |  |  |  |  | 216.8287 | 0
|  | 216.9271 | 0 |  |  |  |  | 216.9271 | 0
|  | 217.0255 | 0 |  |  |  |  | 217.0255 | 0
|  | 217.1238 | 0 |  |  |  |  | 217.1238 | 0
|  | 217.2222 | 0 |  |  |  |  | 217.2222 | 0
|  | 217.3206 | 0 |  |  |  |  | 217.3206 | 0
|  | 217.419 | 0 |  |  |  |  | 217.419 | 0
|  | 217.5173 | 0 |  |  |  |  | 217.5174 | 0
|  | 217.6157 | 0 |  |  |  |  | 217.6157 | 0
|  | 217.7141 | 0 |  |  |  |  | 217.7141 | 0
|  | 217.8125 | 0 |  |  |  |  | 217.8125 | 0
|  | 217.9109 | 55.8971 |  |  |  |  | 217.9109 | 0
|  | 218.0092 | 57.9777 |  |  |  |  | 218.0093 | 0
|  | 218.1076 | 58.938 |  |  |  |  | 218.1076 | 0
|  | 218.206 | 58.3778 |  |  |  |  | 218.206 | 0
|  | 218.3044 | 58.8179 |  |  |  |  | 218.3044 | 0
|  | 218.4028 | 60.5785 |  |  |  |  | 218.4028 | 0
|  | 218.5011 | 62.0989 |  |  |  |  | 218.5012 | 0
|  | 218.5995 | 62.179 |  |  |  |  | 218.5995 | 0
|  | 218.6979 | 61.7388 |  |  |  |  | 218.6979 | 0
|  | 218.7963 | 61.8589 |  |  |  |  | 218.7963 | 0
|  | 218.8947 | 62.0989 |  |  |  |  | 218.8947 | 0
|  | 218.993 | 61.9789 |  |  |  |  | 218.9931 | 0
|  | 219.0914 | 61.8189 |  |  |  |  | 219.0914 | 0
|  | 219.1898 | 61.9389 |  |  |  |  | 219.1898 | 0
|  | 219.2882 | 62.259 |  |  |  |  | 219.2882 | 0
|  | 219.3866 | 62.339 |  |  |  |  | 219.3866 | 0
|  | 219.485 | 62.339 |  |  |  |  | 219.485 | 0
|  | 219.5833 | 62.339 |  |  |  |  | 219.5833 | 0
|  | 219.6817 | 62.339 |  |  |  |  | 219.6817 | 0
|  | 219.7801 | 62.339 |  |  |  |  | 219.7801 | 0
|  | 219.8785 | 60.6985 |  |  |  |  | 219.8785 | 0
|  | 219.9769 | 60.4584 |  |  |  |  | 219.9769 | 0
|  | 220.0752 | 59.2581 |  |  |  |  | 220.0752 | 0
|  | 220.1736 | 58.2178 |  |  |  |  | 220.1736 | 0
|  | 220.272 | 57.0974 |  |  |  |  | 220.272 | 0
|  | 220.3704 | 55.4969 |  |  |  |  | 220.3704 | 0
|  | 220.4688 | 53.6564 |  |  |  |  | 220.4688 | 0
|  | 220.5671 | 51.8958 |  |  |  |  | 220.5671 | 0
|  | 220.6655 | 50.1353 |  |  |  |  | 220.6655 | 0
|  | 220.7639 | 49.7752 |  |  |  |  | 220.7639 | 0
|  | 220.8623 | 49.7752 |  |  |  |  | 220.8623 | 0
|  | 220.9606 | 47.8546 |  |  |  |  | 220.9606 | 0
|  | 221.059 | 44.1735 |  |  |  |  | 221.059 | 0
|  | 221.1574 | 41.6127 |  |  |  |  | 221.1574 | 0
|  | 221.2558 | 41.6127 |  |  |  |  | 221.2558 | 0
|  | 221.3542 | 41.6127 |  |  |  |  | 221.3542 | 0
|  | 221.4525 | 0 |  |  |  |  | 221.4525 | 0
|  | 221.5509 | 0 |  |  |  |  | 221.5509 | 0
|  | 221.6493 | 0 |  |  |  |  | 221.6493 | 0
|  | 221.7477 | 0 |  |  |  |  | 221.7477 | 0
|  | 221.8461 | 0 |  |  |  |  | 221.8461 | 0
|  | 221.9444 | 0 |  |  |  |  | 221.9444 | 0
|  | 222.0428 | 0 |  |  |  |  | 222.0428 | 0
|  | 222.1412 | 0 |  |  |  |  | 222.1412 | 0
|  | 222.2396 | 0 |  |  |  |  | 222.2396 | 0
|  | 222.338 | 0 |  |  |  |  | 222.338 | 0
|  | 222.4363 | 0 |  |  |  |  | 222.4363 | 0
|  | 222.5347 | 0 |  |  |  |  | 222.5347 | 0
|  | 222.6331 | 0 |  |  |  |  | 222.6331 | 0
|  | 222.7315 | 0 |  |  |  |  | 222.7315 | 0
|  | 222.8298 | 0 |  |  |  |  | 222.8299 | 0
|  | 222.9282 | 0 |  |  |  |  | 222.9282 | 0
|  | 223.0266 | 0 |  |  |  |  | 223.0266 | 0
|  | 223.125 | 0 |  |  |  |  | 223.125 | 0
|  | 223.2234 | 0 |  |  |  |  | 223.2234 | 0
|  | 223.3217 | 0 |  |  |  |  | 223.3218 | 0
|  | 223.4201 | 0 |  |  |  |  | 223.4201 | 0
|  | 223.5185 | 0 |  |  |  |  | 223.5185 | 0
|  | 223.6169 | 0 |  |  |  |  | 223.6169 | 0
|  | 223.7153 | 0 |  |  |  |  | 223.7153 | 0
|  | 223.8136 | 0 |  |  |  |  | 223.8137 | 0
|  | 223.912 | 0 |  |  |  |  | 223.912 | 0
|  | 224.0104 | 0 |  |  |  |  | 224.0104 | 0
|  | 224.1088 | 0 |  |  |  |  | 224.1088 | 0
|  | 224.2072 | 0 |  |  |  |  | 224.2072 | 0
|  | 224.3055 | 0 |  |  |  |  | 224.3056 | 0
|  | 224.4039 | 0 |  |  |  |  | 224.4039 | 0
|  | 224.5023 | 0 |  |  |  |  | 224.5023 | 0
|  | 224.6007 | 0 |  |  |  |  | 224.6007 | 0
|  | 224.6991 | 0 |  |  |  |  | 224.6991 | 0
|  | 224.7975 | 0 |  |  |  |  | 224.7975 | 0
|  | 224.8958 | 0 |  |  |  |  | 224.8958 | 0
|  | 224.9942 | 0 |  |  |  |  | 224.9942 | 0
|  | 225.0926 | 0 |  |  |  |  | 225.0926 | 0
|  | 225.191 | 0 |  |  |  |  | 225.191 | 0
|  | 225.2894 | 0 |  |  |  |  | 225.2894 | 0
|  | 225.3877 | 0 |  |  |  |  | 225.3877 | 0
|  | 225.4861 | 0 |  |  |  |  | 225.4861 | 0
|  | 225.5845 | 0 |  |  |  |  | 225.5845 | 0
|  | 225.6829 | 0 |  |  |  |  | 225.6829 | 0
|  | 225.7812 | 0 |  |  |  |  | 225.7812 | 0
|  | 225.8796 | 0 |  |  |  |  | 225.8796 | 0
|  | 225.978 | 0 |  |  |  |  | 225.978 | 0
|  | 226.0764 | 0 |  |  |  |  | 226.0764 | 0
|  | 226.1748 | 0 |  |  |  |  | 226.1748 | 0
|  | 226.2731 | 0 |  |  |  |  | 226.2731 | 0
|  | 226.3715 | 0 |  |  |  |  | 226.3715 | 0
|  | 226.4699 | 0 |  |  |  |  | 226.4699 | 0
|  | 226.5683 | 0 |  |  |  |  | 226.5683 | 0
|  | 226.6667 | 0 |  |  |  |  | 226.6667 | 0
|  | 226.765 | 0 |  |  |  |  | 226.765 | 0
|  | 226.8634 | 0 |  |  |  |  | 226.8634 | 0
|  | 226.9618 | 0 |  |  |  |  | 226.9618 | 0
|  | 227.0602 | 0 |  |  |  |  | 227.0602 | 0
|  | 227.1586 | 0 |  |  |  |  | 227.1586 | 0
|  | 227.2569 | 0 |  |  |  |  | 227.2569 | 0
|  | 227.3553 | 0 |  |  |  |  | 227.3553 | 0
|  | 227.4537 | 0 |  |  |  |  | 227.4537 | 0
|  | 227.5521 | 0 |  |  |  |  | 227.5521 | 0
|  | 227.6505 | 0 |  |  |  |  | 227.6505 | 0
|  | 227.7488 | 0 |  |  |  |  | 227.7488 | 0
|  | 227.8472 | 0 |  |  |  |  | 227.8472 | 0
|  | 227.9456 | 0 |  |  |  |  | 227.9456 | 0
|  | 228.044 | 0 |  |  |  |  | 228.044 | 0
|  | 228.1423 | 0 |  |  |  |  | 228.1424 | 0
|  | 228.2407 | 0 |  |  |  |  | 228.2407 | 0
|  | 228.3391 | 0 |  |  |  |  | 228.3391 | 0
|  | 228.4375 | 0 |  |  |  |  | 228.4375 | 0
|  | 228.5359 | 0 |  |  |  |  | 228.5359 | 0
|  | 228.6342 | 0 |  |  |  |  | 228.6343 | 0
|  | 228.7326 | 0 |  |  |  |  | 228.7326 | 0
|  | 228.831 | 0 |  |  |  |  | 228.831 | 0
|  | 228.9294 | 0 |  |  |  |  | 228.9294 | 0
|  | 229.0278 | 0 |  |  |  |  | 229.0278 | 0
|  | 229.1261 | 0 |  |  |  |  | 229.1262 | 0
|  | 229.2245 | 0 |  |  |  |  | 229.2245 | 0
|  | 229.3229 | 0 |  |  |  |  | 229.3229 | 0
|  | 229.4213 | 0 |  |  |  |  | 229.4213 | 0
|  | 229.5197 | 0 |  |  |  |  | 229.5197 | 0
|  | 229.618 | 0 |  |  |  |  | 229.6181 | 0
|  | 229.7164 | 0 |  |  |  |  | 229.7164 | 0
|  | 229.8148 | 0 |  |  |  |  | 229.8148 | 0
|  | 229.9132 | 0 |  |  |  |  | 229.9132 | 0
|  | 230.0116 | 0 |  |  |  |  | 230.0116 | 0
|  | 230.11 | 0 |  |  |  |  | 230.11 | 0
|  | 230.2083 | 0 |  |  |  |  | 230.2083 | 0
|  | 230.3067 | 0 |  |  |  |  | 230.3067 | 0
|  | 230.4051 | 0 |  |  |  |  | 230.4051 | 0
|  | 230.5035 | 0 |  |  |  |  | 230.5035 | 0
|  | 230.6019 | 0 |  |  |  |  | 230.6019 | 0
|  | 230.7002 | 0 |  |  |  |  | 230.7002 | 0
|  | 230.7986 | 0 |  |  |  |  | 230.7986 | 0
|  | 230.897 | 0 |  |  |  |  | 230.897 | 0
|  | 230.9954 | 0 |  |  |  |  | 230.9954 | 0
|  | 231.0938 | 0 |  |  |  |  | 231.0938 | 0
|  | 231.1921 | 0 |  |  |  |  | 231.1921 | 0
|  | 231.2905 | 0 |  |  |  |  | 231.2905 | 0
|  | 231.3889 | 0 |  |  |  |  | 231.3889 | 0
|  | 231.4873 | 0 |  |  |  |  | 231.4873 | 0
|  | 231.5856 | 0 |  |  |  |  | 231.5856 | 0
|  | 231.684 | 0 |  |  |  |  | 231.684 | 0
|  | 231.7824 | 0 |  |  |  |  | 231.7824 | 0
|  | 231.8808 | 0 |  |  |  |  | 231.8808 | 0
|  | 231.9792 | 0 |  |  |  |  | 231.9792 | 0
|  | 232.0775 | 0 |  |  |  |  | 232.0775 | 0
|  | 232.1759 | 0 |  |  |  |  | 232.1759 | 0
|  | 232.2743 | 0 |  |  |  |  | 232.2743 | 0
|  | 232.3727 | 0 |  |  |  |  | 232.3727 | 0
|  | 232.4711 | 0 |  |  |  |  | 232.4711 | 0
|  | 232.5694 | 0 |  |  |  |  | 232.5694 | 0
|  | 232.6678 | 0 |  |  |  |  | 232.6678 | 0
|  | 232.7662 | 0 |  |  |  |  | 232.7662 | 0
|  | 232.8646 | 0 |  |  |  |  | 232.8646 | 0
|  | 232.963 | 0 |  |  |  |  | 232.963 | 0
|  | 233.0613 | 0 |  |  |  |  | 233.0613 | 0
|  | 233.1597 | 0 |  |  |  |  | 233.1597 | 0
|  | 233.2581 | 0 |  |  |  |  | 233.2581 | 0
|  | 233.3565 | 0 |  |  |  |  | 233.3565 | 29.3289
|  | 233.4548 | 0 |  |  |  |  | 233.4549 | 65.8201
|  | 233.5532 | 0 |  |  |  |  | 233.5532 | 62.139
|  | 233.6516 | 0 |  |  |  |  | 233.6516 | 61.7388
|  | 233.75 | 0 |  |  |  |  | 233.75 | 76.8234
|  | 233.8484 | 0 |  |  |  |  | 233.8484 | 80.3445
|  | 233.9467 | 0 |  |  |  |  | 233.9468 | 79.9444
|  | 234.0451 | 0 |  |  |  |  | 234.0451 | 80.2245
|  | 234.1435 | 0 |  |  |  |  | 234.1435 | 80.6246
|  | 234.2419 | 0 |  |  |  |  | 234.2419 | 80.8647
|  | 234.3403 | 0 |  |  |  |  | 234.3403 | 80.5446
|  | 234.4386 | 0 |  |  |  |  | 234.4387 | 80.5846
|  | 234.537 | 0 |  |  |  |  | 234.537 | 80.5446
|  | 234.6354 | 56.2172 |  |  |  |  | 234.6354 | 80.3045
|  | 234.7338 | 58.2978 |  |  |  |  | 234.7338 | 80.3445
|  | 234.8322 | 59.2581 |  |  |  |  | 234.8322 | 79.9844
|  | 234.9305 | 60.0583 |  |  |  |  | 234.9306 | 79.7843
|  | 235.0289 | 60.1384 |  |  |  |  | 235.0289 | 79.9044
|  | 235.1273 | 60.6185 |  |  |  |  | 235.1273 | 80.3045
|  | 235.2257 | 61.0186 |  |  |  |  | 235.2257 | 80.6646
|  | 235.3241 | 61.6588 |  |  |  |  | 235.3241 | 80.6246
|  | 235.4224 | 62.0589 |  |  |  |  | 235.4225 | 82.3851
|  | 235.5208 | 62.6191 |  |  |  |  | 235.5208 | 84.4658
|  | 235.6192 | 62.7391 |  |  |  |  | 235.6192 | 23.6472
|  | 235.7176 | 62.179 |  |  |  |  | 235.7176 | 23.6472
|  | 235.816 | 62.379 |  |  |  |  | 235.816 | 23.6472
|  | 235.9144 | 62.4591 |  |  |  |  | 235.9144 | 0
|  | 236.0127 | 62.5391 |  |  |  |  | 236.0127 | 0
|  | 236.1111 | 62.8592 |  |  |  |  | 236.1111 | 0
|  | 236.2095 | 62.7391 |  |  |  |  | 236.2095 | 0
|  | 236.3079 | 62.379 |  |  |  |  | 236.3079 | 0
|  | 236.4062 | 62.4591 |  |  |  |  | 236.4062 | 0
|  | 236.5046 | 62.0189 |  |  |  |  | 236.5046 | 0
|  | 236.603 | 61.0986 |  |  |  |  | 236.603 | 0
|  | 236.7014 | 60.8186 |  |  |  |  | 236.7014 | 0
|  | 236.7998 | 60.5785 |  |  |  |  | 236.7998 | 0
|  | 236.8981 | 59.5382 |  |  |  |  | 236.8981 | 0
|  | 236.9965 | 58.6579 |  |  |  |  | 236.9965 | 0
|  | 237.0949 | 57.2975 |  |  |  |  | 237.0949 | 0
|  | 237.1933 | 55.2168 |  |  |  |  | 237.1933 | 0
|  | 237.2917 | 52.9762 |  |  |  |  | 237.2917 | 0
|  | 237.39 | 50.7355 |  |  |  |  | 237.39 | 0
|  | 237.4884 | 48.0147 |  |  |  |  | 237.4884 | 0
|  | 237.5868 | 47.1744 |  |  |  |  | 237.5868 | 0
|  | 237.6852 | 47.5745 |  |  |  |  | 237.6852 | 0
|  | 237.7836 | 45.974 |  |  |  |  | 237.7836 | 0
|  | 237.8819 | 41.8128 |  |  |  |  | 237.8819 | 0
|  | 237.9803 | 0 |  |  |  |  | 237.9803 | 0
|  | 238.0787 | 0 |  |  |  |  | 238.0787 | 0
|  | 238.1771 | 0 |  |  |  |  | 238.1771 | 0
|  | 238.2755 | 0 |  |  |  |  | 238.2755 | 0
|  | 238.3738 | 0 |  |  |  |  | 238.3738 | 0
[8 empty] | 238.4722 | 0
[8 empty] | 238.5706 | 0
[8 empty] | 238.669 | 0
[8 empty] | 238.7674 | 0
[8 empty] | 238.8657 | 0
[8 empty] | 238.9641 | 0
[8 empty] | 239.0625 | 0
[8 empty] | 239.1609 | 0
[8 empty] | 239.2593 | 0
[8 empty] | 239.3576 | 0
[8 empty] | 239.456 | 0
[8 empty] | 239.5544 | 0
[8 empty] | 239.6528 | 0
[8 empty] | 239.7512 | 0
[8 empty] | 239.8495 | 0
[8 empty] | 239.9479 | 0
[8 empty] | 240.0463 | 0
[8 empty] | 240.1447 | 0
[8 empty] | 240.2431 | 0
[8 empty] | 240.3414 | 0
[8 empty] | 240.4398 | 0
[8 empty] | 240.5382 | 0
[8 empty] | 240.6366 | 0
[8 empty] | 240.735 | 0
[8 empty] | 240.8333 | 0
[8 empty] | 240.9317 | 0
[8 empty] | 241.0301 | 0
[8 empty] | 241.1285 | 0
[8 empty] | 241.2269 | 0
[8 empty] | 241.3252 | 0
[8 empty] | 241.4236 | 0
[8 empty] | 241.522 | 0
[8 empty] | 241.6204 | 0
[8 empty] | 241.7188 | 0
[8 empty] | 241.8171 | 0
[8 empty] | 241.9155 | 0
[8 empty] | 242.0139 | 0
[8 empty] | 242.1123 | 0
[8 empty] | 242.2106 | 0
[8 empty] | 242.309 | 0
[8 empty] | 242.4074 | 0
[8 empty] | 242.5058 | 0
[8 empty] | 242.6042 | 0
[8 empty] | 242.7025 | 0
[8 empty] | 242.8009 | 0
[8 empty] | 242.8993 | 0
[8 empty] | 242.9977 | 0
[8 empty] | 243.0961 | 0
[8 empty] | 243.1944 | 0
[8 empty] | 243.2928 | 0
[8 empty] | 243.3912 | 0
[8 empty] | 243.4896 | 0
[8 empty] | 243.588 | 0
[8 empty] | 243.6863 | 0
[8 empty] | 243.7847 | 0
[8 empty] | 243.8831 | 0
[8 empty] | 243.9815 | 0
[8 empty] | 244.0799 | 0
[8 empty] | 244.1782 | 0
[8 empty] | 244.2766 | 0
[8 empty] | 244.375 | 0
[8 empty] | 244.4734 | 0
[8 empty] | 244.5718 | 0
[8 empty] | 244.6701 | 0
[8 empty] | 244.7685 | 0
[8 empty] | 244.8669 | 0
[8 empty] | 244.9653 | 0
[8 empty] | 245.0637 | 0
[8 empty] | 245.162 | 0
[8 empty] | 245.2604 | 0
[8 empty] | 245.3588 | 0
[8 empty] | 245.4572 | 0
[8 empty] | 245.5556 | 0
[8 empty] | 245.6539 | 0
[8 empty] | 245.7523 | 0
[8 empty] | 245.8507 | 0
[8 empty] | 245.9491 | 0
[8 empty] | 246.0475 | 0
[8 empty] | 246.1458 | 0
[8 empty] | 246.2442 | 0
[8 empty] | 246.3426 | 0
[8 empty] | 246.441 | 0
[8 empty] | 246.5394 | 0
[8 empty] | 246.6377 | 0
[8 empty] | 246.7361 | 0
[8 empty] | 246.8345 | 0
[8 empty] | 246.9329 | 0
[8 empty] | 247.0312 | 0
[8 empty] | 247.1296 | 0
[8 empty] | 247.228 | 0
[8 empty] | 247.3264 | 0
[8 empty] | 247.4248 | 0
[8 empty] | 247.5231 | 0
[8 empty] | 247.6215 | 0
[8 empty] | 247.7199 | 0
[8 empty] | 247.8183 | 0
[8 empty] | 247.9167 | 0
[8 empty] | 248.015 | 0
[8 empty] | 248.1134 | 0
[8 empty] | 248.2118 | 0
[8 empty] | 248.3102 | 0
[8 empty] | 248.4086 | 0
[8 empty] | 248.5069 | 0
[8 empty] | 248.6053 | 0
[8 empty] | 248.7037 | 0
[8 empty] | 248.8021 | 0
[8 empty] | 248.9005 | 0
[8 empty] | 248.9988 | 0
[8 empty] | 249.0972 | 0
[8 empty] | 249.1956 | 0
[8 empty] | 249.294 | 0
[8 empty] | 249.3924 | 0
[8 empty] | 249.4907 | 0
[8 empty] | 249.5891 | 0
[8 empty] | 249.6875 | 0
[8 empty] | 249.7859 | 0
[8 empty] | 249.8843 | 0
[8 empty] | 249.9826 | 0
[8 empty] | 250.081 | 0
[8 empty] | 250.1794 | 0
[8 empty] | 250.2778 | 0
[8 empty] | 250.3762 | 0
[8 empty] | 250.4745 | 0
[8 empty] | 250.5729 | 0
[8 empty] | 250.6713 | 0
[8 empty] | 250.7697 | 0
[8 empty] | 250.8681 | 0
[8 empty] | 250.9664 | 0
[8 empty] | 251.0648 | 0
[8 empty] | 251.1632 | 0
[8 empty] | 251.2616 | 0
[8 empty] | 251.36 | 0
[8 empty] | 251.4583 | 0
[8 empty] | 251.5567 | 0
[8 empty] | 251.6551 | 0
[8 empty] | 251.7535 | 0
[8 empty] | 251.8519 | 0
[8 empty] | 251.9502 | 0
[8 empty] | 252.0486 | 0
[8 empty] | 252.147 | 0
[8 empty] | 252.2454 | 0
[8 empty] | 252.3438 | 0
[8 empty] | 252.4421 | 0
[8 empty] | 252.5405 | 0
[8 empty] | 252.6389 | 0
[8 empty] | 252.7373 | 0
[8 empty] | 252.8356 | 0
[8 empty] | 252.934 | 0
[8 empty] | 253.0324 | 0
[8 empty] | 253.1308 | 0
[8 empty] | 253.2292 | 0
[8 empty] | 253.3275 | 0
[8 empty] | 253.4259 | 0
[8 empty] | 253.5243 | 0
[8 empty] | 253.6227 | 0
[8 empty] | 253.7211 | 0
[8 empty] | 253.8194 | 0
[8 empty] | 253.9178 | 0
[8 empty] | 254.0162 | 0
[8 empty] | 254.1146 | 0
[8 empty] | 254.213 | 0
[8 empty] | 254.3113 | 0
[8 empty] | 254.4097 | 0
[8 empty] | 254.5081 | 0
[8 empty] | 254.6065 | 0
[8 empty] | 254.7049 | 0
[8 empty] | 254.8032 | 0
[8 empty] | 254.9016 | 0
[8 empty] | 255 | 0
[8 empty] | 255.0984 | 0
[8 empty] | 255.1968 | 0
[8 empty] | 255.2951 | 0
[8 empty] | 255.3935 | 0
[8 empty] | 255.4919 | 0
[8 empty] | 255.5903 | 0
[8 empty] | 255.6887 | 0
[8 empty] | 255.787 | 0
[8 empty] | 255.8854 | 0
[8 empty] | 255.9838 | 0
[8 empty] | 256.0822 | 0
[8 empty] | 256.1806 | 0
[8 empty] | 256.2789 | 0
[8 empty] | 256.3773 | 0
[8 empty] | 256.4757 | 0
[8 empty] | 256.5741 | 0
[8 empty] | 256.6725 | 0
[8 empty] | 256.7708 | 0
[8 empty] | 256.8692 | 0
[8 empty] | 256.9676 | 0
[8 empty] | 257.066 | 0
[8 empty] | 257.1644 | 0
[8 empty] | 257.2627 | 0
[8 empty] | 257.3611 | 0
[8 empty] | 257.4595 | 0
[8 empty] | 257.5579 | 0
[8 empty] | 257.6562 | 0
[8 empty] | 257.7546 | 0
[8 empty] | 257.853 | 0
[8 empty] | 257.9514 | 0
[8 empty] | 258.0498 | 0
[8 empty] | 258.1481 | 0
[8 empty] | 258.2465 | 0
[8 empty] | 258.3449 | 0
[8 empty] | 258.4433 | 0
[8 empty] | 258.5417 | 0
[8 empty] | 258.64 | 0
[8 empty] | 258.7384 | 0
[8 empty] | 258.8368 | 0
[8 empty] | 258.9352 | 0
[8 empty] | 259.0336 | 0
[8 empty] | 259.1319 | 0
[8 empty] | 259.2303 | 0
[8 empty] | 259.3287 | 0
[8 empty] | 259.4271 | 0
[8 empty] | 259.5255 | 0
[8 empty] | 259.6238 | 0
[8 empty] | 259.7222 | 0
[8 empty] | 259.8206 | 0
[8 empty] | 259.919 | 0
[8 empty] | 260.0174 | 0
[8 empty] | 260.1157 | 0
[8 empty] | 260.2141 | 0
[8 empty] | 260.3125 | 0
[8 empty] | 260.4109 | 0
[8 empty] | 260.5093 | 0
[8 empty] | 260.6076 | 0
[8 empty] | 260.706 | 0
[8 empty] | 260.8044 | 0
[8 empty] | 260.9028 | 0
[8 empty] | 261.0012 | 0
[8 empty] | 261.0995 | 0
[8 empty] | 261.1979 | 0
[8 empty] | 261.2963 | 0
[8 empty] | 261.3947 | 0
[8 empty] | 261.4931 | 0
[8 empty] | 261.5914 | 0
[8 empty] | 261.6898 | 0
[8 empty] | 261.7882 | 0
[8 empty] | 261.8866 | 0
[8 empty] | 261.985 | 0
[8 empty] | 262.0833 | 0
[8 empty] | 262.1817 | 0
[8 empty] | 262.2801 | 0
[8 empty] | 262.3785 | 0
[8 empty] | 262.4769 | 0
[8 empty] | 262.5752 | 0
[8 empty] | 262.6736 | 0
[8 empty] | 262.772 | 0
[8 empty] | 262.8704 | 0
[8 empty] | 262.9688 | 0
[8 empty] | 263.0671 | 0
[8 empty] | 263.1655 | 0
[8 empty] | 263.2639 | 0
[8 empty] | 263.3623 | 0
[8 empty] | 263.4606 | 0
[8 empty] | 263.559 | 0
[8 empty] | 263.6574 | 0
[8 empty] | 263.7558 | 0
[8 empty] | 263.8542 | 0
[8 empty] | 263.9525 | 0
[8 empty] | 264.0509 | 0
[8 empty] | 264.1493 | 0
[8 empty] | 264.2477 | 0
[8 empty] | 264.3461 | 0
[8 empty] | 264.4444 | 0
[8 empty] | 264.5428 | 0
[8 empty] | 264.6412 | 0
[8 empty] | 264.7396 | 0
[8 empty] | 264.838 | 0
[8 empty] | 264.9363 | 0
[8 empty] | 265.0347 | 0
[8 empty] | 265.1331 | 0
[8 empty] | 265.2315 | 0
[8 empty] | 265.3299 | 0
[8 empty] | 265.4282 | 0
[8 empty] | 265.5266 | 0
[8 empty] | 265.625 | 0
[8 empty] | 265.7234 | 0
[8 empty] | 265.8218 | 0
[8 empty] | 265.9201 | 0
[8 empty] | 266.0185 | 0
[8 empty] | 266.1169 | 0
[8 empty] | 266.2153 | 0
[8 empty] | 266.3137 | 0
[8 empty] | 266.412 | 0
[8 empty] | 266.5104 | 0
[8 empty] | 266.6088 | 0
[8 empty] | 266.7072 | 0
[8 empty] | 266.8056 | 0
[8 empty] | 266.9039 | 33.8103
[8 empty] | 267.0023 | 60.0583
[8 empty] | 267.1007 | 79.9044
[8 empty] | 267.1991 | 80.9047
[8 empty] | 267.2975 | 80.3845
[8 empty] | 267.3958 | 80.2645
[8 empty] | 267.4942 | 80.3845
[8 empty] | 267.5926 | 80.8247
[8 empty] | 267.691 | 80.6246
[8 empty] | 267.7894 | 80.3045
[8 empty] | 267.8877 | 80.3845
[8 empty] | 267.9861 | 80.4646
[8 empty] | 268.0845 | 80.0644
[8 empty] | 268.1829 | 79.7043
[8 empty] | 268.2812 | 80.0644
[8 empty] | 268.3796 | 80.2645
[8 empty] | 268.478 | 80.7446
[8 empty] | 268.5764 | 80.7046
[8 empty] | 268.6748 | 79.9044
[8 empty] | 268.7731 | 79.9044
[8 empty] | 268.8715 | 82.3051
[8 empty] | 268.9699 | 85.8262
[8 empty] | 269.0683 | 54.4566
[8 empty] | 269.1667 | 54.4566
[8 empty] | 269.265 | 54.4566
[8 empty] | 269.3634 | 0
[8 empty] | 269.4618 | 0
[8 empty] | 269.5602 | 0
[8 empty] | 269.6586 | 0
[8 empty] | 269.7569 | 0
[8 empty] | 269.8553 | 0
[8 empty] | 269.9537 | 0
[8 empty] | 270.0521 | 0
[8 empty] | 270.1505 | 0
[8 empty] | 270.2488 | 0
[8 empty] | 270.3472 | 0
[8 empty] | 270.4456 | 0
[8 empty] | 270.544 | 0
[8 empty] | 270.6424 | 0
[8 empty] | 270.7407 | 0
[8 empty] | 270.8391 | 0
[8 empty] | 270.9375 | 0
[8 empty] | 271.0359 | 0
[8 empty] | 271.1343 | 0
[8 empty] | 271.2326 | 0
[8 empty] | 271.331 | 0
[8 empty] | 271.4294 | 0
[8 empty] | 271.5278 | 0
[8 empty] | 271.6262 | 0
[8 empty] | 271.7245 | 0
[8 empty] | 271.8229 | 0
[8 empty] | 271.9213 | 0
[8 empty] | 272.0197 | 0
[8 empty] | 272.1181 | 0
[8 empty] | 272.2164 | 0
[8 empty] | 272.3148 | 0
[8 empty] | 272.4132 | 0
[8 empty] | 272.5116 | 0
[8 empty] | 272.61 | 0
[8 empty] | 272.7083 | 0
[8 empty] | 272.8067 | 0
[8 empty] | 272.9051 | 0
[8 empty] | 273.0035 | 0
[8 empty] | 273.1019 | 0
[8 empty] | 273.2002 | 0
[8 empty] | 273.2986 | 0
[8 empty] | 273.397 | 0
[8 empty] | 273.4954 | 0
[8 empty] | 273.5938 | 0
[8 empty] | 273.6921 | 0
[8 empty] | 273.7905 | 0
[8 empty] | 273.8889 | 0
[8 empty] | 273.9873 | 0
[8 empty] | 274.0856 | 0
[8 empty] | 274.184 | 0
[8 empty] | 274.2824 | 0
[8 empty] | 274.3808 | 0
[8 empty] | 274.4792 | 0
[8 empty] | 274.5775 | 0
[8 empty] | 274.6759 | 0
[8 empty] | 274.7743 | 0
[8 empty] | 274.8727 | 0
[8 empty] | 274.9711 | 0
[8 empty] | 275.0694 | 0
[8 empty] | 275.1678 | 0
[8 empty] | 275.2662 | 0
[8 empty] | 275.3646 | 0
[8 empty] | 275.463 | 0
[8 empty] | 275.5613 | 0
[8 empty] | 275.6597 | 0
[8 empty] | 275.7581 | 0
[8 empty] | 275.8565 | 0
[8 empty] | 275.9549 | 0
[8 empty] | 276.0532 | 0
[8 empty] | 276.1516 | 0
[8 empty] | 276.25 | 0
[8 empty] | 276.3484 | 0
[8 empty] | 276.4468 | 0
[8 empty] | 276.5451 | 0
[8 empty] | 276.6435 | 0
[8 empty] | 276.7419 | 0
[8 empty] | 276.8403 | 0
[8 empty] | 276.9387 | 0
[8 empty] | 277.037 | 0
[8 empty] | 277.1354 | 0
[8 empty] | 277.2338 | 0
[8 empty] | 277.3322 | 0
[8 empty] | 277.4306 | 0
[8 empty] | 277.5289 | 0
[8 empty] | 277.6273 | 0
[8 empty] | 277.7257 | 0
[8 empty] | 277.8241 | 0
[8 empty] | 277.9225 | 0
[8 empty] | 278.0208 | 0
[8 empty] | 278.1192 | 0
[8 empty] | 278.2176 | 0
[8 empty] | 278.316 | 0
[8 empty] | 278.4144 | 0
[8 empty] | 278.5127 | 0
[8 empty] | 278.6111 | 0
[8 empty] | 278.7095 | 0
[8 empty] | 278.8079 | 0
[8 empty] | 278.9062 | 0
[8 empty] | 279.0046 | 0
[8 empty] | 279.103 | 0
[8 empty] | 279.2014 | 0
[8 empty] | 279.2998 | 0
[8 empty] | 279.3981 | 0
[8 empty] | 279.4965 | 0
[8 empty] | 279.5949 | 0
[8 empty] | 279.6933 | 0
[8 empty] | 279.7917 | 0
[8 empty] | 279.89 | 0
[8 empty] | 279.9884 | 0
[8 empty] | 280.0868 | 0
[8 empty] | 280.1852 | 0
[8 empty] | 280.2836 | 0
[8 empty] | 280.3819 | 0
[8 empty] | 280.4803 | 0
[8 empty] | 280.5787 | 0
[8 empty] | 280.6771 | 0
[8 empty] | 280.7755 | 0
[8 empty] | 280.8738 | 0
[8 empty] | 280.9722 | 0
[8 empty] | 281.0706 | 0
[8 empty] | 281.169 | 0
[8 empty] | 281.2674 | 0
[8 empty] | 281.3657 | 0
[8 empty] | 281.4641 | 0
[8 empty] | 281.5625 | 0
[8 empty] | 281.6609 | 0
[8 empty] | 281.7593 | 0
[8 empty] | 281.8576 | 0
[8 empty] | 281.956 | 0
[8 empty] | 282.0544 | 0
[8 empty] | 282.1528 | 0
[8 empty] | 282.2512 | 0
[8 empty] | 282.3495 | 0
[8 empty] | 282.4479 | 0
[8 empty] | 282.5463 | 0
[8 empty] | 282.6447 | 0
[8 empty] | 282.7431 | 0
[8 empty] | 282.8414 | 0
[8 empty] | 282.9398 | 0
[8 empty] | 283.0382 | 0
[8 empty] | 283.1366 | 0
[8 empty] | 283.235 | 0
[8 empty] | 283.3333 | 0
[8 empty] | 283.4317 | 0
[8 empty] | 283.5301 | 0
[8 empty] | 283.6285 | 0
[8 empty] | 283.7269 | 0
[8 empty] | 283.8252 | 0
[8 empty] | 283.9236 | 0
[8 empty] | 284.022 | 0
[8 empty] | 284.1204 | 0
[8 empty] | 284.2188 | 0
[8 empty] | 284.3171 | 0
[8 empty] | 284.4155 | 0
[8 empty] | 284.5139 | 0
[8 empty] | 284.6123 | 0
[8 empty] | 284.7106 | 0
[8 empty] | 284.809 | 0
[8 empty] | 284.9074 | 0
[8 empty] | 285.0058 | 0
[8 empty] | 285.1042 | 0
[8 empty] | 285.2025 | 0
[8 empty] | 285.3009 | 0
[8 empty] | 285.3993 | 0
[8 empty] | 285.4977 | 0
[8 empty] | 285.5961 | 0
[8 empty] | 285.6944 | 0
[8 empty] | 285.7928 | 0
[8 empty] | 285.8912 | 0
[8 empty] | 285.9896 | 0
[8 empty] | 286.088 | 0
[8 empty] | 286.1863 | 0
[8 empty] | 286.2847 | 0
[8 empty] | 286.3831 | 0
[8 empty] | 286.4815 | 0
[8 empty] | 286.5799 | 0
[8 empty] | 286.6782 | 0
[8 empty] | 286.7766 | 0
[8 empty] | 286.875 | 0
[8 empty] | 286.9734 | 0
[8 empty] | 287.0718 | 0
[8 empty] | 287.1701 | 0
[8 empty] | 287.2685 | 0
[8 empty] | 287.3669 | 0
[8 empty] | 287.4653 | 0
[8 empty] | 287.5637 | 0
[8 empty] | 287.662 | 0
[8 empty] | 287.7604 | 0
[8 empty] | 287.8588 | 0
[8 empty] | 287.9572 | 0
[8 empty] | 288.0556 | 0
[8 empty] | 288.1539 | 0
[8 empty] | 288.2523 | 0
[8 empty] | 288.3507 | 0
[8 empty] | 288.4491 | 0
[8 empty] | 288.5475 | 0
[8 empty] | 288.6458 | 0
[8 empty] | 288.7442 | 0
[8 empty] | 288.8426 | 0
[8 empty] | 288.941 | 0
[8 empty] | 289.0394 | 0
[8 empty] | 289.1377 | 0
[8 empty] | 289.2361 | 0
[8 empty] | 289.3345 | 0
[8 empty] | 289.4329 | 0
[8 empty] | 289.5312 | 0
[8 empty] | 289.6296 | 0
[8 empty] | 289.728 | 0
[8 empty] | 289.8264 | 0
[8 empty] | 289.9248 | 0
[8 empty] | 290.0231 | 0
[8 empty] | 290.1215 | 0
[8 empty] | 290.2199 | 0
[8 empty] | 290.3183 | 0
[8 empty] | 290.4167 | 0
[8 empty] | 290.515 | 0
[8 empty] | 290.6134 | 0
[8 empty] | 290.7118 | 0
[8 empty] | 290.8102 | 0
[8 empty] | 290.9086 | 0
[8 empty] | 291.0069 | 0
[8 empty] | 291.1053 | 0
[8 empty] | 291.2037 | 0
[8 empty] | 291.3021 | 0
[8 empty] | 291.4005 | 0
[8 empty] | 291.4988 | 0
[8 empty] | 291.5972 | 0
[8 empty] | 291.6956 | 0
[8 empty] | 291.794 | 0
[8 empty] | 291.8924 | 0
[8 empty] | 291.9907 | 0
[8 empty] | 292.0891 | 0
[8 empty] | 292.1875 | 0
[8 empty] | 292.2859 | 0
[8 empty] | 292.3843 | 0
[8 empty] | 292.4826 | 0
[8 empty] | 292.581 | 0
[8 empty] | 292.6794 | 0
[8 empty] | 292.7778 | 0
[8 empty] | 292.8762 | 0
[8 empty] | 292.9745 | 0
[8 empty] | 293.0729 | 0
[8 empty] | 293.1713 | 0
[8 empty] | 293.2697 | 0
[8 empty] | 293.3681 | 0
[8 empty] | 293.4664 | 0
[8 empty] | 293.5648 | 0
[8 empty] | 293.6632 | 0
[8 empty] | 293.7616 | 0
[8 empty] | 293.86 | 0
[8 empty] | 293.9583 | 0
[8 empty] | 294.0567 | 0
[8 empty] | 294.1551 | 0
[8 empty] | 294.2535 | 0
[8 empty] | 294.3519 | 0
[8 empty] | 294.4502 | 0
[8 empty] | 294.5486 | 0
[8 empty] | 294.647 | 0
[8 empty] | 294.7454 | 0
[8 empty] | 294.8438 | 0
[8 empty] | 294.9421 | 0
[8 empty] | 295.0405 | 0
[8 empty] | 295.1389 | 0
[8 empty] | 295.2373 | 0
[8 empty] | 295.3356 | 0
[8 empty] | 295.434 | 0
[8 empty] | 295.5324 | 0
[8 empty] | 295.6308 | 0
[8 empty] | 295.7292 | 0
[8 empty] | 295.8275 | 0
[8 empty] | 295.9259 | 0
[8 empty] | 296.0243 | 0
[8 empty] | 296.1227 | 0
[8 empty] | 296.2211 | 0
[8 empty] | 296.3194 | 0
[8 empty] | 296.4178 | 0
[8 empty] | 296.5162 | 0
[8 empty] | 296.6146 | 0
[8 empty] | 296.713 | 0
[8 empty] | 296.8113 | 0
[8 empty] | 296.9097 | 0
[8 empty] | 297.0081 | 0
[8 empty] | 297.1065 | 0
[8 empty] | 297.2049 | 0
[8 empty] | 297.3032 | 0
[8 empty] | 297.4016 | 0
[8 empty] | 297.5 | 0
[8 empty] | 297.5984 | 0
[8 empty] | 297.6968 | 0
[8 empty] | 297.7951 | 0
[8 empty] | 297.8935 | 0
[8 empty] | 297.9919 | 0
[8 empty] | 298.0903 | 0
[8 empty] | 298.1887 | 0
[8 empty] | 298.287 | 0
[8 empty] | 298.3854 | 0
[8 empty] | 298.4838 | 0
[8 empty] | 298.5822 | 0
[8 empty] | 298.6806 | 0
[8 empty] | 298.7789 | 0
[8 empty] | 298.8773 | 0
[8 empty] | 298.9757 | 0
[8 empty] | 299.0741 | 0
[8 empty] | 299.1725 | 0
[8 empty] | 299.2708 | 0
[8 empty] | 299.3692 | 0
[8 empty] | 299.4676 | 0
[8 empty] | 299.566 | 0
[8 empty] | 299.6644 | 0
[8 empty] | 299.7627 | 0
[8 empty] | 299.8611 | 0
[8 empty] | 299.9595 | 0
[8 empty] | 300.0579 | 0
[8 empty] | 300.1562 | 0
[8 empty] | 300.2546 | 0
[8 empty] | 300.353 | 48.5348
[8 empty] | 300.4514 | 78.624
[8 empty] | 300.5498 | 81.1848
[8 empty] | 300.6481 | 81.2648
[8 empty] | 300.7465 | 80.7847
[8 empty] | 300.8449 | 80.0644
[8 empty] | 300.9433 | 79.9444
[8 empty] | 301.0417 | 80.3845
[8 empty] | 301.14 | 80.1044
[8 empty] | 301.2384 | 80.0644
[8 empty] | 301.3368 | 80.1445
[8 empty] | 301.4352 | 80.4646
[8 empty] | 301.5336 | 80.9447
[8 empty] | 301.6319 | 80.4245
[8 empty] | 301.7303 | 79.9844
[8 empty] | 301.8287 | 80.4245
[8 empty] | 301.9271 | 80.4245
[8 empty] | 302.0255 | 79.9844
[8 empty] | 302.1238 | 80.1445
[8 empty] | 302.2222 | 80.3445
[8 empty] | 302.3206 | 84.5058
[8 empty] | 302.419 | 87.3867
[8 empty] | 302.5174 | 87.3867
[8 empty] | 302.6157 | 87.3867
[8 empty] | 302.7141 | 0
[8 empty] | 302.8125 | 0
[8 empty] | 302.9109 | 0
[8 empty] | 303.0093 | 0
[8 empty] | 303.1076 | 0
[8 empty] | 303.206 | 0
[8 empty] | 303.3044 | 0
[8 empty] | 303.4028 | 0
[8 empty] | 303.5012 | 0
[8 empty] | 303.5995 | 0
[8 empty] | 303.6979 | 0
[8 empty] | 303.7963 | 0
[8 empty] | 303.8947 | 0
[8 empty] | 303.9931 | 0
[8 empty] | 304.0914 | 0
[8 empty] | 304.1898 | 0
[8 empty] | 304.2882 | 0
[8 empty] | 304.3866 | 0
[8 empty] | 304.485 | 0
[8 empty] | 304.5833 | 0
[8 empty] | 304.6817 | 0
[8 empty] | 304.7801 | 0
[8 empty] | 304.8785 | 0
[8 empty] | 304.9769 | 0
[8 empty] | 305.0752 | 0
[8 empty] | 305.1736 | 0
[8 empty] | 305.272 | 0
[8 empty] | 305.3704 | 0
[8 empty] | 305.4688 | 0
[8 empty] | 305.5671 | 0
[8 empty] | 305.6655 | 0
[8 empty] | 305.7639 | 0
[8 empty] | 305.8623 | 0
[8 empty] | 305.9606 | 0
[8 empty] | 306.059 | 0
[8 empty] | 306.1574 | 0
[8 empty] | 306.2558 | 0
[8 empty] | 306.3542 | 0
[8 empty] | 306.4525 | 0
[8 empty] | 306.5509 | 0
[8 empty] | 306.6493 | 0
[8 empty] | 306.7477 | 0
[8 empty] | 306.8461 | 0
[8 empty] | 306.9444 | 0
[8 empty] | 307.0428 | 0
[8 empty] | 307.1412 | 0
[8 empty] | 307.2396 | 0
[8 empty] | 307.338 | 0
[8 empty] | 307.4363 | 0
[8 empty] | 307.5347 | 0
[8 empty] | 307.6331 | 0
[8 empty] | 307.7315 | 0
[8 empty] | 307.8299 | 0
[8 empty] | 307.9282 | 0
[8 empty] | 308.0266 | 0
[8 empty] | 308.125 | 0
[8 empty] | 308.2234 | 0
[8 empty] | 308.3218 | 0
[8 empty] | 308.4201 | 0
[8 empty] | 308.5185 | 0
[8 empty] | 308.6169 | 0
[8 empty] | 308.7153 | 0
[8 empty] | 308.8137 | 0
[8 empty] | 308.912 | 0
[8 empty] | 309.0104 | 0
[8 empty] | 309.1088 | 0
[8 empty] | 309.2072 | 0
[8 empty] | 309.3056 | 0
[8 empty] | 309.4039 | 0
[8 empty] | 309.5023 | 0
[8 empty] | 309.6007 | 0
[8 empty] | 309.6991 | 0
[8 empty] | 309.7975 | 0
[8 empty] | 309.8958 | 0
[8 empty] | 309.9942 | 0
[8 empty] | 310.0926 | 0
[8 empty] | 310.191 | 0
[8 empty] | 310.2894 | 0
[8 empty] | 310.3877 | 0
[8 empty] | 310.4861 | 0
[8 empty] | 310.5845 | 0
[8 empty] | 310.6829 | 0
[8 empty] | 310.7812 | 0
[8 empty] | 310.8796 | 0
[8 empty] | 310.978 | 0
[8 empty] | 311.0764 | 0
[8 empty] | 311.1748 | 0
[8 empty] | 311.2731 | 0
[8 empty] | 311.3715 | 0
[8 empty] | 311.4699 | 0
[8 empty] | 311.5683 | 0
[8 empty] | 311.6667 | 0
[8 empty] | 311.765 | 0
[8 empty] | 311.8634 | 0
[8 empty] | 311.9618 | 0
[8 empty] | 312.0602 | 0
[8 empty] | 312.1586 | 0
[8 empty] | 312.2569 | 0
[8 empty] | 312.3553 | 0
[8 empty] | 312.4537 | 0
[8 empty] | 312.5521 | 0
[8 empty] | 312.6505 | 0
[8 empty] | 312.7488 | 0
[8 empty] | 312.8472 | 0
[8 empty] | 312.9456 | 0
[8 empty] | 313.044 | 0
[8 empty] | 313.1424 | 0
[8 empty] | 313.2407 | 0
[8 empty] | 313.3391 | 0
[8 empty] | 313.4375 | 0
[8 empty] | 313.5359 | 0
[8 empty] | 313.6343 | 0
[8 empty] | 313.7326 | 0
[8 empty] | 313.831 | 0
[8 empty] | 313.9294 | 0
[8 empty] | 314.0278 | 0
[8 empty] | 314.1262 | 0
[8 empty] | 314.2245 | 0
[8 empty] | 314.3229 | 0
[8 empty] | 314.4213 | 0
[8 empty] | 314.5197 | 0
[8 empty] | 314.6181 | 0
[8 empty] | 314.7164 | 0
[8 empty] | 314.8148 | 0
[8 empty] | 314.9132 | 0
[8 empty] | 315.0116 | 0
[8 empty] | 315.11 | 0
[8 empty] | 315.2083 | 0
[8 empty] | 315.3067 | 0
[8 empty] | 315.4051 | 0
[8 empty] | 315.5035 | 0
[8 empty] | 315.6019 | 0
[8 empty] | 315.7002 | 0
[8 empty] | 315.7986 | 0
[8 empty] | 315.897 | 0
[8 empty] | 315.9954 | 0
[8 empty] | 316.0938 | 0
[8 empty] | 316.1921 | 0
[8 empty] | 316.2905 | 0
[8 empty] | 316.3889 | 0
[8 empty] | 316.4873 | 0
[8 empty] | 316.5856 | 0
[8 empty] | 316.684 | 0
[8 empty] | 316.7824 | 0
[8 empty] | 316.8808 | 0
[8 empty] | 316.9792 | 0
[8 empty] | 317.0775 | 0
[8 empty] | 317.1759 | 0
[8 empty] | 317.2743 | 0
[8 empty] | 317.3727 | 0
[8 empty] | 317.4711 | 0
[8 empty] | 317.5694 | 0
[8 empty] | 317.6678 | 0
[8 empty] | 317.7662 | 0
[8 empty] | 317.8646 | 0
[8 empty] | 317.963 | 0
[8 empty] | 318.0613 | 0
[8 empty] | 318.1597 | 0
[8 empty] | 318.2581 | 0
[8 empty] | 318.3565 | 0
[8 empty] | 318.4549 | 0
[8 empty] | 318.5532 | 0
[8 empty] | 318.6516 | 0
[8 empty] | 318.75 | 0
[8 empty] | 318.8484 | 0
[8 empty] | 318.9468 | 0
[8 empty] | 319.0451 | 0
[8 empty] | 319.1435 | 0
[8 empty] | 319.2419 | 0
[8 empty] | 319.3403 | 0
[8 empty] | 319.4387 | 0
[8 empty] | 319.537 | 0
[8 empty] | 319.6354 | 0
[8 empty] | 319.7338 | 0
[8 empty] | 319.8322 | 0
[8 empty] | 319.9306 | 0
[8 empty] | 320.0289 | 0
[8 empty] | 320.1273 | 0
[8 empty] | 320.2257 | 0
[8 empty] | 320.3241 | 0
[8 empty] | 320.4225 | 0
[8 empty] | 320.5208 | 0
[8 empty] | 320.6192 | 0
[8 empty] | 320.7176 | 0
[8 empty] | 320.816 | 0
[8 empty] | 320.9144 | 0
[8 empty] | 321.0127 | 0
[8 empty] | 321.1111 | 0
[8 empty] | 321.2095 | 0
[8 empty] | 321.3079 | 0
[8 empty] | 321.4062 | 0
[8 empty] | 321.5046 | 0
[8 empty] | 321.603 | 0
[8 empty] | 321.7014 | 0
[8 empty] | 321.7998 | 0
[8 empty] | 321.8981 | 0
[8 empty] | 321.9965 | 0
[8 empty] | 322.0949 | 0
[8 empty] | 322.1933 | 0
[8 empty] | 322.2917 | 0
[8 empty] | 322.39 | 0
[8 empty] | 322.4884 | 0
[8 empty] | 322.5868 | 0
[8 empty] | 322.6852 | 0
[8 empty] | 322.7836 | 0
[8 empty] | 322.8819 | 0
[8 empty] | 322.9803 | 0
[8 empty] | 323.0787 | 0
[8 empty] | 323.1771 | 0
[8 empty] | 323.2755 | 0
[8 empty] | 323.3738 | 0
[8 empty] | 323.4722 | 0
[8 empty] | 323.5706 | 0
[8 empty] | 323.669 | 0
[8 empty] | 323.7674 | 0
[8 empty] | 323.8657 | 0
[8 empty] | 323.9641 | 0
[8 empty] | 324.0625 | 0
[8 empty] | 324.1609 | 0
[8 empty] | 324.2593 | 0
[8 empty] | 324.3576 | 0
[8 empty] | 324.456 | 0
[8 empty] | 324.5544 | 0
[8 empty] | 324.6528 | 0
[8 empty] | 324.7512 | 0
[8 empty] | 324.8495 | 0
[8 empty] | 324.9479 | 0
[8 empty] | 325.0463 | 0
[8 empty] | 325.1447 | 0
[8 empty] | 325.2431 | 0
[8 empty] | 325.3414 | 0
[8 empty] | 325.4398 | 0
[8 empty] | 325.5382 | 0
[8 empty] | 325.6366 | 0
[8 empty] | 325.735 | 0
[8 empty] | 325.8333 | 0
[8 empty] | 325.9317 | 0
[8 empty] | 326.0301 | 0
[8 empty] | 326.1285 | 0
[8 empty] | 326.2269 | 0
[8 empty] | 326.3252 | 0
[8 empty] | 326.4236 | 0
[8 empty] | 326.522 | 0
[8 empty] | 326.6204 | 0
[8 empty] | 326.7188 | 0
[8 empty] | 326.8171 | 0
[8 empty] | 326.9155 | 0
[8 empty] | 327.0139 | 0
[8 empty] | 327.1123 | 0
[8 empty] | 327.2106 | 0
[8 empty] | 327.309 | 0
[8 empty] | 327.4074 | 0
[8 empty] | 327.5058 | 0
[8 empty] | 327.6042 | 0
[8 empty] | 327.7025 | 0
[8 empty] | 327.8009 | 0
[8 empty] | 327.8993 | 0
[8 empty] | 327.9977 | 0
[8 empty] | 328.0961 | 0
[8 empty] | 328.1944 | 0
[8 empty] | 328.2928 | 0
[8 empty] | 328.3912 | 0
[8 empty] | 328.4896 | 0
[8 empty] | 328.588 | 0
[8 empty] | 328.6863 | 0
[8 empty] | 328.7847 | 0
[8 empty] | 328.8831 | 0
[8 empty] | 328.9815 | 0
[8 empty] | 329.0799 | 0
[8 empty] | 329.1782 | 0
[8 empty] | 329.2766 | 0
[8 empty] | 329.375 | 0
[8 empty] | 329.4734 | 0
[8 empty] | 329.5718 | 0
[8 empty] | 329.6701 | 0
[8 empty] | 329.7685 | 0
[8 empty] | 329.8669 | 0
[8 empty] | 329.9653 | 0
[8 empty] | 330.0637 | 0
[8 empty] | 330.162 | 0
[8 empty] | 330.2604 | 0
[8 empty] | 330.3588 | 0
[8 empty] | 330.4572 | 0
[8 empty] | 330.5556 | 0
[8 empty] | 330.6539 | 0
[8 empty] | 330.7523 | 0
[8 empty] | 330.8507 | 0
[8 empty] | 330.9491 | 0
[8 empty] | 331.0475 | 0
[8 empty] | 331.1458 | 0
[8 empty] | 331.2442 | 0
[8 empty] | 331.3426 | 0
[8 empty] | 331.441 | 0
[8 empty] | 331.5394 | 0
[8 empty] | 331.6377 | 0
[8 empty] | 331.7361 | 0
[8 empty] | 331.8345 | 0
[8 empty] | 331.9329 | 0
[8 empty] | 332.0312 | 0
[8 empty] | 332.1296 | 0
[8 empty] | 332.228 | 0
[8 empty] | 332.3264 | 0
[8 empty] | 332.4248 | 0
[8 empty] | 332.5231 | 0
[8 empty] | 332.6215 | 0
[8 empty] | 332.7199 | 0
[8 empty] | 332.8183 | 0
[8 empty] | 332.9167 | 0
[8 empty] | 333.015 | 0
[8 empty] | 333.1134 | 0
[8 empty] | 333.2118 | 0
[8 empty] | 333.3102 | 0
[8 empty] | 333.4086 | 0
[8 empty] | 333.5069 | 0
[8 empty] | 333.6053 | 0
[8 empty] | 333.7037 | 0
[8 empty] | 333.8021 | 36.3711
[8 empty] | 333.9005 | 69.0211
[8 empty] | 333.9988 | 79.2642
[8 empty] | 334.0972 | 80.6246
[8 empty] | 334.1956 | 80.5446
[8 empty] | 334.294 | 80.3045
[8 empty] | 334.3924 | 80.1044
[8 empty] | 334.4907 | 80.0244
[8 empty] | 334.5891 | 80.3845
[8 empty] | 334.6875 | 80.2645
[8 empty] | 334.7859 | 80.0644
[8 empty] | 334.8843 | 79.8244
[8 empty] | 334.9826 | 79.7443
[8 empty] | 335.081 | 80.3045
[8 empty] | 335.1794 | 80.5446
[8 empty] | 335.2778 | 80.5046
[8 empty] | 335.3762 | 80.3845
[8 empty] | 335.4745 | 79.9044
[8 empty] | 335.5729 | 79.9844
[8 empty] | 335.6713 | 80.3445
[8 empty] | 335.7697 | 83.2254
[8 empty] | 335.8681 | 87.2266
[8 empty] | 335.9664 | 31.5296
[8 empty] | 336.0648 | 31.5296
[8 empty] | 336.1632 | 31.5296
[8 empty] | 336.2616 | 0
[8 empty] | 336.36 | 0
[8 empty] | 336.4583 | 0
[8 empty] | 336.5567 | 0
[8 empty] | 336.6551 | 0
[8 empty] | 336.7535 | 0
[8 empty] | 336.8519 | 0
[8 empty] | 336.9502 | 0
[8 empty] | 337.0486 | 0
[8 empty] | 337.147 | 0
[8 empty] | 337.2454 | 0
[8 empty] | 337.3438 | 0
[8 empty] | 337.4421 | 0
[8 empty] | 337.5405 | 0
[8 empty] | 337.6389 | 0
[8 empty] | 337.7373 | 0
[8 empty] | 337.8356 | 0
[8 empty] | 337.934 | 0
[8 empty] | 338.0324 | 0
[8 empty] | 338.1308 | 0
[8 empty] | 338.2292 | 0
[8 empty] | 338.3275 | 0
[8 empty] | 338.4259 | 0
[8 empty] | 338.5243 | 0
[8 empty] | 338.6227 | 0
[8 empty] | 338.7211 | 0
[8 empty] | 338.8194 | 0
[8 empty] | 338.9178 | 0
[8 empty] | 339.0162 | 0
[8 empty] | 339.1146 | 0
[8 empty] | 339.213 | 0
[8 empty] | 339.3113 | 0
[8 empty] | 339.4097 | 0
[8 empty] | 339.5081 | 0
[8 empty] | 339.6065 | 0
[8 empty] | 339.7049 | 0
[8 empty] | 339.8032 | 0
[8 empty] | 339.9016 | 0
[8 empty] | 340 | 0
[8 empty] | 340.0984 | 0
[8 empty] | 340.1968 | 0
[8 empty] | 340.2951 | 0
[8 empty] | 340.3935 | 0
[8 empty] | 340.4919 | 0
[8 empty] | 340.5903 | 0
[8 empty] | 340.6887 | 0
[8 empty] | 340.787 | 0
[8 empty] | 340.8854 | 0
[8 empty] | 340.9838 | 0
[8 empty] | 341.0822 | 0
[8 empty] | 341.1806 | 0
[8 empty] | 341.2789 | 0
[8 empty] | 341.3773 | 0
[8 empty] | 341.4757 | 0
[8 empty] | 341.5741 | 0
[8 empty] | 341.6725 | 0
[8 empty] | 341.7708 | 0
[8 empty] | 341.8692 | 0
[8 empty] | 341.9676 | 0
[8 empty] | 342.066 | 0
[8 empty] | 342.1644 | 0
[8 empty] | 342.2627 | 0
[8 empty] | 342.3611 | 0
[8 empty] | 342.4595 | 0
[8 empty] | 342.5579 | 0
[8 empty] | 342.6562 | 0
[8 empty] | 342.7546 | 0
[8 empty] | 342.853 | 0
[8 empty] | 342.9514 | 0
[8 empty] | 343.0498 | 0
[8 empty] | 343.1481 | 0
[8 empty] | 343.2465 | 0
[8 empty] | 343.3449 | 0
[8 empty] | 343.4433 | 0
[8 empty] | 343.5417 | 0
[8 empty] | 343.64 | 0
[8 empty] | 343.7384 | 0
[8 empty] | 343.8368 | 0
[8 empty] | 343.9352 | 0
[8 empty] | 344.0336 | 0
[8 empty] | 344.1319 | 0
[8 empty] | 344.2303 | 0
[8 empty] | 344.3287 | 0
[8 empty] | 344.4271 | 0
[8 empty] | 344.5255 | 0
[8 empty] | 344.6238 | 0
[8 empty] | 344.7222 | 0
[8 empty] | 344.8206 | 0
[8 empty] | 344.919 | 0
[8 empty] | 345.0174 | 0
[8 empty] | 345.1157 | 0
[8 empty] | 345.2141 | 0
[8 empty] | 345.3125 | 0
[8 empty] | 345.4109 | 0
[8 empty] | 345.5093 | 0
[8 empty] | 345.6076 | 0
[8 empty] | 345.706 | 0
[8 empty] | 345.8044 | 0
[8 empty] | 345.9028 | 0
[8 empty] | 346.0012 | 0
[8 empty] | 346.0995 | 0
[8 empty] | 346.1979 | 0
[8 empty] | 346.2963 | 0
[8 empty] | 346.3947 | 0
[8 empty] | 346.4931 | 0
[8 empty] | 346.5914 | 0
[8 empty] | 346.6898 | 0
[8 empty] | 346.7882 | 0
[8 empty] | 346.8866 | 0
[8 empty] | 346.985 | 0
[8 empty] | 347.0833 | 0
[8 empty] | 347.1817 | 0
[8 empty] | 347.2801 | 0
[8 empty] | 347.3785 | 0
[8 empty] | 347.4769 | 0
[8 empty] | 347.5752 | 0
[8 empty] | 347.6736 | 0
[8 empty] | 347.772 | 0
[8 empty] | 347.8704 | 0
[8 empty] | 347.9688 | 0
[8 empty] | 348.0671 | 0
[8 empty] | 348.1655 | 0
[8 empty] | 348.2639 | 0
[8 empty] | 348.3623 | 0
[8 empty] | 348.4606 | 0
[8 empty] | 348.559 | 0
[8 empty] | 348.6574 | 0
[8 empty] | 348.7558 | 0
[8 empty] | 348.8542 | 0
[8 empty] | 348.9525 | 0
[8 empty] | 349.0509 | 0
[8 empty] | 349.1493 | 0
[8 empty] | 349.2477 | 0
[8 empty] | 349.3461 | 0
[8 empty] | 349.4444 | 0
[8 empty] | 349.5428 | 0
[8 empty] | 349.6412 | 0
[8 empty] | 349.7396 | 0
[8 empty] | 349.838 | 0
[8 empty] | 349.9363 | 0
[8 empty] | 350.0347 | 0
[8 empty] | 350.1331 | 0
[8 empty] | 350.2315 | 0
[8 empty] | 350.3299 | 0
[8 empty] | 350.4282 | 0
[8 empty] | 350.5266 | 0
[8 empty] | 350.625 | 0
[8 empty] | 350.7234 | 0
[8 empty] | 350.8218 | 0
[8 empty] | 350.9201 | 0
[8 empty] | 351.0185 | 0
[8 empty] | 351.1169 | 0
[8 empty] | 351.2153 | 0
[8 empty] | 351.3137 | 0
[8 empty] | 351.412 | 0
[8 empty] | 351.5104 | 0
[8 empty] | 351.6088 | 0
[8 empty] | 351.7072 | 0
[8 empty] | 351.8056 | 0
[8 empty] | 351.9039 | 0
[8 empty] | 352.0023 | 0
[8 empty] | 352.1007 | 0
[8 empty] | 352.1991 | 0
[8 empty] | 352.2975 | 0
[8 empty] | 352.3958 | 0
[8 empty] | 352.4942 | 0
[8 empty] | 352.5926 | 0
[8 empty] | 352.691 | 0
[8 empty] | 352.7894 | 0
[8 empty] | 352.8877 | 0
[8 empty] | 352.9861 | 0
[8 empty] | 353.0845 | 0
[8 empty] | 353.1829 | 0
[8 empty] | 353.2812 | 0
[8 empty] | 353.3796 | 0
[8 empty] | 353.478 | 0
[8 empty] | 353.5764 | 0
[8 empty] | 353.6748 | 0
[8 empty] | 353.7731 | 0
[8 empty] | 353.8715 | 0
[8 empty] | 353.9699 | 0
[8 empty] | 354.0683 | 0
[8 empty] | 354.1667 | 0
[8 empty] | 354.265 | 0
[8 empty] | 354.3634 | 0
[8 empty] | 354.4618 | 0
[8 empty] | 354.5602 | 0
[8 empty] | 354.6586 | 0
[8 empty] | 354.7569 | 0
[8 empty] | 354.8553 | 0
[8 empty] | 354.9537 | 0
[8 empty] | 355.0521 | 0
[8 empty] | 355.1505 | 0
[8 empty] | 355.2488 | 0
[8 empty] | 355.3472 | 0
[8 empty] | 355.4456 | 0
[8 empty] | 355.544 | 0
[8 empty] | 355.6424 | 0
[8 empty] | 355.7407 | 0
[8 empty] | 355.8391 | 0
[8 empty] | 355.9375 | 0
[8 empty] | 356.0359 | 0
[8 empty] | 356.1343 | 0
[8 empty] | 356.2326 | 0
[8 empty] | 356.331 | 0
[8 empty] | 356.4294 | 0
[8 empty] | 356.5278 | 0
[8 empty] | 356.6262 | 0
[8 empty] | 356.7245 | 0
[8 empty] | 356.8229 | 0
[8 empty] | 356.9213 | 0
[8 empty] | 357.0197 | 0
[8 empty] | 357.1181 | 0
[8 empty] | 357.2164 | 0
[8 empty] | 357.3148 | 0
[8 empty] | 357.4132 | 0
[8 empty] | 357.5116 | 0
[8 empty] | 357.61 | 0
[8 empty] | 357.7083 | 0
[8 empty] | 357.8067 | 0
[8 empty] | 357.9051 | 0
[8 empty] | 358.0035 | 0
[8 empty] | 358.1019 | 0
[8 empty] | 358.2002 | 0
[8 empty] | 358.2986 | 0
[8 empty] | 358.397 | 0
[8 empty] | 358.4954 | 0
[8 empty] | 358.5938 | 0
[8 empty] | 358.6921 | 0
[8 empty] | 358.7905 | 0
[8 empty] | 358.8889 | 0
[8 empty] | 358.9873 | 0
[8 empty] | 359.0856 | 0
[8 empty] | 359.184 | 0
[8 empty] | 359.2824 | 0
[8 empty] | 359.3808 | 0
[8 empty] | 359.4792 | 0
[8 empty] | 359.5775 | 0
[8 empty] | 359.6759 | 0
[8 empty] | 359.7743 | 0
[8 empty] | 359.8727 | 0
[8 empty] | 359.9711 | 0
[8 empty] | 360.0694 | 0
[8 empty] | 360.1678 | 0
[8 empty] | 360.2662 | 0
[8 empty] | 360.3646 | 0
[8 empty] | 360.463 | 0
[8 empty] | 360.5613 | 0
[8 empty] | 360.6597 | 0
[8 empty] | 360.7581 | 0
[8 empty] | 360.8565 | 0
[8 empty] | 360.9549 | 0
[8 empty] | 361.0532 | 0
[8 empty] | 361.1516 | 0
[8 empty] | 361.25 | 0
[8 empty] | 361.3484 | 0
[8 empty] | 361.4468 | 0
[8 empty] | 361.5451 | 0
[8 empty] | 361.6435 | 0
[8 empty] | 361.7419 | 0
[8 empty] | 361.8403 | 0
[8 empty] | 361.9387 | 0
[8 empty] | 362.037 | 0
[8 empty] | 362.1354 | 0
[8 empty] | 362.2338 | 0
[8 empty] | 362.3322 | 0
[8 empty] | 362.4306 | 0
[8 empty] | 362.5289 | 0
[8 empty] | 362.6273 | 0
[8 empty] | 362.7257 | 0
[8 empty] | 362.8241 | 0
[8 empty] | 362.9225 | 0
[8 empty] | 363.0208 | 0
[8 empty] | 363.1192 | 0
[8 empty] | 363.2176 | 0
[8 empty] | 363.316 | 0
[8 empty] | 363.4144 | 0
[8 empty] | 363.5127 | 0
[8 empty] | 363.6111 | 0
[8 empty] | 363.7095 | 0
[8 empty] | 363.8079 | 0
[8 empty] | 363.9062 | 0
[8 empty] | 364.0046 | 0
[8 empty] | 364.103 | 0
[8 empty] | 364.2014 | 0
[8 empty] | 364.2998 | 0
[8 empty] | 364.3981 | 0
[8 empty] | 364.4965 | 0
[8 empty] | 364.5949 | 0
[8 empty] | 364.6933 | 0
[8 empty] | 364.7917 | 0
[8 empty] | 364.89 | 0
[8 empty] | 364.9884 | 0
[8 empty] | 365.0868 | 0
[8 empty] | 365.1852 | 0
[8 empty] | 365.2836 | 0
[8 empty] | 365.3819 | 0
[8 empty] | 365.4803 | 0
[8 empty] | 365.5787 | 0
[8 empty] | 365.6771 | 0
[8 empty] | 365.7755 | 0
[8 empty] | 365.8738 | 0
[8 empty] | 365.9722 | 0
[8 empty] | 366.0706 | 0
[8 empty] | 366.169 | 0
[8 empty] | 366.2674 | 0
[8 empty] | 366.3657 | 0
[8 empty] | 366.4641 | 0
[8 empty] | 366.5625 | 0
[8 empty] | 366.6609 | 0
[8 empty] | 366.7593 | 0
[8 empty] | 366.8576 | 0
[8 empty] | 366.956 | 0
[8 empty] | 367.0544 | 0
[8 empty] | 367.1528 | 38.2517
[8 empty] | 367.2512 | 64.4997
[8 empty] | 367.3495 | 60.9786
[8 empty] | 367.4479 | 60.7785
[8 empty] | 367.5463 | 75.8632
[8 empty] | 367.6447 | 80.3445
[8 empty] | 367.7431 | 80.6246
[8 empty] | 367.8414 | 80.4245
[8 empty] | 367.9398 | 80.2245
[8 empty] | 368.0382 | 80.1845
[8 empty] | 368.1366 | 80.1445
[8 empty] | 368.235 | 79.9444
[8 empty] | 368.3333 | 79.9444
[8 empty] | 368.4317 | 80.3845
[8 empty] | 368.5301 | 79.9844
[8 empty] | 368.6285 | 79.6243
[8 empty] | 368.7269 | 79.6643
[8 empty] | 368.8252 | 80.0644
[8 empty] | 368.9236 | 80.5846
[8 empty] | 369.022 | 80.1845
[8 empty] | 369.1204 | 80.1845
[8 empty] | 369.2188 | 83.2254
[8 empty] | 369.3171 | 86.1063
[8 empty] | 369.4155 | 22.7269
[8 empty] | 369.5139 | 22.7269
[8 empty] | 369.6123 | 22.7269
[8 empty] | 369.7106 | 0
[8 empty] | 369.809 | 0
[8 empty] | 369.9074 | 0
[8 empty] | 370.0058 | 0
[8 empty] | 370.1042 | 0
[8 empty] | 370.2025 | 0
[8 empty] | 370.3009 | 0
[8 empty] | 370.3993 | 0
[8 empty] | 370.4977 | 0
[8 empty] | 370.5961 | 0
[8 empty] | 370.6944 | 0
[8 empty] | 370.7928 | 0
[8 empty] | 370.8912 | 0
[8 empty] | 370.9896 | 0
[8 empty] | 371.088 | 0
[8 empty] | 371.1863 | 0
[8 empty] | 371.2847 | 0
[8 empty] | 371.3831 | 0
[8 empty] | 371.4815 | 0
[8 empty] | 371.5799 | 0
[8 empty] | 371.6782 | 0
[8 empty] | 371.7766 | 0
[8 empty] | 371.875 | 0
[8 empty] | 371.9734 | 0
[8 empty] | 372.0718 | 0
[8 empty] | 372.1701 | 0
[8 empty] | 372.2685 | 0
[8 empty] | 372.3669 | 0
[8 empty] | 372.4653 | 0
[8 empty] | 372.5637 | 0
[8 empty] | 372.662 | 0
[8 empty] | 372.7604 | 0
[8 empty] | 372.8588 | 0
[8 empty] | 372.9572 | 0
[8 empty] | 373.0556 | 0
[8 empty] | 373.1539 | 0
[8 empty] | 373.2523 | 0
[8 empty] | 373.3507 | 0
[8 empty] | 373.4491 | 0
[8 empty] | 373.5475 | 0
[8 empty] | 373.6458 | 0
[8 empty] | 373.7442 | 0
[8 empty] | 373.8426 | 0
[8 empty] | 373.941 | 0
[8 empty] | 374.0394 | 0
[8 empty] | 374.1377 | 0
[8 empty] | 374.2361 | 0
[8 empty] | 374.3345 | 0
[8 empty] | 374.4329 | 0
[8 empty] | 374.5312 | 0
[8 empty] | 374.6296 | 0
[8 empty] | 374.728 | 0
[8 empty] | 374.8264 | 0
[8 empty] | 374.9248 | 0
[8 empty] | 375.0231 | 0
[8 empty] | 375.1215 | 0
[8 empty] | 375.2199 | 0
[8 empty] | 375.3183 | 0
[8 empty] | 375.4167 | 0
[8 empty] | 375.515 | 0
[8 empty] | 375.6134 | 0
[8 empty] | 375.7118 | 0
[8 empty] | 375.8102 | 0
[8 empty] | 375.9086 | 0
[8 empty] | 376.0069 | 0
[8 empty] | 376.1053 | 0
[8 empty] | 376.2037 | 0
[8 empty] | 376.3021 | 0
[8 empty] | 376.4005 | 0
[8 empty] | 376.4988 | 0
[8 empty] | 376.5972 | 0
[8 empty] | 376.6956 | 0
[8 empty] | 376.794 | 0
[8 empty] | 376.8924 | 0
[8 empty] | 376.9907 | 0
[8 empty] | 377.0891 | 0
[8 empty] | 377.1875 | 0
[8 empty] | 377.2859 | 0
[8 empty] | 377.3843 | 0
[8 empty] | 377.4826 | 0
[8 empty] | 377.581 | 0
[8 empty] | 377.6794 | 0
[8 empty] | 377.7778 | 0
[8 empty] | 377.8762 | 0
[8 empty] | 377.9745 | 0
[8 empty] | 378.0729 | 0
[8 empty] | 378.1713 | 0
[8 empty] | 378.2697 | 0
[8 empty] | 378.3681 | 0
[8 empty] | 378.4664 | 0
[8 empty] | 378.5648 | 0
[8 empty] | 378.6632 | 0
[8 empty] | 378.7616 | 0
[8 empty] | 378.86 | 0
[8 empty] | 378.9583 | 0
[8 empty] | 379.0567 | 0
[8 empty] | 379.1551 | 0
[8 empty] | 379.2535 | 0
[8 empty] | 379.3519 | 0
[8 empty] | 379.4502 | 0
[8 empty] | 379.5486 | 0
[8 empty] | 379.647 | 0
[8 empty] | 379.7454 | 0
[8 empty] | 379.8438 | 0
[8 empty] | 379.9421 | 0
[8 empty] | 380.0405 | 0
[8 empty] | 380.1389 | 0
[8 empty] | 380.2373 | 0
[8 empty] | 380.3356 | 0
[8 empty] | 380.434 | 0
[8 empty] | 380.5324 | 0
[8 empty] | 380.6308 | 0
[8 empty] | 380.7292 | 0
[8 empty] | 380.8275 | 0
[8 empty] | 380.9259 | 0
[8 empty] | 381.0243 | 0
[8 empty] | 381.1227 | 0
[8 empty] | 381.2211 | 0
[8 empty] | 381.3194 | 0
[8 empty] | 381.4178 | 0
[8 empty] | 381.5162 | 0
[8 empty] | 381.6146 | 0
[8 empty] | 381.713 | 0
[8 empty] | 381.8113 | 0
[8 empty] | 381.9097 | 0
[8 empty] | 382.0081 | 0
[8 empty] | 382.1065 | 0
[8 empty] | 382.2049 | 0
[8 empty] | 382.3032 | 0
[8 empty] | 382.4016 | 0
[8 empty] | 382.5 | 0
[8 empty] | 382.5984 | 0
[8 empty] | 382.6968 | 0
[8 empty] | 382.7951 | 0
[8 empty] | 382.8935 | 0
[8 empty] | 382.9919 | 0
[8 empty] | 383.0903 | 0
[8 empty] | 383.1887 | 0
[8 empty] | 383.287 | 0
[8 empty] | 383.3854 | 0
[8 empty] | 383.4838 | 0
[8 empty] | 383.5822 | 0
[8 empty] | 383.6806 | 0
[8 empty] | 383.7789 | 0
[8 empty] | 383.8773 | 0
[8 empty] | 383.9757 | 0
[8 empty] | 384.0741 | 0
[8 empty] | 384.1725 | 0
[8 empty] | 384.2708 | 0
[8 empty] | 384.3692 | 0
[8 empty] | 384.4676 | 0
[8 empty] | 384.566 | 0
[8 empty] | 384.6644 | 0
[8 empty] | 384.7627 | 0
[8 empty] | 384.8611 | 0
[8 empty] | 384.9595 | 0
[8 empty] | 385.0579 | 0
[8 empty] | 385.1562 | 0
[8 empty] | 385.2546 | 0
[8 empty] | 385.353 | 0
[8 empty] | 385.4514 | 0
[8 empty] | 385.5498 | 0
[8 empty] | 385.6481 | 0
[8 empty] | 385.7465 | 0
[8 empty] | 385.8449 | 0
[8 empty] | 385.9433 | 0
[8 empty] | 386.0417 | 0
[8 empty] | 386.14 | 0
[8 empty] | 386.2384 | 0
[8 empty] | 386.3368 | 0
[8 empty] | 386.4352 | 0
[8 empty] | 386.5336 | 0
[8 empty] | 386.6319 | 0
[8 empty] | 386.7303 | 0
[8 empty] | 386.8287 | 0
[8 empty] | 386.9271 | 0
[8 empty] | 387.0255 | 0
[8 empty] | 387.1238 | 0
[8 empty] | 387.2222 | 0
[8 empty] | 387.3206 | 0
[8 empty] | 387.419 | 0
[8 empty] | 387.5174 | 0
[8 empty] | 387.6157 | 0
[8 empty] | 387.7141 | 0
[8 empty] | 387.8125 | 0
[8 empty] | 387.9109 | 0
[8 empty] | 388.0093 | 0
[8 empty] | 388.1076 | 0
[8 empty] | 388.206 | 0
[8 empty] | 388.3044 | 0
[8 empty] | 388.4028 | 0
[8 empty] | 388.5012 | 0
[8 empty] | 388.5995 | 0
[8 empty] | 388.6979 | 0
[8 empty] | 388.7963 | 0
[8 empty] | 388.8947 | 0
[8 empty] | 388.9931 | 0
[8 empty] | 389.0914 | 0
[8 empty] | 389.1898 | 0
[8 empty] | 389.2882 | 0
[8 empty] | 389.3866 | 0
[8 empty] | 389.485 | 0
[8 empty] | 389.5833 | 0
[8 empty] | 389.6817 | 0
[8 empty] | 389.7801 | 0
[8 empty] | 389.8785 | 0
[8 empty] | 389.9769 | 0
[8 empty] | 390.0752 | 0
[8 empty] | 390.1736 | 0
[8 empty] | 390.272 | 0
[8 empty] | 390.3704 | 0
[8 empty] | 390.4688 | 0
[8 empty] | 390.5671 | 0
[8 empty] | 390.6655 | 0
[8 empty] | 390.7639 | 0
[8 empty] | 390.8623 | 0
[8 empty] | 390.9606 | 0
[8 empty] | 391.059 | 0
[8 empty] | 391.1574 | 0
[8 empty] | 391.2558 | 0
[8 empty] | 391.3542 | 0
[8 empty] | 391.4525 | 0
[8 empty] | 391.5509 | 0
[8 empty] | 391.6493 | 0
[8 empty] | 391.7477 | 0
[8 empty] | 391.8461 | 0
[8 empty] | 391.9444 | 0
[8 empty] | 392.0428 | 0
[8 empty] | 392.1412 | 0
[8 empty] | 392.2396 | 0
[8 empty] | 392.338 | 0
[8 empty] | 392.4363 | 0
[8 empty] | 392.5347 | 0
[8 empty] | 392.6331 | 0
[8 empty] | 392.7315 | 0
[8 empty] | 392.8299 | 0
[8 empty] | 392.9282 | 0
[8 empty] | 393.0266 | 0
[8 empty] | 393.125 | 0
[8 empty] | 393.2234 | 0
[8 empty] | 393.3218 | 0
[8 empty] | 393.4201 | 0
[8 empty] | 393.5185 | 0
[8 empty] | 393.6169 | 0
[8 empty] | 393.7153 | 0
[8 empty] | 393.8137 | 0
[8 empty] | 393.912 | 0
[8 empty] | 394.0104 | 0
[8 empty] | 394.1088 | 0
[8 empty] | 394.2072 | 0
[8 empty] | 394.3056 | 0
[8 empty] | 394.4039 | 0
[8 empty] | 394.5023 | 0
[8 empty] | 394.6007 | 0
[8 empty] | 394.6991 | 0
[8 empty] | 394.7975 | 0
[8 empty] | 394.8958 | 0
[8 empty] | 394.9942 | 0
[8 empty] | 395.0926 | 0
[8 empty] | 395.191 | 0
[8 empty] | 395.2894 | 0
[8 empty] | 395.3877 | 0
[8 empty] | 395.4861 | 0
[8 empty] | 395.5845 | 0
[8 empty] | 395.6829 | 0
[8 empty] | 395.7812 | 0
[8 empty] | 395.8796 | 0
[8 empty] | 395.978 | 0
[8 empty] | 396.0764 | 0
[8 empty] | 396.1748 | 0
[8 empty] | 396.2731 | 0
[8 empty] | 396.3715 | 0
[8 empty] | 396.4699 | 0
[8 empty] | 396.5683 | 0
[8 empty] | 396.6667 | 0
[8 empty] | 396.765 | 0
[8 empty] | 396.8634 | 0
[8 empty] | 396.9618 | 0
[8 empty] | 397.0602 | 0
[8 empty] | 397.1586 | 0
[8 empty] | 397.2569 | 0
[8 empty] | 397.3553 | 0
[8 empty] | 397.4537 | 0
[8 empty] | 397.5521 | 0
[8 empty] | 397.6505 | 0
[8 empty] | 397.7488 | 0
[8 empty] | 397.8472 | 0
[8 empty] | 397.9456 | 0
[8 empty] | 398.044 | 0
[8 empty] | 398.1424 | 0
[8 empty] | 398.2407 | 0
[8 empty] | 398.3391 | 0
[8 empty] | 398.4375 | 0
[8 empty] | 398.5359 | 0
[8 empty] | 398.6343 | 0
[8 empty] | 398.7326 | 0
[8 empty] | 398.831 | 0
[8 empty] | 398.9294 | 0
[8 empty] | 399.0278 | 0
[8 empty] | 399.1262 | 0
[8 empty] | 399.2245 | 0
[8 empty] | 399.3229 | 0
[8 empty] | 399.4213 | 0
[8 empty] | 399.5197 | 0
[8 empty] | 399.6181 | 0
[8 empty] | 399.7164 | 0
[8 empty] | 399.8148 | 0
[8 empty] | 399.9132 | 0
[8 empty] | 400.0116 | 0
[8 empty] | 400.11 | 0
[8 empty] | 400.2083 | 0
[8 empty] | 400.3067 | 0
[8 empty] | 400.4051 | 0
[8 empty] | 400.5035 | 0
[8 empty] | 400.6019 | 34.4505
[8 empty] | 400.7002 | 63.2593
[8 empty] | 400.7986 | 60.3784
[8 empty] | 400.897 | 60.2184
[8 empty] | 400.9954 | 76.5834
[8 empty] | 401.0938 | 80.5846
[8 empty] | 401.1921 | 80.9447
[8 empty] | 401.2905 | 80.3845
[8 empty] | 401.3889 | 80.2245
[8 empty] | 401.4873 | 79.9044
[8 empty] | 401.5856 | 79.7843
[8 empty] | 401.684 | 79.5043
[8 empty] | 401.7824 | 79.5843
[8 empty] | 401.8808 | 79.7443
[8 empty] | 401.9792 | 80.4245
[8 empty] | 402.0775 | 80.3445
[8 empty] | 402.1759 | 80.2645
[8 empty] | 402.2743 | 80.1445
[8 empty] | 402.3727 | 79.7843
[8 empty] | 402.4711 | 79.9844
[8 empty] | 402.5694 | 80.4245
[8 empty] | 402.6678 | 82.9853
[8 empty] | 402.7662 | 84.1057
[8 empty] | 402.8646 | 38.6518
[8 empty] | 402.963 | 38.6518
[8 empty] | 403.0613 | 38.6518
[8 empty] | 403.1597 | 0
[8 empty] | 403.2581 | 0
[8 empty] | 403.3565 | 0
[8 empty] | 403.4549 | 0
[8 empty] | 403.5532 | 0
[8 empty] | 403.6516 | 0
[8 empty] | 403.75 | 0
[8 empty] | 403.8484 | 0
[8 empty] | 403.9468 | 0
[8 empty] | 404.0451 | 0
[8 empty] | 404.1435 | 0
[8 empty] | 404.2419 | 0
[8 empty] | 404.3403 | 0
[8 empty] | 404.4387 | 0
[8 empty] | 404.537 | 0
[8 empty] | 404.6354 | 0
[8 empty] | 404.7338 | 0
[8 empty] | 404.8322 | 0
[8 empty] | 404.9306 | 0
[8 empty] | 405.0289 | 0
[8 empty] | 405.1273 | 0
[8 empty] | 405.2257 | 0
[8 empty] | 405.3241 | 0
[8 empty] | 405.4225 | 0
[8 empty] | 405.5208 | 0
[8 empty] | 405.6192 | 0
[8 empty] | 405.7176 | 0
[8 empty] | 405.816 | 0
[8 empty] | 405.9144 | 0
[8 empty] | 406.0127 | 0
[8 empty] | 406.1111 | 0
[8 empty] | 406.2095 | 0
[8 empty] | 406.3079 | 0
[8 empty] | 406.4062 | 0
[8 empty] | 406.5046 | 0
[8 empty] | 406.603 | 0
[8 empty] | 406.7014 | 0
[8 empty] | 406.7998 | 0
[8 empty] | 406.8981 | 0
[8 empty] | 406.9965 | 0
[8 empty] | 407.0949 | 0
[8 empty] | 407.1933 | 0
[8 empty] | 407.2917 | 0
[8 empty] | 407.39 | 0
[8 empty] | 407.4884 | 0
[8 empty] | 407.5868 | 0
[8 empty] | 407.6852 | 0
[8 empty] | 407.7836 | 0
[8 empty] | 407.8819 | 0
[8 empty] | 407.9803 | 0
[8 empty] | 408.0787 | 0
[8 empty] | 408.1771 | 0
[8 empty] | 408.2755 | 0
[8 empty] | 408.3738 | 0
[8 empty] | 408.4722 | 0
[8 empty] | 408.5706 | 0
[8 empty] | 408.669 | 0
[8 empty] | 408.7674 | 0
[8 empty] | 408.8657 | 0
[8 empty] | 408.9641 | 0
[8 empty] | 409.0625 | 0
[8 empty] | 409.1609 | 0
[8 empty] | 409.2593 | 0
[8 empty] | 409.3576 | 0
[8 empty] | 409.456 | 0
[8 empty] | 409.5544 | 0
[8 empty] | 409.6528 | 0
[8 empty] | 409.7512 | 0
[8 empty] | 409.8495 | 0
[8 empty] | 409.9479 | 0
[8 empty] | 410.0463 | 0
[8 empty] | 410.1447 | 0
[8 empty] | 410.2431 | 0
[8 empty] | 410.3414 | 0
[8 empty] | 410.4398 | 0
[8 empty] | 410.5382 | 0
[8 empty] | 410.6366 | 0
[8 empty] | 410.735 | 0
[8 empty] | 410.8333 | 0
[8 empty] | 410.9317 | 0
[8 empty] | 411.0301 | 0
[8 empty] | 411.1285 | 0
[8 empty] | 411.2269 | 0
[8 empty] | 411.3252 | 0
[8 empty] | 411.4236 | 0
[8 empty] | 411.522 | 0
[8 empty] | 411.6204 | 0
[8 empty] | 411.7188 | 0
[8 empty] | 411.8171 | 0
[8 empty] | 411.9155 | 0
[8 empty] | 412.0139 | 0
[8 empty] | 412.1123 | 0
[8 empty] | 412.2106 | 0
[8 empty] | 412.309 | 0
[8 empty] | 412.4074 | 0
[8 empty] | 412.5058 | 0
[8 empty] | 412.6042 | 0
[8 empty] | 412.7025 | 0
[8 empty] | 412.8009 | 0
[8 empty] | 412.8993 | 0
[8 empty] | 412.9977 | 0
[8 empty] | 413.0961 | 0
[8 empty] | 413.1944 | 0
[8 empty] | 413.2928 | 0
[8 empty] | 413.3912 | 0
[8 empty] | 413.4896 | 0
[8 empty] | 413.588 | 0
[8 empty] | 413.6863 | 0
[8 empty] | 413.7847 | 0
[8 empty] | 413.8831 | 0
[8 empty] | 413.9815 | 0
[8 empty] | 414.0799 | 0
[8 empty] | 414.1782 | 0
[8 empty] | 414.2766 | 0
[8 empty] | 414.375 | 0
[8 empty] | 414.4734 | 0
[8 empty] | 414.5718 | 0
[8 empty] | 414.6701 | 0
[8 empty] | 414.7685 | 0
[8 empty] | 414.8669 | 0
[8 empty] | 414.9653 | 0
[8 empty] | 415.0637 | 0
[8 empty] | 415.162 | 0
[8 empty] | 415.2604 | 0
[8 empty] | 415.3588 | 0
[8 empty] | 415.4572 | 0
[8 empty] | 415.5556 | 0
[8 empty] | 415.6539 | 0
[8 empty] | 415.7523 | 0
[8 empty] | 415.8507 | 0
[8 empty] | 415.9491 | 0
[8 empty] | 416.0475 | 0
[8 empty] | 416.1458 | 0
[8 empty] | 416.2442 | 0
[8 empty] | 416.3426 | 0
[8 empty] | 416.441 | 0
[8 empty] | 416.5394 | 0
[8 empty] | 416.6377 | 0
[8 empty] | 416.7361 | 0
[8 empty] | 416.8345 | 0
[8 empty] | 416.9329 | 0
[8 empty] | 417.0312 | 0
[8 empty] | 417.1296 | 0
[8 empty] | 417.228 | 0
[8 empty] | 417.3264 | 0
[8 empty] | 417.4248 | 0
[8 empty] | 417.5231 | 0
[8 empty] | 417.6215 | 0
[8 empty] | 417.7199 | 0
[8 empty] | 417.8183 | 0
[8 empty] | 417.9167 | 0
[8 empty] | 418.015 | 0
[8 empty] | 418.1134 | 0
[8 empty] | 418.2118 | 0
[8 empty] | 418.3102 | 0
[8 empty] | 418.4086 | 0
[8 empty] | 418.5069 | 0
[8 empty] | 418.6053 | 0
[8 empty] | 418.7037 | 0
[8 empty] | 418.8021 | 0
[8 empty] | 418.9005 | 0
[8 empty] | 418.9988 | 0
[8 empty] | 419.0972 | 0
[8 empty] | 419.1956 | 0
[8 empty] | 419.294 | 0
[8 empty] | 419.3924 | 0
[8 empty] | 419.4907 | 0
[8 empty] | 419.5891 | 0
[8 empty] | 419.6875 | 0
[8 empty] | 419.7859 | 0
[8 empty] | 419.8843 | 0
[8 empty] | 419.9826 | 0
[8 empty] | 420.081 | 0
[8 empty] | 420.1794 | 0
[8 empty] | 420.2778 | 0
[8 empty] | 420.3762 | 0
[8 empty] | 420.4745 | 0
[8 empty] | 420.5729 | 0
[8 empty] | 420.6713 | 0
[8 empty] | 420.7697 | 0
[8 empty] | 420.8681 | 0
[8 empty] | 420.9664 | 0
[8 empty] | 421.0648 | 0
[8 empty] | 421.1632 | 0
[8 empty] | 421.2616 | 0
[8 empty] | 421.36 | 0
[8 empty] | 421.4583 | 0
[8 empty] | 421.5567 | 0
[8 empty] | 421.6551 | 0
[8 empty] | 421.7535 | 0
[8 empty] | 421.8519 | 0
[8 empty] | 421.9502 | 0
[8 empty] | 422.0486 | 0
[8 empty] | 422.147 | 0
[8 empty] | 422.2454 | 0
[8 empty] | 422.3438 | 0
[8 empty] | 422.4421 | 0
[8 empty] | 422.5405 | 0
[8 empty] | 422.6389 | 0
[8 empty] | 422.7373 | 0
[8 empty] | 422.8356 | 0
[8 empty] | 422.934 | 0
[8 empty] | 423.0324 | 0
[8 empty] | 423.1308 | 0
[8 empty] | 423.2292 | 0
[8 empty] | 423.3275 | 0
[8 empty] | 423.4259 | 0
[8 empty] | 423.5243 | 0
[8 empty] | 423.6227 | 0
[8 empty] | 423.7211 | 0
[8 empty] | 423.8194 | 0
[8 empty] | 423.9178 | 0
[8 empty] | 424.0162 | 0
[8 empty] | 424.1146 | 0
[8 empty] | 424.213 | 0
[8 empty] | 424.3113 | 0
[8 empty] | 424.4097 | 0
[8 empty] | 424.5081 | 0
[8 empty] | 424.6065 | 0
[8 empty] | 424.7049 | 0
[8 empty] | 424.8032 | 0
[8 empty] | 424.9016 | 0
[8 empty] | 425 | 0
[8 empty] | 425.0984 | 0
[8 empty] | 425.1968 | 0
[8 empty] | 425.2951 | 0
[8 empty] | 425.3935 | 0
[8 empty] | 425.4919 | 0
[8 empty] | 425.5903 | 0
[8 empty] | 425.6887 | 0
[8 empty] | 425.787 | 0
[8 empty] | 425.8854 | 0
[8 empty] | 425.9838 | 0
[8 empty] | 426.0822 | 0
[8 empty] | 426.1806 | 0
[8 empty] | 426.2789 | 0
[8 empty] | 426.3773 | 0
[8 empty] | 426.4757 | 0
[8 empty] | 426.5741 | 0
[8 empty] | 426.6725 | 0
[8 empty] | 426.7708 | 0
[8 empty] | 426.8692 | 0
[8 empty] | 426.9676 | 0
[8 empty] | 427.066 | 0
[8 empty] | 427.1644 | 0
[8 empty] | 427.2627 | 0
[8 empty] | 427.3611 | 0
[8 empty] | 427.4595 | 0
[8 empty] | 427.5579 | 0
[8 empty] | 427.6562 | 0
[8 empty] | 427.7546 | 0
[8 empty] | 427.853 | 0
[8 empty] | 427.9514 | 0
[8 empty] | 428.0498 | 0
[8 empty] | 428.1481 | 0
[8 empty] | 428.2465 | 0
[8 empty] | 428.3449 | 0
[8 empty] | 428.4433 | 0
[8 empty] | 428.5417 | 0
[8 empty] | 428.64 | 0
[8 empty] | 428.7384 | 0
[8 empty] | 428.8368 | 0
[8 empty] | 428.9352 | 0
[8 empty] | 429.0336 | 0
[8 empty] | 429.1319 | 0
[8 empty] | 429.2303 | 0
[8 empty] | 429.3287 | 0
[8 empty] | 429.4271 | 0
[8 empty] | 429.5255 | 0
[8 empty] | 429.6238 | 0
[8 empty] | 429.7222 | 0
[8 empty] | 429.8206 | 0
[8 empty] | 429.919 | 0
[8 empty] | 430.0174 | 0
[8 empty] | 430.1157 | 0
[8 empty] | 430.2141 | 0
[8 empty] | 430.3125 | 0
[8 empty] | 430.4109 | 0
[8 empty] | 430.5093 | 0
[8 empty] | 430.6076 | 0
[8 empty] | 430.706 | 0
[8 empty] | 430.8044 | 0
[8 empty] | 430.9028 | 0
[8 empty] | 431.0012 | 0
[8 empty] | 431.0995 | 0
[8 empty] | 431.1979 | 0
[8 empty] | 431.2963 | 0
[8 empty] | 431.3947 | 0
[8 empty] | 431.4931 | 0
[8 empty] | 431.5914 | 0
[8 empty] | 431.6898 | 0
[8 empty] | 431.7882 | 0
[8 empty] | 431.8866 | 0
[8 empty] | 431.985 | 0
[8 empty] | 432.0833 | 0
[8 empty] | 432.1817 | 0
[8 empty] | 432.2801 | 0
[8 empty] | 432.3785 | 0
[8 empty] | 432.4769 | 0
[8 empty] | 432.5752 | 0
[8 empty] | 432.6736 | 0
[8 empty] | 432.772 | 0
[8 empty] | 432.8704 | 0
[8 empty] | 432.9688 | 0
[8 empty] | 433.0671 | 0
[8 empty] | 433.1655 | 0
[8 empty] | 433.2639 | 0
[8 empty] | 433.3623 | 0
[8 empty] | 433.4606 | 0
[8 empty] | 433.559 | 0
[8 empty] | 433.6574 | 0
[8 empty] | 433.7558 | 0
[8 empty] | 433.8542 | 0
[8 empty] | 433.9525 | 0
[8 empty] | 434.0509 | 0
[8 empty] | 434.1493 | 65.1799
[8 empty] | 434.2477 | 76.7034
[8 empty] | 434.3461 | 81.0247
[8 empty] | 434.4444 | 81.1448
[8 empty] | 434.5428 | 80.7446
[8 empty] | 434.6412 | 80.3845
[8 empty] | 434.7396 | 80.3845
[8 empty] | 434.838 | 80.2245
[8 empty] | 434.9363 | 80.0644
[8 empty] | 435.0347 | 80.1445
[8 empty] | 435.1331 | 80.1845
[8 empty] | 435.2315 | 80.1044
[8 empty] | 435.3299 | 79.8644
[8 empty] | 435.4282 | 79.8644
[8 empty] | 435.5266 | 80.1044
[8 empty] | 435.625 | 80.1044
[8 empty] | 435.7234 | 80.0244
[8 empty] | 435.8218 | 79.7843
[8 empty] | 435.9201 | 79.9044
[8 empty] | 436.0185 | 80.3045
[8 empty] | 436.1169 | 80.3845
[8 empty] | 436.2153 | 80.5046
[8 empty] | 436.3137 | 80.7847
[8 empty] | 436.412 | 80.8247
[8 empty] | 436.5104 | 81.0247
[8 empty] | 436.6088 | 80.8247
[8 empty] | 436.7072 | 80.7446
[8 empty] | 436.8056 | 80.7446
[8 empty] | 436.9039 | 80.7046
[8 empty] | 437.0023 | 80.6246
[8 empty] | 437.1007 | 81.2648
[8 empty] | 437.1991 | 85.266
[8 empty] | 437.2975 | 87.0266
[8 empty] | 437.3958 | 0
[8 empty] | 437.4942 | 0
[8 empty] | 437.5926 | 0
[8 empty] | 437.691 | 0
[8 empty] | 437.7894 | 0

## Test
XLPRO ™ Version 4.06
Operation Settings Area |  | 0 | R/min | 0 | R/min | 0 | R | 4.0244e-05 | R | 0.53052 | R | 11693.7794 | uR/pls
|  | 0 | R/min | 0 | R/min | 0 | R | 2.0404 | R | 1.0614 | R | 11703.8777 | uR/pls
Units |  | 63.6559 | R/min | 32.9397 | R/min | 0.2799 | R | 4.4641 | R |  |  | 11703.8777 | uR/pls
Radiation: | R | 63.6659 | R/min | 32.9397 | R/min | 1.4025 | R | 6.6091 | R |  |  | 11703.8777 | uR/pls
/Time: | /min | 63.6659 | R/min | 32.9397 | R/min | 2.5253 | R | 8.6555 | R |  |  | 11703.8777 | uR/pls
Scale: |  | 63.676 | R/min | 63.6458 | R/min | 3.6542 | R | 9.5932 | R |  |  | 2108.5157 | uR/pls
Pulse Rate: | 15 | 14.4174 | R/min | 63.6458 | R/min | 4.4872 | R | 9.5932 | R
Mode |  |  |  | 63.6458 | R/min | 4.4872 | R | 9.5932 | R
Units Display: | Rescale |  |  | 63.6458 | R/min | 4.4872 | R | [6 empty] | 2.816 | R | 220 | ms | 60.5 | kVp
Cell Update: | Incremental |  |  | 63.6458 | R/min | 4.9752 | R | [6 empty] | 63.03 | R/min | 88.5 | kVp
Auto Zero: | Disabled |  |  | 63.6458 | R/min | 6.098 | R | [6 empty] | 63.04 | R/min | 88.4 | kVp
Capture Mode: | Auto |  |  | 63.6559 | R/min | 7.2218 | R
Screen Tracking: | Yes |  |  | 63.6559 | R/min | 8.3446 | R
kVp |  |  |  | 63.6559 | R/min | 9.0197 | R
Config: | Dose+Time+kVp |  |  | 63.6559 | R/min | 9.0197 | R
Scale: | DAccu-kV |  |  | 63.6559 | R/min | 9.0197 | R
Gain: | 32.77
Offset: | 0.00
Comm |  | 0.187 | R | 195 | ms-9000 | 111.3 | kVp-9000
Index: | 0 | Peak= | 112.7 | kVp-FT1 | PPV= | 110.6 | kV-Prac.
Baud: | 9600 | Avg= | 110.921 | kV | Ripple= | 3.6329 | kV p-p.
Monitor Configuration |  | Pk Freq= | 1498.8914 | Hz | Rate= | 1.5 | Hz
Serial No: | A 1 | Width= | 194.7924 | ms-Mem. | Pulses= | 1
ROM Ver: | 3.25 | ms | kV |  | Hz | kV | LINES
Detector: | 90x6 | 0 | 0 |  | 0 | 1.9582e-15 | 0
Chamb Kq: | 1.00 | 0.09838 | 0 |  | 2.4816 | 0.015498 | 198.3341
Atm Corr: | Automatic | 0.19676 | 0 |  | 4.9632 | 0.022788 | 112.7374
Temp (°C): | 2 | 0.29514 | 0 |  | 7.4448 | 0.029323 | 112.7374
Press (kPa): | 100.5 | 0.39352 | 0 |  | 9.9264 | 0.051629 | 110.921
|  | 0.4919 | 40.8525 |  | 12.408 | 0.078418 | 110.921
|  | 0.59028 | 40.8925 |  | 14.8896 | 0.091295
|  | 0.68866 | 51.1356 |  | 17.3713 | 0.080697
|  | 0.78704 | 51.0556 |  | 19.8529 | 0.055884 | 2042
|  | 0.88542 | 51.2156 |  | 22.3345 | 0.062529
|  | 0.9838 | 64.0195 |  | 24.8161 | 0.10175
|  | 1.0822 | 95.3891 |  | 27.2977 | 0.12928
|  | 1.1806 | 103.0715 |  | 29.7793 | 0.12822
|  | 1.2789 | 107.0727 |  | 32.2609 | 0.098886
|  | 1.3773 | 106.7526 |  | 34.7425 | 0.053949
|  | 1.4757 | 105.7123 |  | 37.2241 | 0.024019
|  | 1.5741 | 105.6322 |  | 39.7057 | 0.043963
|  | 1.6725 | 106.9926 |  | 42.1873 | 0.059408
|  | 1.7708 | 109.3934 |  | 44.6689 | 0.062117
|  | 1.8692 | 110.7538 |  | 47.1506 | 0.060991
|  | 1.9676 | 109.6335 |  | 49.6322 | 0.061018
|  | 2.066 | 108.7532 |  | 52.1138 | 0.058906
|  | 2.1644 | 109.2333 |  | 54.5954 | 0.055763
|  | 2.2627 | 111.1539 |  | 57.077 | 0.063605
|  | 2.3611 | 112.4343 |  | 59.5586 | 0.084505
|  | 2.4595 | 111.7941 |  | 62.0402 | 0.10412
|  | 2.5579 | 110.6738 |  | 64.5218 | 0.11025
|  | 2.6563 | 110.4737 |  | 67.0034 | 0.098822
|  | 2.7546 | 111.9942 |  | 69.485 | 0.074651
|  | 2.853 | 113.9148 |  | 71.9666 | 0.052004
|  | 2.9514 | 114.4349 |  | 74.4482 | 0.04757
|  | 3.0498 | 113.2346 |  | 76.9299 | 0.051186
|  | 3.1482 | 112.1142 |  | 79.4115 | 0.043725
|  | 3.2465 | 111.0739 |  | 81.8931 | 0.021624
|  | 3.3449 | 110.2737 |  | 84.3747 | 0.012326
|  | 3.4433 | 110.3137 |  | 86.8563 | 0.040891
|  | 3.5417 | 112.3943 |  | 89.3379 | 0.059135
|  | 3.6401 | 113.8347 |  | 91.8195 | 0.061363
|  | 3.7384 | 112.9545 |  | 94.3011 | 0.049956
|  | 3.8368 | 112.0742 |  | 96.7827 | 0.034044
|  | 3.9352 | 111.0339 |  | 99.2643 | 0.024773
|  | 4.0336 | 109.9135 |  | 101.7459 | 0.025943
|  | 4.132 | 109.7535 |  | 104.2275 | 0.040389
|  | 4.2303 | 111.514 |  | 106.7092 | 0.066779
|  | 4.3287 | 113.1945 |  | 109.1908 | 0.092057
|  | 4.4271 | 113.0345 |  | 111.6724 | 0.10193
|  | 4.5255 | 111.6741 |  | 114.154 | 0.088508
|  | 4.6239 | 110.7938 |  | 116.6356 | 0.056591
|  | 4.7222 | 109.7535 |  | 119.1172 | 0.036969
|  | 4.8206 | 109.6735 |  | 121.5988 | 0.056178
|  | 4.919 | 111.5941 |  | 124.0804 | 0.064209
|  | 5.0174 | 113.1145 |  | 126.562 | 0.040509
|  | 5.1158 | 112.7544 |  | 129.0436 | 0.022876
|  | 5.2141 | 111.7141 |  | 131.5252 | 0.088414
|  | 5.3125 | 110.8338 |  | 134.0068 | 0.15055
|  | 5.4109 | 109.9936 |  | 136.4885 | 0.18778
|  | 5.5093 | 109.6335 |  | 138.9701 | 0.19246
|  | 5.6077 | 111.394 |  | 141.4517 | 0.17077
|  | 5.706 | 112.8344 |  | 143.9333 | 0.13929
|  | 5.8044 | 112.5543 |  | 146.4149 | 0.11617
|  | 5.9028 | 111.5941 |  | 148.8965 | 0.10995
|  | 6.0012 | 110.3137 |  | 151.3781 | 0.11777
|  | 6.0996 | 109.3534 |  | 153.8597 | 0.13213
|  | 6.1979 | 109.7935 |  | 156.3413 | 0.14174
|  | 6.2963 | 111.8741 |  | 158.8229 | 0.13532
|  | 6.3947 | 112.9145 |  | 161.3045 | 0.10884
|  | 6.4931 | 112.6344 |  | 163.7861 | 0.069578
|  | 6.5915 | 111.1939 |  | 166.2678 | 0.036455
|  | 6.6898 | 110.1536 |  | 168.7494 | 0.030248
|  | 6.7882 | 109.1933 |  | 171.231 | 0.027104
|  | 6.8866 | 109.7535 |  | 173.7126 | 0.021196
|  | 6.985 | 111.6741 |  | 176.1942 | 0.050937
|  | 7.0834 | 112.9545 |  | 178.6758 | 0.088736
|  | 7.1817 | 112.2342 |  | 181.1574 | 0.11258
|  | 7.2801 | 111.0339 |  | 183.639 | 0.11299
|  | 7.3785 | 110.0736 |  | 186.1206 | 0.090664
|  | 7.4769 | 109.1933 |  | 188.6022 | 0.054947
|  | 7.5753 | 109.7535 |  | 191.0838 | 0.019504
|  | 7.6736 | 111.8341 |  | 193.5654 | 0.0037111
|  | 7.772 | 112.7944 |  | 196.0471 | 0.0080713
|  | 7.8704 | 111.6741 |  | 198.5287 | 0.0051008
|  | 7.9688 | 110.8738 |  | 201.0103 | 0.024926
|  | 8.0672 | 109.8335 |  | 203.4919 | 0.042538
|  | 8.1655 | 108.7132 |  | 205.9735 | 0.050395
|  | 8.2639 | 109.8335 |  | 208.4551 | 0.050815
|  | 8.3623 | 111.9142 |  | 210.9367 | 0.056441
|  | 8.4607 | 112.5543 |  | 213.4183 | 0.071326
|  | 8.5591 | 112.1142 |  | 215.8999 | 0.082248
|  | 8.6574 | 110.5137 |  | 218.3815 | 0.077333
|  | 8.7558 | 109.3934 |  | 220.8631 | 0.054894
|  | 8.8542 | 108.8732 |  | 223.3447 | 0.033424
|  | 8.9526 | 110.3137 |  | 225.8264 | 0.055484
|  | 9.051 | 112.2342 |  | 228.308 | 0.088897
|  | 9.1493 | 112.5944 |  | 230.7896 | 0.10806
|  | 9.2477 | 111.474 |  | 233.2712 | 0.10749
|  | 9.3461 | 110.3537 |  | 235.7528 | 0.090942
|  | 9.4445 | 109.7135 |  | 238.2344 | 0.06882
|  | 9.5429 | 109.3534 |  | 240.716 | 0.05422
|  | 9.6412 | 110.9539 |  | 243.1976 | 0.051597
|  | 9.7396 | 112.4743 |  | 245.6792 | 0.052169
|  | 9.838 | 112.3943 |  | 248.1608 | 0.050899
|  | 9.9364 | 111.1139 |  | 250.6424 | 0.050664
|  | 10.0348 | 110.1536 |  | 253.124 | 0.0544
|  | 10.1331 | 109.2733 |  | 255.6057 | 0.060995
|  | 10.2315 | 109.2333 |  | 258.0873 | 0.068826
|  | 10.3299 | 111.6741 |  | 260.5689 | 0.076681
|  | 10.4283 | 112.8744 |  | 263.0505 | 0.081554
|  | 10.5267 | 112.0742 |  | 265.5321 | 0.080238
|  | 10.625 | 111.434 |  | 268.0137 | 0.074889
|  | 10.7234 | 110.0736 |  | 270.4953 | 0.075977
|  | 10.8218 | 109.2733 |  | 272.9769 | 0.089221
|  | 10.9202 | 109.9135 |  | 275.4585 | 0.10301
|  | 11.0186 | 112.1542 |  | 277.9401 | 0.1034
|  | 11.1169 | 112.7144 |  | 280.4217 | 0.087232
|  | 11.2153 | 112.0742 |  | 282.9033 | 0.071055
|  | 11.3137 | 110.9939 |  | 285.385 | 0.085914
|  | 11.4121 | 109.6335 |  | 287.8666 | 0.11946
|  | 11.5105 | 109.1533 |  | 290.3482 | 0.14234
|  | 11.6088 | 110.4337 |  | 292.8298 | 0.14273
|  | 11.7072 | 111.9542 |  | 295.3114 | 0.12267
|  | 11.8056 | 112.4743 |  | 297.793 | 0.09357
|  | 11.904 | 111.9542 |  | 300.2746 | 0.068705
|  | 12.0024 | 110.5937 |  | 302.7562 | 0.051433
|  | 12.1007 | 109.7135 |  | 305.2378 | 0.033901
|  | 12.1991 | 109.1933 |  | 307.7194 | 0.015313
|  | 12.2975 | 110.3537 |  | 310.201 | 0.020144
|  | 12.3959 | 112.2743 |  | 312.6826 | 0.031424
|  | 12.4943 | 112.5143 |  | 315.1643 | 0.032903
|  | 12.5926 | 111.394 |  | 317.6459 | 0.039653
|  | 12.691 | 110.5137 |  | 320.1275 | 0.065357
|  | 12.7894 | 109.1533 |  | 322.6091 | 0.094028
|  | 12.8878 | 109.3134 |  | 325.0907 | 0.11014
|  | 12.9862 | 111.394 |  | 327.5723 | 0.10719
|  | 13.0845 | 112.9145 |  | 330.0539 | 0.0881
|  | 13.1829 | 112.3943 |  | 332.5355 | 0.063682
|  | 13.2813 | 111.274 |  | 335.0171 | 0.047489
|  | 13.3797 | 110.0736 |  | 337.4987 | 0.043809
|  | 13.4781 | 109.0333 |  | 339.9803 | 0.046872
|  | 13.5764 | 109.5934 |  | 342.4619 | 0.057078
|  | 13.6748 | 111.514 |  | 344.9436 | 0.073104
|  | 13.7732 | 112.6344 |  | 347.4252 | 0.087091
|  | 13.8716 | 112.0742 |  | 349.9068 | 0.093582
|  | 13.97 | 111.1139 |  | 352.3884 | 0.093866
|  | 14.0683 | 109.8335 |  | 354.87 | 0.094009
|  | 14.1667 | 109.0333 |  | 357.3516 | 0.098971
|  | 14.2651 | 109.4734 |  | 359.8332 | 0.10968
|  | 14.3635 | 111.554 |  | 362.3148 | 0.12591
|  | 14.4619 | 112.9145 |  | 364.7964 | 0.14603
|  | 14.5602 | 111.9542 |  | 367.278 | 0.1636
|  | 14.6586 | 110.7538 |  | 369.7596 | 0.16869
|  | 14.757 | 109.9536 |  | 372.2412 | 0.15308
|  | 14.8554 | 108.8332 |  | 374.7229 | 0.11463
|  | 14.9538 | 109.6335 |  | 377.2045 | 0.058951
|  | 15.0521 | 111.554 |  | 379.6861 | 0.0067098
|  | 15.1505 | 112.5944 |  | 382.1677 | 0.055257
|  | 15.2489 | 111.7941 |  | 384.6493 | 0.08585
|  | 15.3473 | 110.5937 |  | 387.1309 | 0.088785
|  | 15.4457 | 109.6335 |  | 389.6125 | 0.06622
|  | 15.544 | 108.9132 |  | 392.0941 | 0.028874
|  | 15.6424 | 110.4337 |  | 394.5757 | 0.018877
|  | 15.7408 | 112.5143 |  | 397.0573 | 0.045887
|  | 15.8392 | 112.3943 |  | 399.5389 | 0.056838
|  | 15.9376 | 111.434 |  | 402.0205 | 0.051633
|  | 16.0359 | 110.3937 |  | 404.5022 | 0.041547
|  | 16.1343 | 109.2733 |  | 406.9838 | 0.038942
|  | 16.2327 | 109.1933 |  | 409.4654 | 0.038706
|  | 16.3311 | 110.7138 |  | 411.947 | 0.032539
|  | 16.4295 | 112.3943 |  | 414.4286 | 0.0304
|  | 16.5278 | 112.5543 |  | 416.9102 | 0.042139
|  | 16.6262 | 111.434 |  | 419.3918 | 0.050368
|  | 16.7246 | 110.2336 |  | 421.8734 | 0.042082
|  | 16.823 | 109.3534 |  | 424.355 | 0.01832
|  | 16.9214 | 109.2733 |  | 426.8366 | 0.024858
|  | 17.0197 | 111.354 |  | 429.3182 | 0.053108
|  | 17.1181 | 112.5543 |  | 431.7998 | 0.064058
|  | 17.2165 | 112.2342 |  | 434.2815 | 0.050801
|  | 17.3149 | 111.1939 |  | 436.7631 | 0.020068
|  | 17.4133 | 110.2336 |  | 439.2447 | 0.030379
|  | 17.5116 | 109.2733 |  | 441.7263 | 0.062035
|  | 17.61 | 109.5534 |  | 444.2079 | 0.074667
|  | 17.7084 | 111.7941 |  | 446.6895 | 0.064637
|  | 17.8068 | 113.0745 |  | 449.1711 | 0.040951
|  | 17.9052 | 111.9542 |  | 451.6527 | 0.026479
|  | 18.0035 | 110.9939 |  | 454.1343 | 0.032493
|  | 18.1019 | 109.7935 |  | 456.6159 | 0.0317
|  | 18.2003 | 108.9933 |  | 459.0975 | 0.017891
|  | 18.2987 | 110.2737 |  | 461.5791 | 0.0094013
|  | 18.3971 | 112.1942 |  | 464.0608 | 0.019086
|  | 18.4954 | 112.5944 |  | 466.5424 | 0.017017
|  | 18.5938 | 111.8741 |  | 469.024 | 0.025727
|  | 18.6922 | 110.7538 |  | 471.5056 | 0.065046
|  | 18.7906 | 109.5534 |  | 473.9872 | 0.11016
|  | 18.889 | 109.1533 |  | 476.4688 | 0.145
|  | 18.9873 | 110.4337 |  | 478.9504 | 0.15879
|  | 19.0857 | 112.2743 |  | 481.432 | 0.14815
|  | 19.1841 | 112.3543 |  | 483.9136 | 0.11827
|  | 19.2825 | 111.394 |  | 486.3952 | 0.083212
|  | 19.3809 | 110.1936 |  | 488.8768 | 0.064606
|  | 19.4792 | 109.3134 |  | 491.3584 | 0.068525
|  | 19.5776 | 109.1933 |  | 493.8401 | 0.071868
|  | 19.676 | 111.1139 |  | 496.3217 | 0.061455
|  | 19.7744 | 112.7944 |  | 498.8033 | 0.039536
|  | 19.8728 | 112.4743 |  | 501.2849 | 0.019473
|  | 19.9711 | 111.1139 |  | 503.7665 | 0.023357
|  | 20.0695 | 110.1536 |  | 506.2481 | 0.030475
|  | 20.1679 | 109.2333 |  | 508.7297 | 0.026419
|  | 20.2663 | 109.3134 |  | 511.2113 | 0.013041
|  | 20.3647 | 111.7141 |  | 513.6929 | 0.0098982
|  | 20.463 | 112.7544 |  | 516.1745 | 0.029513
|  | 20.5614 | 112.0342 |  | 518.6561 | 0.048806
|  | 20.6598 | 110.9939 |  | 521.1377 | 0.06281
|  | 20.7582 | 109.9536 |  | 523.6194 | 0.065737
|  | 20.8566 | 108.9933 |  | 526.101 | 0.052825
|  | 20.9549 | 109.8735 |  | 528.5826 | 0.025375
|  | 21.0533 | 111.7941 |  | 531.0642 | 0.026417
|  | 21.1517 | 112.6344 |  | 533.5458 | 0.065487
|  | 21.2501 | 111.8341 |  | 536.0274 | 0.095056
|  | 21.3485 | 110.7938 |  | 538.509 | 0.10337
|  | 21.4468 | 109.5934 |  | 540.9906 | 0.086894
|  | 21.5452 | 108.8732 |  | 543.4722 | 0.049997
|  | 21.6436 | 110.0736 |  | 545.9538 | 0.0056079
|  | 21.742 | 111.9942 |  | 548.4354 | 0.041371
|  | 21.8404 | 112.7944 |  | 550.917 | 0.070953
|  | 21.9387 | 111.7541 |  | 553.3987 | 0.081032
|  | 22.0371 | 110.3937 |  | 555.8803 | 0.073203
|  | 22.1355 | 109.4734 |  | 558.3619 | 0.054558
|  | 22.2339 | 109.0333 |  | 560.8435 | 0.033527
|  | 22.3323 | 110.6338 |  | 563.3251 | 0.020185
|  | 22.4306 | 112.5543 |  | 565.8067 | 0.033544
|  | 22.529 | 112.5143 |  | 568.2883 | 0.060168
|  | 22.6274 | 111.1539 |  | 570.7699 | 0.083635
|  | 22.7258 | 110.3537 |  | 573.2515 | 0.093077
|  | 22.8242 | 109.3134 |  | 575.7331 | 0.082168
|  | 22.9225 | 109.1933 |  | 578.2147 | 0.052165
|  | 23.0209 | 111.434 |  | 580.6963 | 0.014015
|  | 23.1193 | 112.4743 |  | 583.178 | 0.030293
|  | 23.2177 | 112.2342 |  | 585.6596 | 0.054837
|  | 23.3161 | 111.1139 |  | 588.1412 | 0.060582
|  | 23.4144 | 110.0736 |  | 590.6228 | 0.049477
|  | 23.5128 | 109.0333 |  | 593.1044 | 0.030886
|  | 23.6112 | 109.5934 |  | 595.586 | 0.018841
|  | 23.7096 | 111.8341 |  | 598.0676 | 0.020614
|  | 23.808 | 112.9545 |  | 600.5492 | 0.026974
|  | 23.9063 | 112.2342 |  | 603.0308 | 0.036597
|  | 24.0047 | 111.274 |  | 605.5124 | 0.045901
|  | 24.1031 | 109.8735 |  | 607.994 | 0.04754
|  | 24.2015 | 109.0733 |  | 610.4756 | 0.03793
|  | 24.2999 | 109.7135 |  | 612.9573 | 0.022149
|  | 24.3982 | 111.7941 |  | 615.4389 | 0.021642
|  | 24.4966 | 112.9945 |  | 617.9205 | 0.033117
|  | 24.595 | 112.0342 |  | 620.4021 | 0.033322
|  | 24.6934 | 110.6738 |  | 622.8837 | 0.017373
|  | 24.7918 | 109.8735 |  | 625.3653 | 0.018505
|  | 24.8901 | 109.2333 |  | 627.8469 | 0.051547
|  | 24.9885 | 110.4337 |  | 630.3285 | 0.076337
|  | 25.0869 | 112.3543 |  | 632.8101 | 0.080624
|  | 25.1853 | 112.5944 |  | 635.2917 | 0.059782
|  | 25.2837 | 111.6341 |  | 637.7733 | 0.019659
|  | 25.382 | 110.5937 |  | 640.2549 | 0.035581
|  | 25.4804 | 109.7135 |  | 642.7366 | 0.0788
|  | 25.5788 | 109.1533 |  | 645.2182 | 0.10324
|  | 25.6772 | 110.8338 |  | 647.6998 | 0.10446
|  | 25.7756 | 112.7544 |  | 650.1814 | 0.089952
|  | 25.8739 | 112.4343 |  | 652.663 | 0.075516
|  | 25.9723 | 111.474 |  | 655.1446 | 0.069847
|  | 26.0707 | 110.3137 |  | 657.6262 | 0.06319
|  | 26.1691 | 109.1133 |  | 660.1078 | 0.046905
|  | 26.2675 | 109.3934 |  | 662.5894 | 0.029355
|  | 26.3658 | 111.474 |  | 665.071 | 0.031628
|  | 26.4642 | 112.5143 |  | 667.5526 | 0.038093
|  | 26.5626 | 112.4343 |  | 670.0342 | 0.029038
|  | 26.661 | 110.9939 |  | 672.5159 | 0.0045932
|  | 26.7594 | 109.7935 |  | 674.9975 | 0.028676
|  | 26.8577 | 109.2733 |  | 677.4791 | 0.055483
|  | 26.9561 | 109.6335 |  | 679.9607 | 0.067249
|  | 27.0545 | 110.6738 |  | 682.4423 | 0.061139
|  | 27.1529 | 112.7544 |  | 684.9239 | 0.042512
|  | 27.2513 | 112.0342 |  | 687.4055 | 0.024116
|  | 27.3496 | 110.8338 |  | 689.8871 | 0.020351
|  | 27.448 | 110.7138 |  | 692.3687 | 0.019787
|  | 27.5464 | 109.0333 |  | 694.8503 | 0.0095608
|  | 27.6448 | 108.8732 |  | 697.3319 | 0.0082864
|  | 27.7432 | 111.7541 |  | 699.8135 | 0.026467
|  | 27.8415 | 112.1942 |  | 702.2952 | 0.037843
|  | 27.9399 | 111.9142 |  | 704.7768 | 0.038779
|  | 28.0383 | 111.9542 |  | 707.2584 | 0.030579
|  | 28.1367 | 109.8735 |  | 709.74 | 0.018972
|  | 28.2351 | 109.3934 |  | 712.2216 | 0.011843
|  | 28.3334 | 109.9536 |  | 714.7032 | 0.013298
|  | 28.4318 | 110.3137 |  | 717.1848 | 0.025738
|  | 28.5302 | 112.0742 |  | 719.6664 | 0.04951
|  | 28.6286 | 112.5543 |  | 722.148 | 0.07773
|  | 28.727 | 111.514 |  | 724.6296 | 0.10081
|  | 28.8253 | 110.3137 |  | 727.1112 | 0.10974
|  | 28.9237 | 109.3534 |  | 729.5928 | 0.099378
|  | 29.0221 | 109.3134 |  | 732.0744 | 0.071444
|  | 29.1205 | 110.9138 |  | 734.5561 | 0.040945
|  | 29.2189 | 112.9945 |  | 737.0377 | 0.048974
|  | 29.3172 | 112.4343 |  | 739.5193 | 0.077971
|  | 29.4156 | 111.1539 |  | 742.0009 | 0.093576
|  | 29.514 | 109.9536 |  | 744.4825 | 0.088437
|  | 29.6124 | 109.4334 |  | 746.9641 | 0.066915
|  | 29.7108 | 109.7535 |  | 749.4457 | 0.044013
|  | 29.8091 | 111.5941 |  | 751.9273 | 0.04253
|  | 29.9075 | 112.8744 |  | 754.4089 | 0.054235
|  | 30.0059 | 111.9142 |  | 756.8905 | 0.058985
|  | 30.1043 | 110.9539 |  | 759.3721 | 0.055254
|  | 30.2027 | 109.9135 |  | 761.8537 | 0.051656
|  | 30.301 | 108.9532 |  | 764.3354 | 0.056696
|  | 30.3994 | 110.0736 |  | 766.817 | 0.068173
|  | 30.4978 | 112.1542 |  | 769.2986 | 0.078574
|  | 30.5962 | 112.5543 |  | 771.7802 | 0.08169
|  | 30.6946 | 111.8341 |  | 774.2618 | 0.0733
|  | 30.7929 | 110.6338 |  | 776.7434 | 0.053363
|  | 30.8913 | 109.8335 |  | 779.225 | 0.033186
|  | 30.9897 | 109.3134 |  | 781.7066 | 0.041771
|  | 31.0881 | 110.5137 |  | 784.1882 | 0.063352
|  | 31.1865 | 112.3543 |  | 786.6698 | 0.073883
|  | 31.2848 | 112.5143 |  | 789.1514 | 0.070012
|  | 31.3832 | 111.554 |  | 791.633 | 0.061845
|  | 31.4816 | 110.2737 |  | 794.1147 | 0.064833
|  | 31.58 | 109.5534 |  | 796.5963 | 0.075566
|  | 31.6784 | 109.3534 |  | 799.0779 | 0.078484
|  | 31.7767 | 110.7138 |  | 801.5595 | 0.066986
|  | 31.8751 | 112.4743 |  | 804.0411 | 0.045631
|  | 31.9735 | 112.5944 |  | 806.5227 | 0.028524
|  | 32.0719 | 111.1539 |  | 809.0043 | 0.027354
|  | 32.1703 | 110.1136 |  | 811.4859 | 0.026016
|  | 32.2686 | 109.0733 |  | 813.9675 | 0.017042
|  | 32.367 | 109.2733 |  | 816.4491 | 0.024249
|  | 32.4654 | 111.354 |  | 818.9307 | 0.044306
|  | 32.5638 | 112.7944 |  | 821.4123 | 0.055077
|  | 32.6622 | 111.9942 |  | 823.894 | 0.048055
|  | 32.7605 | 111.1139 |  | 826.3756 | 0.024028
|  | 32.8589 | 109.6735 |  | 828.8572 | 0.02189
|  | 32.9573 | 108.9933 |  | 831.3388 | 0.057997
|  | 33.0557 | 110.0336 |  | 833.8204 | 0.085493
|  | 33.1541 | 111.9542 |  | 836.302 | 0.094663
|  | 33.2524 | 112.5944 |  | 838.7836 | 0.084004
|  | 33.3508 | 111.9542 |  | 841.2652 | 0.058903
|  | 33.4492 | 110.6738 |  | 843.7468 | 0.029225
|  | 33.5476 | 109.7135 |  | 846.2284 | 0.00674
|  | 33.646 | 108.9933 |  | 848.71 | 0.010473
|  | 33.7443 | 110.4337 |  | 851.1916 | 0.014843
|  | 33.8427 | 112.1142 |  | 853.6733 | 0.023946
|  | 33.9411 | 112.5143 |  | 856.1549 | 0.038923
|  | 34.0395 | 111.7141 |  | 858.6365 | 0.051354
|  | 34.1379 | 110.2737 |  | 861.1181 | 0.056762
|  | 34.2362 | 109.5534 |  | 863.5997 | 0.056269
|  | 34.3346 | 109.0333 |  | 866.0813 | 0.054775
|  | 34.433 | 110.6338 |  | 868.5629 | 0.057306
|  | 34.5314 | 112.5543 |  | 871.0445 | 0.067149
|  | 34.6298 | 112.5543 |  | 873.5261 | 0.0851
|  | 34.7281 | 112.5543 |  | 876.0077 | 0.10644
|  | 34.8265 | 112.5543 |  | 878.4893 | 0.12155
|  | 34.9249 | 109.1533 |  | 880.9709 | 0.12084
|  | 35.0233 | 109.1533 |  | 883.4526 | 0.099779
|  | 35.1217 | 109.1533 |  | 885.9342 | 0.061758
|  | 35.22 | 109.1533 |  | 888.4158 | 0.019902
|  | 35.3184 | 109.1533 |  | 890.8974 | 0.028276
|  | 35.4168 | 109.1533 |  | 893.379 | 0.04789
|  | 35.5152 | 109.1533 |  | 895.8606 | 0.046172
|  | 35.6136 | 109.1533 |  | 898.3422 | 0.027953
|  | 35.7119 | 109.1533 |  | 900.8238 | 0.023669
|  | 35.8103 | 109.1533 |  | 903.3054 | 0.046613
|  | 35.9087 | 112.6344 |  | 905.787 | 0.06358
|  | 36.0071 | 111.9942 |  | 908.2686 | 0.068869
|  | 36.1055 | 110.9539 |  | 910.7502 | 0.067755
|  | 36.2038 | 109.7535 |  | 913.2319 | 0.067102
|  | 36.3022 | 109.1933 |  | 915.7135 | 0.066979
|  | 36.4006 | 110.1536 |  | 918.1951 | 0.063657
|  | 36.499 | 112.0742 |  | 920.6767 | 0.058305
|  | 36.5974 | 112.7144 |  | 923.1583 | 0.056799
|  | 36.6957 | 111.6741 |  | 925.6399 | 0.059516
|  | 36.7941 | 110.7138 |  | 928.1215 | 0.058975
|  | 36.8925 | 109.6335 |  | 930.6031 | 0.049503
|  | 36.9909 | 108.8332 |  | 933.0847 | 0.031505
|  | 37.0893 | 110.6738 |  | 935.5663 | 0.010233
|  | 37.1876 | 112.4343 |  | 938.0479 | 0.0083472
|  | 37.286 | 112.7144 |  | 940.5295 | 0.017568
|  | 37.3844 | 111.6741 |  | 943.0112 | 0.018648
|  | 37.4828 | 110.7138 |  | 945.4928 | 0.021338
|  | 37.5812 | 109.4334 |  | 947.9744 | 0.032243
|  | 37.6795 | 109.0733 |  | 950.456 | 0.041853
|  | 37.7779 | 110.7538 |  | 952.9376 | 0.041999
|  | 37.8763 | 112.5543 |  | 955.4192 | 0.030799
|  | 37.9747 | 112.3943 |  | 957.9008 | 0.013011
|  | 38.0731 | 111.354 |  | 960.3824 | 0.011029
|  | 38.1714 | 110.4737 |  | 962.864 | 0.020835
|  | 38.2698 | 109.1933 |  | 965.3456 | 0.019963
|  | 38.3682 | 109.3534 |  | 967.8272 | 0.0097161
|  | 38.4666 | 111.274 |  | 970.3088 | 0.0075618
|  | 38.565 | 112.6344 |  | 972.7905 | 0.013639
|  | 38.6633 | 112.1942 |  | 975.2721 | 0.006467
|  | 38.7617 | 111.0739 |  | 977.7537 | 0.018305
|  | 38.8601 | 110.0336 |  | 980.2353 | 0.054376
|  | 38.9585 | 109.3934 |  | 982.7169 | 0.093167
|  | 39.0569 | 109.4734 |  | 985.1985 | 0.12375
|  | 39.1552 | 111.554 |  | 987.6801 | 0.13799
|  | 39.2536 | 112.6744 |  | 990.1617 | 0.13312
|  | 39.352 | 112.0342 |  | 992.6433 | 0.11224
|  | 39.4504 | 110.8338 |  | 995.1249 | 0.082723
|  | 39.5488 | 109.9536 |  | 997.6065 | 0.053811
|  | 39.6471 | 108.9933 |  | 1000.0881 | 0.034091
|  | 39.7455 | 109.9536 |  | 1002.5698 | 0.026753
|  | 39.8439 | 112.1142 |  | 1005.0514 | 0.025363
|  | 39.9423 | 112.9145 |  | 1007.533 | 0.025611
|  | 40.0407 | 111.394 |  | 1010.0146 | 0.028048
|  | 40.139 | 110.3537 |  | 1012.4962 | 0.031042
|  | 40.2374 | 109.4734 |  | 1014.9778 | 0.032196
|  | 40.3358 | 108.9132 |  | 1017.4594 | 0.031981
|  | 40.4342 | 110.7538 |  | 1019.941 | 0.033316
|  | 40.5326 | 112.5143 |  | 1022.4226 | 0.038206
|  | 40.6309 | 112.5543 |  | 1024.9042 | 0.045861
|  | 40.7293 | 111.434 |  | 1027.3858 | 0.053874
|  | 40.8277 | 110.2336 |  | 1029.8674 | 0.058947
|  | 40.9261 | 109.2733 |  | 1032.3491 | 0.057727
|  | 41.0245 | 109.4334 |  | 1034.8307 | 0.048469
|  | 41.1228 | 110.9539 |  | 1037.3123 | 0.032185
|  | 41.2212 | 112.4743 |  | 1039.7939 | 0.01221
|  | 41.3196 | 112.1142 |  | 1042.2755 | 0.0074689
|  | 41.418 | 111.0739 |  | 1044.7571 | 0.023828
|  | 41.5164 | 109.9536 |  | 1047.2387 | 0.035409
|  | 41.6147 | 108.9933 |  | 1049.7203 | 0.042387
|  | 41.7131 | 109.2733 |  | 1052.2019 | 0.047214
|  | 41.8115 | 111.274 |  | 1054.6835 | 0.053864
|  | 41.9099 | 112.6344 |  | 1057.1651 | 0.062698
|  | 42.0083 | 112.2743 |  | 1059.6467 | 0.068354
|  | 42.1066 | 110.9138 |  | 1062.1284 | 0.065458
|  | 42.205 | 109.7935 |  | 1064.61 | 0.054709
|  | 42.3034 | 108.9132 |  | 1067.0916 | 0.044707
|  | 42.4018 | 109.6335 |  | 1069.5732 | 0.042055
|  | 42.5002 | 112.0342 |  | 1072.0548 | 0.038389
|  | 42.5985 | 112.7544 |  | 1074.5364 | 0.030657
|  | 42.6969 | 111.7141 |  | 1077.018 | 0.043228
|  | 42.7953 | 110.6338 |  | 1079.4996 | 0.073334
|  | 42.8937 | 109.5934 |  | 1081.9812 | 0.095499
|  | 42.9921 | 108.8732 |  | 1084.4628 | 0.096184
|  | 43.0904 | 110.5537 |  | 1086.9444 | 0.074188
|  | 43.1888 | 112.4743 |  | 1089.426 | 0.044978
|  | 43.2872 | 112.7544 |  | 1091.9077 | 0.04705
|  | 43.3856 | 111.6341 |  | 1094.3893 | 0.068398
|  | 43.484 | 110.5137 |  | 1096.8709 | 0.07528
|  | 43.5823 | 109.3134 |  | 1099.3525 | 0.063098
|  | 43.6807 | 109.0733 |  | 1101.8341 | 0.040077
|  | 43.7791 | 111.0339 |  | 1104.3157 | 0.021331
|  | 43.8775 | 112.6344 |  | 1106.7973 | 0.022514
|  | 43.9759 | 112.6744 |  | 1109.2789 | 0.030114
|  | 44.0742 | 111.554 |  | 1111.7605 | 0.033904
|  | 44.1726 | 110.1936 |  | 1114.2421 | 0.031132
|  | 44.271 | 109.2333 |  | 1116.7237 | 0.019147
|  | 44.3694 | 109.4334 |  | 1119.2053 | 0.0043169
|  | 44.4678 | 111.0339 |  | 1121.687 | 0.025436
|  | 44.5661 | 112.7944 |  | 1124.1686 | 0.042229
|  | 44.6645 | 112.3943 |  | 1126.6502 | 0.046008
|  | 44.7629 | 111.0339 |  | 1129.1318 | 0.036138
|  | 44.8613 | 110.1536 |  | 1131.6134 | 0.022948
|  | 44.9597 | 109.1133 |  | 1134.095 | 0.029
|  | 45.058 | 109.6335 |  | 1136.5766 | 0.042486
|  | 45.1564 | 112.0342 |  | 1139.0582 | 0.046929
|  | 45.2548 | 112.6744 |  | 1141.5398 | 0.040081
|  | 45.3532 | 112.0342 |  | 1144.0214 | 0.026166
|  | 45.4516 | 110.6738 |  | 1146.503 | 0.016599
|  | 45.5499 | 109.7135 |  | 1148.9846 | 0.025777
|  | 45.6483 | 108.9933 |  | 1151.4663 | 0.040679
|  | 45.7467 | 109.9135 |  | 1153.9479 | 0.053931
|  | 45.8451 | 111.8341 |  | 1156.4295 | 0.063446
|  | 45.9435 | 112.3943 |  | 1158.9111 | 0.066903
|  | 46.0418 | 111.7541 |  | 1161.3927 | 0.061953
|  | 46.1402 | 110.5537 |  | 1163.8743 | 0.048207
|  | 46.2386 | 109.1933 |  | 1166.3559 | 0.028883
|  | 46.337 | 109.1533 |  | 1168.8375 | 0.011025
|  | 46.4354 | 111.0739 |  | 1171.3191 | 0.0086272
|  | 46.5337 | 112.6744 |  | 1173.8007 | 0.015751
|  | 46.6321 | 112.6744 |  | 1176.2823 | 0.027042
|  | 46.7305 | 111.354 |  | 1178.7639 | 0.040692
|  | 46.8289 | 110.1536 |  | 1181.2456 | 0.046656
|  | 46.9273 | 109.2733 |  | 1183.7272 | 0.036512
|  | 47.0256 | 109.2733 |  | 1186.2088 | 0.0091653
|  | 47.124 | 109.2733 |  | 1188.6904 | 0.028635
|  | 47.2224 | 109.2733 |  | 1191.172 | 0.063288
|  | 47.3208 | 109.2733 |  | 1193.6536 | 0.082575
|  | 47.4192 | 109.2733 |  | 1196.1352 | 0.080561
|  | 47.5175 | 109.2733 |  | 1198.6168 | 0.062262
|  | 47.6159 | 109.2733 |  | 1201.0984 | 0.046407
|  | 47.7143 | 109.5534 |  | 1203.58 | 0.052458
|  | 47.8127 | 111.7941 |  | 1206.0616 | 0.065337
|  | 47.9111 | 112.9945 |  | 1208.5432 | 0.070101
|  | 48.0094 | 112.1142 |  | 1211.0249 | 0.06548
|  | 48.1078 | 111.2339 |  | 1213.5065 | 0.053709
|  | 48.2062 | 110.1136 |  | 1215.9881 | 0.038358
|  | 48.3046 | 108.9132 |  | 1218.4697 | 0.032073
|  | 48.403 | 109.6335 |  | 1220.9513 | 0.04326
|  | 48.5013 | 111.7141 |  | 1223.4329 | 0.052192
|  | 48.5997 | 112.5944 |  | 1225.9145 | 0.045492
|  | 48.6981 | 112.2342 |  | 1228.3961 | 0.022838
|  | 48.7965 | 110.9539 |  | 1230.8777 | 0.0178
|  | 48.8949 | 109.5934 |  | 1233.3593 | 0.045775
|  | 48.9932 | 109.1533 |  | 1235.8409 | 0.061185
|  | 49.0916 | 110.5937 |  | 1238.3225 | 0.056706
|  | 49.19 | 112.2743 |  | 1240.8042 | 0.035636
|  | 49.2884 | 112.7144 |  | 1243.2858 | 0.019405
|  | 49.3868 | 111.274 |  | 1245.7674 | 0.040164
|  | 49.4851 | 110.2336 |  | 1248.249 | 0.059961
|  | 49.5835 | 109.3534 |  | 1250.7306 | 0.065176
|  | 49.6819 | 109.2733 |  | 1253.2122 | 0.054827
|  | 49.7803 | 111.0339 |  | 1255.6938 | 0.03425
|  | 49.8787 | 112.6344 |  | 1258.1754 | 0.016172
|  | 49.977 | 112.4343 |  | 1260.657 | 0.016341
|  | 50.0754 | 111.2339 |  | 1263.1386 | 0.01356
|  | 50.1738 | 110.2737 |  | 1265.6202 | 0.015175
|  | 50.2722 | 109.3934 |  | 1268.1018 | 0.047188
|  | 50.3706 | 109.5934 |  | 1270.5835 | 0.082183
|  | 50.4689 | 111.6741 |  | 1273.0651 | 0.1054
|  | 50.5673 | 112.7144 |  | 1275.5467 | 0.1088
|  | 50.6657 | 112.0742 |  | 1278.0283 | 0.093688
|  | 50.7641 | 111.1139 |  | 1280.5099 | 0.069989
|  | 50.8625 | 109.8335 |  | 1282.9915 | 0.049561
|  | 50.9608 | 108.9532 |  | 1285.4731 | 0.036225
|  | 51.0592 | 109.8335 |  | 1287.9547 | 0.032225
|  | 51.1576 | 112.0742 |  | 1290.4363 | 0.048513
|  | 51.256 | 112.7944 |  | 1292.9179 | 0.073827
|  | 51.3544 | 112.2342 |  | 1295.3995 | 0.092166
|  | 51.4527 | 110.8738 |  | 1297.8811 | 0.09564
|  | 51.5511 | 109.8335 |  | 1300.3628 | 0.083848
|  | 51.6495 | 108.9933 |  | 1302.8444 | 0.063067
|  | 51.7479 | 110.2737 |  | 1305.326 | 0.046462
|  | 51.8463 | 112.1942 |  | 1307.8076 | 0.046707
|  | 51.9446 | 112.6744 |  | 1310.2892 | 0.052247
|  | 52.043 | 111.554 |  | 1312.7708 | 0.046839
|  | 52.1414 | 110.4337 |  | 1315.2524 | 0.02872
|  | 52.2398 | 109.4734 |  | 1317.734 | 0.022036
|  | 52.3382 | 109.1933 |  | 1320.2156 | 0.042501
|  | 52.4365 | 110.4737 |  | 1322.6972 | 0.054632
|  | 52.5349 | 112.3943 |  | 1325.1788 | 0.050833
|  | 52.6333 | 112.1542 |  | 1327.6604 | 0.044174
|  | 52.7317 | 111.0339 |  | 1330.1421 | 0.054105
|  | 52.8301 | 110.1536 |  | 1332.6237 | 0.06727
|  | 52.9284 | 109.3534 |  | 1335.1053 | 0.066158
|  | 53.0268 | 109.4734 |  | 1337.5869 | 0.050834
|  | 53.1252 | 111.554 |  | 1340.0685 | 0.038273
|  | 53.2236 | 112.7544 |  | 1342.5501 | 0.044356
|  | 53.322 | 112.1142 |  | 1345.0317 | 0.053635
|  | 53.4203 | 111.2339 |  | 1347.5133 | 0.05661
|  | 53.5187 | 110.1936 |  | 1349.9949 | 0.054787
|  | 53.6171 | 109.3534 |  | 1352.4765 | 0.047931
|  | 53.7155 | 109.7535 |  | 1354.9581 | 0.03804
|  | 53.8139 | 111.8341 |  | 1357.4397 | 0.040012
|  | 53.9122 | 112.6344 |  | 1359.9214 | 0.051168
|  | 54.0106 | 111.9942 |  | 1362.403 | 0.047657
|  | 54.109 | 110.8738 |  | 1364.8846 | 0.018613
|  | 54.2074 | 109.7535 |  | 1367.3662 | 0.030739
|  | 54.3058 | 109.3534 |  | 1369.8478 | 0.080274
|  | 54.4041 | 110.2336 |  | 1372.3294 | 0.10835
|  | 54.5025 | 111.6741 |  | 1374.811 | 0.10106
|  | 54.6009 | 112.5143 |  | 1377.2926 | 0.060628
|  | 54.6993 | 111.8741 |  | 1379.7742 | 0.0042371
|  | 54.7977 | 110.7538 |  | 1382.2558 | 0.046244
|  | 54.896 | 110.5937 |  | 1384.7374 | 0.077065
|  | 54.9944 | 109.2333 |  | 1387.219 | 0.090908
|  | 55.0928 | 108.9933 |  | 1389.7007 | 0.096422
|  | 55.1912 | 112.0342 |  | 1392.1823 | 0.089228
|  | 55.2896 | 112.1542 |  | 1394.6639 | 0.059555
|  | 55.3879 | 111.9942 |  | 1397.1455 | 0.037513
|  | 55.4863 | 111.9142 |  | 1399.6271 | 0.08549
|  | 55.5847 | 109.6335 |  | 1402.1087 | 0.12659
|  | 55.6831 | 109.4334 |  | 1404.5903 | 0.12812
|  | 55.7815 | 109.1533 |  | 1407.0719 | 0.091457
|  | 55.8798 | 111.2339 |  | 1409.5535 | 0.049959
|  | 55.9782 | 112.5944 |  | 1412.0351 | 0.048466
|  | 56.0766 | 112.3943 |  | 1414.5167 | 0.043575
|  | 56.175 | 111.354 |  | 1416.9983 | 0.052318
|  | 56.2734 | 110.0736 |  | 1419.48 | 0.12177
|  | 56.3717 | 109.2733 |  | 1421.9616 | 0.18414
|  | 56.4701 | 109.6335 |  | 1424.4432 | 0.19696
|  | 56.5685 | 111.314 |  | 1426.9248 | 0.15001
|  | 56.6669 | 112.7544 |  | 1429.4064 | 0.078291
|  | 56.7653 | 112.0342 |  | 1431.888 | 0.084618
|  | 56.8636 | 110.9138 |  | 1434.3696 | 0.11783
|  | 56.962 | 110.1136 |  | 1436.8512 | 0.097208
|  | 57.0604 | 108.9933 |  | 1439.3328 | 0.036549
|  | 57.1588 | 109.6335 |  | 1441.8144 | 0.06235
|  | 57.2572 | 111.8741 |  | 1444.296 | 0.098714
|  | 57.3555 | 112.9145 |  | 1446.7776 | 0.074277
|  | 57.4539 | 111.7141 |  | 1449.2592 | 0.015912
|  | 57.5523 | 110.7938 |  | 1451.7409 | 0.096089
|  | 57.6507 | 109.5134 |  | 1454.2225 | 0.15332
|  | 57.7491 | 108.7932 |  | 1456.7041 | 0.1609
|  | 57.8474 | 109.9936 |  | 1459.1857 | 0.14966
|  | 57.9458 | 112.0742 |  | 1461.6673 | 0.15992
|  | 58.0442 | 112.7144 |  | 1464.1489 | 0.15846
|  | 58.1426 | 111.8341 |  | 1466.6305 | 0.11214
|  | 58.241 | 110.5537 |  | 1469.1121 | 0.10313
|  | 58.3393 | 109.2733 |  | 1471.5937 | 0.17516
|  | 58.4377 | 109.0733 |  | 1474.0753 | 0.18926
|  | 58.5361 | 110.7538 |  | 1476.5569 | 0.11232
|  | 58.6345 | 112.3543 |  | 1479.0385 | 0.17854
|  | 58.7329 | 112.5143 |  | 1481.5202 | 0.35723
|  | 58.8312 | 111.2339 |  | 1484.0018 | 0.40984
|  | 58.9296 | 110.0336 |  | 1486.4834 | 0.22592
|  | 59.028 | 109.3934 |  | 1488.965 | 0.23119
|  | 59.1264 | 109.3934 |  | 1491.4466 | 0.83335
|  | 59.2248 | 109.3934 |  | 1493.9282 | 1.4222
|  | 59.3231 | 109.3934 |  | 1496.4098 | 1.796
|  | 59.4215 | 109.3934 |  | 1498.8914 | 1.8182
|  | 59.5199 | 111.0739 |  | 1501.373 | 1.4736
|  | 59.6183 | 110.1936 |  | 1503.8546 | 0.87713
|  | 59.7167 | 109.2333 |  | 1506.3362 | 0.22978
|  | 59.815 | 109.7935 |  | 1508.8178 | 0.2746
|  | 59.9134 | 112.1942 |  | 1511.2995 | 0.4952
|  | 60.0118 | 112.9145 |  | 1513.7811 | 0.43346
|  | 60.1102 | 111.8741 |  | 1516.2627 | 0.19808
|  | 60.2086 | 110.6738 |  | 1518.7443 | 0.10664
|  | 60.3069 | 109.7135 |  | 1521.2259 | 0.23289
|  | 60.4053 | 108.7532 |  | 1523.7075 | 0.19683
|  | 60.5037 | 110.1936 |  | 1526.1891 | 0.089129
|  | 60.6021 | 112.2743 |  | 1528.6707 | 0.26636
|  | 60.7005 | 112.4743 |  | 1531.1523 | 0.43124
|  | 60.7988 | 111.8341 |  | 1533.6339 | 0.46376
|  | 60.8972 | 110.7138 |  | 1536.1155 | 0.36279
|  | 60.9956 | 109.4334 |  | 1538.5971 | 0.19597
|  | 61.094 | 109.3134 |  | 1541.0788 | 0.081463
|  | 61.1924 | 110.7538 |  | 1543.5604 | 0.091365
|  | 61.2907 | 112.6744 |  | 1546.042 | 0.060592
|  | 61.3891 | 112.6744 |  | 1548.5236 | 0.024046
|  | 61.4875 | 111.474 |  | 1551.0052 | 0.10638
|  | 61.5859 | 110.2737 |  | 1553.4868 | 0.13684
|  | 61.6843 | 109.5534 |  | 1555.9684 | 0.10182
|  | 61.7826 | 109.0733 |  | 1558.45 | 0.033155
|  | 61.881 | 110.9138 |  | 1560.9316 | 0.055851
|  | 61.9794 | 112.7544 |  | 1563.4132 | 0.086486
|  | 62.0778 | 112.6344 |  | 1565.8948 | 0.088613
|  | 62.1762 | 111.434 |  | 1568.3764 | 0.11234
|  | 62.2745 | 110.0736 |  | 1570.8581 | 0.15022
|  | 62.3729 | 109.1933 |  | 1573.3397 | 0.15335
|  | 62.4713 | 109.4334 |  | 1575.8213 | 0.10728
|  | 62.5697 | 111.6741 |  | 1578.3029 | 0.038757
|  | 62.6681 | 112.9545 |  | 1580.7845 | 0.045167
|  | 62.7664 | 111.9942 |  | 1583.2661 | 0.05964
|  | 62.8648 | 110.8738 |  | 1585.7477 | 0.021941
|  | 62.9632 | 109.8335 |  | 1588.2293 | 0.050157
|  | 63.0616 | 109.0333 |  | 1590.7109 | 0.1151
|  | 63.16 | 109.9936 |  | 1593.1925 | 0.13649
|  | 63.2583 | 111.9142 |  | 1595.6741 | 0.10077
|  | 63.3567 | 112.7144 |  | 1598.1557 | 0.040885
|  | 63.4551 | 111.9942 |  | 1600.6374 | 0.090589
|  | 63.5535 | 110.7138 |  | 1603.119 | 0.14889
|  | 63.6519 | 109.6735 |  | 1605.6006 | 0.16057
|  | 63.7502 | 109.1933 |  | 1608.0822 | 0.12917
|  | 63.8486 | 110.3137 |  | 1610.5638 | 0.081681
|  | 63.947 | 112.2342 |  | 1613.0454 | 0.051575
|  | 64.0454 | 112.5543 |  | 1615.527 | 0.048631
|  | 64.1438 | 111.6741 |  | 1618.0086 | 0.06102
|  | 64.2421 | 110.3137 |  | 1620.4902 | 0.081412
|  | 64.3405 | 109.1933 |  | 1622.9718 | 0.090157
|  | 64.4389 | 108.9532 |  | 1625.4534 | 0.073284
|  | 64.5373 | 110.4737 |  | 1627.935 | 0.035682
|  | 64.6357 | 112.5543 |  | 1630.4167 | 0.018162
|  | 64.734 | 112.5944 |  | 1632.8983 | 0.041221
|  | 64.8324 | 111.2339 |  | 1635.3799 | 0.050379
|  | 64.9308 | 110.1136 |  | 1637.8615 | 0.065656
|  | 65.0292 | 108.8332 |  | 1640.3431 | 0.095777
|  | 65.1276 | 109.1933 |  | 1642.8247 | 0.11667
|  | 65.2259 | 111.274 |  | 1645.3063 | 0.112
|  | 65.3243 | 112.4743 |  | 1647.7879 | 0.085141
|  | 65.4227 | 112.2342 |  | 1650.2695 | 0.055352
|  | 65.5211 | 111.1139 |  | 1652.7511 | 0.03917
|  | 65.6195 | 110.1536 |  | 1655.2327 | 0.023766
|  | 65.7178 | 109.1133 |  | 1657.7143 | 0.0093839
|  | 65.8162 | 109.5534 |  | 1660.196 | 0.040542
|  | 65.9146 | 111.6341 |  | 1662.6776 | 0.062033
|  | 66.013 | 112.8344 |  | 1665.1592 | 0.061198
|  | 66.1114 | 111.8741 |  | 1667.6408 | 0.041241
|  | 66.2097 | 110.7538 |  | 1670.1224 | 0.020263
|  | 66.3081 | 109.7135 |  | 1672.604 | 0.018011
|  | 66.4065 | 108.9132 |  | 1675.0856 | 0.019903
|  | 66.5049 | 110.3537 |  | 1677.5672 | 0.042735
|  | 66.6033 | 112.2743 |  | 1680.0488 | 0.079317
|  | 66.7016 | 112.9945 |  | 1682.5304 | 0.10679
|  | 66.8 | 111.7141 |  | 1685.012 | 0.11201
|  | 66.8984 | 110.4337 |  | 1687.4936 | 0.094576
|  | 66.9968 | 109.5534 |  | 1689.9753 | 0.064512
|  | 67.0952 | 108.7532 |  | 1692.4569 | 0.033635
|  | 67.1935 | 110.3537 |  | 1694.9385 | 0.016784
|  | 67.2919 | 112.5944 |  | 1697.4201 | 0.042275
|  | 67.3903 | 112.3943 |  | 1699.9017 | 0.075205
|  | 67.4887 | 111.434 |  | 1702.3833 | 0.098092
|  | 67.5871 | 110.4737 |  | 1704.8649 | 0.10239
|  | 67.6854 | 109.6735 |  | 1707.3465 | 0.089186
|  | 67.7838 | 109.0333 |  | 1709.8281 | 0.068662
|  | 67.8822 | 110.9539 |  | 1712.3097 | 0.054016
|  | 67.9806 | 112.5543 |  | 1714.7913 | 0.054115
|  | 68.079 | 112.5944 |  | 1717.2729 | 0.068319
|  | 68.1773 | 111.6341 |  | 1719.7546 | 0.085923
|  | 68.2757 | 110.5137 |  | 1722.2362 | 0.093424
|  | 68.3741 | 109.1933 |  | 1724.7178 | 0.083337
|  | 68.4725 | 109.0733 |  | 1727.1994 | 0.05998
|  | 68.5709 | 111.1539 |  | 1729.681 | 0.04028
|  | 68.6692 | 112.4343 |  | 1732.1626 | 0.03631
|  | 68.7676 | 112.3943 |  | 1734.6442 | 0.029618
|  | 68.866 | 111.0339 |  | 1737.1258 | 0.0086394
|  | 68.9644 | 109.7535 |  | 1739.6074 | 0.020086
|  | 69.0628 | 109.1133 |  | 1742.089 | 0.044059
|  | 69.1611 | 109.9135 |  | 1744.5706 | 0.05497
|  | 69.2595 | 111.8341 |  | 1747.0522 | 0.054043
|  | 69.3579 | 112.6744 |  | 1749.5339 | 0.048445
|  | 69.4563 | 112.6744 |  | 1752.0155 | 0.039746
|  | 69.5547 | 112.6744 |  | 1754.4971 | 0.023965
|  | 69.653 | 112.6744 |  | 1756.9787 | 0.019101
|  | 69.7514 | 112.6744 |  | 1759.4603 | 0.041553
|  | 69.8498 | 112.6744 |  | 1761.9419 | 0.055167
|  | 69.9482 | 112.6744 |  | 1764.4235 | 0.047833
|  | 70.0466 | 112.6744 |  | 1766.9051 | 0.021626
|  | 70.1449 | 112.6744 |  | 1769.3867 | 0.010792
|  | 70.2433 | 112.6744 |  | 1771.8683 | 0.033511
|  | 70.3417 | 109.3934 |  | 1774.3499 | 0.038404
|  | 70.4401 | 108.9532 |  | 1776.8315 | 0.037992
|  | 70.5385 | 110.8738 |  | 1779.3132 | 0.054504
|  | 70.6368 | 112.7144 |  | 1781.7948 | 0.071749
|  | 70.7352 | 112.2342 |  | 1784.2764 | 0.07011
|  | 70.8336 | 111.434 |  | 1786.758 | 0.053108
|  | 70.932 | 110.2336 |  | 1789.2396 | 0.056619
|  | 71.0304 | 109.0333 |  | 1791.7212 | 0.084333
|  | 71.1287 | 109.3134 |  | 1794.2028 | 0.098765
|  | 71.2271 | 111.1539 |  | 1796.6844 | 0.087447
|  | 71.3255 | 112.5543 |  | 1799.166 | 0.063069
|  | 71.4239 | 112.3543 |  | 1801.6476 | 0.058526
|  | 71.5223 | 110.8338 |  | 1804.1292 | 0.071493
|  | 71.6206 | 109.8735 |  | 1806.6108 | 0.069891
|  | 71.719 | 109.0733 |  | 1809.0925 | 0.047464
|  | 71.8174 | 109.7935 |  | 1811.5741 | 0.023182
|  | 71.9158 | 111.7141 |  | 1814.0557 | 0.034461
|  | 72.0142 | 112.7544 |  | 1816.5373 | 0.045784
|  | 72.1125 | 111.8741 |  | 1819.0189 | 0.039501
|  | 72.2109 | 110.6738 |  | 1821.5005 | 0.021585
|  | 72.3093 | 109.8335 |  | 1823.9821 | 0.0038324
|  | 72.4077 | 109.1933 |  | 1826.4637 | 0.007776
|  | 72.5061 | 110.0736 |  | 1828.9453 | 0.022149
|  | 72.6044 | 111.9142 |  | 1831.4269 | 0.044787
|  | 72.7028 | 112.3943 |  | 1833.9085 | 0.06703
|  | 72.8012 | 111.434 |  | 1836.3901 | 0.07898
|  | 72.8996 | 110.5537 |  | 1838.8718 | 0.0762
|  | 72.998 | 109.4334 |  | 1841.3534 | 0.061092
|  | 73.0963 | 109.0333 |  | 1843.835 | 0.040477
|  | 73.1947 | 110.9539 |  | 1846.3166 | 0.021857
|  | 73.2931 | 112.6744 |  | 1848.7982 | 0.013077
|  | 73.3915 | 112.6744 |  | 1851.2798 | 0.016992
|  | 73.4899 | 112.6744 |  | 1853.7614 | 0.022072
|  | 73.5882 | 112.6744 |  | 1856.243 | 0.024133
|  | 73.6866 | 112.6744 |  | 1858.7246 | 0.0237
|  | 73.785 | 112.6744 |  | 1861.2062 | 0.0215
|  | 73.8834 | 112.6744 |  | 1863.6878 | 0.016486
|  | 73.9818 | 112.6744 |  | 1866.1694 | 0.010148
|  | 74.0801 | 112.6744 |  | 1868.6511 | 0.016104
|  | 74.1785 | 112.6744 |  | 1871.1327 | 0.029011
|  | 74.2769 | 109.8335 |  | 1873.6143 | 0.038427
|  | 74.3753 | 109.1133 |  | 1876.0959 | 0.041359
|  | 74.4737 | 109.9135 |  | 1878.5775 | 0.038541
|  | 74.572 | 111.8341 |  | 1881.0591 | 0.032239
|  | 74.6704 | 112.7944 |  | 1883.5407 | 0.024598
|  | 74.7688 | 112.0742 |  | 1886.0223 | 0.018899
|  | 74.8672 | 110.7938 |  | 1888.5039 | 0.018021
|  | 74.9656 | 109.9135 |  | 1890.9855 | 0.016465
|  | 75.0639 | 108.8732 |  | 1893.4671 | 0.0088142
|  | 75.1623 | 109.8335 |  | 1895.9487 | 0.013666
|  | 75.2607 | 111.9142 |  | 1898.4304 | 0.031034
|  | 75.3591 | 112.7944 |  | 1900.912 | 0.040975
|  | 75.4575 | 111.9142 |  | 1903.3936 | 0.035209
|  | 75.5558 | 110.8738 |  | 1905.8752 | 0.013607
|  | 75.6542 | 109.7535 |  | 1908.3568 | 0.025341
|  | 75.7526 | 109.2333 |  | 1910.8384 | 0.059049
|  | 75.851 | 110.4337 |  | 1913.32 | 0.081836
|  | 75.9494 | 112.3543 |  | 1915.8016 | 0.087399
|  | 76.0477 | 112.3543 |  | 1918.2832 | 0.078304
|  | 76.1461 | 112.3543 |  | 1920.7648 | 0.06366
|  | 76.2445 | 110.4337 |  | 1923.2464 | 0.051341
|  | 76.3429 | 109.0733 |  | 1925.728 | 0.040283
|  | 76.4413 | 109.2733 |  | 1928.2097 | 0.02643
|  | 76.5396 | 110.9539 |  | 1930.6913 | 0.01364
|  | 76.638 | 112.4743 |  | 1933.1729 | 0.012517
|  | 76.7364 | 112.4743 |  | 1935.6545 | 0.01123
|  | 76.8348 | 112.4743 |  | 1938.1361 | 0.002636
|  | 76.9332 | 112.4743 |  | 1940.6177 | 0.020179
|  | 77.0315 | 112.4743 |  | 1943.0993 | 0.03977
|  | 77.1299 | 112.4743 |  | 1945.5809 | 0.053249
|  | 77.2283 | 111.434 |  | 1948.0625 | 0.05808
|  | 77.3267 | 112.8744 |  | 1950.5441 | 0.05636
|  | 77.4251 | 112.1542 |  | 1953.0257 | 0.05197
|  | 77.5234 | 110.6338 |  | 1955.5073 | 0.046808
|  | 77.6218 | 110.0736 |  | 1957.989 | 0.0405
|  | 77.7202 | 108.9532 |  | 1960.4706 | 0.032925
|  | 77.8186 | 109.5134 |  | 1962.9522 | 0.025038
|  | 77.917 | 111.9142 |  | 1965.4338 | 0.019285
|  | 78.0153 | 112.6344 |  | 1967.9154 | 0.021548
|  | 78.1137 | 111.8341 |  | 1970.397 | 0.031428
|  | 78.2121 | 110.9138 |  | 1972.8786 | 0.041472
|  | 78.3105 | 109.6335 |  | 1975.3602 | 0.047017
|  | 78.4089 | 108.9933 |  | 1977.8418 | 0.046926
|  | 78.5072 | 110.4337 |  | 1980.3234 | 0.042634
|  | 78.6056 | 112.3543 |  | 1982.805 | 0.036222
|  | 78.704 | 112.6744 |  | 1985.2866 | 0.028661
|  | 78.8024 | 111.6341 |  | 1987.7683 | 0.020014
|  | 78.9008 | 110.5937 |  | 1990.2499 | 0.010923
|  | 78.9991 | 109.3934 |  | 1992.7315 | 0.0040169
|  | 79.0975 | 108.9132 |  | 1995.2131 | 0.0068198
|  | 79.1959 | 110.5537 |  | 1997.6947 | 0.015108
|  | 79.2943 | 112.4743 |  | 2000.1763 | 0.030222
|  | 79.3927 | 112.5543 |  | 2002.6579 | 0.052514
|  | 79.491 | 111.434 |  | 2005.1395 | 0.077387
|  | 79.5894 | 109.9936 |  | 2007.6211 | 0.097377
|  | 79.6878 | 109.2733 |  | 2010.1027 | 0.1052
|  | 79.7862 | 109.4334 |  | 2012.5843 | 0.097218
|  | 79.8846 | 111.6741 |  | 2015.0659 | 0.075796
|  | 79.9829 | 112.8744 |  | 2017.5476 | 0.049375
|  | 80.0813 | 112.3943 |  | 2020.0292 | 0.0298
|  | 80.1797 | 111.2339 |  | 2022.5108 | 0.022416
|  | 80.2781 | 110.1936 |  | 2024.9924 | 0.016813
|  | 80.3765 | 109.0733 |  | 2027.474 | 0.01121
|  | 80.4748 | 109.8735 |  | 2029.9556 | 0.017533
|  | 80.5732 | 111.9542 |  | 2032.4372 | 0.026816
|  | 80.6716 | 112.7544 |  | 2034.9188 | 0.034819
|  | 80.77 | 111.9542 |  | 2037.4004 | 0.043374
|  | 80.8684 | 110.9939 |  | 2039.882 | 0.049171
|  | 80.9667 | 109.9536 |  | 2042.3636 | 0.045247
|  | 81.0651 | 109.0733 |  | 2044.8452 | 0.028859
|  | 81.1635 | 109.9936 |  | 2047.3269 | 0.012738
|  | 81.2619 | 112.0742 |  | 2049.8085 | 0.031126
|  | 81.3603 | 112.7144 |  | 2052.2901 | 0.046894
|  | 81.4586 | 111.9142 |  | 2054.7717 | 0.046717
|  | 81.557 | 110.8738 |  | 2057.2533 | 0.031301
|  | 81.6554 | 109.5134 |  | 2059.7349 | 0.011576
|  | 81.7538 | 109.1933 |  | 2062.2165 | 0.020562
|  | 81.8522 | 110.3937 |  | 2064.6981 | 0.033789
|  | 81.9505 | 110.3537 |  | 2067.1797 | 0.038938
|  | 82.0489 | 112.4343 |  | 2069.6613 | 0.039506
|  | 82.1473 | 111.6741 |  | 2072.1429 | 0.037503
|  | 82.2457 | 110.5537 |  | 2074.6245 | 0.030462
|  | 82.3441 | 110.5937 |  | 2077.1062 | 0.019051
|  | 82.4424 | 109.2333 |  | 2079.5878 | 0.018661
|  | 82.5408 | 109.1933 |  | 2082.0694 | 0.030519
|  | 82.6392 | 112.5543 |  | 2084.551 | 0.036072
|  | 82.7376 | 112.7144 |  | 2087.0326 | 0.029975
|  | 82.836 | 111.274 |  | 2089.5142 | 0.01427
|  | 82.9343 | 111.2339 |  | 2091.9958 | 0.0078782
|  | 83.0327 | 109.3934 |  | 2094.4774 | 0.025743
|  | 83.1311 | 109.0733 |  | 2096.959 | 0.037804
|  | 83.2295 | 109.8735 |  | 2099.4406 | 0.040209
|  | 83.3279 | 111.7941 |  | 2101.9222 | 0.031618
|  | 83.4262 | 112.5143 |  | 2104.4038 | 0.017268
|  | 83.5246 | 112.1542 |  | 2106.8855 | 0.026537
|  | 83.623 | 110.7138 |  | 2109.3671 | 0.049356
|  | 83.7214 | 109.9135 |  | 2111.8487 | 0.063773
|  | 83.8198 | 109.1933 |  | 2114.3303 | 0.064054
|  | 83.9181 | 110.1536 |  | 2116.8119 | 0.054712
|  | 84.0165 | 112.2342 |  | 2119.2935 | 0.05032
|  | 84.1149 | 112.8744 |  | 2121.7751 | 0.05757
|  | 84.2133 | 111.8341 |  | 2124.2567 | 0.063418
|  | 84.3117 | 110.4737 |  | 2126.7383 | 0.05908
|  | 84.41 | 109.6735 |  | 2129.2199 | 0.047208
|  | 84.5084 | 109.1533 |  | 2131.7015 | 0.035925
|  | 84.6068 | 110.4337 |  | 2134.1831 | 0.02767
|  | 84.7052 | 112.3543 |  | 2136.6648 | 0.016411
|  | 84.8036 | 112.5543 |  | 2139.1464 | 0.0086578
|  | 84.9019 | 111.5941 |  | 2141.628 | 0.025951
|  | 85.0003 | 110.3937 |  | 2144.1096 | 0.040537
|  | 85.0987 | 109.4734 |  | 2146.5912 | 0.044177
|  | 85.1971 | 109.0733 |  | 2149.0728 | 0.035778
|  | 85.2955 | 110.9939 |  | 2151.5544 | 0.018841
|  | 85.3938 | 112.8344 |  | 2154.036 | 0.0037175
|  | 85.4922 | 112.5543 |  | 2156.5176 | 0.020971
|  | 85.5906 | 111.274 |  | 2158.9992 | 0.034897
|  | 85.689 | 110.1536 |  | 2161.4808 | 0.041548
|  | 85.7874 | 109.5134 |  | 2163.9624 | 0.039872
|  | 85.8857 | 109.5134 |  | 2166.444 | 0.032997
|  | 85.9841 | 109.5134 |  | 2168.9257 | 0.030933
|  | 86.0825 | 112.7144 |  | 2171.4073 | 0.038429
|  | 86.1809 | 112.4743 |  | 2173.8889 | 0.044993
|  | 86.2793 | 111.354 |  | 2176.3705 | 0.043022
|  | 86.3776 | 109.9936 |  | 2178.8521 | 0.032588
|  | 86.476 | 109.2733 |  | 2181.3337 | 0.020217
|  | 86.5744 | 110.1536 |  | 2183.8153 | 0.017588
|  | 86.6728 | 112.2342 |  | 2186.2969 | 0.022843
|  | 86.7712 | 113.0345 |  | 2188.7785 | 0.026125
|  | 86.8695 | 111.9942 |  | 2191.2601 | 0.02712
|  | 86.9679 | 110.7938 |  | 2193.7417 | 0.028089
|  | 87.0663 | 110.0736 |  | 2196.2233 | 0.029139
|  | 87.1647 | 109.2733 |  | 2198.705 | 0.028468
|  | 87.2631 | 110.5537 |  | 2201.1866 | 0.024788
|  | 87.3614 | 112.6344 |  | 2203.6682 | 0.019213
|  | 87.4598 | 112.7944 |  | 2206.1498 | 0.016278
|  | 87.5582 | 111.8341 |  | 2208.6314 | 0.017524
|  | 87.6566 | 110.8738 |  | 2211.113 | 0.016929
|  | 87.755 | 109.8335 |  | 2213.5946 | 0.015267
|  | 87.8533 | 109.0333 |  | 2216.0762 | 0.025588
|  | 87.9517 | 110.9539 |  | 2218.5578 | 0.042974
|  | 88.0501 | 112.7144 |  | 2221.0394 | 0.055712
|  | 88.1485 | 112.5944 |  | 2223.521 | 0.057321
|  | 88.2469 | 111.554 |  | 2226.0026 | 0.047206
|  | 88.3452 | 110.5937 |  | 2228.4843 | 0.030574
|  | 88.4436 | 109.3934 |  | 2230.9659 | 0.016366
|  | 88.542 | 109.3534 |  | 2233.4475 | 0.015398
|  | 88.6404 | 111.1139 |  | 2235.9291 | 0.027667
|  | 88.7388 | 112.6344 |  | 2238.4107 | 0.046113
|  | 88.8371 | 112.5543 |  | 2240.8923 | 0.063168
|  | 88.9355 | 111.434 |  | 2243.3739 | 0.070409
|  | 89.0339 | 109.9135 |  | 2245.8555 | 0.063528
|  | 89.1323 | 109.1133 |  | 2248.3371 | 0.045424
|  | 89.2307 | 109.6335 |  | 2250.8187 | 0.026631
|  | 89.329 | 111.554 |  | 2253.3003 | 0.020769
|  | 89.4274 | 112.7544 |  | 2255.7819 | 0.019991
|  | 89.5258 | 111.9542 |  | 2258.2636 | 0.012722
|  | 89.6242 | 110.5937 |  | 2260.7452 | 0.015232
|  | 89.7226 | 109.7935 |  | 2263.2268 | 0.030632
|  | 89.8209 | 108.9132 |  | 2265.7084 | 0.041495
|  | 89.9193 | 110.0336 |  | 2268.19 | 0.044059
|  | 90.0177 | 112.1142 |  | 2270.6716 | 0.040438
|  | 90.1161 | 112.6744 |  | 2273.1532 | 0.033692
|  | 90.2145 | 111.7141 |  | 2275.6348 | 0.023674
|  | 90.3128 | 110.6738 |  | 2278.1164 | 0.0092831
|  | 90.4112 | 109.4734 |  | 2280.598 | 0.011865
|  | 90.5096 | 108.9933 |  | 2283.0796 | 0.028777
|  | 90.608 | 110.9138 |  | 2285.5612 | 0.038893
|  | 90.7064 | 112.8344 |  | 2288.0429 | 0.038538
|  | 90.8047 | 112.1942 |  | 2290.5245 | 0.029622
|  | 90.9031 | 111.314 |  | 2293.0061 | 0.022787
|  | 91.0015 | 110.1936 |  | 2295.4877 | 0.030881
|  | 91.0999 | 109.2333 |  | 2297.9693 | 0.043408
|  | 91.1983 | 109.1933 |  | 2300.4509 | 0.050644
|  | 91.2966 | 111.1139 |  | 2302.9325 | 0.050059
|  | 91.395 | 112.3943 |  | 2305.4141 | 0.042427
|  | 91.4934 | 112.2743 |  | 2307.8957 | 0.03068
|  | 91.5918 | 110.9939 |  | 2310.3773 | 0.019852
|  | 91.6902 | 109.7935 |  | 2312.8589 | 0.016189
|  | 91.7885 | 109.5134 |  | 2315.3405 | 0.01762
|  | 91.8869 | 109.5534 |  | 2317.8222 | 0.016938
|  | 91.9853 | 111.5941 |  | 2320.3038 | 0.013462
|  | 92.0837 | 112.8744 |  | 2322.7854 | 0.012321
|  | 92.1821 | 112.3943 |  | 2325.267 | 0.015829
|  | 92.2804 | 110.7938 |  | 2327.7486 | 0.017963
|  | 92.3788 | 109.6735 |  | 2330.2302 | 0.015248
|  | 92.4772 | 109.1933 |  | 2332.7118 | 0.0099458
|  | 92.5756 | 109.6735 |  | 2335.1934 | 0.011998
|  | 92.674 | 112.0742 |  | 2337.675 | 0.018366
|  | 92.7723 | 112.5543 |  | 2340.1566 | 0.021258
|  | 92.8707 | 111.354 |  | 2342.6382 | 0.022948
|  | 92.9691 | 110.5937 |  | 2345.1198 | 0.029727
|  | 93.0675 | 109.3134 |  | 2347.6015 | 0.038158
|  | 93.1659 | 108.9532 |  | 2350.0831 | 0.03986
|  | 93.2642 | 111.0339 |  | 2352.5647 | 0.032289
|  | 93.3626 | 112.7944 |  | 2355.0463 | 0.027262
|  | 93.461 | 112.3143 |  | 2357.5279 | 0.040811
|  | 93.5594 | 111.514 |  | 2360.0095 | 0.056315
|  | 93.6578 | 110.3937 |  | 2362.4911 | 0.059454
|  | 93.7561 | 109.2733 |  | 2364.9727 | 0.04755
|  | 93.8545 | 109.5134 |  | 2367.4543 | 0.027954
|  | 93.9529 | 111.1939 |  | 2369.9359 | 0.020828
|  | 94.0513 | 112.7144 |  | 2372.4175 | 0.029316
|  | 94.1497 | 112.4743 |  | 2374.8991 | 0.033259
|  | 94.248 | 111.514 |  | 2377.3808 | 0.03553
|  | 94.3464 | 110.1536 |  | 2379.8624 | 0.043651
|  | 94.4448 | 109.1133 |  | 2382.344 | 0.051831
|  | 94.5432 | 109.4734 |  | 2384.8256 | 0.052461
|  | 94.6416 | 111.7141 |  | 2387.3072 | 0.046775
|  | 94.7399 | 112.9945 |  | 2389.7888 | 0.042464
|  | 94.8383 | 112.2743 |  | 2392.2704 | 0.04091
|  | 94.9367 | 110.7538 |  | 2394.752 | 0.033863
|  | 95.0351 | 109.7935 |  | 2397.2336 | 0.017318
|  | 95.1335 | 109.0733 |  | 2399.7152 | 0.010412
|  | 95.2318 | 109.6335 |  | 2402.1968 | 0.029045
|  | 95.3302 | 111.8741 |  | 2404.6784 | 0.037964
|  | 95.4286 | 112.7544 |  | 2407.1601 | 0.032033
|  | 95.527 | 111.7141 |  | 2409.6417 | 0.01382
|  | 95.6254 | 110.6738 |  | 2412.1233 | 0.0098032
|  | 95.7237 | 109.6335 |  | 2414.6049 | 0.030046
|  | 95.8221 | 108.8332 |  | 2417.0865 | 0.041919
|  | 95.9205 | 110.3537 |  | 2419.5681 | 0.044152
|  | 96.0189 | 112.0342 |  | 2422.0497 | 0.03902
|  | 96.1173 | 112.4343 |  | 2424.5313 | 0.03064
|  | 96.2156 | 111.554 |  | 2427.0129 | 0.024102
|  | 96.314 | 110.5137 |  | 2429.4945 | 0.022622
|  | 96.4124 | 109.4734 |  | 2431.9761 | 0.02235
|  | 96.5108 | 109.3934 |  | 2434.4577 | 0.01822
|  | 96.6092 | 111.0739 |  | 2436.9394 | 0.010596
|  | 96.7075 | 112.3543 |  | 2439.421 | 0.0077086
|  | 96.8059 | 112.3943 |  | 2441.9026 | 0.010377
|  | 96.9043 | 111.394 |  | 2444.3842 | 0.0072093
|  | 97.0027 | 110.0336 |  | 2446.8658 | 0.003582
|  | 97.1011 | 109.3134 |  | 2449.3474 | 0.017161
|  | 97.1994 | 109.3134 |  | 2451.829 | 0.028044
|  | 97.2978 | 109.3134 |  | 2454.3106 | 0.031696
|  | 97.3962 | 109.3134 |  | 2456.7922 | 0.027926
|  | 97.4946 | 109.3134 |  | 2459.2738 | 0.020645
|  | 97.593 | 109.3134 |  | 2461.7554 | 0.014033
|  | 97.6913 | 109.3134 |  | 2464.237 | 0.010236
|  | 97.7897 | 109.3134 |  | 2466.7187 | 0.017369
|  | 97.8881 | 109.5534 |  | 2469.2003 | 0.031369
|  | 97.9865 | 111.9542 |  | 2471.6819 | 0.042264
|  | 98.0849 | 112.6744 |  | 2474.1635 | 0.044065
|  | 98.1832 | 111.9542 |  | 2476.6451 | 0.036024
|  | 98.2816 | 110.7538 |  | 2479.1267 | 0.024053
|  | 98.38 | 109.7935 |  | 2481.6083 | 0.020234
|  | 98.4784 | 108.9132 |  | 2484.0899 | 0.02655
|  | 98.5768 | 110.1936 |  | 2486.5715 | 0.035263
|  | 98.6751 | 112.4343 |  | 2489.0531 | 0.045575
|  | 98.7735 | 112.5143 |  | 2491.5347 | 0.054315
|  | 98.8719 | 111.8341 |  | 2494.0163 | 0.055453
|  | 98.9703 | 110.6338 |  | 2496.498 | 0.046088
|  | 99.0687 | 109.1933 |  | 2498.9796 | 0.029701
|  | 99.167 | 109.0333 |  | 2501.4612 | 0.016983
|  | 99.2654 | 110.5537 |  | 2503.9428 | 0.018234
|  | 99.3638 | 112.4743 |  | 2506.4244 | 0.023276
|  | 99.4622 | 112.7544 |  | 2508.906 | 0.032038
|  | 99.5606 | 111.474 |  | 2511.3876 | 0.046465
|  | 99.6589 | 110.1136 |  | 2513.8692 | 0.059161
|  | 99.7573 | 109.3134 |  | 2516.3508 | 0.06307
|  | 99.8557 | 109.3934 |  | 2518.8324 | 0.056497
|  | 99.9541 | 111.1539 |  | 2521.314 | 0.043038
|  | 100.0525 | 112.9145 |  | 2523.7956 | 0.028783
|  | 100.1508 | 112.2743 |  | 2526.2773 | 0.018179
|  | 100.2492 | 111.0739 |  | 2528.7589 | 0.012081
|  | 100.3476 | 110.1136 |  | 2531.2405 | 0.011843
|  | 100.446 | 109.3134 |  | 2533.7221 | 0.015645
|  | 100.5444 | 109.7135 |  | 2536.2037 | 0.016979
|  | 100.6427 | 111.7941 |  | 2538.6853 | 0.012378
|  | 100.7411 | 112.9145 |  | 2541.1669 | 0.0021573
|  | 100.8395 | 112.1142 |  | 2543.6485 | 0.012982
|  | 100.9379 | 111.0739 |  | 2546.1301 | 0.027177
|  | 101.0363 | 109.9536 |  | 2548.6117 | 0.036944
|  | 101.1346 | 108.8332 |  | 2551.0933 | 0.039436
|  | 101.233 | 110.0336 |  | 2553.5749 | 0.035414
|  | 101.3314 | 112.1142 |  | 2556.0566 | 0.030804
|  | 101.4298 | 112.7544 |  | 2558.5382 | 0.031911
|  | 101.5282 | 111.9542 |  | 2561.0198 | 0.035152
|  | 101.6265 | 110.5137 |  | 2563.5014 | 0.03582
|  | 101.7249 | 109.7135 |  | 2565.983 | 0.038005
|  | 101.8233 | 109.3534 |  | 2568.4646 | 0.045917
|  | 101.9217 | 110.7138 |  | 2570.9462 | 0.052557
|  | 102.0201 | 112.3943 |  | 2573.4278 | 0.049112
|  | 102.1184 | 112.6344 |  | 2575.9094 | 0.03342
|  | 102.2168 | 111.5941 |  | 2578.391 | 0.010764
|  | 102.3152 | 110.4737 |  | 2580.8726 | 0.0097892
|  | 102.4136 | 109.3534 |  | 2583.3542 | 0.019083
|  | 102.512 | 109.5134 |  | 2585.8359 | 0.016666
|  | 102.6103 | 111.0339 |  | 2588.3175 | 0.019674
|  | 102.7087 | 112.4743 |  | 2590.7991 | 0.035274
|  | 102.8071 | 112.2342 |  | 2593.2807 | 0.045178
|  | 102.9055 | 110.9539 |  | 2595.7623 | 0.041108
|  | 103.0039 | 110.3137 |  | 2598.2439 | 0.025479
|  | 103.1022 | 109.3534 |  | 2600.7255 | 0.017731
|  | 103.2006 | 109.5534 |  | 2603.2071 | 0.029889
|  | 103.299 | 111.7941 |  | 2605.6887 | 0.035231
|  | 103.3974 | 112.7544 |  | 2608.1703 | 0.027725
|  | 103.4958 | 111.8741 |  | 2610.6519 | 0.014375
|  | 103.5941 | 110.9939 |  | 2613.1335 | 0.015689
|  | 103.6925 | 109.7935 |  | 2615.6152 | 0.024028
|  | 103.7909 | 109.2333 |  | 2618.0968 | 0.026441
|  | 103.8893 | 110.2737 |  | 2620.5784 | 0.027064
|  | 103.9877 | 112.3543 |  | 2623.06 | 0.030235
|  | 104.086 | 112.7544 |  | 2625.5416 | 0.031824
|  | 104.1844 | 111.9542 |  | 2628.0232 | 0.027691
|  | 104.2828 | 110.5937 |  | 2630.5048 | 0.019379
|  | 104.3812 | 109.4734 |  | 2632.9864 | 0.01184
|  | 104.4796 | 108.8332 |  | 2635.468 | 0.0077566
|  | 104.5779 | 110.4337 |  | 2637.9496 | 0.010062
|  | 104.6763 | 112.1142 |  | 2640.4312 | 0.020115
|  | 104.7747 | 112.4743 |  | 2642.9128 | 0.030108
|  | 104.8731 | 111.5941 |  | 2645.3945 | 0.03415
|  | 104.9715 | 110.1536 |  | 2647.8761 | 0.03014
|  | 105.0698 | 109.5134 |  | 2650.3577 | 0.020312
|  | 105.1682 | 109.1133 |  | 2652.8393 | 0.0098083
|  | 105.2666 | 110.6338 |  | 2655.3209 | 0.0038875
|  | 105.365 | 112.5543 |  | 2657.8025 | 0.0043701
|  | 105.4634 | 112.3543 |  | 2660.2841 | 0.010978
|  | 105.5617 | 111.314 |  | 2662.7657 | 0.020954
|  | 105.6601 | 110.3537 |  | 2665.2473 | 0.030535
|  | 105.7585 | 109.1533 |  | 2667.7289 | 0.03693
|  | 105.8569 | 109.5134 |  | 2670.2105 | 0.038812
|  | 105.9553 | 111.5941 |  | 2672.6921 | 0.037418
|  | 106.0536 | 112.8744 |  | 2675.1738 | 0.037824
|  | 106.152 | 112.0742 |  | 2677.6554 | 0.043581
|  | 106.2504 | 111.0339 |  | 2680.137 | 0.048898
|  | 106.3488 | 109.8335 |  | 2682.6186 | 0.045349
|  | 106.4472 | 109.1133 |  | 2685.1002 | 0.030327
|  | 106.5455 | 110.3137 |  | 2687.5818 | 0.017656
|  | 106.6439 | 112.3943 |  | 2690.0634 | 0.035134
|  | 106.7423 | 112.7544 |  | 2692.545 | 0.052402
|  | 106.8407 | 111.7941 |  | 2695.0266 | 0.054839
|  | 106.9391 | 110.6738 |  | 2697.5082 | 0.0415
|  | 107.0374 | 109.3934 |  | 2699.9898 | 0.019802
|  | 107.1358 | 109.1133 |  | 2702.4714 | 0.011378
|  | 107.2342 | 110.3937 |  | 2704.9531 | 0.020963
|  | 107.3326 | 112.1542 |  | 2707.4347 | 0.021443
|  | 107.431 | 112.5543 |  | 2709.9163 | 0.014892
|  | 107.5293 | 111.6741 |  | 2712.3979 | 0.010812
|  | 107.6277 | 110.4737 |  | 2714.8795 | 0.010643
|  | 107.7261 | 109.6335 |  | 2717.3611 | 0.013479
|  | 107.8245 | 108.9132 |  | 2719.8427 | 0.026023
|  | 107.9229 | 110.7538 |  | 2722.3243 | 0.039847
|  | 108.0212 | 112.5143 |  | 2724.8059 | 0.045867
|  | 108.1196 | 112.6344 |  | 2727.2875 | 0.040563
|  | 108.218 | 111.274 |  | 2729.7691 | 0.026982
|  | 108.3164 | 110.4737 |  | 2732.2507 | 0.012881
|  | 108.4148 | 109.1933 |  | 2734.7324 | 0.0063868
|  | 108.5131 | 109.1533 |  | 2737.214 | 0.011569
|  | 108.6115 | 111.0739 |  | 2739.6956 | 0.026172
|  | 108.7099 | 112.6344 |  | 2742.1772 | 0.042677
|  | 108.8083 | 112.1142 |  | 2744.6588 | 0.052469
|  | 108.9067 | 110.9939 |  | 2747.1404 | 0.050452
|  | 109.005 | 110.0336 |  | 2749.622 | 0.037973
|  | 109.1034 | 108.9132 |  | 2752.1036 | 0.023143
|  | 109.2018 | 109.8735 |  | 2754.5852 | 0.017567
|  | 109.3002 | 111.7941 |  | 2757.0668 | 0.018143
|  | 109.3986 | 112.5143 |  | 2759.5484 | 0.016
|  | 109.4969 | 112.0342 |  | 2762.03 | 0.016384
|  | 109.5953 | 111.474 |  | 2764.5117 | 0.021482
|  | 109.6937 | 109.8335 |  | 2766.9933 | 0.024108
|  | 109.7921 | 109.9135 |  | 2769.4749 | 0.02261
|  | 109.8905 | 110.0736 |  | 2771.9565 | 0.020877
|  | 109.9888 | 110.2336 |  | 2774.4381 | 0.020473
|  | 110.0872 | 112.9545 |  | 2776.9197 | 0.01852
|  | 110.1856 | 112.8344 |  | 2779.4013 | 0.017462
|  | 110.284 | 111.554 |  | 2781.8829 | 0.022935
|  | 110.3824 | 110.5137 |  | 2784.3645 | 0.028274
|  | 110.4807 | 109.2333 |  | 2786.8461 | 0.025227
|  | 110.5791 | 109.1133 |  | 2789.3277 | 0.011976
|  | 110.6775 | 111.274 |  | 2791.8093 | 0.012468
|  | 110.7759 | 112.7944 |  | 2794.291 | 0.033405
|  | 110.8743 | 112.2342 |  | 2796.7726 | 0.047552
|  | 110.9726 | 111.0339 |  | 2799.2542 | 0.049239
|  | 111.071 | 110.2336 |  | 2801.7358 | 0.03822
|  | 111.1694 | 109.0333 |  | 2804.2174 | 0.019338
|  | 111.2678 | 109.3534 |  | 2806.699 | 0.0040624
|  | 111.3662 | 111.434 |  | 2809.1806 | 0.012664
|  | 111.4645 | 112.4743 |  | 2811.6622 | 0.011833
|  | 111.5629 | 112.1542 |  | 2814.1438 | 0.0032019
|  | 111.6613 | 111.314 |  | 2816.6254 | 0.020729
|  | 111.7597 | 110.3537 |  | 2819.107 | 0.037833
|  | 111.8581 | 109.3934 |  | 2821.5886 | 0.045479
|  | 111.9564 | 109.6735 |  | 2824.0703 | 0.040789
|  | 112.0548 | 111.9142 |  | 2826.5519 | 0.028094
|  | 112.1532 | 112.7944 |  | 2829.0335 | 0.020169
|  | 112.2516 | 111.9142 |  | 2831.5151 | 0.023774
|  | 112.35 | 110.7138 |  | 2833.9967 | 0.026908
|  | 112.4483 | 109.8335 |  | 2836.4783 | 0.029919
|  | 112.5467 | 109.1133 |  | 2838.9599 | 0.037694
|  | 112.6451 | 109.9936 |  | 2841.4415 | 0.043486
|  | 112.7435 | 112.2342 |  | 2843.9231 | 0.038827
|  | 112.8419 | 112.8744 |  | 2846.4047 | 0.022853
|  | 112.9402 | 111.6741 |  | 2848.8863 | 0.0080794
|  | 113.0386 | 110.4737 |  | 2851.3679 | 0.019748
|  | 113.137 | 109.4334 |  | 2853.8496 | 0.024599
|  | 113.2354 | 109.3134 |  | 2856.3312 | 0.021301
|  | 113.3338 | 111.2339 |  | 2858.8128 | 0.029778
|  | 113.4321 | 112.7544 |  | 2861.2944 | 0.047097
|  | 113.5305 | 112.3943 |  | 2863.776 | 0.056079
|  | 113.6289 | 111.274 |  | 2866.2576 | 0.05085
|  | 113.7273 | 110.2336 |  | 2868.7392 | 0.036006
|  | 113.8257 | 109.4334 |  | 2871.2208 | 0.023653
|  | 113.924 | 109.3934 |  | 2873.7024 | 0.01888
|  | 114.0224 | 111.314 |  | 2876.184 | 0.0095806
|  | 114.1208 | 112.9145 |  | 2878.6656 | 0.013307
|  | 114.2192 | 112.2743 |  | 2881.1472 | 0.037444
|  | 114.3176 | 111.0739 |  | 2883.6288 | 0.057099
|  | 114.4159 | 110.1936 |  | 2886.1105 | 0.065556
|  | 114.5143 | 109.0733 |  | 2888.5921 | 0.062591
|  | 114.6127 | 109.7135 |  | 2891.0737 | 0.052763
|  | 114.7111 | 111.9542 |  | 2893.5553 | 0.040781
|  | 114.8095 | 112.9945 |  | 2896.0369 | 0.028403
|  | 114.9078 | 112.1142 |  | 2898.5185 | 0.016682
|  | 115.0062 | 110.7538 |  | 2901.0001 | 0.0090635
|  | 115.1046 | 109.6335 |  | 2903.4817 | 0.007265
|  | 115.203 | 108.9132 |  | 2905.9633 | 0.007558
|  | 115.3014 | 109.6335 |  | 2908.4449 | 0.010679
|  | 115.3997 | 112.0342 |  | 2910.9265 | 0.016517
|  | 115.4981 | 112.7544 |  | 2913.4081 | 0.029405
|  | 115.5965 | 111.554 |  | 2915.8898 | 0.049694
|  | 115.6949 | 110.4337 |  | 2918.3714 | 0.067296
|  | 115.7933 | 109.6335 |  | 2920.853 | 0.069249
|  | 115.8916 | 108.9933 |  | 2923.3346 | 0.049252
|  | 115.99 | 110.6738 |  | 2925.8162 | 0.031545
|  | 116.0884 | 112.3543 |  | 2928.2978 | 0.069274
|  | 116.1868 | 112.5543 |  | 2930.7794 | 0.10814
|  | 116.2852 | 111.434 |  | 2933.261 | 0.11864
|  | 116.3835 | 110.4737 |  | 2935.7426 | 0.094283
|  | 116.4819 | 109.3534 |  | 2938.2242 | 0.045838
|  | 116.5803 | 109.3134 |  | 2940.7058 | 0.011656
|  | 116.6787 | 111.0739 |  | 2943.1874 | 0.038437
|  | 116.7771 | 112.5143 |  | 2945.6691 | 0.035562
|  | 116.8754 | 112.5143 |  | 2948.1507 | 0.0045288
|  | 116.9738 | 112.5143 |  | 2950.6323 | 0.043572
|  | 117.0722 | 112.5143 |  | 2953.1139 | 0.084302
|  | 117.1706 | 112.5143 |  | 2955.5955 | 0.10974
|  | 117.269 | 112.5143 |  | 2958.0771 | 0.12113
|  | 117.3673 | 112.5143 |  | 2960.5587 | 0.11962
|  | 117.4657 | 112.5143 |  | 2963.0403 | 0.097278
|  | 117.5641 | 112.0342 |  | 2965.5219 | 0.0497
|  | 117.6625 | 110.6738 |  | 2968.0035 | 0.042264
|  | 117.7609 | 109.6335 |  | 2970.4851 | 0.1043
|  | 117.8592 | 108.8332 |  | 2972.9667 | 0.13991
|  | 117.9576 | 109.6335 |  | 2975.4484 | 0.13267
|  | 118.056 | 112.0342 |  | 2977.93 | 0.10475
|  | 118.1544 | 112.5543 |  | 2980.4116 | 0.098416
|  | 118.2528 | 111.8341 |  | 2982.8932 | 0.085981
|  | 118.3511 | 110.7138 |  | 2985.3748 | 0.029545
|  | 118.4495 | 109.8335 |  | 2987.8564 | 0.1666
|  | 118.5479 | 109.0333 |  | 2990.338 | 0.37545
|  | 118.6463 | 110.5537 |  | 2992.8196 | 0.56558
|  | 118.7447 | 112.4743 |  | 2995.3012 | 0.66603
|  | 118.843 | 112.5944 |  | 2997.7828 | 0.63418
|  | 118.9414 | 111.6341 |  | 3000.2644 | 0.47535
|  | 119.0398 | 110.4337 |  | 3002.746 | 0.24298
|  | 119.1382 | 109.2333 |  | 3005.2277 | 0.042553
|  | 119.2366 | 109.1933 |  | 3007.7093 | 0.15257
|  | 119.3349 | 110.9539 |  | 3010.1909 | 0.19796
|  | 119.4333 | 112.3143 |  | 3012.6725 | 0.15711
|  | 119.5317 | 112.4743 |  | 3015.1541 | 0.11066
|  | 119.6301 | 111.354 |  | 3017.6357 | 0.12599
|  | 119.7285 | 109.9936 |  | 3020.1173 | 0.11937
|  | 119.8268 | 109.5934 |  | 3022.5989 | 0.050594
|  | 119.9252 | 109.7535 |  | 3025.0805 | 0.097428
|  | 120.0236 | 111.6741 |  | 3027.5621 | 0.21946
|  | 120.122 | 113.0345 |  | 3030.0437 | 0.29307
|  | 120.2204 | 112.0742 |  | 3032.5253 | 0.29079
|  | 120.3187 | 110.9539 |  | 3035.007 | 0.22173
|  | 120.4171 | 110.0736 |  | 3037.4886 | 0.12313
|  | 120.5155 | 109.3134 |  | 3039.9702 | 0.048016
|  | 120.6139 | 110.1136 |  | 3042.4518 | 0.044882
|  | 120.7123 | 112.3543 |  | 3044.9334 | 0.043923
|  | 120.8106 | 113.1545 |  | 3047.415 | 0.039501
|  | 120.909 | 111.9542 |  | 3049.8966 | 0.058363
|  | 121.0074 | 110.9138 |  | 3052.3782 | 0.068684
|  | 121.1058 | 109.8735 |  | 3054.8598 | 0.054151
|  | 121.2042 | 108.9132 |  | 3057.3414 | 0.021806
|  | 121.3025 | 110.6738 |  | 3059.823 | 0.021816
|  | 121.4009 | 112.5143 |  | 3062.3046 | 0.052324
|  | 121.4993 | 112.5543 |  | 3064.7863 | 0.073572
|  | 121.5977 | 111.9142 |  | 3067.2679 | 0.082676
|  | 121.6961 | 110.4737 |  | 3069.7495 | 0.075352
|  | 121.7944 | 109.2733 |  | 3072.2311 | 0.052148
|  | 121.8928 | 109.2333 |  | 3074.7127 | 0.037802
|  | 121.9912 | 111.0739 |  | 3077.1943 | 0.057652
|  | 122.0896 | 112.6744 |  | 3079.6759 | 0.068978
|  | 122.188 | 112.6344 |  | 3082.1575 | 0.050388
|  | 122.2863 | 111.354 |  | 3084.6391 | 0.013219
|  | 122.3847 | 110.1536 |  | 3087.1207 | 0.053453
|  | 122.4831 | 109.2733 |  | 3089.6023 | 0.090474
|  | 122.5815 | 109.5134 |  | 3092.0839 | 0.094089
|  | 122.6799 | 111.434 |  | 3094.5656 | 0.064856
|  | 122.7782 | 112.8744 |  | 3097.0472 | 0.042114
|  | 122.8766 | 112.4743 |  | 3099.5288 | 0.075918
|  | 122.975 | 111.0339 |  | 3102.0104 | 0.10414
|  | 123.0734 | 109.9135 |  | 3104.492 | 0.10219
|  | 123.1718 | 109.1133 |  | 3106.9736 | 0.075029
|  | 123.2701 | 109.5934 |  | 3109.4552 | 0.0422
|  | 123.3685 | 111.8341 |  | 3111.9368 | 0.025906
|  | 123.4669 | 112.6344 |  | 3114.4184 | 0.018587
|  | 123.5653 | 111.6741 |  | 3116.9 | 0.0010407
|  | 123.6637 | 110.7138 |  | 3119.3816 | 0.022611
|  | 123.762 | 109.6735 |  | 3121.8632 | 0.038746
|  | 123.8604 | 109.0333 |  | 3124.3449 | 0.040657
|  | 123.9588 | 110.5537 |  | 3126.8265 | 0.036454
|  | 124.0572 | 112.4743 |  | 3129.3081 | 0.045121
|  | 124.1556 | 112.6344 |  | 3131.7897 | 0.060643
|  | 124.2539 | 111.7541 |  | 3134.2713 | 0.06793
|  | 124.3523 | 110.5537 |  | 3136.7529 | 0.063998
|  | 124.4507 | 109.5134 |  | 3139.2345 | 0.052846
|  | 124.5491 | 109.1133 |  | 3141.7161 | 0.037483
|  | 124.6475 | 110.7938 |  | 3144.1977 | 0.019638
|  | 124.7458 | 112.4743 |  | 3146.6793 | 0.021151
|  | 124.8442 | 112.5543 |  | 3149.1609 | 0.040962
|  | 124.9426 | 111.514 |  | 3151.6425 | 0.051692
|  | 125.041 | 110.3137 |  | 3154.1242 | 0.046088
|  | 125.1394 | 109.5134 |  | 3156.6058 | 0.028493
|  | 125.2377 | 109.5534 |  | 3159.0874 | 0.022234
|  | 125.3361 | 111.2339 |  | 3161.569 | 0.03624
|  | 125.4345 | 112.7544 |  | 3164.0506 | 0.042867
|  | 125.5329 | 112.2743 |  | 3166.5322 | 0.03674
|  | 125.6313 | 111.0739 |  | 3169.0138 | 0.023736
|  | 125.7296 | 109.9536 |  | 3171.4954 | 0.014843
|  | 125.828 | 109.3134 |  | 3173.977 | 0.01852
|  | 125.9264 | 109.7535 |  | 3176.4586 | 0.028684
|  | 126.0248 | 111.9942 |  | 3178.9402 | 0.039744
|  | 126.1232 | 112.7944 |  | 3181.4218 | 0.046679
|  | 126.2215 | 111.7541 |  | 3183.9035 | 0.047146
|  | 126.3199 | 110.7938 |  | 3186.3851 | 0.043949
|  | 126.4183 | 109.6735 |  | 3188.8667 | 0.039839
|  | 126.5167 | 108.7932 |  | 3191.3483 | 0.031033
|  | 126.6151 | 110.1536 |  | 3193.8299 | 0.014447
|  | 126.7134 | 112.2342 |  | 3196.3115 | 0.020981
|  | 126.8118 | 112.5143 |  | 3198.7931 | 0.045915
|  | 126.9102 | 111.8741 |  | 3201.2747 | 0.059928
|  | 127.0086 | 110.7538 |  | 3203.7563 | 0.055873
|  | 127.107 | 109.3934 |  | 3206.2379 | 0.038943
|  | 127.2053 | 109.1933 |  | 3208.7195 | 0.032142
|  | 127.3037 | 111.1139 |  | 3211.2011 | 0.043079
|  | 127.4021 | 112.5143 |  | 3213.6828 | 0.044525
|  | 127.5005 | 112.4743 |  | 3216.1644 | 0.02739
|  | 127.5989 | 111.1139 |  | 3218.646 | 0.010111
|  | 127.6972 | 110.1536 |  | 3221.1276 | 0.03895
|  | 127.7956 | 109.4334 |  | 3223.6092 | 0.060618
|  | 127.894 | 109.5134 |  | 3226.0908 | 0.066247
|  | 127.9924 | 111.434 |  | 3228.5724 | 0.058367
|  | 128.0908 | 112.7944 |  | 3231.054 | 0.045941
|  | 128.1891 | 112.3543 |  | 3233.5356 | 0.040183
|  | 128.2875 | 110.9939 |  | 3236.0172 | 0.047533
|  | 128.3859 | 110.1936 |  | 3238.4988 | 0.062623
|  | 128.4843 | 109.3134 |  | 3240.9804 | 0.075152
|  | 128.5827 | 109.3134 |  | 3243.4621 | 0.076356
|  | 128.681 | 109.3134 |  | 3245.9437 | 0.062715
|  | 128.7794 | 109.3134 |  | 3248.4253 | 0.038309
|  | 128.8778 | 109.3134 |  | 3250.9069 | 0.016567
|  | 128.9762 | 109.3134 |  | 3253.3885 | 0.019948
|  | 129.0746 | 109.3134 |  | 3255.8701 | 0.023872
|  | 129.1729 | 109.3134 |  | 3258.3517 | 0.015301
|  | 129.2713 | 109.3134 |  | 3260.8333 | 0.0082188
|  | 129.3697 | 112.1142 |  | 3263.3149 | 0.028182
|  | 129.4681 | 112.3943 |  | 3265.7965 | 0.048192
|  | 129.5665 | 111.6741 |  | 3268.2781 | 0.062375
|  | 129.6648 | 110.5537 |  | 3270.7597 | 0.067362
|  | 129.7632 | 109.5934 |  | 3273.2414 | 0.06119
|  | 129.8616 | 109.2333 |  | 3275.723 | 0.045233
|  | 129.96 | 110.8338 |  | 3278.2046 | 0.025177
|  | 130.0584 | 112.4343 |  | 3280.6862 | 0.0089394
|  | 130.1567 | 112.4343 |  | 3283.1678 | 0.0051944
|  | 130.2551 | 112.4343 |  | 3285.6494 | 0.015615
|  | 130.3535 | 110.3137 |  | 3288.131 | 0.031131
|  | 130.4519 | 109.4334 |  | 3290.6126 | 0.043296
|  | 130.5503 | 109.6735 |  | 3293.0942 | 0.044083
|  | 130.6486 | 111.434 |  | 3295.5758 | 0.032251
|  | 130.747 | 112.7944 |  | 3298.0574 | 0.019443
|  | 130.8454 | 112.3143 |  | 3300.539 | 0.026488
|  | 130.9438 | 111.0339 |  | 3303.0207 | 0.036917
|  | 131.0422 | 110.1536 |  | 3305.5023 | 0.044171
|  | 131.1405 | 109.1933 |  | 3307.9839 | 0.055155
|  | 131.2389 | 109.2733 |  | 3310.4655 | 0.064875
|  | 131.3373 | 111.354 |  | 3312.9471 | 0.059847
|  | 131.4357 | 112.7144 |  | 3315.4287 | 0.037293
|  | 131.5341 | 111.6741 |  | 3317.9103 | 0.039258
|  | 131.6324 | 110.7938 |  | 3320.3919 | 0.080288
|  | 131.7308 | 109.9135 |  | 3322.8735 | 0.10971
|  | 131.8292 | 108.8732 |  | 3325.3551 | 0.10948
|  | 131.9276 | 109.9135 |  | 3327.8367 | 0.079065
|  | 132.026 | 111.9942 |  | 3330.3183 | 0.033122
|  | 132.1243 | 112.7944 |  | 3332.8 | 0.021826
|  | 132.2227 | 111.9142 |  | 3335.2816 | 0.042909
|  | 132.3211 | 110.3937 |  | 3337.7632 | 0.040338
|  | 132.4195 | 109.5134 |  | 3340.2448 | 0.019373
|  | 132.5179 | 109.1533 |  | 3342.7264 | 0.0094047
|  | 132.6162 | 110.6738 |  | 3345.208 | 0.018787
|  | 132.7146 | 112.1942 |  | 3347.6896 | 0.0099685
|  | 132.813 | 112.5143 |  | 3350.1712 | 0.014485
|  | 132.9114 | 111.394 |  | 3352.6528 | 0.036409
|  | 133.0098 | 110.3537 |  | 3355.1344 | 0.041775
|  | 133.1081 | 109.3134 |  | 3357.616 | 0.023842
|  | 133.2065 | 109.7935 |  | 3360.0976 | 0.011219
|  | 133.3049 | 111.554 |  | 3362.5793 | 0.042937
|  | 133.4033 | 112.7544 |  | 3365.0609 | 0.05337
|  | 133.5017 | 112.3943 |  | 3367.5425 | 0.033349
|  | 133.6 | 111.0339 |  | 3370.0241 | 0.023728
|  | 133.6984 | 110.2336 |  | 3372.5057 | 0.06947
|  | 133.7968 | 109.2733 |  | 3374.9873 | 0.094832
|  | 133.8952 | 109.4334 |  | 3377.4689 | 0.084017
|  | 133.9936 | 111.9942 |  | 3379.9505 | 0.085399
|  | 134.0919 | 112.7944 |  | 3382.4321 | 0.16713
|  | 134.1903 | 111.7541 |  | 3384.9137 | 0.26622
|  | 134.2887 | 110.9138 |  | 3387.3953 | 0.3299
|  | 134.3871 | 109.5534 |  | 3389.8769 | 0.33167
|  | 134.4855 | 109.0733 |  | 3392.3586 | 0.27028
|  | 134.5838 | 110.1936 |  | 3394.8402 | 0.16768
|  | 134.6822 | 112.2743 |  | 3397.3218 | 0.059253
|  | 134.7806 | 112.6744 |  | 3399.8034 | 0.02414
|  | 134.879 | 111.8741 |  | 3402.285 | 0.057704
|  | 134.9774 | 110.5937 |  | 3404.7666 | 0.049279
|  | 135.0757 | 109.5534 |  | 3407.2482 | 0.041153
|  | 135.1741 | 109.0733 |  | 3409.7298 | 0.076339
|  | 135.2725 | 110.3937 |  | 3412.2114 | 0.10815
|  | 135.3709 | 112.0742 |  | 3414.693 | 0.11589
|  | 135.4693 | 112.4343 |  | 3417.1746 | 0.1019
|  | 135.5676 | 111.7141 |  | 3419.6562 | 0.082056
|  | 135.666 | 110.5137 |  | 3422.1379 | 0.075172
|  | 135.7644 | 109.4734 |  | 3424.6195 | 0.080315
|  | 135.8628 | 109.0733 |  | 3427.1011 | 0.081456
|  | 135.9612 | 110.5137 |  | 3429.5827 | 0.070053
|  | 136.0595 | 112.4343 |  | 3432.0643 | 0.046305
|  | 136.1579 | 112.4743 |  | 3434.5459 | 0.016751
|  | 136.2563 | 111.474 |  | 3437.0275 | 0.020447
|  | 136.3547 | 110.4337 |  | 3439.5091 | 0.043249
|  | 136.4531 | 109.1533 |  | 3441.9907 | 0.05665
|  | 136.5514 | 109.1933 |  | 3444.4723 | 0.064329
|  | 136.6498 | 110.9539 |  | 3446.9539 | 0.075669
|  | 136.7482 | 112.5543 |  | 3449.4355 | 0.092507
|  | 136.8466 | 112.4743 |  | 3451.9172 | 0.10632
|  | 136.945 | 111.1139 |  | 3454.3988 | 0.10857
|  | 137.0433 | 110.0736 |  | 3456.8804 | 0.096963
|  | 137.1417 | 108.8732 |  | 3459.362 | 0.078012
|  | 137.2401 | 108.9532 |  | 3461.8436 | 0.069411
|  | 137.3385 | 111.8341 |  | 3464.3252 | 0.085076
|  | 137.4369 | 111.7141 |  | 3466.8068 | 0.11182
|  | 137.5352 | 111.9542 |  | 3469.2884 | 0.1343
|  | 137.6336 | 111.9542 |  | 3471.77 | 0.1455
|  | 137.732 | 111.9542 |  | 3474.2516 | 0.14288
|  | 137.8304 | 111.9542 |  | 3476.7332 | 0.12577
|  | 137.9288 | 111.9542 |  | 3479.2148 | 0.094351
|  | 138.0271 | 111.9542 |  | 3481.6965 | 0.050706
|  | 138.1255 | 111.9542 |  | 3484.1781 | 0.015905
|  | 138.2239 | 112.6744 |  | 3486.6597 | 0.060946
|  | 138.3223 | 112.6744 |  | 3489.1413 | 0.10435
|  | 138.4207 | 112.6744 |  | 3491.6229 | 0.12855
|  | 138.519 | 112.6744 |  | 3494.1045 | 0.12792
|  | 138.6174 | 112.6744 |  | 3496.5861 | 0.10604
|  | 138.7158 | 112.6744 |  | 3499.0677 | 0.080781
|  | 138.8142 | 112.6744
|  | 138.9126 | 112.6744
|  | 139.0109 | 112.6744
|  | 139.1093 | 112.6744
|  | 139.2077 | 109.3134
|  | 139.3061 | 109.6735
|  | 139.4045 | 111.5941
|  | 139.5028 | 112.6344
|  | 139.6012 | 112.1942
|  | 139.6996 | 111.1539
|  | 139.798 | 109.7135
|  | 139.8964 | 109.1533
|  | 139.9947 | 109.7135
|  | 140.0931 | 111.6341
|  | 140.1915 | 112.9145
|  | 140.2899 | 111.9542
|  | 140.3883 | 110.6738
|  | 140.4866 | 109.8735
|  | 140.585 | 108.9933
|  | 140.6834 | 110.2737
|  | 140.7818 | 112.3543
|  | 140.8802 | 112.7144
|  | 140.9785 | 111.514
|  | 141.0769 | 110.3937
|  | 141.1753 | 109.5934
|  | 141.2737 | 108.7932
|  | 141.3721 | 110.7138
|  | 141.4704 | 112.6344
|  | 141.5688 | 112.4743
|  | 141.6672 | 111.354
|  | 141.7656 | 110.3137
|  | 141.864 | 109.2733
|  | 141.9623 | 109.2333
|  | 142.0607 | 111.1539
|  | 142.1591 | 112.4743
|  | 142.2575 | 112.2342
|  | 142.3559 | 111.1139
|  | 142.4542 | 109.9936
|  | 142.5526 | 109.2733
|  | 142.651 | 109.2733
|  | 142.7494 | 109.2733
|  | 142.8478 | 109.2733
|  | 142.9461 | 109.2733
|  | 143.0445 | 109.2733
|  | 143.1429 | 109.8335
|  | 143.2413 | 109.0333
|  | 143.3397 | 109.9936
|  | 143.438 | 112.2342
|  | 143.5364 | 112.5944
|  | 143.6348 | 111.554
|  | 143.7332 | 110.6738
|  | 143.8316 | 109.5534
|  | 143.9299 | 109.0333
|  | 144.0283 | 110.9539
|  | 144.1267 | 112.6744
|  | 144.2251 | 112.6744
|  | 144.3235 | 112.6744
|  | 144.4218 | 112.6744
|  | 144.5202 | 112.6744
|  | 144.6186 | 112.6744
|  | 144.717 | 112.6744
|  | 144.8154 | 112.6744
|  | 144.9137 | 112.6744
|  | 145.0121 | 112.6744
|  | 145.1105 | 110.1536
|  | 145.2089 | 109.1133
|  | 145.3073 | 109.3934
|  | 145.4056 | 111.474
|  | 145.504 | 112.8344
|  | 145.6024 | 112.3143
|  | 145.7008 | 110.8738
|  | 145.7992 | 109.9135
|  | 145.8975 | 108.8732
|  | 145.9959 | 109.7535
|  | 146.0943 | 111.9142
|  | 146.1927 | 112.6344
|  | 146.2911 | 111.5941
|  | 146.3894 | 110.7138
|  | 146.4878 | 109.8335
|  | 146.5862 | 109.3934
|  | 146.6846 | 110.9138
|  | 146.783 | 112.1942
|  | 146.8813 | 112.5143
|  | 146.9797 | 111.474
|  | 147.0781 | 110.4737
|  | 147.1765 | 109.5134
|  | 147.2749 | 109.1933
|  | 147.3732 | 111.1139
|  | 147.4716 | 112.6344
|  | 147.57 | 112.5543
|  | 147.6684 | 111.434
|  | 147.7668 | 110.0736
|  | 147.8651 | 109.1933
|  | 147.9635 | 109.3134
|  | 148.0619 | 111.474
|  | 148.1603 | 112.7544
|  | 148.2587 | 112.0342
|  | 148.357 | 110.9939
|  | 148.4554 | 110.1136
|  | 148.5538 | 109.3134
|  | 148.6522 | 109.6735
|  | 148.7506 | 111.7541
|  | 148.8489 | 112.5543
|  | 148.9473 | 111.9942
|  | 149.0457 | 110.7938
|  | 149.1441 | 109.7535
|  | 149.2425 | 108.7132
|  | 149.3408 | 109.9936
|  | 149.4392 | 112.2342
|  | 149.5376 | 112.6744
|  | 149.636 | 111.8741
|  | 149.7344 | 110.5137
|  | 149.8327 | 109.1533
|  | 149.9311 | 108.8332
|  | 150.0295 | 110.5137
|  | 150.1279 | 112.4343
|  | 150.2263 | 112.5944
|  | 150.3246 | 111.314
|  | 150.423 | 109.9536
|  | 150.5214 | 109.2333
|  | 150.6198 | 109.5534
|  | 150.7182 | 111.6341
|  | 150.8165 | 112.9945
|  | 150.9149 | 112.2743
|  | 151.0133 | 111.1539
|  | 151.1117 | 110.1136
|  | 151.2101 | 109.3134
|  | 151.3084 | 109.6735
|  | 151.4068 | 111.7541
|  | 151.5052 | 112.6344
|  | 151.6036 | 111.7541
|  | 151.702 | 110.7938
|  | 151.8003 | 109.7535
|  | 151.8987 | 108.8732
|  | 151.9971 | 110.1936
|  | 152.0955 | 112.2743
|  | 152.1939 | 112.2743
|  | 152.2922 | 112.2743
|  | 152.3906 | 112.2743
|  | 152.489 | 112.2743
|  | 152.5874 | 112.2743
|  | 152.6858 | 112.2743
|  | 152.7841 | 112.2743
|  | 152.8825 | 112.2743
|  | 152.9809 | 111.1539
|  | 153.0793 | 110.1136
|  | 153.1777 | 109.0733
|  | 153.276 | 109.5934
|  | 153.3744 | 111.8341
|  | 153.4728 | 112.9545
|  | 153.5712 | 112.2342
|  | 153.6696 | 110.8738
|  | 153.7679 | 109.9936
|  | 153.8663 | 109.1133
|  | 153.9647 | 109.5534
|  | 154.0631 | 111.7941
|  | 154.1615 | 112.8344
|  | 154.2598 | 111.8741
|  | 154.3582 | 110.9138
|  | 154.4566 | 109.6335
|  | 154.555 | 108.6732
|  | 154.6534 | 110.3537
|  | 154.7517 | 112.5944
|  | 154.8501 | 112.6744
|  | 154.9485 | 111.8741
|  | 155.0469 | 110.6738
|  | 155.1453 | 109.4734
|  | 155.2436 | 108.6732
|  | 155.342 | 110.3537
|  | 155.4404 | 112.2743
|  | 155.5388 | 112.5143
|  | 155.6372 | 111.394
|  | 155.7355 | 110.1936
|  | 155.8339 | 109.4734
|  | 155.9323 | 109.0733
|  | 156.0307 | 110.5137
|  | 156.1291 | 112.5944
|  | 156.2274 | 112.4343
|  | 156.3258 | 111.1539
|  | 156.4242 | 110.2737
|  | 156.5226 | 109.1533
|  | 156.621 | 109.5134
|  | 156.7193 | 111.9142
|  | 156.8177 | 112.8744
|  | 156.9161 | 111.9142
|  | 157.0145 | 111.0339
|  | 157.1129 | 109.8335
|  | 157.2112 | 109.1133
|  | 157.3096 | 110.0736
|  | 157.408 | 112.3143
|  | 157.5064 | 112.8344
|  | 157.6048 | 112.0342
|  | 157.7031 | 110.7538
|  | 157.8015 | 109.3934
|  | 157.8999 | 109.1933
|  | 157.9983 | 110.3937
|  | 158.0967 | 112.3143
|  | 158.195 | 112.6744
|  | 158.2934 | 111.474
|  | 158.3918 | 110.3537
|  | 158.4902 | 109.4734
|  | 158.5886 | 109.2333
|  | 158.6869 | 110.9939
|  | 158.7853 | 112.9145
|  | 158.8837 | 112.8344
|  | 158.9821 | 111.2339
|  | 159.0805 | 110.2737
|  | 159.1788 | 109.1533
|  | 159.2772 | 109.2733
|  | 159.3756 | 111.354
|  | 159.474 | 112.7944
|  | 159.5724 | 112.1542
|  | 159.6707 | 110.9539
|  | 159.7691 | 109.9135
|  | 159.8675 | 108.8732
|  | 159.9659 | 109.9135
|  | 160.0643 | 111.8341
|  | 160.1626 | 112.5543
|  | 160.261 | 111.8341
|  | 160.3594 | 110.7138
|  | 160.4578 | 109.5934
|  | 160.5562 | 109.2333
|  | 160.6545 | 110.3537
|  | 160.7529 | 112.2743
|  | 160.8513 | 112.7144
|  | 160.9497 | 111.7541
|  | 161.0481 | 110.3937
|  | 161.1464 | 109.5134
|  | 161.2448 | 109.1133
|  | 161.3432 | 110.9539
|  | 161.4416 | 112.7144
|  | 161.54 | 112.5143
|  | 161.6383 | 111.0739
|  | 161.7367 | 110.1936
|  | 161.8351 | 109.2733
|  | 161.9335 | 109.2733
|  | 162.0319 | 109.2733
|  | 162.1302 | 109.2733
|  | 162.2286 | 109.2733
|  | 162.327 | 109.2733
|  | 162.4254 | 109.2733
|  | 162.5238 | 109.2733
|  | 162.6221 | 109.2733
|  | 162.7205 | 109.2733
|  | 162.8189 | 112.8344
|  | 162.9173 | 112.0342
|  | 163.0157 | 110.7538
|  | 163.114 | 109.5534
|  | 163.2124 | 109.1133
|  | 163.3108 | 110.3137
|  | 163.4092 | 112.3943
|  | 163.5076 | 112.4743
|  | 163.6059 | 111.6741
|  | 163.7043 | 110.4737
|  | 163.8027 | 109.3934
|  | 163.9011 | 109.3534
|  | 163.9995 | 110.8738
|  | 164.0978 | 111.6741
|  | 164.1962 | 112.5543
|  | 164.2946 | 111.6741
|  | 164.393 | 110.3937
|  | 164.4914 | 109.9135
|  | 164.5897 | 109.4334
|  | 164.6881 | 109.5534
|  | 164.7865 | 112.5543
|  | 164.8849 | 112.7944
|  | 164.9833 | 111.7541
|  | 165.0816 | 111.474
|  | 165.18 | 110.0336
|  | 165.2784 | 109.3134
|  | 165.3768 | 111.394
|  | 165.4752 | 111.8741
|  | 165.5735 | 112.9145
|  | 165.6719 | 112.1942
|  | 165.7703 | 110.5937
|  | 165.8687 | 109.7135
|  | 165.9671 | 109.2733
|  | 166.0654 | 110.0736
|  | 166.1638 | 112.3143
|  | 166.2622 | 112.8344
|  | 166.3606 | 111.7141
|  | 166.459 | 110.4337
|  | 166.5573 | 109.3134
|  | 166.6557 | 109.1533
|  | 166.7541 | 111.0739
|  | 166.8525 | 112.4343
|  | 166.9509 | 112.5543
|  | 167.0492 | 111.514
|  | 167.1476 | 110.3137
|  | 167.246 | 109.0333
|  | 167.3444 | 109.1933
|  | 167.4428 | 111.1139
|  | 167.5411 | 112.5543
|  | 167.6395 | 112.3943
|  | 167.7379 | 111.2339
|  | 167.8363 | 109.6335
|  | 167.9347 | 108.9933
|  | 168.033 | 109.6335
|  | 168.1314 | 111.7141
|  | 168.2298 | 113.0745
|  | 168.3282 | 112.2743
|  | 168.4266 | 110.7538
|  | 168.5249 | 109.8735
|  | 168.6233 | 109.2333
|  | 168.7217 | 109.6735
|  | 168.8201 | 112.2342
|  | 168.9185 | 112.4743
|  | 169.0168 | 111.7541
|  | 169.1152 | 110.9539
|  | 169.2136 | 109.9936
|  | 169.312 | 108.9532
|  | 169.4104 | 110.8738
|  | 169.5087 | 112.3943
|  | 169.6071 | 112.2342
|  | 169.7055 | 111.5941
|  | 169.8039 | 110.5537
|  | 169.9023 | 109.2733
|  | 170.0006 | 109.2733
|  | 170.099 | 109.2733
|  | 170.1974 | 109.2733
|  | 170.2958 | 109.2733
|  | 170.3942 | 109.2733
|  | 170.4925 | 109.2733
|  | 170.5909 | 109.2733
|  | 170.6893 | 109.9135
|  | 170.7877 | 111.9942
|  | 170.8861 | 112.9545
|  | 170.9844 | 111.9942
|  | 171.0828 | 110.8738
|  | 171.1812 | 109.9135
|  | 171.2796 | 109.2733
|  | 171.378 | 109.9936
|  | 171.4763 | 112.0742
|  | 171.5747 | 112.5143
|  | 171.6731 | 111.6741
|  | 171.7715 | 110.7938
|  | 171.8699 | 109.7535
|  | 171.9682 | 108.8732
|  | 172.0666 | 110.3137
|  | 172.165 | 112.5543
|  | 172.2634 | 112.4343
|  | 172.3618 | 111.6341
|  | 172.4601 | 110.4337
|  | 172.5585 | 109.3134
|  | 172.6569 | 109.4334
|  | 172.7553 | 111.1939
|  | 172.8537 | 112.5543
|  | 172.952 | 112.3943
|  | 173.0504 | 110.8738
|  | 173.1488 | 109.9135
|  | 173.2472 | 109.1933
|  | 173.3456 | 109.9135
|  | 173.4439 | 111.8341
|  | 173.5423 | 112.5543
|  | 173.6407 | 111.9142
|  | 173.7391 | 110.6338
|  | 173.8375 | 109.5934
|  | 173.9358 | 108.9532
|  | 174.0342 | 109.7535
|  | 174.1326 | 111.8341
|  | 174.231 | 112.6344
|  | 174.3294 | 111.5941
|  | 174.4277 | 110.7938
|  | 174.5261 | 109.8335
|  | 174.6245 | 109.1133
|  | 174.7229 | 110.5537
|  | 174.8213 | 112.4743
|  | 174.9196 | 112.7144
|  | 175.018 | 111.514
|  | 175.1164 | 110.1536
|  | 175.2148 | 109.3534
|  | 175.3132 | 109.1133
|  | 175.4115 | 110.7938
|  | 175.5099 | 112.7144
|  | 175.6083 | 112.3143
|  | 175.7067 | 111.1939
|  | 175.8051 | 109.9936
|  | 175.9034 | 109.2733
|  | 176.0018 | 109.3934
|  | 176.1002 | 111.0739
|  | 176.1986 | 112.5944
|  | 176.297 | 112.3543
|  | 176.3953 | 110.9138
|  | 176.4937 | 110.1936
|  | 176.5921 | 108.9933
|  | 176.6905 | 109.6335
|  | 176.7889 | 112.1942
|  | 176.8872 | 112.9945
|  | 176.9856 | 111.9542
|  | 177.084 | 110.9939
|  | 177.1824 | 109.6335
|  | 177.2808 | 109.0733
|  | 177.3791 | 110.4337
|  | 177.4775 | 112.3543
|  | 177.5759 | 112.5543
|  | 177.6743 | 111.6741
|  | 177.7727 | 110.6338
|  | 177.871 | 109.2733
|  | 177.9694 | 108.9933
|  | 178.0678 | 110.6738
|  | 178.1662 | 112.1142
|  | 178.2646 | 112.3543
|  | 178.3629 | 111.394
|  | 178.4613 | 110.2737
|  | 178.5597 | 109.3934
|  | 178.6581 | 109.3134
|  | 178.7565 | 111.2339
|  | 178.8548 | 112.8344
|  | 178.9532 | 112.4743
|  | 179.0516 | 111.274
|  | 179.15 | 110.0736
|  | 179.2484 | 108.9532
|  | 179.3467 | 109.6735
|  | 179.4451 | 111.5941
|  | 179.5435 | 112.5543
|  | 179.6419 | 111.8341
|  | 179.7403 | 110.7938
|  | 179.8386 | 109.6735
|  | 179.937 | 109.0333
|  | 180.0354 | 110.3137
|  | 180.1338 | 112.2342
|  | 180.2322 | 112.3543
|  | 180.3305 | 111.6341
|  | 180.4289 | 110.5137
|  | 180.5273 | 109.4334
|  | 180.6257 | 109.3934
|  | 180.7241 | 110.9939
|  | 180.8224 | 112.3543
|  | 180.9208 | 112.3143
|  | 181.0192 | 111.434
|  | 181.1176 | 110.3137
|  | 181.216 | 109.2733
|  | 181.3143 | 109.2333
|  | 181.4127 | 111.314
|  | 181.5111 | 112.6744
|  | 181.6095 | 112.6744
|  | 181.7079 | 112.6744
|  | 181.8062 | 112.6744
|  | 181.9046 | 112.6744
|  | 182.003 | 112.6744
|  | 182.1014 | 112.6744
|  | 182.1998 | 112.6744
|  | 182.2981 | 112.6744
|  | 182.3965 | 112.6744
|  | 182.4949 | 109.8735
|  | 182.5933 | 109.0733
|  | 182.6917 | 110.3537
|  | 182.79 | 112.2743
|  | 182.8884 | 112.4743
|  | 182.9868 | 111.8341
|  | 183.0852 | 110.7938
|  | 183.1836 | 109.6735
|  | 183.2819 | 109.3534
|  | 183.3803 | 110.7138
|  | 183.4787 | 112.3943
|  | 183.5771 | 112.3543
|  | 183.6755 | 111.394
|  | 183.7738 | 110.1136
|  | 183.8722 | 109.2333
|  | 183.9706 | 109.3534
|  | 184.069 | 111.274
|  | 184.1674 | 112.9545
|  | 184.2657 | 112.5143
|  | 184.3641 | 111.0739
|  | 184.4625 | 110.1136
|  | 184.5609 | 109.0733
|  | 184.6593 | 109.4734
|  | 184.7576 | 111.8741
|  | 184.856 | 112.8344
|  | 184.9544 | 111.9542
|  | 185.0528 | 111.1539
|  | 185.1512 | 109.9536
|  | 185.2495 | 108.8332
|  | 185.3479 | 110.1136
|  | 185.4463 | 112.2342
|  | 185.5447 | 112.7144
|  | 185.6431 | 111.5941
|  | 185.7414 | 110.6338
|  | 185.8398 | 109.4334
|  | 185.9382 | 108.8732
|  | 186.0366 | 110.1536
|  | 186.135 | 112.2342
|  | 186.2333 | 112.5944
|  | 186.3317 | 111.474
|  | 186.4301 | 110.1136
|  | 186.5285 | 109.2333
|  | 186.6269 | 109.1933
|  | 186.7252 | 111.1139
|  | 186.8236 | 112.8744
|  | 186.922 | 112.2342
|  | 187.0204 | 110.9539
|  | 187.1188 | 109.9936
|  | 187.2171 | 109.1133
|  | 187.3155 | 109.6335
|  | 187.4139 | 111.8741
|  | 187.5123 | 112.9145
|  | 187.6107 | 112.1142
|  | 187.709 | 110.9138
|  | 187.8074 | 109.9536
|  | 187.9058 | 109.0733
|  | 188.0042 | 109.9536
|  | 188.1026 | 112.0342
|  | 188.2009 | 112.3543
|  | 188.2993 | 111.9542
|  | 188.3977 | 110.7938
|  | 188.4961 | 109.6735
|  | 188.5945 | 109.0333
|  | 188.6928 | 110.3937
|  | 188.7912 | 112.3143
|  | 188.8896 | 112.8744
|  | 188.988 | 111.9942
|  | 189.0864 | 110.8738
|  | 189.1847 | 109.8335
|  | 189.2831 | 109.4334
|  | 189.3815 | 111.0739
|  | 189.4799 | 112.8344
|  | 189.5783 | 112.4343
|  | 189.6766 | 111.394
|  | 189.775 | 110.4337
|  | 189.8734 | 109.2333
|  | 189.9718 | 109.5134
|  | 190.0702 | 111.5941
|  | 190.1685 | 112.7944
|  | 190.2669 | 111.8341
|  | 190.3653 | 111.1139
|  | 190.4637 | 109.9936
|  | 190.5621 | 109.1933
|  | 190.6604 | 110.3137
|  | 190.7588 | 112.2342
|  | 190.8572 | 112.4343
|  | 190.9556 | 111.7941
|  | 191.054 | 110.5937
|  | 191.1523 | 109.6335
|  | 191.2507 | 109.1533
|  | 191.3491 | 110.3537
|  | 191.4475 | 112.2743
|  | 191.5459 | 112.9145
|  | 191.6442 | 111.7141
|  | 191.7426 | 111.5941
|  | 191.841 | 109.5134
|  | 191.9394 | 109.4734
|  | 192.0378 | 110.5137
|  | 192.1361 | 110.9539
|  | 192.2345 | 112.4743
|  | 192.3329 | 112.1942
|  | 192.4313 | 110.8338
|  | 192.5297 | 110.1936
|  | 192.628 | 109.1533
|  | 192.7264 | 109.5934
|  | 192.8248 | 111.6741
|  | 192.9232 | 112.5543
|  | 193.0216 | 112.0342
|  | 193.1199 | 110.8338
|  | 193.2183 | 109.7135
|  | 193.3167 | 109.1933
|  | 193.4151 | 110.5537
|  | 193.5135 | 112.2342
|  | 193.6118 | 112.6344
|  | 193.7102 | 111.514
|  | 193.8086 | 110.3937
|  | 193.907 | 109.4334
|  | 194.0054 | 108.3931
|  | 194.1037 | 107.2727
|  | 194.2021 | 106.3925
|  | 194.3005 | 105.0721
|  | 194.3989 | 103.9517
|  | 194.4973 | 102.9114
|  | 194.5956 | 101.711
|  | 194.694 | 100.5907
|  | 194.7924 | 99.4704
|  | 194.8908 | 98.35
|  | 194.9892 | 97.2297
|  | 195.0875 | 96.1894
|  | 195.1859 | 94.5889
|  | 195.2843 | 93.4685
|  | 195.3827 | 91.868
|  | 195.4811 | 90.5876
|  | 195.5794 | 88.9071
|  | 195.6778 | 87.7068
|  | 195.7762 | 85.9462
|  | 195.8746 | 84.0256
|  | 195.973 | 82.4252
|  | 196.0713 | 80.5846
|  | 196.1697 | 78.504
|  | 196.2681 | 75.7831
|  | 196.3665 | 72.9022
|  | 196.4649 | 68.901
|  | 196.5632 | 63.4594
|  | 196.6616 | 61.0586
|  | 196.76 | 61.9389
|  | 196.8584 | 60.7385
|  | 196.9568 | 58.6579
|  | 197.0551 | 56.0971
|  | 197.1535 | 53.8564
|  | 197.2519 | 52.8561
|  | 197.3503 | 53.9765
|  | 197.4487 | 60.3784
|  | 197.547 | 59.6582
|  | 197.6454 | 57.5776
|  | 197.7438 | 56.0571
|  | 197.8422 | 48.3748
|  | 197.9406 | 0
|  | 198.0389 | 0
|  | 198.1373 | 0
|  | 198.2357 | 0
|  | 198.3341 | 0

## Settings
XLPRO ™ Version 4.06 |  | BW-1 | DEFAULT
Operation Settings Area |  | 3500 | 3500 | Max. Freq. (Hz)
|  | 20 | 20 | Initial Delay (ms)
Units |  | 7 | 7 | End Delay (ms)
Radiation: | R | 100 | 100 | Width (ms)
/Time: | /min | 7 | 7 | Pulses in window
Scale: |  | 3 | 3 | Initial Interval (ms)
Pulse Rate: | 15 | 20 | 20 | Noise threshold (kV)
Mode |  | 0.9 | 0.9 | Upper threshold
Units Display: | Rescale | 0.7 | 0.7 | Lower threshold
Cell Update: | Incremental | 3 | 3 | FT-resolution control
Auto Zero: | Disabled |  | INSTRUMENT PARAMETERS
Capture Mode: | Auto | 0.09838 | 0.09838 | 9000 Sample Interval (ms)
Screen Tracking: | Yes | 38.08 | 38.08 | 9000 Input Filter (us) |  |  | 4081/4082
kVp |  | 37.3 | 37.3 | 37.3 | 37.3 | Sample Interval (us)
Config: | Dose+Time+kV | 2.9764 | 2.9764 | 2.9764 | 2.9764 | Input Filter -3dB (kHz)
Scale: | DAccu-kV | 16.2 | 16.2 | 16.2 | 16.2 | Output Filter (us)
Gain: | 32.77 | 4 | 4 | 2.6 | 2.6 | Averaging Intervals
Offset: | 0.00 | 0.75 | 0.75 | Width fraction of kVp
Comm |  | 0.17 | 0.17 | Duplicate width (ms)
Index: | 0 | 0.0768 | 0.0768 | 9095 Sample Interval (ms)
Baud: | 9600
Monitor Configuration |  |  |  | COEFFICIENTS FOR PRACTICAL kV
Serial No: | N/A | Power | Values
ROM Ver: | N/A | Min. kV | 20 | 36 | 150 | 161
Detector: | N/A | 0 | -18.6452 | -0.017472 | -0.017472 | 0
Chamb Kq: | N/A | 1 | 0.85068 | 1.0308e-05 | 1.0308e-05 | 0
Atm Corr: | Automatic | 2 | -0.019681 | 2.3082e-05 | 2.3082e-05 | 0
Temp (°C): | 22 | 3 | 0.00025663 | -1.662e-07 | -1.662e-07 | 0
Press (kPa): | 101.3 | 4 | -2.1424e-06 | 4.3106e-10 | 4.3106e-10 | 0
|  | 5 | 0 | 0 | 0 | 0
|  | EXP? | 1 | 0 | 0 | 0
NON-LINEAR CORRECTION FACTORS
|  | 1030 | 2700 | 0.12
|  | 0.343 | -0.569 | 0.2809 | -0.0331
|  | 2.9804 | -5.4582 | 4.6956 | -1.4557