## Readme
General Notes:
These worksheets in this workbook were derived from the old "Examp9010.xls" workbook. They have
been updated for use with 9095 XLPRO and provide automatic kVp measurment capture and
analysis. In some cases where the kVp data might be erroneous (e.g. a hvl data point with added filtration)
the kVp values are hidden by coloring the font white.
Font colors have been used to help identify different types of data. In general, all user entry data is in a medium-blue color.
All text that should not need to be changed in is black. Dose data is dark-blue, time data is medium-green, and kVp data is dark-red
| EXAMPLE |  | Measured
|  | SET | Dose | Time | HV
|  | mA | mR | ms | kVp
|  | 25 | 6.68 | 44.3 | 58.5
To protect formatting and foumulas, all cells that should not need to be changed are locked and protected.
Those cells that contain formulas have also been hidden. You may of course unlock them and edit as you see fit.
Before unlocking or making changes, be sure you have retained a copy of the original workbook. To remove the protection,
Select "Unprotect Sheet" from the Tools\Protection menu.
The worksheets have been formatted to print a report and exclude comments and extraneous data. To view how a
sheet will look when printed, select "File" and then "Print Preview".
For additional information see the Radcal "XLPROManual.pdf" file either on the install CD or
in your 9000PXL directory.
Note on E13: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.

## Worksheet Information
Sample Worksheets in "9095 Sample Template.xls"
The three sample worksheets in the workbook 9095 Sample Template.xls delivered with
9095 XLPRO have been designed to demonstrate the ease of use of XLPRO with Excel and
a formatted worksheet. These worksheets capture dose and rate & kVp from a 9095 and
implement typical QA measurements and data-collection requirements for a radiology department.
The degree of macro automation used in these worksheets is minimal, which makes it very easy
for users to modify the sheets to fit their individual data-collection needs.
Most of the usefulness of these worksheets comes from the formulas and graphs that are in the
worksheets themselves. The user places the cursor at the cell into which a measurement is to be
inserted and then initiates a measurement through a control-key sequence. This control-key
sequence initiates communication with the 9095 through the XLPRO control interface.
Once XLPRO inserts the measurement into the cell, control is returned to the user either automatically
.Or by pressing ESC. The user then ensures that the cursor is in the correct cell for the next measurement
and initiates a new measurement. This sequence continues until all the required measurements are completed.
For each section of these worksheets, there is a setup area to input machine parameters, a data collection
area for either manual or automatic data collection, and a comment area to outline the
results of the test and the action to be taken. All three of these worksheets come with the
protection property of Excel turned on so that only the data-entry areas of the worksheet are
unprotected and modifiable by the user. You may turn the protection property off so that you can
modify other areas easily. No password is required.
You should keep in mind when modifying these worksheets that there are items such as the
Settings Area and intermediate calculation areas which are located outside of the main worksheet.
These areas must not be broken by inserting rows into the middle of them.
Macros
The workbook contains the macros listed below. The Auto_Open subroutine initializes
combination CONTROL-SHIFT keys to run the subroutines that they are associated with. These
subroutines first determine the cell at the top-left corner of the Settings Area and use these coordinates to
record the initial settings for Scale and Cell Update. Then they modify the settings area cells as needed to
execute the command, and execute the appropriate 9000XLMacro.xls macro (e.g.MacroRunMaxRate),
which does not exit if no argument is passed to the macro. The subroutines end after restoring the
Settings Area to its initial configuration and advancing the selection by one row. The Auto_Close
subroutine allows you to delete the Auto_Activate commands when the workbook closes.
FOR EXAMPLE
Sub Auto_Open()
Application.OnKey "+^o", "RateOverwrite"
Application.OnKey "+^f", "FluorokVpOverwrite"
End Sub
Sub RateOverwrite()
On Error GoTo Quit 'Trap no Settings Area or its write protection.
'Find the coordinates of the top-left cell of the Settings Area.
r1 = ActiveSheet.Names("_RCTL").RefersToRange.Row
c1 = ActiveSheet.Names("_RCTL").RefersToRange.Column
'Store the current values from the Settings Area.
incr = Cells(r1 + 10, c1 + 2)
units = Cells(r1 + 6, c1 + 2)
'Modify the Settings Area.
Cells(r1 + 10, c1 + 2) = "Overwrite"
Cells(r1 + 6, c1 + 2) = ""
'Run the MaxRate macro.
Application.Run "9000XLMacro.xls!MacroRunMaxRate"
'Restore the Settings Area to its old values and advance the selection by one row.
Cells(r1 + 10, c1 + 2) = incr
Cells(r1 + 6, c1 + 2) = units
ActiveCell.Offset(1, 0).Select
Quit:
If (Err) Then Call MsgBox("No recorded Settings Area or it is protected.", _
vbCritical, "Local Error")
On Error GoTo 0
End Sub
Sub Auto_Close()
If (Right(ActiveWorkbook.Name, 3) = "xlt") Then Exit Sub
i = 0
On Error Resume Next
For j = 1 To ActiveWorkbook.Sheets.Count 'Look for Auto_Activates.
ws = ActiveWorkbook.Sheets(j).Name
If (ws <> "Settings") Then i = Sheets(j).Names("Auto_Activate").Index
Next j
On Error GoTo 0
If (i = 0) Then Exit Sub 'Avoid query if no Auto_Activates are present.
If (MsgBox("Do you want to keep auto activation?", vbYesNo) <> vbNo) Then Exit Sub
On Error Resume Next
For j = 1 To ActiveWorkbook.Sheets.Count 'Delete Auto_Activates.
ws = ActiveWorkbook.Sheets(j).Name
If (ws <> "Settings") Then Sheets(j).Names("Auto_Activate").Delete
Next j
On Error GoTo 0
End Sub
Sheet named Radiographic
The worksheet Radiographic is set up for a strictly radiographic room or station. The red comment indicators,
in the yellow cells indicate comments that describe the steps required for executing the Measure Pulse
| macro (CONTROL-SHIFT P). |  |  | Example | Note: these comments will not be visible when you print.
A. Facility Information
Facility information contains data areas for equipment location, x-ray equipment type and
equipment operational ratings. There are no formulas associated with this section of the
worksheet.
B. kVp Accuracy and Exposure Linearity
kVp Accuracy and Exposure Linearity provides automatic input of the Dose Time & kVp
data from the radiation monitor. Formulas for this section are for kV error
calculation & Pass/Fail criteria ($24:G37), linear-regression calculation of the best-fit
straight line to the measured dose data, and the error of the measured dose data
relative to the best-fit dose data (H24:I37). Linear-regression calculations are performed just to
the right of the graph in an area of the worksheet (T21:AA37) which is not in the print area.
Formulas in this area must be modified by the user if the full data area is not used.
C. mA and Exposure Linearity
mA and Exposure Linearity allows for the automatic input of exposure data from the radiation
monitor. The Set mA data area is unprotected, and users may change these numbers to reflect the
settings on their x-ray machines. Formulas for this section are for the linear-regression calculation
of the best-fit straight line to the measured dose data (F57:F63), the error of the measured dose
data relative to the best-fit dose data (G57:G63), and average dose/mAs (B65). Again, the linefitting
calculations are performed just to the right of the graph in an area of the worksheet
(X56:AB64) that is not in the print area. Formulas in this area must be modified by the user if the
full data area is not used.
D. Timer and Exposure Linearity
Timer and Exposure Linearity allows for the automatic input of exposure data from the radiation
monitor. The Set Time data area is unprotected, and users may change these numbers
to reflect the settings on their x-ray machines. Formulas for this section are for the calculation
of timer error (F79:F94), the linear regression calculation of the best-fit straight
line to the measured dose data (G79:G94), and the error of the measured dose data relative to
the best-fit dose data (H79:H94). Again, the linefitting calculations are performed just to the right
of the graph in an area of the worksheet (X79:AB87) that is not in the print area.
Formulas in this area must be modified by the user if the full data area is not used.
E. Exposure Reproducibility
Exposure Reproducibility is a summary area where the user may enter information about the x-ray
techniques used, a summary of the results, and comments about the action required.
F. Half Value Layer
Formulas for this section are for the calculation of mR/s (E134:F139), the normalization of the mR/s values
(G134:G139), and the calculation of the HVL (D141) done in cells T132:X139. The HVL is calculated
using the mR/s values. The formulas for mR/s calculation use the time in the cell
just to the left of the mR/s cell. Again, the two-point log interpolation for
determining the calculated HVL is done just to the right of this section in a non-printing area of
the worksheet. If the user wishes to do fewer measurements in determining HVL, then the
interpolation formula must be changed also.
G. Phototimer
Phototimer allows capture of data for three different cell areas of the phototimer area.
The last four sections (H., I., J. and K.) are text-entry areas for users to insert their observations
and comments about various non-radiation related aspects of the x-ray machine performance.
Sheet named Special
The "Special" worksheet is set up for a fluoroscopic or radiographic room. The red indicators
in the yellow cells indicate comments that describe the steps required for executing the MeasureFluorokVp
| macro (CONTROL-SHIFT F). |  |  | Example | Note: these comments will not be visible when you print.
A. Facility Information
Facility information contains data areas for equipment location, x-ray equipment type and
equipment operational ratings. There are no formulas associated with this section of the
worksheet.
B. Fluoroscopy
The fluoroscopy section has data-entry areas for machine-setup parameters, for dose-rate or
exposure information, and result and action comments. There are no formulas associated with this
section of the worksheet.
1) Maximum Table Top Exposure Rate
2) Automatic Dose Control
3) Automatic Brightness Control
4) Resolution
5) Timer
6) kV Accuracy
7) Half Value Layer
8) Spot Films
9) Exposure Reproducibility
10) Mechanical
C. Radiographic Overhead X-Ray Tube
See Radiographic Sheet discussion above
Sheet named Special
The "Mobile" worksheet is set up for a mobile x-ray machine. The red comment indicators,
in the yellow cells indicate comments that describe the steps required for executing the Measure kVp
| macro (CONTROL-SHIFT K). |  |  | Example | Note: these comments will not be visible when you print.
This sheet follows the format and measurement principles found in the Radiographic sheet
discussed above.
Sheet named Waveforms
This sheet contains a number of waveforms that have been captured from different types of generators.
In all cases, the data was collected by selecting Measure kVp from the 9000XL menu (CONTROL-SHIFT K)
Sheet named Test
This sheet is provided as a place to make simple "scratch measurments" as you learn how
to use 9000XL.
Sheet named Settings
This sheet is found in all 9000XL workbooks and contains the basic settings information needed to start
making measurments. For further information regarding the Settings sheet see the Radcal "XLPROManual.pdf"
file either on the install CD or in your 9000PXL directory.
Note on F89: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on F147: 1. Place cursor in this cell and press "Control-Shift f" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make a fluoro exposure. 3. Wait for kV and Rate values to settle to stable numbers. 4. Press Escape to end.
Note on F175: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.

## Radiographic
2026 ANNUAL CALIBRATION REPORT | [25 empty] | XLPRO ™ Version 4.06
RADIOLOGY DEPARTMENT | [25 empty] | Operation Settings Area
[7 empty] | DATE: | 2026-07-30 14:06
Units
A. | FACILITY |  |  | Almost anywhere USA | [6 empty] | Room No: | B243 |  | Other: | [11 empty] | Radiation: | R
|  | State Reg. No: |  | CA423-619 |  |  |  | UNIT: | Radiographic | [16 empty] | /Time: | /min
[26 empty] | Scale: | milli
| EQUIPMENT |  |  | Manufacturer | [5 empty] | Model |  |  |  | Serial No. | [11 empty] | Pulse Rate: | 0.25-9-6-5
|  |  | CONSOLE: | General Electric |  |  |  |  | DXR-750 |  |  |  | A417176 | [12 empty] | Mode
|  |  | TUBE: | Machlett |  |  |  |  | Dynamax 69 B |  |  |  | 99828288 | [12 empty] | Units Display: | Omit
|  |  | RATINGS: |  |  | Maximum tube kV: | 150 | kVp. |  |  | Maximum tube mA: | 700 | mA. | Phase: | 3 | [10 empty] | Cell Update: | Incremental
[6 empty] | Inherent filtration: | 1.2 | mm Al |  |  | Total Filtration | 3.1 | mm Al | [12 empty] | Auto Zero: | Disabled
[26 empty] | Capture Mode: | Auto
[26 empty] | Screen Tracking: | Yes
B. | kVp ACCURACY AND EXPOSURE LINEARITY | [24 empty] | kVp | 9000
|  | Focal Spot: | lg |  | Field Size: | Max | [19 empty] | Config: | Dose+Time+kV
|  | 60 | mA | 100 | ms | 6 | mAs | [18 empty] | Scale: | 9095
| Radcal Accu-kV diagnostic sensor placed on table . | [24 empty] | Start delay (ms): | 0
|  | @ | 20.5 | inches from target, | [21 empty] | Width (ms): | 5000
| Radcal Chamber at: |  |  | 15.25 | inches from target | [20 empty] | Filtration (mm-Al): | 2.5
|  | Measured |  |  | kV Accuracy |  | Linearity | [10 empty] | Exposure time: | 0.1 | sec | [5 empty] | Dental threshold: | 0
| Set | Dose | Time | HV | Error | Pass/ | mR | % | [8 empty] | Ln Values |  | Reg. |  |  | Regression Output: |  |  |  | Device Index | 0
| kVp | mR | ms | kVp | kV | Fail | regrs. | Error | [8 empty] | kVp | Dose | Dose |  |  |  | b | a |  | Monitor Configuration
| 50 | 198 | 99.9 | 49.1 | -0.9 |  | 192.6169 | -2.795% | [8 empty] | 3.8939 | 5.2883 | 192.6169 |  |  | Coeff Value | 2.079 | -2.8348 |  | Serial No: | 00008
| 55 | 225 | 99.5 | 52 | -3 |  | 217.0242 | -3.675% | [8 empty] | 3.9512 | 5.4161 | 217.0242 |  |  | Std Error | 0.029373 | 0.12877 |  | ROM Ver: | 4.01
| 60 | 260 | 99.7 | 58 | -2 |  | 272.3367 | 4.53% | [8 empty] | 4.0604 | 5.5607 | 272.3367 |  |  | R squared | 0.99761 | 0.032073 |  | Detector: | 6 cc
| 65 | 291 | 100 | 61 | -4 | FAIL | 302.4414 | 3.783% | [8 empty] | 4.1109 | 5.6733 | 302.4414 | [6 empty] | Chamb Kq: | 1.02
| 70 | 381 | 101 | 69 | -1 |  | 390.7602 | 2.498% | [8 empty] | 4.2341 | 5.9428 | 390.7602 |  | y = a + (b * x) |  |  |  |  | Atm Corr: | Automatic
| 75 | 421 | 101 | 72 | -3 |  | 426.9119 | 1.385% | [8 empty] | 4.2767 | 6.0426 | 426.9119 |  | a= | -2.8348 |  |  |  | Temp (°C): | 24.1
| 80 | 525 | 101 | 79 | -1 |  | 517.7412 | -1.402% | [8 empty] | 4.3694 | 6.2634 | 517.7412 |  | b= | 2.079 |  |  |  | Press (kPa): | 100.1
| 85 | 560 | 100 | 83 | -2 |  | 573.7337 | 2.394% | [8 empty] | 4.4188 | 6.3279 | 573.7337
| 90 | 677 | 101 | 88 | -2 |  | 647.9295 | -4.487% | [8 empty] | 4.4773 | 6.5177 | 647.9295
| 95 | 731 | 102 | 91 | -4 | FAIL | 694.698 | -5.226% | [8 empty] | 4.5109 | 6.5944 | 694.698
| 100 | 832 | 100 | 99 | -1 |  | 827.7064 | -0.5187% | [8 empty] | 4.5951 | 6.7238 | 827.7064
| 110 | 1030 | 101 | 110 | 0 |  | 1030.406 | 0.0394% | [8 empty] | 4.7005 | 6.9373 | 1030.406
| 120 | 1210 | 102 | 120 | 0 |  | 1234.7318 | 2.003% | [8 empty] | 4.7875 | 7.0984 | 1234.7318
| 130 | 1400 | 103 | 128 | -2 |  | 1412.0359 | 0.8524% | [8 empty] | 4.852 | 7.2442 | 1412.0359
|  |  |  | Max.: | -4 | kV | Max: | -5.226%
|  |  |  | Limit ± | 3 | kV
Comments:
|  |  |  | kV Accuracy: | Service recommended for the 65 & 95 kV stations.
GE's specifications for kV accuracy need to be checked.
|  |  |  | Exposure Linearity: | Exposures linear to within 6% at all kV stations tested
Acceptable results
C. | mA AND EXPOSURE LINEARITY
|  | Technique: | 60 | kV | 50 | ms
| 21 | in. target to chamber distance
Measured
| SET | Dose | Time | HV | Regression
| mA | mR | ms | kVp | mR | Error | [16 empty] | Regression Output:
| 25 | 6.68 | 44.3 | 58.5 | 8.4324 | -20.78% | [14 empty] | 6.7459 |  |  | b | a
| 50 | 13.2 | 45 | 59 | 17.3399 | -23.88% | [14 empty] | 6.936 |  | Coeff Value | 0.3563 | -0.47518
| 100 | 33.7 | 47.5 | 59.1 | 35.155 | -4.139% | [14 empty] | 7.031 |  | Std Error | 0.015556 | 4.7879
| 200 | 70.9 | 49.2 | 59.4 | 70.7853 | 0.1621% | [14 empty] | 7.0785 |  | R squared | 0.99056 | 7.9673
| 300 | 111 | 48.9 | 59.4 | 106.4155 | 4.308% | [14 empty] | 7.0944
| 400 | 155 | 48.6 | 59.5 | 142.0457 | 9.12% | [14 empty] | 7.1023 | y = a + (b * x)
| 600 | 203 | 48.3 | 59.4 | 213.3062 | -4.832% | [14 empty] | 7.1102 | a= | -0.47518
| Regression Avg.: |  |  |  | Max. Error: | -23.88% | [15 empty] | b= | 0.3563
| 7.014 | mR/mAs @ 21 inches
Comments:
|  |  | Results: | Exposures should be linear to within +/-10%
At 25 and 50 mA the exposures are too low
|  |  | Action: | SERVICE REQUIRED
D. | TIMER AND EXPOSURE LINEARITY
|  | Technique: | 75 | kV | 300 | mA
| 21 | in. target to chamber distance
Measured
| SET | Dose | Time | HV | Timer | Regression
| Time | mR | ms | kVp | Error | mR | Error
| 17 | 78 | 18 | 74.5 | 5.882% | 79.2031 | -1.519% | [14 empty] | Regression Output:
| 25 | 89.8 | 24 | 74 | -4% | 106.4557 | -15.65% | [15 empty] | b | a
| 33 | 118 | 35 | 73.7 | 6.061% | 133.7083 | -11.75% | [14 empty] | Coeff Value | 3.4066 | 21.2914
| 50 | 178 | 52 | 73.6 | 4% | 191.62 | -7.108% | [14 empty] | Std Error | 0.019324 | 7.8785
| 67 | 233 | 65 | 74.2 | -2.985% | 249.5317 | -6.625% | [14 empty] | R squared | 0.99955 | 22.2629
| 100 | 352 | 102 | 74.17 | 2% | 361.9485 | -2.749%
| 125 | 448 | 128 | 74.14 | 2.4% | 447.1128 | 0.1984% | [13 empty] | y = a + (b * x)
| 150 | 535 | 155 | 74.11 | 3.333% | 532.277 | 0.5116% | [13 empty] | a= | 21.2914
| 200 | 704 | 205 | 74.08 | 2.5% | 702.6056 | 0.1985% | [13 empty] | b= | 3.4066
| 250 | 913 | 255 | 74.05 | 2% | 872.9341 | 4.59%
| 300 | 1060 | 306 | 74.02 | 2% | 1043.2627 | 1.604%
| 400 | 1400 | 408 | 73.99 | 2% | 1383.9197 | 1.162%
| 500 | 1760 | 508 | 73.96 | 1.6% | 1724.5768 | 2.054%
| 600 | 2080 | 610 | 73.93 | 1.667% | 2065.2339 | 0.715%
| 800 | 2700 | 815 | 73.9 | 1.875% | 2746.548 | -1.695%
| 1000 | 3420 | 1008 | 73.87 | 0.8% | 3427.8622 | -0.2294%
Comments:
|  |  | Results: | Exposures should be linear to within +/-10%
At .0167 and .025 seconds the exposures are too low
|  |  | Action: | SERVICE REQUIRED
E. | EXPOSURE REPRODUCIBILITY
|  |  |  | Dose | Time | kVp | mR/mAs
|  |  | PASS/FAIL Criteria: | 5% | 5% | 5% | 5%
|  |  | mA: | 200 | ms: | 100 |  | mA: | 200 | ms: | 100 |  | mA: | 200 | ms: | 100
|  |  | kVp: | 60 | mAs: | 20 |  | kVp: | 60 | mAs: | 20 |  | kVp: | 60 | mAs: | 20
|  |  | Measured |  |  |  |  | Measured |  |  |  |  | Measured
|  |  | Dose | Time | HV | mR/ |  | Dose | Time | HV | mR/ |  | Dose | Time | HV | mR/
|  |  | mR | ms | kVp | mAs |  | mR | ms | kVp | mAs |  | mR | ms | kVp | mAs
|  | X1 | 100 | 100 | 59.4 | 5 | X1 | 100 | 100 | 59.4 | 5 | X1 | 100 | 100 | 59.4 | 5
|  | X2 | 99 | 98 | 60 | 5.051 | X2 | 99 | 98 | 60 | 5.051 | X2 | 99 | 98 | 60 | 5.051
|  | X3 | 98 | 98 | 60.1 | 5 | X3 | 98 | 98 | 60.1 | 5 | X3 | 98 | 98 | 60.1 | 5
|  | X4 | 99 | 98.5 | 59.8 | 5.0254 | X4 | 99 | 98.5 | 59.8 | 5.0254 | X4 | 99 | 98.5 | 59.8 | 5.0254
|  | X5 | 94 | 99.6 | 60.3 | 4.7189 | X5 | 94 | 99.6 | 60.3 | 4.7189 | X5 | 94 | 99.6 | 60.3 | 4.7189
|  | Avg | 98 | 98.82 | 59.92 | 4.9591 | Avg | 98 | 98.82 | 59.92 | 4.9591 | Avg | 98 | 98.82 | 59.92 | 4.9591
|  | CV | 2.393% | 0.9395% | 0.5708% | 2.741% | CV | 2.393% | 0.9395% | 0.5708% | 2.741% | CV | 2.393% | 0.9395% | 0.5708% | 2.741%
|  | Pass/Fail | Pass | Pass | Pass | Pass |  | Pass | Pass | Pass | Pass |  | Pass | Pass | Pass | Pass
|  |  | RESULTS: | Reproducibility was within the required +/- 5%.
|  |  | COMMENT: | Acceptable results
F. | HALF VALUE LAYER
Technique
| kV | mA | ms | added filtration |  |  | 74.5 kV (75 kV setting) 200 mA, 2.0 mmAl Added
| 75 | 200 | 100 | 2 | mmAl |  | Calculated HVL @75.0 kV setting is 2.5 mmAl
| 15.25 | in. target to chamber distance
Measured
| mm | Dose | Time | HV | Calculated | [12 empty] | HVL (2 point Log interpolation)
| Al | mR | ms | kVp | mR/s | Atten | [13 empty] | (a)Just > | (b)Just < |  | 0.5 line data
| 0 | 218 | 100.1 | 74.5 | 2177.8222 | 1 | [10 empty] | 1 |  | Target Value | 0.5 |  |  | 0 | 0.5
| 1 | 158 | 99.8 | 74.8 | 1583.1663 | 0.72695 | [10 empty] | 0.72695 |  | Row | 3 | 4 |  | 5 | 0.5
| 2 | 122 | 99.9 | 75 | 1221.2212 | 0.56075 | [10 empty] | 0.56075 |  | Exposure Rate | 1221.2212 | 970
| 3 | 97 | 100 | 74.1 | 970 | 0.4454 | [10 empty] | 0.4454 |  | Thickness (mm) | 2 | 3
| 4 | 78.1 | 99.9 | 84.9 | 781.7818 | 0.35897 | [10 empty] | 0.35897
| 5 | 65.1 | 99.9 | 75 | 651.6517 | 0.29922 | [10 empty] | 0.29922 |  | Calc. HVL | 2.4979 | mmAl
|  | Calculated HVL: | 2.4979 | mmAl
Comments:
|  |  | Results: | Acceptable
|  |  | Action: | None Required
G. | PHOTOTIMER
Technique
| kV | mA | ms | cassette |  |  | 72 | inch SFD |  | 67.5 | inch SCD, Lead Apron
| 121 | 200 | 500 | 14 x 17 | in. |  |  | Results: | Chest Stand + Lead Apron
|  | Center Cell |  |  |  |  | Right Cell |  |  |  |  | Left Cell
|  | Measured |  |  |  |  | Measured |  |  |  |  | Measured
|  | Dose | Time | HV | Dose |  | Dose | Time | HV | Dose |  | Dose | Time | HV | Dose
|  | mR | ms | kVp | Ratio |  | mR | ms | kVp | Ratio |  | mR | ms | kVp | Ratio
| NORM | 44 | 77 | 121.1 | N/A |  | 31 | 52 | 121 | N/A |  | 48 | 85 | 120.8 | N/A
| 1.25 | 52 | 94 | 120.8 | 1.1818 |  | 39 | 68 | 120.8667 | 1.2581 |  | 56 | 101 | 121 | 1.1667
| 1.5 | 65 | 118 | 121 | 1.4773 |  | 47 | 77 | 120.8167 | 1.5161 |  | 72 | 134 | 120.8667 | 1.5
| 2 | 73 | 134 | 120.8667 | 1.6591 |  | 56 | 101 | 120.7667 | 1.8065 |  | 85 | 159 | 121 | 1.7708
| 0.75 | 30 | 51 | 120.8167 | 0.68182 |  | 22 | 35 | 120.7167 | 0.70968 |  | 39 | 68 | 120.8667 | 0.8125
| 0.5 | 22 | 35 | 120.7667 | 0.5 |  | 17 | 27 | 120.8667 | 0.54839 |  | 26 | 43 | 120.8167 | 0.54167
| 0.25 | 13 | 18 | 120.7167 | 0.29545 |  | 9 | 10 | 120.8167 | 0.29032 |  | 13 | 18 | 120.7667 | 0.27083
Comments:
|  |  | Results: | On both normal and 3/4 density settings, Exposures are not balanced
for the three cells. There should be better agreement in % variation
among the three cells.
|  |  | Action: | SERVICE RECOMMENDED
H. | LIGHT FIELD/X-RAY FIELD ALIGNMENT
|  | TABLE: | At 37 inch SFD - Alignment was not within 2% of SFD. The x-ray field is 1/2 inch smaller
and 1/4 inch larger than the light field (cross table), and shifted ~ 1/4 inch longitudinally.
|  | CHEST STAND: | At 72 inch SFD, the alignment is poor: 3/4 inch of the x-ray field is cut off (or shifted)
on the left side (facing the chest unit) and ~3/4 inch too large on the right side.
|  | Comments: | SERVICE REQUIRED
I. | CENTERING OF X-RAY FIELD TO CASSETTE IN BUCKY TRAY
|  | TABLE: | At 40 inch SFD, Center of x-ray field and cassette are within 2% of SFD.
|  | Comments: | Acceptable Results
|  | WALL: | Center of x-ray field and cassette were not within 2% of SFD
|  | Comments: | SERVICE REQUIRED
J. | AUTOMATIC COLLIMATION
SERVICE REQUIRED, this system was not operational.
K. | MECHANICAL
| 1 | Ready light on control panel did not light
| 2 | Chest stand cassette size indicator light did not light.
| 3 | When using low technique (~ 50 - 55 kV, 5 mAs) the RESET indicator came on.
| 4 | Several times, during testing no exposure occurred when the exposure switch was engaged.
L. | IN AIR EXPOSURE FROM CLINICAL PROTOCOLS
Measured
|  |  |  | Generator Settings |  |  | Dose | Time | HV
|  | STUDY |  | KV | mA | ms | mR | ms | kVp
|  | ABDOMEN |  | 75 | 200 | 300 | 704 | 305 | 74.8
|  |  |  | 105 | 200 | 50 | 242 | 48 | 105.1
|  |  |  | 68 | 300 | 100 | 278 | 97 | 68.7
CHEST
|  |  | a) Erect | 75 | 300 | 25 | 89 | 22 | 74.3
|  |  |  | 85 | 400 | 100 | 667 | 98 | 84.2
|  |  |  | 88 | 400 | 100 | 712 | 7 | 87.2
|  |  | b) Supine | 63 | 300 | 25 | 64 | 22 | 62.8
|  |  |  | 65 | 300 | 17 | 42 | 13 | 64.9
|  |  |  | 76 | 300 | 33 | 124 | 31 | 75.8
|  |  | LAT | 95 | 600 | 25 | 3050 | 256 | 93.7
|  |  |  | 110 | 300 | 33 | 278 | 32 | 109.6
c) PHOTOTIMED - on 3/4 Density
|  |  |  | 80 | 300 | 100 | 433 | 97 | 79.7
|  |  |  | 100 | 200 | 500 | 2240 | 506 | 100.1
HIP & PELVIS
|  |  |  | 69 | 100 | 200 | 183 | 203 | 69.2
|  |  |  | 70 | 300 | 200 | 635 | 197 | 69.9
|  |  |  | 74 | 200 | 250 | 608 | 255 | 74.2
|  |  |  | 90 | 300 | 250 | 1400 | 256 | 89.9
|  | L-SPINE |  | 75 | 200 | 250 | 585 | 255 | 75
|  |  |  | 80 | 200 | 250 | 719 | 255 | 80.1
|  |  |  | 85 | 200 | 800 | 2330 | 797 | 85
|  |  |  | 90 | 200 | 800 | 2810 | 797 | 89.8
OBSTRUCTION SERIES
|  |  |  | 78 | 300 | 150 | 604 | 147 | 77.3
|  |  |  | 80 | 300 | 400 | 1680 | 397 | 79.2
|  |  |  | 85 | 300 | 400 | 1770 | 397 | 84.4
CALIBRATION PERFORMED BY:
CERTIFIED RADIOLOGICAL PHYSICIST
Note on C24: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on C57: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on C79: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on C134: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on C157: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on D111: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on H157: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on M157: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.

## Special
2026 ANNUAL CALIBRATION REPORT | [22 empty] | XLPRO ™ Version 4.06
RADIOLOGY DEPARTMENT | [22 empty] | Operation Settings Area
[6 empty] | DATE: | 2026-07-30 14:06
Units
A. | FACILITY |  |  | Almost anywhere USA | [6 empty] | Room No: | 453B |  | Other: | [8 empty] | Radiation: |  | R
|  | State Reg. No: |  | No. 705544 (Fluro) |  |  |  | UNIT: | Special Procedures | [13 empty] | /Time: |  | /min
|  |  |  | No 703190 (Radiographic) | [18 empty] | Scale:
| EQUIPMENT |  |  | Manufacturer | [5 empty] | Model |  |  |  | Serial No. | [8 empty] | Pulse Rate: |  | 15-9-6-5
|  |  | CONSOLE: | General Electric |  |  |  |  | DXR-750 |  |  |  | A417176 | [9 empty] | Mode
|  |  | TUBE: | Machlett |  |  |  |  | Dynamax 69 B |  |  |  | 99828288 | [9 empty] | Units Display: |  | Omit
|  |  | RATINGS: |  |  | Maximum tube kV: | 165 | kVp. |  |  | Maximum tube mA: | 100 | mA. | Phase: | 1 | [7 empty] | Cell Update: |  | Overwrite
[6 empty] | Inherent filtration: |  | mm Al |  |  | Total Filtration |  | mm Al | [9 empty] | Auto Zero: |  | Disabled
[23 empty] | Capture Mode: |  | Auto
B. | FLUOROSCOPY | [21 empty] | Screen Tracking: |  | Yes
| 1) | Maximum Table Top Exposure Rate | [20 empty] | kVp |  | 9000F
|  | R/min | kV: | [19 empty] | Config: |  | Dose+kV
|  | 8.6 | 120.1 | [19 empty] | Scale: |  | 9095
|  | mA: | None seen | [19 empty] | Start delay (ms): |  | 0
[23 empty] | Width (ms): |  | 5000
|  | COMMENT: |  | Acceptable. This complies with federal and state regulations | [18 empty] | Filtration (mm-Al): |  | 2.5
[23 empty] | Dental threshold: |  | 0
[23 empty] | Device Index |  | 0
| 2) | Automatic Dose Control | [20 empty] | Monitor Configuration
|  | Technique: | [20 empty] | Serial No: |  | 00008
|  | kV | mA |  | Auto mode will override manual mode if phantom's thickness and density | [17 empty] | ROM Ver: |  | 4.01
|  | 70 | MAX |  | drive system beyond the selected kV (via dial on image intensifier) | [17 empty] | Detector: |  | 6 cc
[23 empty] | Chamb Kq: |  | 1.02
[5 empty] | Table Top Exposures | [17 empty] | Atm Corr: |  | Automatic
|  |  | Phantom |  | kV |  | R/min | kVp | [14 empty] | Temp (°C): |  | 23.9
|  |  | Copper |  | 75 |  | 1.65 | 74.3 | [14 empty] | Press (kPa): |  | 100
[5 empty] | 80 |  | 0.73 | 80.1
[5 empty] | 90 |  | 0.46 | 91
|  |  | 0.75" Al + Cu |  | 72 |  | 7.11 | 70.4
[5 empty] | 80 |  | 2.97 | 80.4
[5 empty] | 120 |  | 0.63 | 121.3
|  |  | 1.5" Al + Cu |  | 86.5 |  | 6.955 | 80.8
[5 empty] | 121 |  | 1.6 | 120.2
|  |  | Lead + Cu + Al |  | 121 |  | 8 | 119.7
|  | COMMENT: |  | This unit is performing as per GE's Design. It will operate on the kV selected
(ie 70 kV) unless the patient/phantom's thickness or density drives the system
to maximum mA first. Then it will increase the kV. The above data illustrates
how patient dose is lowered at higher kV's. Conversely, as kV increases
contrast decreases.
| 3) | Automatic Brightness Control
|  | Technique: |  | Different thicknesses of different materials were placed on the table in the beam
The ion chamber was placed as close as possible to the outside entrance to image
intensifier assembly
[5 empty] | Phantom Material |  |  | R/min | kV
[5 empty] | 0.75 " Aluminum |  |  | 0.0065 | 70
[5 empty] | 1.5 " Aluminum |  |  | 0.006 | 70
[5 empty] | Copper |  |  | 0.005 | 70
[5 empty] | Cu + 0.75"Al |  |  | 0.005 | 70
[5 empty] | 1.5" Al + Cu |  |  | 0.006 | 85
|  | COMMENT: |  | Acceptable Results
| 4) | Resolution
|  |  | Technique: | Line pair phantom was taped to the outside entrance of the image intensifier.
|  |  | Results: | 1.1 line pairs per millimeter
|  | COMMENT: |  | Acceptable Results. This is not as good as in Room 1
| 5) | Timer
|  |  | Results: | Has audible indication at the completion of nominal 5 minutes of Fluoro time.
|  |  | COMMENT: | Acceptable Results.
| 6) | kV Accuracy
|  | Technique: |  | Radcal AccukV sensor was placed on the table at 90 degree angle with the
anode/cathode of the x-ray tube.
[6 empty] | Radcal |  | Apparent
|  |  | SET | kV (TV) |  | R/min | kVp | kV Error
|  |  | 75 | 75 |  |  | 76.1 | 1.1
|  |  | 80 | 80 |  |  | 80 | 0
|  |  | 85 | 85,85 |  |  | 82.5 | -2.5
|  |  | 90 | 89 |  |  | 93 | 3
|  |  | 100 | 99,100 |  |  | 95 | -5
|  |  | 110 |  |  |  | 103 | -7
|  |  | 120 | 121 |  |  | 115 | -5
|  |  | COMMENT: | Service Required for the high kV's
NOTE: The kV meter used to measure the kV drove the automatic system to a
minimum kV of 75.
| 7) | Half Value Layer
|  |  | Result: | 2.9 mm of Al at 75 kV
|  |  | COMMENT: | Acceptable Results
| 8) | Spot Films
|  | Technique: |  | Routine clinical settings of 110 kV, 1.5 seconds back-up for phototimer, phantom.
a) Spacing
No everlap noted on the 4/1 and 2/1 films.
b) Film Density
Variation in density corresponded with exposure variations and phantoms.
See patient exposures below.
c) Patient Exposures
|  |  |  | Technique: | Automatic Mode, Copper, Copper plus Aluminum phantoms.
70 kV on Fluoro, 95 kV, 400 mA, 0.5 sec backup.
|  |  |  | Results: | [5 empty] | Dose | Time
|  |  |  | Phantom |  |  | Spacing |  | Grid | mR | ms
[6 empty] | Copper + 1.5" Al |  | 1/1 | No | 1330 | 77
[9 empty] | Yes | 1670 | 93
[6 empty] | Copper + 3/4" Al |  | 2/1 | No | 917 | 56
[9 empty] | Yes | 874 | 56
[6 empty] | Copper + 1.5" Al |  | 2/1 | No | 2010 | 124
[9 empty] | Yes | 1920 | 119
[6 empty] | Copper + 1.5" Al |  | 4/1 | No | 2670 | 1610
[9 empty] | Yes | 2310 | 1450
[6 empty] | Copper +3/4" Al |  | 4/1 | No | 954 | 64
[6 empty] | Copper Alone |  | 4/1 | Yes | 309 | 26
| 9) | Exposure Reproducibility
|  |  | Result: | Exposures were not within the required 5% for the Fluoro mode using Copper
+ 1.5 in Aluminum phantom. The fluoro kV was 70 kV selected via the kV dial on
the image intensifier assembly. Table top exposures varried as follows:
7.51, 6.7, 6.92, 6.7, 6.38, 6.95, 6.91 R/min
|  |  | COMMENT: | Service Needed
| 10) | Mechanical
|  |  | NOTE: | This system appears to have only the 9 inch fluoro field of view
in spite of 6 and 4.5 inch FOV indicators on the equipment.
C. | RADIOGRAPHIC OVERHEAD X-RAY TUBE
| 1). | kVp Accuracy and Exposure Linearity
|  | Focal Spot: | lg |  | Field Size: | Max
Radcal Accu-kV diagnostic sensor placed on table .
| @ | 20.5 | inches from target, |  |  | 6 | mAs
| Radcal Chamber at: |  |  | 15.25 | inches from target
|  | Measured |  |  | kVp Accuracy |  | Linearity | [9 empty] | Exposure time: | 0.999 | sec
| Set | Dose | Time | HV | Error | Pass/ | mR | % | [8 empty] | Ln Values |  | Reg. |  |  | Regression Output:
| kVp | mR | ms | kVp | kV | Fail | regrs. | Error | [8 empty] | kVp | Dose | Dose |  |  |  | b | a
| 50 | 185 | 99.9 | 49 | -1 |  | 181.9213 | -1.692% | [8 empty] | 3.8918 | 5.2204 | 181.9213 |  |  | Coeff Value | 2.1598 | -3.2019
| 55 | 195 | 99.5 | 52 | -3 |  | 206.8338 | 5.721% | [8 empty] | 3.9512 | 5.273 | 206.8338 |  |  | Std Error | 0.029362 | 0.12873
| 60 | 260 | 99.7 | 58 | -2 |  | 261.8475 | 0.7056% | [8 empty] | 4.0604 | 5.5607 | 261.8475 |  |  | R squared | 0.99779 | 0.032187
| 65 | 291 | 100 | 61 | -4 | FAIL | 291.979 | 0.3353% | [8 empty] | 4.1109 | 5.6733 | 291.979
| 70 | 381 | 101 | 69 | -1 |  | 381.0146 | 0.00384% | [8 empty] | 4.2341 | 5.9428 | 381.0146 |  | y = a + (b * x)
| 75 | 421 | 101 | 72 | -3 |  | 417.6974 | -0.7907% | [8 empty] | 4.2767 | 6.0426 | 417.6974 |  | a= | -3.2019
| 80 | 525 | 101 | 79 | -1 |  | 510.375 | -2.866% | [8 empty] | 4.3694 | 6.2634 | 510.375 |  | b= | 2.1598
| 85 | 560 | 100 | 83 | -2 |  | 567.8307 | 1.379% | [8 empty] | 4.4188 | 6.3279 | 567.8307
| 90 | 677 | 101 | 88 | -2 |  | 644.2987 | -5.075% | [8 empty] | 4.4773 | 6.5177 | 644.2987
| 95 | 731 | 102 | 91 | -4 | FAIL | 692.6772 | -5.533% | [8 empty] | 4.5109 | 6.5944 | 692.6772
| 100 | 832 | 100 | 99 | -1 |  | 830.9322 | -0.1285% | [8 empty] | 4.5951 | 6.7238 | 830.9322
| 110 | 1030 | 101 | 110 | 0 |  | 1043.2582 | 1.271% | [8 empty] | 4.7005 | 6.9373 | 1043.2582
| 120 | 1210 | 102 | 120 | 0 |  | 1258.9454 | 3.888% | [8 empty] | 4.7875 | 7.0984 | 1258.9454
| 130 | 1440 | 103 | 129 | -1 |  | 1471.7781 | 2.159% | [8 empty] | 4.8598 | 7.2724 | 1471.7781
|  |  |  | Max.: | -4 | kV | Max: | 5.721%
Comments:
|  |  |  | kV Accuracy: | Service recommended for the 65 & 95 kV stations.
GE's specifications for kV accuracy need to be checked.
|  |  |  | Exposure Linearity: | Exposures linear to within 6% at all kV stations tested
Acceptable results
| 2). | mA and Exposure Linearity
|  | Technique: | 60 | kV | 50 | ms
| 21 | in. target to chamber distance
Measured
| SET | Dose | Time | HV | Regression
| mA | mR | ms | kVp | mR | Error | [15 empty] | Regression Output:
| 25 | 6.68 | 55 | 58.5 | 8.4324 | -20.78% | [13 empty] | 6.7459 |  |  | b | a
| 50 | 13.2 | 52 | 59 | 17.3399 | -23.88% | [13 empty] | 6.936 |  | Coeff Value | 0.3563 | -0.47518
| 100 | 33.7 | 54 | 59.1 | 35.155 | -4.139% | [13 empty] | 7.031 |  | Std Error | 0.015556 | 4.7879
| 200 | 70.9 | 53 | 59.4 | 70.7853 | 0.1621% | [13 empty] | 7.0785 |  | R squared | 0.99056 | 7.9673
| 300 | 111 | 52 | 59.4 | 106.4155 | 4.308% | [13 empty] | 7.0944
| 400 | 155 | 53 | 59.5 | 142.0457 | 9.12% | [13 empty] | 7.1023 | y = a + (b * x)
| 600 | 203 | 55 | 59.4 | 213.3062 | -4.832% | [13 empty] | 7.1102 | a= | -0.47518
[5 empty] | Max. Error: | -23.88% | [14 empty] | b= | 0.3563
Regression Avg.:
| 7.014 | mR/mAs @ 21 inches
Comments:
|  |  | Results: | Exposures should be linear to within +/-10%
At 25 and 50 mA the exposures are too low
|  |  | Action: | SERVICE REQUIRED
| 3). | Timer and Exposure Linearity
|  | Technique: | 75 | kV | 300 | mA
| 21 | in. target to chamber distance
Measured
| SET | Dose | Time | HV | Timer | Regression
| Time | mR | ms | kVp | Error | mR | Error
| 17 | 59.5 | 18 | 74.5 | 5.882% | 77.019 | -22.75% | [12 empty] | Regression Output:
| 25 | 89.8 | 22 | 74 | -12% | 104.3019 | -13.9% | [13 empty] | b | a
| 33 | 118 | 35 | 73.7 | 6.061% | 131.5847 | -10.32% | [12 empty] | Coeff Value | 3.4104 | 19.043
| 50 | 178 | 52 | 73.6 | 4% | 189.5608 | -6.099% | [12 empty] | Std Error | 0.0198 | 8.0727
| 67 | 233 | 65 | 74.2 | -2.985% | 247.5368 | -5.873% | [12 empty] | R squared | 0.99953 | 22.8117
| 100 | 352 | 102 | 74.17 | 2% | 360.0785 | -2.244%
| 125 | 448 | 128 | 74.14 | 2.4% | 445.3374 | 0.5979% | [11 empty] | y = a + (b * x)
| 150 | 535 | 155 | 74.11 | 3.333% | 530.5963 | 0.8299% | [11 empty] | a= | 19.043
| 200 | 704 | 205 | 74.08 | 2.5% | 701.1141 | 0.4116% | [11 empty] | b= | 3.4104
| 250 | 913 | 255 | 74.05 | 2% | 871.6319 | 4.746%
| 300 | 1060 | 306 | 74.02 | 2% | 1042.1497 | 1.713%
| 400 | 1400 | 408 | 73.99 | 2% | 1383.1853 | 1.216%
| 500 | 1760 | 508 | 73.96 | 1.6% | 1724.2208 | 2.075%
| 600 | 2080 | 610 | 73.93 | 1.667% | 2065.2564 | 0.7139%
| 800 | 2700 | 815 | 73.9 | 1.875% | 2747.3276 | -1.723%
| 1000 | 3420 | 1008 | 73.87 | 0.8% | 3429.3987 | -0.2741%
Comments:
|  |  | Results: | Exposures should be linear to within +/-10%
At .0167 and .025 seconds the exposures are too low
|  |  | Action: | SERVICE REQUIRED
| 4). | Exposure Reproducibility
|  |  | TECHNIQUES USED = | 55, 65, 75, 85, 95 kV, 200 Ma, 0.1s
75 kV, 300 Ma, 4/5 s
|  |  | RESULTS= | Reproducibility was within the required +/- 5%.
Exposures varied by 8%
|  |  | COMMENT: | Acceptable results
| 5). | Half Value Layer
Technique
| kV | mA | ms | added filtration |  |  | 74.5 kV (75 kV setting) 200 mA, 2.0 mmAl Added
| 75 | 200 | 100 | 2 | mmAl |  | Calculated HVL @75.0 kV setting is 2.5 mmAl
| 15.25 | in. target to chamber distance
Measured
| mm | Dose | Time | HV | Calculated | [11 empty] | HVL (2 point Log interpolation)
| Al | mR | ms | kVp | mR/s | Atten | [12 empty] | (a)Just > | (b)Just < |  | 0.5 line data
| 0 | 218 | 100.1 | 74.5 | 2177.8222 | 1 | [9 empty] | 1 |  | Target Value | 0.5 |  |  | 0 | 0.5
| 1 | 158 | 99.8 | 74.8 | 1583.1663 | 0.72695 | [9 empty] | 0.72695 |  | Row | 3 | 4 |  | 5 | 0.5
| 2 | 122 | 99.9 | 75 | 1221.2212 | 0.56075 | [9 empty] | 0.56075 |  | Exposure Rate | 1221.2212 | 970
| 3 | 97 | 100 | 74.1 | 970 | 0.4454 | [9 empty] | 0.4454 |  | Thickness (mm) | 2 | 3
| 4 | 78.1 | 99.9 | 84.9 | 781.7818 | 0.35897 | [9 empty] | 0.35897
| 5 | 65.1 | 99.9 | 75 | 651.6517 | 0.29922 | [9 empty] | 0.29922 |  | Calc. HVL | 2.4979 | mmAl
|  | Calculated HVL: | 2.4979 | mmAl
Comments:
|  |  | Results: | Acceptable
|  |  | Action: | None Required
| 6). | Phototimer
Technique
| kV | mA | ms | cassette |  |  | 72 | inch SFD |  | 67.5 | inch SCD, Lead Apron
| 121 | 200 | 500 | 14 x 17 | in. |  |  | Results: | Chest Stand + Lead Apron
|  | Center Cell |  |  |  |  | Right Cell |  |  |  |  | Left Cell
|  | Measured |  |  |  |  | Measured |  |  |  |  | Measured
|  | Dose | Time | HV | Dose |  | Dose | Time | HV | Dose |  | Dose | Time | HV | Dose
|  | mR | ms | kVp | Ratio |  | mR | ms | kVp | Ratio |  | mR | ms | kVp | Ratio
| NORM | 44 | 77 | 121.1 | N/A |  | 31 | 52 | 121 | N/A |  | 48 | 85 | 120.8 | N/A
| 1.25 | 52 | 94 | 120.8 | 1.1818 |  | 39 | 68 | 120.8667 | 1.2581 |  | 56 | 101 | 121 | 1.1667
| 1.5 | 65 | 118 | 121 | 1.4773 |  | 47 | 77 | 120.8167 | 1.5161 |  | 72 | 134 | 120.8667 | 1.5
| 2 | 73 | 134 | 120.8667 | 1.6591 |  | 56 | 101 | 120.7667 | 1.8065 |  | 85 | 159 | 121 | 1.7708
| 0.75 | 30 | 51 | 120.8167 | 0.68182 |  | 22 | 35 | 120.7167 | 0.70968 |  | 39 | 68 | 120.8667 | 0.8125
| 0.5 | 22 | 35 | 120.7667 | 0.5 |  | 17 | 27 | 120.8667 | 0.54839 |  | 26 | 43 | 120.8167 | 0.54167
| 0.25 | 13 | 18 | 120.7167 | 0.29545 |  | 9 | 10 | 120.8167 | 0.29032 |  | 13 | 18 | 120.7667 | 0.27083
|  |  | Results: | On both normal and 3/4 density settings, Exposures are not balanced
for the three cells. There should be better agreement in % variation
among the three cells.
|  |  | Action: | SERVICE RECOMMENDED
| 7). | Light Field/X-ray Field Alignment
|  | TABLE: | At 37 inch SFD - Alignment was not within 2% of SFD. The x-ray field is 1/2 inch smaller
and 1/4 inch larger than the light field (cross table), and shifted ~ 1/4 inch longitudinally.
|  | CHEST STAND: | At 72 inch SFD, the alignment is poor: 3/4 inch of the x-ray field is cut off (or shifted)
on the left side (facing the chest unit) and ~ 3/4 inch too large on the right side.
|  | Comments: | SERVICE REQUIRED
| 8). | Centering of X-ray field to Cassette in Bucky Tray
|  | TABLE: | At 40 inch SFD, Center of x-ray field and cassette are within 2% of SFD.
|  | Comments: | Acceptable Results
|  | WALL: | Center of x-ray field and cassette were not within 2% of SFD
|  | Comments: | SERVICE REQUIRED
| 9). | Automatic Collimation
SERVICE REQUIRED, this system was not operational.
| 10). | Mechanical
| 1 | Ready light on control panel did not light
| 2 | Chest stand cassette size indicator light did not light.
| 3 | When using low technique (~ 50 - 55 kV, 5 mAs) the RESET indicator came on.
| 4 | Several times, during testing no exposure occurred when the exposure switch was engaged.
| 11). | In air exposure from clinical protocols
Measured
|  |  |  | Generator Settings |  |  | Dose | Time | HV
|  | STUDY |  | KV | mA | ms | mR | ms | kVp
|  | ABDOMEN |  | 75 | 200 | 300 | 704 | 305 | 74.8
|  |  |  | 105 | 200 | 50 | 242 | 48 | 105.1
|  |  |  | 68 | 300 | 100 | 278 | 97 | 68.7
CHEST
|  |  | a) Erect | 75 | 300 | 25 | 89 | 22 | 74.3
|  |  |  | 85 | 400 | 100 | 667 | 98 | 84.2
|  |  |  | 88 | 400 | 100 | 712 | 7 | 87.2
|  |  | b) Supine | 63 | 300 | 25 | 64 | 22 | 62.8
|  |  |  | 65 | 300 | 17 | 42 | 13 | 64.9
|  |  |  | 76 | 300 | 33 | 124 | 31 | 75.8
|  |  | LAT | 95 | 600 | 25 | 3050 | 256 | 93.7
|  |  |  | 110 | 300 | 33 | 278 | 32 | 109.6
c) PHOTOTIMED - on 3/4 Density
|  |  |  | 80 | 300 | 100 | 433 | 97 | 79.7
|  |  |  | 100 | 200 | 500 | 2240 | 506 | 100.1
HIP & PELVIS
|  |  |  | 69 | 100 | 200 | 183 | 203 | 69.2
|  |  |  | 70 | 300 | 200 | 635 | 197 | 69.9
|  |  |  | 74 | 200 | 250 | 608 | 255 | 74.2
|  |  |  | 90 | 300 | 250 | 1400 | 256 | 89.9
|  | L-SPINE |  | 75 | 200 | 250 | 585 | 255 | 75
|  |  |  | 80 | 200 | 250 | 719 | 255 | 80.1
|  |  |  | 85 | 200 | 800 | 2330 | 797 | 85
|  |  |  | 90 | 200 | 800 | 2810 | 797 | 89.8
OBSTRUCTION SERIES
|  |  |  | 78 | 300 | 150 | 604 | 147 | 77.3
|  |  |  | 80 | 300 | 400 | 1680 | 397 | 79.2
|  |  |  | 85 | 300 | 400 | 1770 | 397 | 84.4
CALIBRATION PERFORMED BY:
CERTIFIED RADIOLOGICAL PHYSICIST
Note on C17: 1. Place cursor in this cell and press "Control-Shift f" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make a fluoro exposure. 3. Wait for kV and Rate values to settle to stable numbers. 4. Press Escape to end.
Note on C159: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on C187: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on C209: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on C249: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on C271: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on G87: To Make a Fluoro Rate & kVp Measurement 1. Place cursor in this cell and press "Control-Shift-f" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make a fluoro exposure. 3. Wait for kV and Rate values to settle to stable numbers. 4. Press Escape to end. 5. Move cursor to next cell
Note on H30: To Make a Fluoro Rate & kVp Measurement 1. Place cursor in this cell and press "Control-Shift-f" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make a fluoro exposure. 3. Wait for kV and Rate values to settle to stable numbers. 4. Press Escape to end. 5. Move cursor to next cell
Note on H271: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on H325: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on I55: To Make a Fluoro Rate & kVp Measurement 1. Place cursor in this cell and press "Control-Shift-f" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make a fluoro exposure. 3. Wait for kV and Rate values to settle to stable numbers. 4. Press Escape to end. 5. Move cursor to next cell
Note on K121: To make a dose & time measurement (BE SURE THE kVp SENSOR IS IN THE BEAM): 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Press Escape to end.
Note on M271: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.

## Mobile
2026 ANNUAL CALIBRATION REPORT | [24 empty] | XLPRO ™ Version 4.06
RADIOLOGY DEPARTMENT | [24 empty] | Operation Settings Area
[7 empty] | DATE: | 2026-07-30 14:06
Units
A. | FACILITY |  |  | Almost anywhere USA | [6 empty] | Room No: | 223 |  | Other: | [10 empty] | Radiation: | R
|  | State Reg. No: |  | CA 324-916 |  |  |  | UNIT: | Radiographic Mobile | [15 empty] | /Time: | /min
[25 empty] | Scale: | milli
| EQUIPMENT |  |  | Manufacturer | [5 empty] | Model |  |  |  | Serial No. | [10 empty] | Pulse Rate: | 15-9-6-5
|  |  | CONSOLE: | General Electric |  |  |  |  | AMX-4 |  |  |  | 371294 | [11 empty] | Mode
|  |  | TUBE: | General Electric |  |  |  |  | GE MX75 |  |  |  | MX9184 | [11 empty] | Units Display: | Omit
|  |  | RATINGS: |  |  | Maximum tube kV: | 75 | kVp. Maximum tube mA: |  |  |  | 40 | mA. | Phase: | 1 | [9 empty] | Cell Update: | Incremental
[25 empty] | Auto Zero: | Disabled
[25 empty] | Capture Mode: | Manual
B. | kVp ACCURACY AND EXPOSURE LINEARITY | [23 empty] | Screen Tracking: | Yes
|  | Focal Spot: | lg |  | Field Size: | Max | [18 empty] | kVp | 9000
| Radcal Accu-kV diagnostic sensor placed on table . | [23 empty] | Config: | Dose+Time+kV
| @ | 20.5 | inches from target, |  |  | 6 | mAs | [17 empty] | Scale: | 9095
| Radcal Chamber at: |  |  | 15 | inches from target | [19 empty] | Start delay (ms): | 0
| Radcal Chamber at: |  |  | 15.25 | inches from target | [19 empty] | Width (ms): | 4914.6
|  | Measured |  |  | kV Accuracy |  | Linearity | [9 empty] | Exposure time: | 0.0999 | sec | [5 empty] | Filtration (mm-Al): | 2.5
| Set | Dose | Time | HV | Error | Pass/ | mR | % | [7 empty] | Ln Values |  | Reg. |  |  | Regression Output: |  |  |  | Dental threshold: | 0
| kVp | mR | ms | kVp | kV | Fail | regrs. | Error | [7 empty] | kVp | Dose | Dose |  |  |  | b | a |  | Device Index | 0
| 50 | 81 | 99.9 | 50.5 | 0.5 |  | 83.8508 | 3.4% | [7 empty] | 3.922 | 4.3944 | 83.8508 |  |  | Coeff Value | 2.1689 | -4.0772 |  | Monitor Configuration
| 55 | 100 | 99.5 | 55 | 0 |  | 100.9043 | 0.8962% | [7 empty] | 4.0073 | 4.6052 | 100.9043 |  |  | Std Error | 0.037879 | 0.16778 |  | Serial No: | 00008
| 60 | 125 | 99.7 | 61 | 1 |  | 126.3098 | 1.037% | [7 empty] | 4.1109 | 4.8283 | 126.3098 |  |  | R squared | 0.99635 | 0.039822 |  | ROM Ver: | 4.01
| 65 | 157 | 100 | 68 | 3 |  | 159.868 | 1.794% | [7 empty] | 4.2195 | 5.0562 | 159.868 | [6 empty] | Detector: | 6 cc
| 70 | 181 | 101 | 73 | 3 |  | 186.4629 | 2.93% | [7 empty] | 4.2905 | 5.1985 | 186.4629 |  | y = a + (b * x) |  |  |  |  | Chamb Kq: | 1.02
| 75 | 218 | 101 | 77 | 2 |  | 209.3342 | -4.14% | [7 empty] | 4.3438 | 5.3845 | 209.3342 |  | a= | -4.0772 |  |  |  | Atm Corr: | Automatic
| 80 | 248 | 101 | 83 | 3 |  | 246.3301 | -0.6779% | [7 empty] | 4.4188 | 5.5134 | 246.3301 |  | b= | 2.1689 |  |  |  | Temp (°C): | 24.1
| 85 | 284 | 100 | 86.5 | 1.5 |  | 269.4155 | -5.413% | [7 empty] | 4.4601 | 5.649 | 269.4155 | [6 empty] | Press (kPa): | 100
| 90 | 325 | 101 | 91 | 1 |  | 300.7407 | -8.067% | [7 empty] | 4.5109 | 5.7838 | 300.7407
| 95 | 330 | 102 | 95 | 0 |  | 330.1499 | 0.04541% | [7 empty] | 4.5539 | 5.7991 | 330.1499
| 100 | 404 | 100 | 103 | 3 |  | 393.4301 | -2.687% | [7 empty] | 4.6347 | 6.0014 | 393.4301
| 110 | 481 | 101 | 113 | 3 |  | 480.9999 | -2.893e-05% | [7 empty] | 4.7274 | 6.1759 | 480.9999
| 120 | 572 | 102 | 125 | 5 | FAIL | 598.7 | 4.46% | [7 empty] | 4.8283 | 6.3491 | 598.7
| 124 | 606 | 103 | 129 | 5 | FAIL | 641.0303 | 5.465% | [7 empty] | 4.8598 | 6.4069 | 641.0303
|  |  |  | Max.: | 5 | kV | Max: | -8.067%
|  |  |  | Limit ± | 3 | kV
Comments:
|  |  |  | kV Accuracy: | Results acceptable below 120kV.
SERVICE REQUIRED FOR kV STATIONS ABOVE 110kV
|  |  |  | Exposure Linearity: | Exposures linear to within 6% at all kV stations tested
Very Good Results, which are better than the other Capmobile
C. | mAs AND EXPOSURE LINEARITY
|  | Technique: | 70 | kV | 50 | ms
| 20 | in. target to chamber distance
Measured
| SET | Data | Time | HV | Regression |  | mR/
| mAs | mR | ms | kVp | mR | Error | mAs | [13 empty] | Regression Output:
| 1 | 33.4 | 55 | 70.3 | 52.6483 | -36.56% | 52.6483 | [14 empty] | b | a
| 2 | 34.6 | 52 | 70.2 | 63.9551 | -45.9% | 31.9776 | [13 empty] | Coeff Value | 11.3069 | 41.3414
| 3 | 51.2 | 54 | 70.1 | 75.262 | -31.97% | 25.0873 | [13 empty] | Std Error | 0.7334 | 10.9405
| 4 | 65.7 | 53 | 70.2 | 86.5688 | -24.11% | 21.6422 | [13 empty] | R squared | 0.94814 | 27.1034
| 5 | 82 | 52 | 70.4 | 97.8757 | -16.22% | 19.5751
| 6 | 112 | 53 | 79.2 | 109.1825 | 2.581% | 18.1971 | [12 empty] | y = a + (b * x)
| 7 | 128 | 55 | 70 | 120.4894 | 6.233% | 17.2128 | [12 empty] | a= | 41.3414
| 8 | 157 | 55 | 70.3 | 131.7962 | 19.12% | 16.4745 | [12 empty] | b= | 11.3069
| 10 | 170 | 55 | 70 | 154.4099 | 10.1% | 15.441
| 12 | 202 | 55 | 70.1 | 177.0237 | 14.11% | 14.752
| 14 | 228 | 55 | 70 | 199.6374 | 14.21% | 14.2598
| 16 | 260 | 55 | 69.9 | 222.2511 | 16.98% | 13.8907
| 22 | 311 | 55 | 70 | 290.0922 | 7.207% | 13.186
| 27 | 346 | 55 | 69.9 | 346.6265 | -0.1807% | 12.838
| 35 | 384 | 55 | 70 | 437.0813 | -12.14% | 12.488
[5 empty] | Max. Error: | -45.9%
Regression Avg.:
| 19.978 | mR/mAs @ 20 inches
Comments:
|  |  | Results: | These results are not as good as a the other Capmobile.
Even though this is a capacitor discharge unit.
Exposures should be linear to within +/- 10% from 1 to 10 or 14 mAs
|  |  | Action: | SERVICE REQUIRED
D. | EXPOSURE REPRODUCIBILITY
|  |  | TECHNIQUES USED= | 55, 70,85,100 kV, 6 mAs
75 kV, 8 mAs
70 kV, 3, 5, 10 mAs
|  |  | RESULTS= | Reproducibility was within the required +/- 5%.
Exposures varied by 8%
|  |  | Comment: | Acceptable results
E. | HALF VALUE LAYER
Technique
| kV | mAs | ms | added filtration |  |  | 75.7 kV (75 kV setting) 8 mA, 2.0 mmAl Added
| 75 | 8 | 100 | 2 | mmAl |  | Calculated HVL @75.0 kV setting is 2.7 mmAl
| 15.25 | in. target to chamber distance
HVL (2 point Log interpolation)
|  | Measured | [17 empty] | (a)Just > | (b)Just < |  | 0.5 line data
| mm | Dose | Time | HV | Calculated | [11 empty] | 1 |  | Target Value | 0.5 |  |  | 0 | 0.5
| Al | mR | ms | kVp | mR/s | Atten | [10 empty] | 0.74854 |  | Row | 3 | 4 |  | 5 | 0.5
| 0 | 171 | 100 | 75.7 | 171 | 1 | [10 empty] | 0.5848 |  | Exposure Rate | 100 | 81
| 1 | 128 | 100 | 75.9 | 128 | 0.74854 | [10 empty] | 0.47368 |  | Thickness (mm) | 2 | 3
| 2 | 100 | 100 | 76.3 | 100 | 0.5848 | [10 empty] | 0.39708
| 3 | 81 | 100 | 76.5 | 81 | 0.47368 | [10 empty] | 0.3345 |  | Calc. HVL | 2.7434 | mmAl
| 4 | 67.9 | 100 | 76.8 | 67.9 | 0.39708
| 5 | 57.2 | 100 | 76.9 | 57.2 | 0.3345
|  | Calculated HVL: | 2.7434 | mmAl
Comments:
|  |  | Results: | Acceptable
|  |  | Action: | None Required
F. | LIGHT FIELD/X-RAY FIELD ALIGNMENT
|  | Findings: | At 27 inch SFD - Alignment was not within 2% of SFD. The x-ray field is 1/2 inch smaller
than the light field, as well as shifted/
|  | Comments: | While this is in compliance, it is at the limit. Service recommended to prevent
it from being out of compliance.
G. | MECHANICAL
| 1 | Locks and light fields were working porperly.
| 2 | Audible beam on indicator needs to be louder.
H. | IN AIR EXPOSURE FROM CLINICAL PROTOCOLS
|  | Measurement Distance from Source: |  |  |  |  | 15 | inches
Measured
|  |  |  | Generator Settings |  |  | Dose | Time | HV
|  | STUDY |  | KV | mAs |  | mR | ms | kVp
|  | ABDOMEN |  | 80 | 22 |  | 756 | 274 | 78.6
|  | & C-Spine |  | 80 | 22 |  | 756 | 274 | 79
CHEST
|  |  |  | 70 | 5 |  | 142 | 14 | 69.8
|  |  |  | 75 | 8 |  | 172 | 980 | 74 | at 20 inches
|  | COLANG. |  | 90 | 35 |  | 1520 | 500 | 89.5
|  | HIP |  | 108 | 27 |  | 1720 | 461 | 107.5
|  |  |  | 85 | 27 |  | 999 | 360 | 84
|  | L-SPINE |  | 120 | 35 |  | 2590 | 618 | 119
|  |  |  | 120 | 50 |  | 3230 | 900 | 118.5
|  | PELVIS |  | 75 | 22 |  | 629 | 242 | 74.5
CALIBRATION PERFORMED BY:
CERTIFIED RADIOLOGICAL PHYSICIST
Note on C23: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on C55: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on C102: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.
Note on H135: 1. Place cursor in this cell and press "Control-Shift K" to arm the 9000. 2. When "Capturing data … appears at the bottom of the screen, make an x-ray exposure. 3. Repeat exposure as needed. 4. Press Escape to end.

## Waveforms
| Waveform Examples |  |  |  |  | Belmont Dental Single Phase 60 Hz | [5 empty] | GE DXR-750 3Phase 12-pulse
[6 empty] | 70 kVp, 10 mA, 15 pulses, SSD 55 cm | [5 empty] | 86 kVp, 300 mA, 500 ms, SSD 75 cm
XLPRO ™ Version 4.06 |  |  | 0.012603 | R | 219 | ms | 61.9 | kVp |  |  | 0.13177 | R | 495 | ms | 81.7 | kVp | 157 | mAs
Operation Settings Area |  |  | Peak= | 61.925 | kVp | PPV= | 55.9 | kV | Control= | 32834 | Peak= | 81.745 | kVp | PPV= | 78.6 | kV | Control= | 32835
|  |  | Rel Dose= | 0.55552 | /s | Average= | 60.18 | kV |  |  | Rel Dose= | 62.6202 | /s | Average= | 79.01 | kV
Units |  |  | Pk Freq= | 60.002 | Hz | Rate= | 65.241 | Hz | Avg. mA= | N/A | Pk Freq= | 120.0039 | Hz | Rate= | 0.606 | Hz | Avg. mA= | 317.3844
Radiation: |  | R | Width= | 219.1872 | ms | Pulses= | 15 |  | mAs= | N/A | Width= | 495.0528 | ms | Pulses= | 1 |  | mAs= | 157.1263
/Time: |  | /min | ms | kV | Dose/s |  |  | LINES | Hz | kV | ms | kV | Dose/s | mA |  | LINES | Hz | kV
Scale: |  |  | -0.3072 | 0 | 0 |  |  | -0.3072 | 0 | 4.4759e-16 | -0.3072 | 0 | 0 | 0 |  | -0.3072 | 0 | 2.5759e-15
Pulse Rate: |  | 0.25-9-6-5 | -0.2304 | 0 | 0 |  |  | 237.0816 | 2.1263 | 0.21054 | -0.2304 | 0 | 0 | 0 |  | 496.2816 | 3.1789 | 0.01253
Mode |  |  | -0.1536 | 0 | 0 |  |  | 61.925 | 4.2525 | 0.3573 | -0.1536 | 0 | 0 | 0 |  | 81.745 | 6.3578 | 0.036175
Units Display: |  | Rescale | -0.0768 | 0 | 0 |  |  | 61.925 | 6.3788 | 0.39401 | -0.0768 | 0 | 0 | 0 |  | 81.745 | 9.5367 | 0.059318
Cell Update: |  | Incremental | 0 | 35.91 | 0.896 |  |  | 60.18 | 8.505 | 0.31113 | 0 | 35.155 | 4.521 | 382.4518 |  | 79.01 | 12.7157 | 0.066648
Auto Zero: |  | Disabled | 0.0768 | 36.24 | 1.0803 |  |  | 60.18 | 10.6313 | 0.16067 | 0.0768 | 36.76 | 7.4957 | 393.0241 |  | 79.01 | 15.8946 | 0.045592
Capture Mode: |  | Auto | 0.1536 | 36.9 | 1.0854 |  |  | 0.55552 | 12.7575 | 0.15513 | 0.1536 | 38.775 | 10.7827 | 402.6717 |  | 62.6202 | 19.0735 | 0.006538
Screen Tracking: |  | Yes | 0.2304 | 37.345 | 1.1622 |  |  | 0.55552 | 14.8838 | 0.26778 | 0.2304 | 41.295 | 13.6448 | 411.4254 |  | 62.6202 | 22.2524 | 0.059748
kVp |  | Full | 0.3072 | 37.6 | 1.3773 |  |  | 3101 | 17.01 | 0.29222 | 0.3072 | 43.925 | 16.1382 | 419.2545 |  | 6477 | 25.4313 | 0.08359
Config: | kVSDmA | Dose+Time+kV | 0.384 | 37.685 | 1.5462 |  |  |  | 19.1363 | 0.21309 | 0.384 | 46.435 | 19.0003 | 426.1588 |  |  | 28.6102 | 0.057224
Scale: |  | 9095 | 0.4608 | 39.06 | 1.3722 |  |  | 0 | 21.2625 | 0.098789 | 0.4608 | 48.78 | 22.4051 | 432.1693 |  | 317.3844 | 31.7891 | 0.03034
Start delay (ms): |  | 0 | 0.5376 | 40.085 | 1.28 |  |  | 0 | 23.3888 | 0.096478 | 0.5376 | 51.14 | 25.6922 | 437.2551 |  | 317.3844 | 34.9681 | 0.05535
Width (ms): |  | 4914.6 | 0.6144 | 39.65 | 1.5206 |  |  |  | 25.515 | 0.12049 | 0.6144 | 53.47 | 29.0406 | 441.3853 |  |  | 38.147 | 0.031641
Filtration (mm-Al): |  | 2.5 | 0.6912 | 39.365 | 1.7818 |  |  |  | 27.6413 | 0.10041 | 0.6912 | 55.64 | 32.4762 | 444.5601 |  |  | 41.3259 | 0.067409
Dental threshold: |  | 0 | 0.768 | 40.445 | 1.5514 |  |  |  | 29.7675 | 0.12476 | 0.768 | 57.86 | 35.369 | 446.9335 |  |  | 44.5048 | 0.12564
Device Index |  | 0 | 0.8448 | 41.49 | 1.3619 |  |  |  | 31.8938 | 0.16331 | 0.8448 | 59.895 | 38.4358 | 448.5363 |  |  | 47.6837 | 0.088821
Monitor Configuration |  |  | 0.9216 | 41.995 | 1.321 |  |  |  | 34.02 | 0.14997 | 0.9216 | 61.975 | 41.0931 | 449.4302 |  |  | 50.8626 | 0.059999
Serial No: |  | 00008 | 0.9984 | 41.715 | 1.4438 |  |  |  | 36.1463 | 0.15472 | 0.9984 | 63.79 | 44.1958 | 449.7076 |  |  | 54.0415 | 0.23965
ROM Ver: |  | 4.01 | 1.0752 | 42.01 | 1.4029 |  |  |  | 38.2725 | 0.22565 | 1.0752 | 66.18 | 44.6003 | 449.3994 |  |  | 57.2205 | 0.3458
Detector: |  | 6 cc | 1.152 | 42.39 | 1.3517 |  |  |  | 40.3988 | 0.23464 | 1.152 | 67.775 | 47.4778 | 448.5363 |  |  | 60.3994 | 0.31896
Chamb Kq: |  | 1.02 | 1.2288 | 42.03 | 1.4797 |  |  |  | 42.525 | 0.13512 | 1.2288 | 69.36 | 50.2426 | 447.0568 |  |  | 63.5783 | 0.19815
Atm Corr: |  | Automatic | 1.3056 | 41.375 | 1.6845 |  |  |  | 44.6513 | 0.22144 | 1.3056 | 70.72 | 53.3043 | 445.0533 |  |  | 66.7572 | 0.08048
Temp (°C): |  | 24.9 | 1.3824 | 41.14 | 1.7459 |  |  |  | 46.7775 | 0.43879 | 1.3824 | 72.075 | 55.9974 | 442.5258 |  |  | 69.9361 | 0.014576
Press (kPa): |  | 100 | 1.4592 | 41.425 | 1.6282 |  |  |  | 48.9038 | 0.50854 | 1.4592 | 73.205 | 58.9875 | 439.5668 |  |  | 73.115 | 0.055015
|  |  | 1.536 | 41.73 | 1.5616 |  |  |  | 51.0301 | 0.4433 | 1.536 | 74.11 | 61.865 | 436.2379 |  |  | 76.2939 | 0.075071
|  |  | 1.6128 | 41.77 | 1.5616 |  |  |  | 53.1563 | 0.76886 | 1.6128 | 75.23 | 63.273 | 432.6008 |  |  | 79.4729 | 0.01271
|  |  | 1.6896 | 41.15 | 1.6794 |  |  |  | 55.2826 | 1.4813 | 1.6896 | 76.93 | 62.4384 | 428.6555 |  |  | 82.6518 | 0.1095
|  |  | 1.7664 | 41.335 | 1.5974 |  |  |  | 57.4088 | 2.1596 | 1.7664 | 77.635 | 64.7014 | 424.4019 |  |  | 85.8307 | 0.15647
|  |  | 1.8432 | 41.04 | 1.6077 |  |  |  | 59.5351 | 2.5106 | 1.8432 | 78.245 | 66.8211 | 419.8709 |  |  | 89.0096 | 0.055932
|  |  | 1.92 | 40.79 | 1.6026 |  |  |  | 61.6613 | 2.381 | 1.92 | 78.955 | 67.5174 | 415.0009 |  |  | 92.1885 | 0.14574
|  |  | 1.9968 | 40.46 | 1.5923 |  |  |  | 63.7876 | 1.8024 | 1.9968 | 79.48 | 68.3418 | 409.8226 |  |  | 95.3674 | 0.25736
|  |  | 2.0736 | 39.77 | 1.7101 |  |  |  | 65.9138 | 0.9829 | 2.0736 | 79.75 | 69.3248 | 404.3978 |  |  | 98.5463 | 0.12957
|  |  | 2.1504 | 39.685 | 1.6282 |  |  |  | 68.0401 | 0.30711 | 2.1504 | 80 | 69.7651 | 398.788 |  |  | 101.7253 | 0.2073
|  |  | 2.2272 | 39.31 | 1.5002 |  |  |  | 70.1663 | 0.4846 | 2.2272 | 80.125 | 70.2874 | 393.0241 |  |  | 104.9042 | 0.47737
|  |  | 2.304 | 38.99 | 1.2851 |  |  |  | 72.2926 | 0.60231 | 2.304 | 80.08 | 70.7686 | 387.1985 |  |  | 108.0831 | 0.34605
|  |  | 2.3808 | 39.16 | 1.0035 |  |  |  | 74.4188 | 0.40943 | 2.3808 | 80.12 | 70.4512 | 381.2805 |  |  | 111.262 | 0.33573
|  |  | 2.4576 | 38.305 | 0.90624 |  |  |  | 76.5451 | 0.071841 | 2.4576 | 80.005 | 70.1286 | 375.3317 |  |  | 114.4409 | 1.3548
|  |  | 2.5344 | 36.525 | 0.93696 |  |  |  | 78.6713 | 0.24258 | 2.5344 | 79.725 | 70.4051 | 369.2596 |  |  | 117.6198 | 2.2184
|  |  | 2.6112 | 36.83 | 0.67072 |  |  |  | 80.7976 | 0.35729 | 2.6112 | 79.355 | 70.6253 | 363.0641 |  |  | 120.7987 | 2.4489
|  |  | 2.688 | 36.17 | 0.60928 |  |  |  | 82.9238 | 0.27182 | 2.688 | 79.235 | 68.9306 | 356.807 |  |  | 123.9777 | 1.9022
|  |  | 2.7648 | 0 | 0 |  |  |  | 85.0501 | 0.092059 | 2.7648 | 79.04 | 67.0054 | 350.5192 |  |  | 127.1566 | 0.87349
|  |  | 2.8416 | 0 | 0 |  |  |  | 87.1763 | 0.1534 | 2.8416 | 78.38 | 66.4064 | 344.3854 |  |  | 130.3355 | 0.097375
|  |  | 2.9184 | 0 | 0 |  |  |  | 89.3026 | 0.22816 | 2.9184 | 77.725 | 65.7869 | 339.8852 |  |  | 133.5144 | 0.55312
|  |  | 2.9952 | 0 | 0 |  |  |  | 91.4288 | 0.18391 | 2.9952 | 77.525 | 64.809 | 338.7448 |  |  | 136.6933 | 0.42777
|  |  | 3.072 | 0 | 0 |  |  |  | 93.5551 | 0.095433 | 3.072 | 77.215 | 66.135 | 341.7346 |  |  | 139.8722 | 0.0066734
|  |  | 3.1488 | 0 | 0 |  |  |  | 95.6813 | 0.1062 | 3.1488 | 77.07 | 66.9594 | 349.1938 |  |  | 143.0511 | 0.31899
|  |  | 3.2256 | 0 | 0 |  |  |  | 97.8076 | 0.11849 | 3.2256 | 77.14 | 65.7971 | 359.9818 |  |  | 146.2301 | 0.31413
|  |  | 3.3024 | 0 | 0 |  |  |  | 99.9339 | 0.068473 | 3.3024 | 76.99 | 65.4541 | 371.3247 |  |  | 149.409 | 0.057166
|  |  | 3.3792 | 0 | 0 |  |  |  | 102.0601 | 0.049674 | 3.3792 | 76.795 | 64.553 | 380.1092 |  |  | 152.5879 | 0.19659
|  |  | 3.456 | 0 | 0 |  |  |  | 104.1864 | 0.080128 | 3.456 | 76.28 | 64.6144 | 384.3628 |  |  | 155.7668 | 0.23732
|  |  | 3.5328 | 0 | 0 |  |  |  | 106.3126 | 0.085295 | 3.5328 | 75.9 | 64.2867 | 383.7463 |  |  | 158.9457 | 0.067972
|  |  | 3.6096 | 0 | 0 |  |  |  | 108.4389 | 0.16651 | 3.6096 | 75.88 | 63.0323 | 379.277 |  |  | 162.1246 | 0.13534
|  |  | 3.6864 | 0 | 0 |  |  |  | 110.5651 | 0.2999 | 3.6864 | 75.95 | 62.2592 | 373.1741 |  |  | 165.3035 | 0.19404
|  |  | 3.7632 | 0 | 0 |  |  |  | 112.6914 | 0.4459 | 3.7632 | 76.015 | 61.9008 | 367.5334 |  |  | 168.4825 | 0.078917
|  |  | 3.84 | 0 | 0 |  |  |  | 114.8176 | 0.63682 | 3.84 | 75.91 | 61.9827 | 363.5573 |  |  | 171.6614 | 0.087333
|  |  | 3.9168 | 0 | 0 |  |  |  | 116.9439 | 0.87713 | 3.9168 | 75.65 | 62.0493 | 361.3997 |  |  | 174.8403 | 0.17262
|  |  | 3.9936 | 0 | 0 |  |  |  | 119.0701 | 1.0777 | 3.9936 | 75.225 | 61.3478 | 360.3825 |  |  | 178.0192 | 0.21283
|  |  | 4.0704 | 0 | 0 |  |  |  | 121.1964 | 1.1199 | 4.0704 | 74.915 | 58.8954 | 358.7181 |  |  | 181.1981 | 0.30131
|  |  | 4.1472 | 0 | 0 |  |  |  | 123.3226 | 0.94029 | 4.1472 | 74.29 | 57.0726 | 354.8652 |  |  | 184.377 | 0.31104
|  |  | 4.224 | 0 | 0 |  |  |  | 125.4489 | 0.57565 | 4.224 | 73.6 | 55.4445 | 348.2691 |  |  | 187.5559 | 0.16232
|  |  | 4.3008 | 0 | 0 |  |  |  | 127.5751 | 0.18889 | 4.3008 | 73.135 | 52.9971 | 339.053 |  |  | 190.7349 | 0.05944
|  |  | 4.3776 | 0 | 0 |  |  |  | 129.7014 | 0.29555 | 4.3776 | 72.47 | 52.6746 | 328.4499 |  |  | 193.9138 | 0.1932
|  |  | 4.4544 | 0 | 0 |  |  |  | 131.8276 | 0.44813 | 4.4544 | 71.775 | 52.4646 | 317.7851 |  |  | 197.0927 | 0.14852
|  |  | 4.5312 | 0 | 0 |  |  |  | 133.9539 | 0.39694 | 4.5312 | 71.415 | 52.7514 | 309.3396 |  |  | 200.2716 | 0.035891
|  |  | 4.608 | 0 | 0 |  |  |  | 136.0801 | 0.2019 | 4.608 | 71.53 | 53.1149 | 308.5382 |  |  | 203.4505 | 0.16622
|  |  | 4.6848 | 0 | 0 |  |  |  | 138.2064 | 0.05542 | 4.6848 | 71.62 | 54.2874 | 306.6888 |  |  | 206.6294 | 0.15846
|  |  | 4.7616 | 0 | 0 |  |  |  | 140.3326 | 0.17007 | 4.7616 | 71.74 | 53.9443 | 299.9078 |  |  | 209.8083 | 0.016224
|  |  | 4.8384 | 0 | 0 |  |  |  | 142.4589 | 0.16929 | 4.8384 | 71.57 | 53.5603 | 293.5891 |  |  | 212.9873 | 0.13219
|  |  | 4.9152 | 0 | 0 |  |  |  | 144.5851 | 0.073797 | 4.9152 | 71.485 | 51.6659 | 288.6266 |  |  | 216.1662 | 0.15557
|  |  | 4.992 | 0 | 0 |  |  |  | 146.7114 | 0.095647 | 4.992 | 70.82 | 52.3418 | 284.9278 |  |  | 219.3451 | 0.032815
|  |  | 5.0688 | 0 | 0 |  |  |  | 148.8377 | 0.16257 | 5.0688 | 70.535 | 52.0243 | 282.5236 |  |  | 222.524 | 0.13193
|  |  | 5.1456 | 0 | 0 |  |  |  | 150.9639 | 0.13698 | 5.1456 | 70.525 | 51.4662 | 281.1058 |  |  | 225.7029 | 0.18862
|  |  | 5.2224 | 0 | 0 |  |  |  | 153.0902 | 0.036516 | 5.2224 | 70.655 | 51.2614 | 280.4893 |  |  | 228.8818 | 0.067605
|  |  | 5.2992 | 0 | 0 |  |  |  | 155.2164 | 0.078231 | 5.2992 | 70.94 | 51.4202 | 280.3352 |  |  | 232.0608 | 0.1746
|  |  | 5.376 | 0 | 0 |  |  |  | 157.3427 | 0.14121 | 5.376 | 71.135 | 52.1421 | 280.705 |  |  | 235.2397 | 0.39648
|  |  | 5.4528 | 0 | 0 |  |  |  | 159.4689 | 0.1324 | 5.4528 | 71.15 | 52.9254 | 281.5681 |  |  | 238.4186 | 0.48114
|  |  | 5.5296 | 0 | 0 |  |  |  | 161.5952 | 0.090996 | 5.5296 | 71.115 | 52.8589 | 282.7394 |  |  | 241.5975 | 0.40604
|  |  | 5.6064 | 0 | 0 |  |  |  | 163.7214 | 0.082737 | 5.6064 | 71.01 | 52.2906 | 286.3457 |  |  | 244.7764 | 0.23915
|  |  | 5.6832 | 0 | 0 |  |  |  | 165.8477 | 0.067969 | 5.6832 | 70.765 | 51.8246 | 300.3701 |  |  | 247.9553 | 0.074543
|  |  | 5.76 | 0 | 0 |  |  |  | 167.9739 | 0.055696 | 5.76 | 70.54 | 51.3741 | 321.6996 |  |  | 251.1342 | 0.032313
|  |  | 5.8368 | 0 | 0 |  |  |  | 170.1002 | 0.16741 | 5.8368 | 70.4 | 51.8605 | 339.2688 |  |  | 254.3132 | 0.064253
|  |  | 5.9136 | 0 | 0 |  |  |  | 172.2264 | 0.29564 | 5.9136 | 70.46 | 53.289 | 350.2418 |  |  | 257.4921 | 0.040581
|  |  | 5.9904 | 0 | 0 |  |  |  | 174.3527 | 0.40419 | 5.9904 | 71.005 | 53.9085 | 355.7591 |  |  | 260.671 | 0.015381
|  |  | 6.0672 | 0 | 0 |  |  |  | 176.4789 | 0.50149 | 6.0672 | 72.09 | 53.6166 | 358.3482 |  |  | 263.8499 | 0.054044
|  |  | 6.144 | 0 | 0 |  |  |  | 178.6052 | 0.58619 | 6.144 | 73.115 | 53.673 | 360.7832 |  |  | 267.0288 | 0.056289
|  |  | 6.2208 | 0 | 0 |  |  |  | 180.7314 | 0.61098 | 6.2208 | 73.69 | 56.8576 | 364.9443 |  |  | 270.2077 | 0.028918
|  |  | 6.2976 | 0 | 0 |  |  |  | 182.8577 | 0.52148 | 6.2976 | 74.465 | 57.6154 | 371.2014 |  |  | 273.3866 | 0.029022
|  |  | 6.3744 | 0 | 0 |  |  |  | 184.9839 | 0.3139 | 6.3744 | 75.07 | 57.344 | 378.3831 |  |  | 276.5656 | 0.030793
|  |  | 6.4512 | 0 | 0 |  |  |  | 187.1102 | 0.10326 | 6.4512 | 75.195 | 57.9533 | 384.7327 |  |  | 279.7445 | 0.013294
|  |  | 6.528 | 0 | 0 |  |  |  | 189.2364 | 0.2535 | 6.528 | 75.015 | 58.8134 | 388.2157 |  |  | 282.9234 | 0.034445
|  |  | 6.6048 | 0 | 0 |  |  |  | 191.3627 | 0.37317 | 6.6048 | 74.6 | 60.375 | 387.5992 |  |  | 286.1023 | 0.038606
|  |  | 6.6816 | 0 | 0 |  |  |  | 193.4889 | 0.34372 | 6.6816 | 74.605 | 60.0166 | 383.0066 |  |  | 289.2812 | 0.017653
|  |  | 6.7584 | 0 | 0 |  |  |  | 195.6152 | 0.20358 | 6.7584 | 74.76 | 59.4227 | 375.6399 |  |  | 292.4601 | 0.021067
|  |  | 6.8352 | 0 | 0 |  |  |  | 197.7415 | 0.10813 | 6.8352 | 74.82 | 59.5251 | 367.2252 |  |  | 295.639 | 0.018121
|  |  | 6.912 | 0 | 0 |  |  |  | 199.8677 | 0.17716 | 6.912 | 74.745 | 59.8221 | 359.6428 |  |  | 298.818 | 0.030488
|  |  | 6.9888 | 0 | 0 |  |  |  | 201.994 | 0.18154 | 6.9888 | 74.395 | 60.841 | 353.8789 |  |  | 301.9969 | 0.071057
|  |  | 7.0656 | 0 | 0 |  |  |  | 204.1202 | 0.089341 | 7.0656 | 74.245 | 59.9347 | 349.9027 |  |  | 305.1758 | 0.078121
|  |  | 7.1424 | 0 | 0 |  |  |  | 206.2465 | 0.061103 | 7.1424 | 74.005 | 58.327 | 346.7279 |  |  | 308.3547 | 0.037977
|  |  | 7.2192 | 0 | 0 |  |  |  | 208.3727 | 0.15649 | 7.2192 | 73.725 | 56.3456 | 342.9984 |  |  | 311.5336 | 0.022609
|  |  | 7.296 | 0 | 0 |  |  |  | 210.499 | 0.16993 | 7.296 | 73.3 | 55.1731 | 337.6043 |  |  | 314.7125 | 0.056333
|  |  | 7.3728 | 0 | 0 |  |  |  | 212.6252 | 0.10019 | 7.3728 | 73.015 | 54.8915 | 331.1623 |  |  | 317.8914 | 0.037916
|  |  | 7.4496 | 0 | 0 |  |  |  | 214.7515 | 0.035828 | 7.4496 | 73.17 | 55.7466 | 323.7648 |  |  | 321.0704 | 0.019547
|  |  | 7.5264 | 0 | 0 |  |  |  | 216.8777 | 0.10203 | 7.5264 | 73.6 | 57.8662 | 316.0282 |  |  | 324.2493 | 0.059607
|  |  | 7.6032 | 0 | 0 |  |  |  | 219.004 | 0.11366 | 7.6032 | 74.595 | 57.9891 | 309.3396 |  |  | 327.4282 | 0.047415
|  |  | 7.68 | 0 | 0 |  |  |  | 221.1302 | 0.062654 | 7.68 | 75.43 | 58.6803 | 304.0073 |  |  | 330.6071 | 0.0168
|  |  | 7.7568 | 0 | 0 |  |  |  | 223.2565 | 0.041467 | 7.7568 | 75.805 | 60.5184 | 299.9078 |  |  | 333.786 | 0.075002
|  |  | 7.8336 | 0 | 0 |  |  |  | 225.3827 | 0.083315 | 7.8336 | 76.155 | 61.2301 | 296.8255 |  |  | 336.9649 | 0.069554
|  |  | 7.9104 | 0 | 0 |  |  |  | 227.509 | 0.082099 | 7.9104 | 76.485 | 60.6669 | 294.5754 |  |  | 340.1438 | 0.019359
|  |  | 7.9872 | 0 | 0 |  |  |  | 229.6352 | 0.10271 | 7.9872 | 76.435 | 60.9997 | 292.9726 |  |  | 343.3228 | 0.13127
|  |  | 8.064 | 0 | 0 |  |  |  | 231.7615 | 0.19313 | 8.064 | 76.14 | 61.7779 | 291.863 |  |  | 346.5017 | 0.15904
|  |  | 8.1408 | 0 | 0 |  |  |  | 233.8877 | 0.2766 | 8.1408 | 76.1 | 61.1328 | 290.9691 |  |  | 349.6806 | 0.024772
|  |  | 8.2176 | 0 | 0 |  |  |  | 236.014 | 0.31785 | 8.2176 | 76.15 | 60.201 | 290.1061 |  |  | 352.8595 | 0.29653
|  |  | 8.2944 | 0 | 0 |  |  |  | 238.1402 | 0.32485 | 8.2944 | 76.185 | 59.991 | 289.0889 |  |  | 356.0384 | 0.64248
|  |  | 8.3712 | 0 | 0 |  |  |  | 240.2665 | 0.31218 | 8.3712 | 76.4 | 59.5558 | 288.0409 |  |  | 359.2173 | 0.84815
|  |  | 8.448 | 0 | 0 |  |  |  | 242.3927 | 0.26643 | 8.448 | 76.475 | 59.392 | 287.2395 |  |  | 362.3962 | 0.79225
|  |  | 8.5248 | 0 | 0 |  |  |  | 244.519 | 0.17006 | 8.5248 | 76.415 | 59.863 | 287.8868 |  |  | 365.5752 | 0.49523
|  |  | 8.6016 | 0 | 0 |  |  |  | 246.6452 | 0.078487 | 8.6016 | 76.64 | 60.1344 | 291.678 |  |  | 368.7541 | 0.11307
|  |  | 8.6784 | 0 | 0 |  |  |  | 248.7715 | 0.1609 | 8.6784 | 76.915 | 61.1789 | 299.2913 |  |  | 371.933 | 0.16019
|  |  | 8.7552 | 0 | 0 |  |  |  | 250.8978 | 0.2326 | 8.7552 | 77.16 | 61.6397 | 311.0965 |  |  | 375.1119 | 0.21515
|  |  | 8.832 | 0 | 0 |  |  |  | 253.024 | 0.21344 | 8.832 | 77.235 | 62.2234 | 325.984 |  |  | 378.2908 | 0.088511
|  |  | 8.9088 | 0 | 0 |  |  |  | 255.1503 | 0.11907 | 8.9088 | 77.53 | 62.4794 | 341.4264 |  |  | 381.4697 | 0.078676
|  |  | 8.9856 | 0 | 0 |  |  |  | 257.2765 | 0.080382 | 8.9856 | 77.705 | 63.2269 | 354.2487 |  |  | 384.6486 | 0.1507
|  |  | 9.0624 | 0 | 0 |  |  |  | 259.4028 | 0.15336 | 9.0624 | 77.475 | 64.8755 | 362.1086 |  |  | 387.8276 | 0.090829
|  |  | 9.1392 | 0 | 0 |  |  |  | 261.529 | 0.16315 | 9.1392 | 77.585 | 64.7168 | 364.3895 |  |  | 391.0065 | 0.029267
|  |  | 9.216 | 0 | 0 |  |  |  | 263.6553 | 0.092266 | 9.216 | 77.795 | 63.959 | 362.2627 |  |  | 394.1854 | 0.10021
|  |  | 9.2928 | 0 | 0 |  |  |  | 265.7815 | 0.020221 | 9.2928 | 77.88 | 64.2202 | 357.7934 |  |  | 397.3643 | 0.06963
|  |  | 9.3696 | 0 | 0 |  |  |  | 267.9078 | 0.10562 | 9.3696 | 77.95 | 64.9011 | 353.1699 |  |  | 400.5432 | 0.024585
|  |  | 9.4464 | 0 | 0 |  |  |  | 270.034 | 0.12404 | 9.4464 | 78.135 | 64.3994 | 349.8411 |  |  | 403.7221 | 0.095013
|  |  | 9.5232 | 0 | 0 |  |  |  | 272.1603 | 0.070025 | 9.5232 | 78.13 | 63.8515 | 348.3307 |  |  | 406.901 | 0.083052
|  |  | 9.6 | 0 | 0 |  |  |  | 274.2865 | 0.028161 | 9.6 | 77.88 | 63.3498 | 347.9608 |  |  | 410.08 | 0.0073997
|  |  | 9.6768 | 0 | 0 |  |  |  | 276.4128 | 0.094558 | 9.6768 | 77.545 | 62.4384 | 348.2074 |  |  | 413.2589 | 0.068854
|  |  | 9.7536 | 0 | 0 |  |  |  | 278.539 | 0.10187 | 9.7536 | 77.11 | 61.4554 | 347.8376 |  |  | 416.4378 | 0.078787
|  |  | 9.8304 | 0 | 0 |  |  |  | 280.6653 | 0.046539 | 9.8304 | 76.865 | 59.5507 | 344.7244 |  |  | 419.6167 | 0.032989
|  |  | 9.9072 | 0 | 0 |  |  |  | 282.7915 | 0.062059 | 9.9072 | 76.48 | 58.7571 | 338.6215 |  |  | 422.7956 | 0.04971
|  |  | 9.984 | 0 | 0 |  |  |  | 284.9178 | 0.12854 | 9.984 | 76.21 | 59.0029 | 330.2684 |  |  | 425.9745 | 0.061374
|  |  | 10.0608 | 0 | 0 |  |  |  | 287.044 | 0.13463 | 10.0608 | 76.46 | 60.0013 | 320.1893 |  |  | 429.1534 | 0.020499
|  |  | 10.1376 | 0 | 0 |  |  |  | 289.1703 | 0.085561 | 10.1376 | 77.23 | 60.7027 | 310.665 |  |  | 432.3324 | 0.040564
|  |  | 10.2144 | 0 | 0 |  |  |  | 291.2965 | 0.10359 | 10.2144 | 78.065 | 60.7642 | 310.2951 |  |  | 435.5113 | 0.058558
|  |  | 10.2912 | 0 | 0 |  |  |  | 293.4228 | 0.19109 | 10.2912 | 78.495 | 61.1174 | 311.9288 |  |  | 438.6902 | 0.025361
|  |  | 10.368 | 0 | 0 |  |  |  | 295.549 | 0.24278 | 10.368 | 78.38 | 62.6176 | 306.4423 |  |  | 441.8691 | 0.027792
|  |  | 10.4448 | 0 | 0 |  |  |  | 297.6753 | 0.24368 | 10.4448 | 78.155 | 63.8771 | 300.0002 |  |  | 445.048 | 0.049764
|  |  | 10.5216 | 0 | 0 |  |  |  | 299.8016 | 0.21807 | 10.5216 | 78.215 | 63.1398 | 293.9898 |  |  | 448.2269 | 0.024588
|  |  | 10.5984 | 0 | 0 |  |  |  | 301.9278 | 0.18413 | 10.5984 | 78.35 | 62.1773 | 288.6882 |  |  | 451.4058 | 0.022134
|  |  | 10.6752 | 0 | 0 |  |  |  | 304.0541 | 0.12778 | 10.6752 | 78.475 | 61.8394 | 284.3113 |  |  | 454.5848 | 0.04753
|  |  | 10.752 | 0 | 0 |  |  |  | 306.1803 | 0.050104 | 10.752 | 78.485 | 62.7046 | 280.9517 |  |  | 457.7637 | 0.032544
|  |  | 10.8288 | 0 | 0 |  |  |  | 308.3066 | 0.094263 | 10.8288 | 78.645 | 63.0426 | 278.4242 |  |  | 460.9426 | 0.0048003
|  |  | 10.9056 | 0 | 0 |  |  |  | 310.4328 | 0.16559 | 10.9056 | 78.72 | 63.488 | 276.5439 |  |  | 464.1215 | 0.024305
|  |  | 10.9824 | 0 | 0 |  |  |  | 312.5591 | 0.17019 | 10.9824 | 78.66 | 63.8259 | 275.2494 |  |  | 467.3004 | 0.016761
|  |  | 11.0592 | 0 | 0 |  |  |  | 314.6853 | 0.10075 | 11.0592 | 78.58 | 63.0374 | 274.448 |  |  | 470.4793 | 0.0084211
|  |  | 11.136 | 0 | 0 |  |  |  | 316.8116 | 0.034829 | 11.136 | 78.145 | 62.5715 | 273.8007 |  |  | 473.6582 | 0.01635
|  |  | 11.2128 | 0 | 0 |  |  |  | 318.9378 | 0.12426 | 11.2128 | 77.795 | 61.2659 | 273.8315 |  |  | 476.8372 | 0.073132
|  |  | 11.2896 | 0 | 0 |  |  |  | 321.0641 | 0.16589 | 11.2896 | 77.355 | 59.8733 | 280.6742 |  |  | 480.0161 | 0.14087
|  |  | 11.3664 | 0 | 0 |  |  |  | 323.1903 | 0.13304 | 11.3664 | 76.795 | 59.3715 | 297.6577 |  |  | 483.195 | 0.16529
|  |  | 11.4432 | 0 | 0 |  |  |  | 325.3166 | 0.048716 | 11.4432 | 76.495 | 58.8595 | 314.3329 |  |  | 486.3739 | 0.11809
|  |  | 11.52 | 0 | 0 |  |  |  | 327.4428 | 0.040206 | 11.52 | 76.305 | 59.1411 | 325.4601 |  |  | 489.5528 | 0.024031
|  |  | 11.5968 | 0 | 0 |  |  |  | 329.5691 | 0.087646 | 11.5968 | 76.455 | 59.8477 | 331.7788 |  |  | 492.7317 | 0.057145
|  |  | 11.6736 | 0 | 0 |  |  |  | 331.6953 | 0.08049 | 11.6736 | 76.96 | 60.6822 | 335.3234 |  |  | 495.9106 | 0.074451
|  |  | 11.7504 | 0 | 0 |  |  |  | 333.8216 | 0.056489 | 11.7504 | 77.61 | 61.6653 | 338.4057 |  |  | 499.0896 | 0.031345
|  |  | 11.8272 | 0 | 0 |  |  |  | 335.9478 | 0.077746 | 11.8272 | 78.195 | 62.6125 | 342.8134 |  |  | 502.2685 | 0.040113
|  |  | 11.904 | 0 | 0 |  |  |  | 338.0741 | 0.089932 | 11.904 | 78.64 | 63.2832 | 349.1629 |  |  | 505.4474 | 0.064223
|  |  | 11.9808 | 0 | 0 |  |  |  | 340.2003 | 0.06074 | 11.9808 | 78.84 | 63.8925 | 356.6221 |  |  | 508.6263 | 0.037549
|  |  | 12.0576 | 0 | 0 |  |  |  | 342.3266 | 0.041701 | 12.0576 | 78.735 | 64.3891 | 363.4648 |  |  | 511.8052 | 0.01317
|  |  | 12.1344 | 0 | 0 |  |  |  | 344.4528 | 0.080533 | 12.1344 | 78.59 | 63.9283 | 367.9341 |  |  | 514.9841 | 0.040613
|  |  | 12.2112 | 0 | 0 |  |  |  | 346.5791 | 0.083719 | 12.2112 | 78.295 | 63.6621 | 368.7972 |  |  | 518.163 | 0.027241
|  |  | 12.288 | 0 | 0 |  |  |  | 348.7054 | 0.027553 | 12.288 | 78.025 | 63.2115 | 365.8382 |  |  | 521.342 | 0.0098326
|  |  | 12.3648 | 0 | 0 |  |  |  | 350.8316 | 0.068984 | 12.3648 | 78.15 | 61.4042 | 359.8893 |  |  | 524.5209 | 0.036679
|  |  | 12.4416 | 0 | 0 |  |  |  | 352.9579 | 0.16396 | 12.4416 | 77.94 | 61.7882 | 352.4918 |  |  | 527.6998 | 0.032815
|  |  | 12.5184 | 0 | 0 |  |  |  | 355.0841 | 0.21768 | 12.5184 | 77.645 | 62.3258 | 345.3717 |  |  | 530.8787 | 0.0072963
|  |  | 12.5952 | 0 | 0 |  |  |  | 357.2104 | 0.21457 | 12.5952 | 77.275 | 62.9094 | 339.577 |  |  | 534.0576 | 0.069171
|  |  | 12.672 | 0 | 0 |  |  |  | 359.3366 | 0.17018 | 12.672 | 76.715 | 63.1501 | 335.2926 |  |  | 537.2365 | 0.14779
|  |  | 12.7488 | 0 | 0 |  |  |  | 361.4629 | 0.11656 | 12.7488 | 76.29 | 61.312 | 331.8096 |  |  | 540.4154 | 0.20244
|  |  | 12.8256 | 0 | 0 |  |  |  | 363.5891 | 0.074497 | 12.8256 | 75.45 | 60.6515 | 327.7101 |  |  | 543.5944 | 0.19095
|  |  | 12.9024 | 0 | 0 |  |  |  | 365.7154 | 0.060364 | 12.9024 | 74.675 | 59.7965 | 322.0079 |  |  | 546.7733 | 0.10663
|  |  | 12.9792 | 0 | 0 |  |  |  | 367.8416 | 0.092641 | 12.9792 | 74.765 | 56.3046 | 315.2268 |  |  | 549.9522 | 0.0068948
|  |  | 13.056 | 0 | 0 |  |  |  | 369.9679 | 0.11969 | 13.056 | 74.92 | 55.9565 | 307.6443 |  |  | 553.1311 | 0.078257
|  |  | 13.1328 | 0 | 0 |  |  |  | 372.0941 | 0.10078 | 13.1328 | 75.375 | 56.4736 | 299.9078 |  |  | 556.31 | 0.070476
|  |  | 13.2096 | 0 | 0 |  |  |  | 374.2204 | 0.041456 | 13.2096 | 75.905 | 57.1238 | 293.4041 |  |  | 559.4889 | 0.012496
|  |  | 13.2864 | 0 | 0 |  |  |  | 376.3466 | 0.088151 | 13.2864 | 76.125 | 59.1309 | 288.5033 |  |  | 562.6678 | 0.068285
|  |  | 13.3632 | 0 | 0 |  |  |  | 378.4729 | 0.16224 | 13.3632 | 76.205 | 60.9178 | 285.1128 |  |  | 565.8468 | 0.079474
|  |  | 13.44 | 0 | 0 |  |  |  | 380.5991 | 0.17595 | 13.44 | 76.225 | 62.0493 | 282.986 |  |  | 569.0257 | 0.032166
|  |  | 13.5168 | 0 | 0 |  |  |  | 382.7254 | 0.11681 | 13.5168 | 76.325 | 61.3837 | 281.8147 |  |  | 572.2046 | 0.027192
|  |  | 13.5936 | 0 | 0 |  |  |  | 384.8516 | 0.033577 | 13.5936 | 76.35 | 59.8323 | 281.3832 |  |  | 575.3835 | 0.045338
|  |  | 13.6704 | 0 | 0 |  |  |  | 386.9779 | 0.10475 | 13.6704 | 76.285 | 58.1581 | 281.3832 |  |  | 578.5624 | 0.012363
|  |  | 13.7472 | 0 | 0 |  |  |  | 389.1041 | 0.14904 | 13.7472 | 76.035 | 57.4464 | 281.6606 |  |  | 581.7413 | 0.04575
|  |  | 13.824 | 0 | 0 |  |  |  | 391.2304 | 0.11606 | 13.824 | 75.745 | 57.5283 | 282.0921 |  |  | 584.9202 | 0.065897
|  |  | 13.9008 | 0 | 0 |  |  |  | 393.3566 | 0.027302 | 13.9008 | 75.315 | 59.0131 | 282.6777 |  |  | 588.0992 | 0.03532
|  |  | 13.9776 | 0 | 0 |  |  |  | 395.4829 | 0.087187 | 13.9776 | 75.2 | 59.1718 | 283.6024 |  |  | 591.2781 | 0.019046
|  |  | 14.0544 | 0 | 0 |  |  |  | 397.6092 | 0.1506 | 14.0544 | 75.325 | 58.5728 | 286.0374 |  |  | 594.457 | 0.051528
|  |  | 14.1312 | 0 | 0 |  |  |  | 399.7354 | 0.1429 | 14.1312 | 75.585 | 59.2026 | 291.7397 |  |  | 597.6359 | 0.06507
|  |  | 14.208 | 0 | 0 |  |  |  | 401.8617 | 0.072017 | 14.208 | 76.025 | 60.544 | 301.1407 |  |  | 600.8148 | 0.095763
|  |  | 14.2848 | 0 | 0 |  |  |  | 403.9879 | 0.02137 | 14.2848 | 76.58 | 60.9331 | 314.5487 |  |  | 603.9937 | 0.10839
|  |  | 14.3616 | 0 | 0 |  |  |  | 406.1142 | 0.087299 | 14.3616 | 76.81 | 61.568 | 330.9774 |  |  | 607.1726 | 0.066553
|  |  | 14.4384 | 0 | 0 |  |  |  | 408.2404 | 0.095694 | 14.4384 | 76.605 | 64.3482 | 347.93 |  |  | 610.3516 | 0.017845
|  |  | 14.5152 | 0 | 0 |  |  |  | 410.3667 | 0.05963 | 14.5152 | 76.66 | 65.4797 | 362.4785 |  |  | 613.5305 | 0.068504
|  |  | 14.592 | 0 | 0 |  |  |  | 412.4929 | 0.06998 | 14.592 | 76.795 | 64.7834 | 372.311 |  |  | 616.7094 | 0.060411
|  |  | 14.6688 | 0 | 0 |  |  |  | 414.6192 | 0.11377 | 14.6688 | 76.74 | 65.2032 | 376.5338 |  |  | 619.8883 | 0.0029991
|  |  | 14.7456 | 0 | 0 |  |  |  | 416.7454 | 0.11497 | 14.7456 | 76.745 | 66.2221 | 376.1022 |  |  | 623.0672 | 0.058527
|  |  | 14.8224 | 0 | 0 |  |  |  | 418.8717 | 0.066857 | 14.8224 | 77.215 | 66.2477 | 372.8967 |  |  | 626.2461 | 0.064498
|  |  | 14.8992 | 0 | 0 |  |  |  | 420.9979 | 0.012956 | 14.8992 | 77.93 | 65.6384 | 369.1979 |  |  | 629.425 | 0.015312
|  |  | 14.976 | 0 | 0 |  |  |  | 423.1242 | 0.065385 | 14.976 | 78.18 | 66.6573 | 366.7321 |  |  | 632.604 | 0.044193
|  |  | 15.0528 | 0 | 0 |  |  |  | 425.2504 | 0.09119 | 15.0528 | 78.34 | 66.9133 | 366.1772 |  |  | 635.7829 | 0.058342
|  |  | 15.1296 | 0 | 0 |  |  |  | 427.3767 | 0.10011 | 15.1296 | 78.265 | 66.8416 | 366.9786 |  |  | 638.9618 | 0.014944
|  |  | 15.2064 | 0 | 0 |  |  |  | 429.5029 | 0.11667 | 15.2064 | 77.755 | 67.2256 | 367.8109 |  |  | 642.1407 | 0.047051
|  |  | 15.2832 | 0 | 0 |  |  |  | 431.6292 | 0.12254 | 15.2832 | 77.135 | 66.7955 | 366.9478 |  |  | 645.3196 | 0.071553
|  |  | 15.36 | 0 | 0 |  |  |  | 433.7554 | 0.088688 | 15.36 | 76.83 | 64.4813 | 363.2799 |  |  | 648.4985 | 0.036746
|  |  | 15.4368 | 0 | 0 |  |  |  | 435.8817 | 0.01587 | 15.4368 | 76.475 | 62.6125 | 356.6837 |  |  | 651.6774 | 0.02585
|  |  | 15.5136 | 0 | 0 |  |  |  | 438.0079 | 0.067106 | 15.5136 | 76.04 | 61.7523 | 347.8376 |  |  | 654.8564 | 0.058007
|  |  | 15.5904 | 0 | 0 |  |  |  | 440.1342 | 0.1195 | 15.5904 | 76.01 | 61.527 | 337.6968 |  |  | 658.0353 | 0.030905
|  |  | 15.6672 | 0 | 0 |  |  |  | 442.2604 | 0.11561 | 15.6672 | 76.285 | 62.7251 | 328.3266 |  |  | 661.2142 | 0.035297
|  |  | 15.744 | 0 | 0 |  |  |  | 444.3867 | 0.05845 | 15.744 | 76.86 | 63.36 | 328.1108 |  |  | 664.3931 | 0.072848
|  |  | 15.8208 | 0 | 0 |  |  |  | 446.513 | 0.024737 | 15.8208 | 77.205 | 64.5376 | 330.1143 |  |  | 667.572 | 0.047996
|  |  | 15.8976 | 0 | 0 |  |  |  | 448.6392 | 0.085383 | 15.8976 | 77.445 | 64.7936 | 325.0902 |  |  | 670.7509 | 0.030825
|  |  | 15.9744 | 0 | 0 |  |  |  | 450.7655 | 0.099243 | 15.9744 | 77.515 | 64.8397 | 319.4496 |  |  | 673.9299 | 0.083128
|  |  | 16.0512 | 0 | 0 |  |  |  | 452.8917 | 0.064358 | 16.0512 | 77.325 | 65.2237 | 314.5795 |  |  | 677.1088 | 0.068916
|  |  | 16.128 | 0 | 0 |  |  |  | 455.018 | 0.025271 | 16.128 | 77.12 | 65.6486 | 310.7575 |  |  | 680.2877 | 0.0083044
|  |  | 16.2048 | 0 | 0 |  |  |  | 457.1442 | 0.057628 | 16.2048 | 77.265 | 65.408 | 308.0759 |  |  | 683.4666 | 0.080854
|  |  | 16.2816 | 36.305 | 0.68608 |  |  |  | 459.2705 | 0.066055 | 16.2816 | 77.845 | 64.384 | 306.3806 |  |  | 686.6455 | 0.083609
|  |  | 16.3584 | 37.635 | 0.86528 |  |  |  | 461.3967 | 0.040898 | 16.3584 | 78.19 | 65.2288 | 305.3943 |  |  | 689.8244 | 0.0080778
|  |  | 16.4352 | 39.02 | 0.95232 |  |  |  | 463.523 | 0.048396 | 16.4352 | 78.35 | 67.625 | 304.8395 |  |  | 693.0033 | 0.084692
|  |  | 16.512 | 41.385 | 0.9216 |  |  |  | 465.6492 | 0.082022 | 16.512 | 78.76 | 68.4851 | 304.6854 |  |  | 696.1823 | 0.10059
|  |  | 16.5888 | 42.035 | 1.1571 |  |  |  | 467.7755 | 0.08514 | 16.5888 | 79.145 | 68.2752 | 304.8395 |  |  | 699.3612 | 0.018753
|  |  | 16.6656 | 42.9 | 1.3312 |  |  |  | 469.9017 | 0.066581 | 16.6656 | 79.21 | 67.9936 | 304.9628 |  |  | 702.5401 | 0.13954
|  |  | 16.7424 | 43.52 | 1.4797 |  |  |  | 472.028 | 0.08043 | 16.7424 | 78.9 | 68.3213 | 304.9628 |  |  | 705.719 | 0.20027
|  |  | 16.8192 | 44.235 | 1.5616 |  |  |  | 474.1542 | 0.10923 | 16.8192 | 78.49 | 68.3469 | 307.7368 |  |  | 708.8979 | 0.077813
|  |  | 16.896 | 44.66 | 1.6794 |  |  |  | 476.2805 | 0.10373 | 16.896 | 78.43 | 66.6778 | 321.1448 |  |  | 712.0768 | 0.22396
|  |  | 16.9728 | 44.91 | 1.8125 |  |  |  | 478.4067 | 0.056989 | 16.9728 | 78.285 | 66.2989 | 340.5017 |  |  | 715.2557 | 0.57285
|  |  | 17.0496 | 44.845 | 2.0019 |  |  |  | 480.533 | 0.013503 | 17.0496 | 78.3 | 66.519 | 356.0056 |  |  | 718.4347 | 0.79238
|  |  | 17.1264 | 45.225 | 2.048 |  |  |  | 482.6592 | 0.057876 | 17.1264 | 78.725 | 66.816 | 365.9307 |  |  | 721.6136 | 0.76602
|  |  | 17.2032 | 45.865 | 1.9968 |  |  |  | 484.7855 | 0.063261 | 17.2032 | 79.29 | 67.9526 | 371.5713 |  |  | 724.7925 | 0.51217
|  |  | 17.28 | 46.555 | 1.9405 |  |  |  | 486.9117 | 0.035563 | 17.28 | 80.2 | 68.9254 | 374.746 |  |  | 727.9714 | 0.17011
|  |  | 17.3568 | 46.965 | 1.9456 |  |  |  | 489.038 | 0.056175 | 17.3568 | 80.835 | 71.7312 | 377.6126 |  |  | 731.1503 | 0.090285
|  |  | 17.4336 | 47.545 | 1.9149 |  |  |  | 491.1642 | 0.091642 | 17.4336 | 81.3 | 74.0864 | 381.4963 |  |  | 734.3292 | 0.16967
|  |  | 17.5104 | 47.695 | 1.9558 |  |  |  | 493.2905 | 0.083878 | 17.5104 | 81.785 | 74.5165 | 386.6437 |  |  | 737.5081 | 0.090094
|  |  | 17.5872 | 47.565 | 2.048 |  |  |  | 495.4168 | 0.028606 | 17.5872 | 81.98 | 74.3834 | 392.1302 |  |  | 740.6871 | 0.040773
|  |  | 17.664 | 47.23 | 2.1862 |  |  |  | 497.543 | 0.048586 | 17.664 | 81.905 | 73.8048 | 396.4146 |  |  | 743.866 | 0.11182
|  |  | 17.7408 | 46.905 | 2.345 |  |  |  | 499.6693 | 0.10639 | 17.7408 | 81.505 | 73.9277 | 398.0174 |  |  | 747.0449 | 0.083055
|  |  | 17.8176 | 47.49 | 2.2221 |  |  |  | 501.7955 | 0.11396 | 17.8176 | 81.165 | 73.9072 | 396.1064 |  |  | 750.2238 | 0.0090202
|  |  | 17.8944 | 48 | 2.135 |  |  |  | 503.9218 | 0.070053 | 17.8944 | 81.02 | 72.9446 | 390.7123 |  |  | 753.4027 | 0.069905
|  |  | 17.9712 | 47.73 | 2.2374 |  |  |  | 506.048 | 0.049349 | 17.9712 | 80.94 | 71.8029 | 382.7908 |  |  | 756.5816 | 0.06627
|  |  | 18.048 | 47.51 | 2.2989 |  |  |  | 508.1743 | 0.10554 | 18.048 | 80.885 | 70.5946 | 373.8522 |  |  | 759.7605 | 0.0088579
|  |  | 18.1248 | 47.855 | 2.1709 |  |  |  | 510.3005 | 0.12703 | 18.1248 | 80.59 | 70.5075 | 365.3759 |  |  | 762.9395 | 0.051295
|  |  | 18.2016 | 48.01 | 2.1146 |  |  |  | 512.4268 | 0.092151 | 18.2016 | 80.125 | 70.1798 | 358.1633 |  |  | 766.1184 | 0.061456
|  |  | 18.2784 | 47.735 | 2.1811 |  |  |  | 514.553 | 0.027186 | 18.2784 | 79.565 | 68.8947 | 352.1836 |  |  | 769.2973 | 0.02606
|  |  | 18.3552 | 47.82 | 2.1453 |  |  |  | 516.6793 | 0.063524 | 18.3552 | 79.05 | 66.3603 | 346.6971 |  |  | 772.4762 | 0.036414
|  |  | 18.432 | 48.16 | 2.0122 |  |  |  | 518.8055 | 0.09735 | 18.432 | 78.33 | 64.855 | 340.625 |  |  | 775.6551 | 0.054384
|  |  | 18.5088 | 47.44 | 2.1402 |  |  |  | 520.9318 | 0.078783 | 18.5088 | 77.77 | 63.2883 | 333.3816 |  |  | 778.834 | 0.046462
|  |  | 18.5856 | 47.155 | 2.1299 |  |  |  | 523.058 | 0.027061 | 18.5856 | 77.225 | 62.9709 | 325.7375 |  |  | 782.0129 | 0.053236
|  |  | 18.6624 | 47.09 | 2.0736 |  |  |  | 525.1843 | 0.052921 | 18.6624 | 77.13 | 64.2816 | 317.5694 |  |  | 785.1919 | 0.051388
|  |  | 18.7392 | 46.85 | 2.0378 |  |  |  | 527.3105 | 0.08331 | 18.7392 | 77.715 | 64.4403 | 309.4013 |  |  | 788.3708 | 0.013657
|  |  | 18.816 | 46.585 | 1.9712 |  |  |  | 529.4368 | 0.065474 | 18.816 | 78.305 | 64.8141 | 302.2195 |  |  | 791.5497 | 0.039323
|  |  | 18.8928 | 46.125 | 1.8944 |  |  |  | 531.563 | 0.020638 | 18.8928 | 78.715 | 65.7408 | 296.2707 |  |  | 794.7286 | 0.065055
|  |  | 18.9696 | 45.125 | 1.8483 |  |  |  | 533.6893 | 0.070123 | 18.9696 | 78.96 | 66.4013 | 291.4315 |  |  | 797.9075 | 0.045267
|  |  | 19.0464 | 44.22 | 1.6998 |  |  |  | 535.8155 | 0.11261 | 19.0464 | 79.04 | 66.1811 | 287.5169 |  |  | 801.0864 | 0.0017153
|  |  | 19.1232 | 43.425 | 1.4746 |  |  |  | 537.9418 | 0.10971 | 19.1232 | 78.59 | 66.6573 | 284.4655 |  |  | 804.2653 | 0.033752
|  |  | 19.2 | 42.15 | 1.4131 |  |  |  | 540.068 | 0.065527 | 19.2 | 77.89 | 67.2461 | 282.2462 |  |  | 807.4443 | 0.027288
|  |  | 19.2768 | 40.865 | 1.3414 |  |  |  | 542.1943 | 0.019149 | 19.2768 | 77.735 | 64.6144 | 280.705 |  |  | 810.6232 | 0.011097
|  |  | 19.3536 | 40.04 | 1.1725 |  |  |  | 544.3206 | 0.053306 | 19.3536 | 77.255 | 63.7952 | 279.7495 |  |  | 813.8021 | 0.038635
|  |  | 19.4304 | 39.575 | 0.96256 |  |  |  | 546.4468 | 0.070447 | 19.4304 | 76.865 | 62.976 | 279.0406 |  |  | 816.981 | 0.032573
|  |  | 19.5072 | 38.215 | 0.93184 |  |  |  | 548.5731 | 0.062857 | 19.5072 | 76.545 | 62.4845 | 278.5475 |  |  | 820.1599 | 0.0066744
|  |  | 19.584 | 36.88 | 0.83456 |  |  |  | 550.6993 | 0.057463 | 19.584 | 76.255 | 62.1414 | 278.3625 |  |  | 823.3388 | 0.03928
|  |  | 19.6608 | 35.495 | 0.7424 |  |  |  | 552.8256 | 0.063088 | 19.6608 | 76.135 | 61.4042 | 279.6571 |  |  | 826.5177 | 0.040877
|  |  | 19.7376 | 0 | 0 |  |  |  | 554.9518 | 0.057933 | 19.7376 | 76.175 | 61.3427 | 284.2497 |  |  | 829.6967 | 0.0086516
|  |  | 19.8144 | 0 | 0 |  |  |  | 557.0781 | 0.038873 | 19.8144 | 76.425 | 61.9315 | 292.8185 |  |  | 832.8756 | 0.028935
|  |  | 19.8912 | 0 | 0 |  |  |  | 559.2043 | 0.02947 | 19.8912 | 76.53 | 62.6125 | 305.8258 |  |  | 836.0545 | 0.038486
|  |  | 19.968 | 0 | 0 |  |  |  | 561.3306 | 0.035206 | 19.968 | 76.545 | 62.1312 | 322.1928 |  |  | 839.2334 | 0.015101
|  |  | 20.0448 | 0 | 0 |  |  |  | 563.4568 | 0.03293 | 20.0448 | 76.635 | 61.5168 | 339.0838 |  |  | 842.4123 | 0.0245
|  |  | 20.1216 | 0 | 0 |  |  |  | 565.5831 | 0.033684 | 20.1216 | 76.615 | 62.2694 | 353.2624 |  |  | 845.5912 | 0.03488
|  |  | 20.1984 | 0 | 0 |  |  |  | 567.7093 | 0.045535 | 20.1984 | 76.275 | 63.9488 | 362.571 |  |  | 848.7701 | 0.012352
|  |  | 20.2752 | 0 | 0 |  |  |  | 569.8356 | 0.046008 | 20.2752 | 76.365 | 63.7594 | 366.5163 |  |  | 851.9491 | 0.029132
|  |  | 20.352 | 0 | 0 |  |  |  | 571.9618 | 0.027666 | 20.352 | 76.67 | 63.1654 | 366.1772 |  |  | 855.128 | 0.044723
|  |  | 20.4288 | 0 | 0 |  |  |  | 574.0881 | 0.023128 | 20.4288 | 76.82 | 63.913 | 363.5573 |  |  | 858.3069 | 0.022979
|  |  | 20.5056 | 0 | 0 |  |  |  | 576.2143 | 0.04237 | 20.5056 | 77.12 | 64.0307 | 360.8448 |  |  | 861.4858 | 0.019842
|  |  | 20.5824 | 0 | 0 |  |  |  | 578.3406 | 0.039764 | 20.5824 | 77.4 | 63.7338 | 359.4886 |  |  | 864.6647 | 0.042927
|  |  | 20.6592 | 0 | 0 |  |  |  | 580.4668 | 0.020806 | 20.6592 | 77.265 | 64.3686 | 359.7044 |  |  | 867.8436 | 0.025393
|  |  | 20.736 | 0 | 0 |  |  |  | 582.5931 | 0.056887 | 20.736 | 76.875 | 64.3942 | 360.8448 |  |  | 871.0225 | 0.021195
|  |  | 20.8128 | 0 | 0 |  |  |  | 584.7193 | 0.098819 | 20.8128 | 76.485 | 63.0118 | 361.2764 |  |  | 874.2015 | 0.055462
|  |  | 20.8896 | 0 | 0 |  |  |  | 586.8456 | 0.11193 | 20.8896 | 76.135 | 60.8973 | 359.3037 |  |  | 877.3804 | 0.047973
|  |  | 20.9664 | 0 | 0 |  |  |  | 588.9718 | 0.091651 | 20.9664 | 75.64 | 59.3408 | 354.1255 |  |  | 880.5593 | 0.018968
|  |  | 21.0432 | 0 | 0 |  |  |  | 591.0981 | 0.06258 | 21.0432 | 75.01 | 58.5011 | 346.1423 |  |  | 883.7382 | 0.05183
|  |  | 21.12 | 0 | 0 |  |  |  | 593.2244 | 0.06382 | 21.12 | 74.36 | 58.2605 | 336.3098 |  |  | 886.9171 | 0.052019
|  |  | 21.1968 | 0 | 0 |  |  |  | 595.3506 | 0.069122 | 21.1968 | 74.375 | 57.2774 | 325.7375 |  |  | 890.096 | 0.0031522
|  |  | 21.2736 | 0 | 0 |  |  |  | 597.4769 | 0.046275 | 21.2736 | 74.61 | 57.7843 | 318.3399 |  |  | 893.2749 | 0.084344
|  |  | 21.3504 | 0 | 0 |  |  |  | 599.6031 | 0.0068955 | 21.3504 | 74.87 | 59.0131 | 319.7886 |  |  | 896.4539 | 0.14255
|  |  | 21.4272 | 0 | 0 |  |  |  | 601.7294 | 0.042485 | 21.4272 | 75.05 | 59.4586 | 317.8468 |  |  | 899.6328 | 0.14566
|  |  | 21.504 | 0 | 0 |  |  |  | 603.8556 | 0.057478 | 21.504 | 74.805 | 60.5542 | 311.5281 |  |  | 902.8117 | 0.099437
|  |  | 21.5808 | 0 | 0 |  |  |  | 605.9819 | 0.041902 | 21.5808 | 74.62 | 59.8221 | 305.795 |  |  | 905.9906 | 0.03865
|  |  | 21.6576 | 0 | 0 |  |  |  | 608.1081 | 0.047301 | 21.6576 | 74.17 | 59.52 | 301.2332 |  |  | 909.1695 | 0.0095729
|  |  | 21.7344 | 0 | 0 |  |  |  | 610.2344 | 0.086966 | 21.7344 | 74.02 | 58.3731 | 297.8118 |  |  | 912.3484 | 0.017344
|  |  | 21.8112 | 0 | 0 |  |  |  | 612.3606 | 0.10328 | 21.8112 | 74.37 | 56.0742 | 295.6542 |  |  | 915.5273 | 0.010802
|  |  | 21.888 | 0 | 0 |  |  |  | 614.4869 | 0.077942 | 21.888 | 74.33 | 56.9293 | 294.3905 |  |  | 918.7063 | 0.0053786
|  |  | 21.9648 | 0 | 0 |  |  |  | 616.6131 | 0.020896 | 21.9648 | 74.445 | 58.1069 | 293.7124 |  |  | 921.8852 | 0.0057056
|  |  | 22.0416 | 0 | 0 |  |  |  | 618.7394 | 0.041987 | 22.0416 | 74.875 | 58.1018 | 293.435 |  |  | 925.0641 | 0.010387
|  |  | 22.1184 | 0 | 0 |  |  |  | 620.8656 | 0.077731 | 22.1184 | 75.005 | 58.7059 | 293.5274 |  |  | 928.243 | 0.011229
|  |  | 22.1952 | 0 | 0 |  |  |  | 622.9919 | 0.071884 | 22.1952 | 75.095 | 58.24 | 293.8973 |  |  | 931.4219 | 0.013632
|  |  | 22.272 | 0 | 0 |  |  |  | 625.1181 | 0.033981 | 22.272 | 74.685 | 58.9056 | 294.4521 |  |  | 934.6008 | 0.016377
|  |  | 22.3488 | 0 | 0 |  |  |  | 627.2444 | 0.034106 | 22.3488 | 74.24 | 58.665 | 298.644 |  |  | 937.7797 | 0.0082821
|  |  | 22.4256 | 0 | 0 |  |  |  | 629.3706 | 0.064304 | 22.4256 | 73.86 | 57.728 | 314.3638 |  |  | 940.9587 | 0.0084319
|  |  | 22.5024 | 0 | 0 |  |  |  | 631.4969 | 0.062204 | 22.5024 | 73.495 | 57.6819 | 335.3542 |  |  | 944.1376 | 0.020969
|  |  | 22.5792 | 0 | 0 |  |  |  | 633.6231 | 0.030374 | 22.5792 | 73.395 | 57.8611 | 351.3206 |  |  | 947.3165 | 0.021749
|  |  | 22.656 | 0 | 0 |  |  |  | 635.7494 | 0.02074 | 22.656 | 73.725 | 57.9942 | 360.5983 |  |  | 950.4954 | 0.023409
|  |  | 22.7328 | 0 | 0 |  |  |  | 637.8756 | 0.042424 | 22.7328 | 74.43 | 58.5216 | 364.7902 |  |  | 953.6743 | 0.030537
|  |  | 22.8096 | 0 | 0 |  |  |  | 640.0019 | 0.033788 | 22.8096 | 75.44 | 59.1718 | 366.4547 |  |  | 956.8532 | 0.024825
|  |  | 22.8864 | 0 | 0 |  |  |  | 642.1281 | 0.025213 | 22.8864 | 76.19 | 61.2966 | 368.5198 |  |  | 960.0321 | 0.0053318
|  |  | 22.9632 | 0 | 0 |  |  |  | 644.2544 | 0.070725 | 22.9632 | 76.69 | 63.8106 | 372.6501 |  |  | 963.2111 | 0.01375
|  |  | 23.04 | 0 | 0 |  |  |  | 646.3807 | 0.10613 | 23.04 | 77.195 | 64.7578 | 378.9996 |  |  | 966.39 | 0.015345
|  |  | 23.1168 | 0 | 0 |  |  |  | 648.5069 | 0.11016 | 23.1168 | 77.78 | 63.8413 | 386.1505 |  |  | 969.5689 | 0.0015439
|  |  | 23.1936 | 0 | 0 |  |  |  | 650.6332 | 0.08748 | 23.1936 | 77.925 | 63.5238 | 392.0069 |  |  | 972.7478 | 0.016032
|  |  | 23.2704 | 0 | 0 |  |  |  | 652.7594 | 0.064176 | 23.2704 | 77.68 | 64.2458 | 394.6269 |  |  | 975.9267 | 0.016201
|  |  | 23.3472 | 0 | 0 |  |  |  | 654.8857 | 0.058078 | 23.3472 | 77.305 | 65.0701 | 392.9624 |  |  | 979.1056 | 0.016883
|  |  | 23.424 | 0 | 0 |  |  |  | 657.0119 | 0.046302 | 23.424 | 77.255 | 64.1741 | 387.2602 |  |  | 982.2845 | 0.037709
|  |  | 23.5008 | 0 | 0 |  |  |  | 659.1382 | 0.015306 | 23.5008 | 77.2 | 63.8618 | 378.8455 |  |  | 985.4635 | 0.041999
|  |  | 23.5776 | 0 | 0 |  |  |  | 661.2644 | 0.022249 | 23.5776 | 77.005 | 64.6502 | 369.7835 |  |  | 988.6424 | 0.020439
|  |  | 23.6544 | 0 | 0 |  |  |  | 663.3907 | 0.04475 | 23.6544 | 76.835 | 65.1981 | 361.9545 |  |  | 991.8213 | 0.01537
|  |  | 23.7312 | 0 | 0 |  |  |  | 665.5169 | 0.038257 | 23.7312 | 76.84 | 64.3174 | 356.1598 |  |  | 995.0002 | 0.0341
|  |  | 23.808 | 0 | 0 |  |  |  | 667.6432 | 0.0064835 | 23.808 | 76.745 | 62.6995 | 352.2761 |  |  | 998.1791 | 0.026055
|  |  | 23.8848 | 0 | 0 |  |  |  | 669.7694 | 0.034637 | 23.8848 | 76.445 | 61.1686 | 349.0705 |  |  | 1001.358 | 0.0091382
|  |  | 23.9616 | 0 | 0 |  |  |  | 671.8957 | 0.059329 | 23.9616 | 76.045 | 59.9962 | 344.7861 |  |  | 1004.5369 | 0.024242
|  |  | 24.0384 | 0 | 0 |  |  |  | 674.0219 | 0.055642 | 24.0384 | 75.74 | 59.3459 | 339.1455 |  |  | 1007.7159 | 0.019268
|  |  | 24.1152 | 0 | 0 |  |  |  | 676.1482 | 0.03155 | 24.1152 | 75.885 | 59.4381 | 332.7035 |  |  | 1010.8948 | 0.0091898
|  |  | 24.192 | 0 | 0 |  |  |  | 678.2744 | 0.03173 | 24.192 | 76.455 | 61.0253 | 325.121 |  |  | 1014.0737 | 0.034197
|  |  | 24.2688 | 0 | 0 |  |  |  | 680.4007 | 0.047896 | 24.2688 | 77.345 | 62.4026 | 318.0009 |  |  | 1017.2526 | 0.038322
|  |  | 24.3456 | 0 | 0 |  |  |  | 682.5269 | 0.040064 | 24.3456 | 78.38 | 62.7814 | 312.1137 |  |  | 1020.4315 | 0.018299
|  |  | 24.4224 | 0 | 0 |  |  |  | 684.6532 | 0.011721 | 24.4224 | 79.09 | 64.041 | 307.4902 |  |  | 1023.6104 | 0.0083089
|  |  | 24.4992 | 0 | 0 |  |  |  | 686.7794 | 0.028935 | 24.4992 | 79.53 | 65.3619 | 303.9456 |  |  | 1026.7893 | 0.019388
|  |  | 24.576 | 0 | 0 |  |  |  | 688.9057 | 0.044106 | 24.576 | 79.92 | 65.2442 | 301.3565 |  |  | 1029.9683 | 0.0081563
|  |  | 24.6528 | 0 | 0 |  |  |  | 691.0319 | 0.02827 | 24.6528 | 79.855 | 66.0224 | 299.4763 |  |  | 1033.1472 | 0.016353
|  |  | 24.7296 | 0 | 0 |  |  |  | 693.1582 | 0.019532 | 24.7296 | 79.625 | 66.6419 | 298.1817 |  |  | 1036.3261 | 0.030303
|  |  | 24.8064 | 0 | 0 |  |  |  | 695.2845 | 0.062158 | 24.8064 | 79.48 | 66.6112 | 297.2878 |  |  | 1039.505 | 0.025844
|  |  | 24.8832 | 0 | 0 |  |  |  | 697.4107 | 0.082521 | 24.8832 | 79.605 | 65.6384 | 296.4865 |  |  | 1042.6839 | 0.012213
|  |  | 24.96 | 0 | 0 |  |  |  | 699.537 | 0.064837 | 24.96 | 79.7 | 65.5923 | 295.5926 |  |  | 1045.8628 | 0.014131
|  |  | 25.0368 | 0 | 0 |  |  |  | 701.6632 | 0.032468 | 25.0368 | 79.745 | 66.1504 | 294.6987 |  |  | 1049.0417 | 0.015178
|  |  | 25.1136 | 0 | 0 |  |  |  | 703.7895 | 0.071687 | 25.1136 | 79.865 | 66.2272 | 294.2672 |  |  | 1052.2207 | 0.0083076
|  |  | 25.1904 | 0 | 0 |  |  |  | 705.9157 | 0.11218 | 25.1904 | 80.02 | 66.3552 | 296.1165 |  |  | 1055.3996 | 0.0074625
|  |  | 25.2672 | 0 | 0 |  |  |  | 708.042 | 0.11286 | 25.2672 | 80.48 | 66.6931 | 301.6647 |  |  | 1058.5785 | 0.013093
|  |  | 25.344 | 0 | 0 |  |  |  | 710.1682 | 0.073621 | 25.344 | 81.035 | 67.5738 | 311.4972 |  |  | 1061.7574 | 0.016358
|  |  | 25.4208 | 0 | 0 |  |  |  | 712.2945 | 0.027522 | 25.4208 | 81.285 | 68.992 | 325.5525 |  |  | 1064.9363 | 0.012964
|  |  | 25.4976 | 0 | 0 |  |  |  | 714.4207 | 0.04887 | 25.4976 | 81.38 | 69.8829 | 342.0428 |  |  | 1068.1152 | 0.0030284
|  |  | 25.5744 | 0 | 0 |  |  |  | 716.547 | 0.060698 | 25.5744 | 81.645 | 70.999 | 358.0092 |  |  | 1071.2941 | 0.009958
|  |  | 25.6512 | 0 | 0 |  |  |  | 718.6732 | 0.040623 | 25.6512 | 82.04 | 71.1885 | 370.2767 |  |  | 1074.4731 | 0.026883
|  |  | 25.728 | 0 | 0 |  |  |  | 720.7995 | 0.04234 | 25.728 | 82.32 | 70.7891 | 376.9345 |  |  | 1077.652 | 0.051703
|  |  | 25.8048 | 0 | 0 |  |  |  | 722.9257 | 0.078851 | 25.8048 | 82.465 | 71.0246 | 378.0441 |  |  | 1080.8309 | 0.07018
|  |  | 25.8816 | 0 | 0 |  |  |  | 725.052 | 0.092378 | 25.8816 | 82.63 | 71.4854 | 375.0543 |  |  | 1084.0098 | 0.063357
|  |  | 25.9584 | 0 | 0 |  |  |  | 727.1782 | 0.070339 | 25.9584 | 82.84 | 72.2944 | 370.5849 |  |  | 1087.1887 | 0.028163
|  |  | 26.0352 | 0 | 0 |  |  |  | 729.3045 | 0.025471 | 26.0352 | 82.91 | 73.4054 | 366.6704 |  |  | 1090.3676 | 0.016933
|  |  | 26.112 | 0 | 0 |  |  |  | 731.4307 | 0.024978 | 26.112 | 82.945 | 74.0198 | 364.2662 |  |  | 1093.5465 | 0.040652
|  |  | 26.1888 | 0 | 0 |  |  |  | 733.557 | 0.049541 | 26.1888 | 83.075 | 72.9139 | 363.5265 |  |  | 1096.7255 | 0.028442
|  |  | 26.2656 | 0 | 0 |  |  |  | 735.6832 | 0.049787 | 26.2656 | 83 | 71.3523 | 363.1566 |  |  | 1099.9044 | 0.0087246
|  |  | 26.3424 | 0 | 0 |  |  |  | 737.8095 | 0.041718 | 26.3424 | 82.58 | 70.1082 | 361.4922 |  |  | 1103.0833 | 0.040866
|  |  | 26.4192 | 0 | 0 |  |  |  | 739.9357 | 0.04347 | 26.4192 | 81.925 | 69.4118 | 357.4235 |  |  | 1106.2622 | 0.043685
|  |  | 26.496 | 0 | 0 |  |  |  | 742.062 | 0.040222 | 26.496 | 81.27 | 68.4902 | 350.2418 |  |  | 1109.4411 | 0.018528
|  |  | 26.5728 | 0 | 0 |  |  |  | 744.1883 | 0.020896 | 26.5728 | 80.655 | 66.9901 | 340.4709 |  |  | 1112.62 | 0.013561
|  |  | 26.6496 | 0 | 0 |  |  |  | 746.3145 | 0.017047 | 26.6496 | 80.105 | 66.1914 | 329.1896 |  |  | 1115.799 | 0.021152
|  |  | 26.7264 | 0 | 0 |  |  |  | 748.4408 | 0.035192 | 26.7264 | 79.835 | 66.8723 | 317.2303 |  |  | 1118.9779 | 0.006894
|  |  | 26.8032 | 0 | 0 |  |  |  | 750.567 | 0.034713 | 26.8032 | 80.04 | 67.712 | 306.658 |  |  | 1122.1568 | 0.024477
|  |  | 26.88 | 0 | 0 |  |  |  | 752.6933 | 0.03314 | 26.88 | 80.33 | 68.4442 | 305.0244 |  |  | 1125.3357 | 0.030511
|  |  | 26.9568 | 0 | 0 |  |  |  | 754.8195 | 0.061395 | 26.9568 | 80.575 | 68.0038 | 304.1306 |  |  | 1128.5146 | 0.01235
|  |  | 27.0336 | 0 | 0 |  |  |  | 756.9458 | 0.085375 | 27.0336 | 80.48 | 67.4048 | 296.733 |  |  | 1131.6935 | 0.02183
|  |  | 27.1104 | 0 | 0 |  |  |  | 759.072 | 0.081465 | 27.1104 | 80.06 | 66.8672 | 289.0581 |  |  | 1134.8724 | 0.03138
|  |  | 27.1872 | 0 | 0 |  |  |  | 761.1983 | 0.053531 | 27.1872 | 79.755 | 65.239 | 282.2154 |  |  | 1138.0514 | 0.011187
|  |  | 27.264 | 0 | 0 |  |  |  | 763.3245 | 0.049859 | 27.264 | 79.62 | 63.3139 | 276.3898 |  |  | 1141.2303 | 0.024188
|  |  | 27.3408 | 0 | 0 |  |  |  | 765.4508 | 0.080003 | 27.3408 | 79.14 | 63.9693 | 271.7972 |  |  | 1144.4092 | 0.045858
|  |  | 27.4176 | 0 | 0 |  |  |  | 767.577 | 0.089403 | 27.4176 | 78.915 | 64.5427 | 268.4067 |  |  | 1147.5881 | 0.03708
|  |  | 27.4944 | 0 | 0 |  |  |  | 769.7033 | 0.065497 | 27.4944 | 78.925 | 64.7014 | 266.0333 |  |  | 1150.767 | 0.0094301
|  |  | 27.5712 | 0 | 0 |  |  |  | 771.8295 | 0.021746 | 27.5712 | 79.03 | 64.512 | 264.4922 |  |  | 1153.9459 | 0.01517
|  |  | 27.648 | 0 | 0 |  |  |  | 773.9558 | 0.024684 | 27.648 | 79.135 | 63.4061 | 263.6908 |  |  | 1157.1248 | 0.0098175
|  |  | 27.7248 | 0 | 0 |  |  |  | 776.082 | 0.047422 | 27.7248 | 78.88 | 62.7507 | 263.475 |  |  | 1160.3038 | 0.019795
|  |  | 27.8016 | 0 | 0 |  |  |  | 778.2083 | 0.057116 | 27.8016 | 78.38 | 62.4947 | 263.5058 |  |  | 1163.4827 | 0.039227
|  |  | 27.8784 | 0 | 0 |  |  |  | 780.3345 | 0.075523 | 27.8784 | 77.84 | 61.2659 | 264.9237 |  |  | 1166.6616 | 0.029203
|  |  | 27.9552 | 0 | 0 |  |  |  | 782.4608 | 0.096369 | 27.9552 | 77.305 | 59.561 | 274.8487 |  |  | 1169.8405 | 0.0094427
|  |  | 28.032 | 0 | 0 |  |  |  | 784.587 | 0.095532 | 28.032 | 76.725 | 58.5318 | 293.2809 |  |  | 1173.0194 | 0.042928
|  |  | 28.1088 | 0 | 0 |  |  |  | 786.7133 | 0.064268 | 28.1088 | 76.2 | 58.6189 | 309.3705 |  |  | 1176.1983 | 0.047698
|  |  | 28.1856 | 0 | 0 |  |  |  | 788.8395 | 0.014275 | 28.1856 | 76.305 | 57.9072 | 319.6962 |  |  | 1179.3772 | 0.020337
|  |  | 28.2624 | 0 | 0 |  |  |  | 790.9658 | 0.031295 | 28.2624 | 76.5 | 58.7878 | 325.6142 |  |  | 1182.5562 | 0.015287
|  |  | 28.3392 | 0 | 0 |  |  |  | 793.0921 | 0.050364 | 28.3392 | 76.825 | 60.4416 | 329.3746 |  |  | 1185.7351 | 0.025313
|  |  | 28.416 | 0 | 0 |  |  |  | 795.2183 | 0.037009 | 28.416 | 77.33 | 61.9213 | 333.2891 |  |  | 1188.914 | 0.0042472
|  |  | 28.4928 | 0 | 0 |  |  |  | 797.3446 | 0.016516 | 28.4928 | 77.97 | 62.7046 | 338.8372 |  |  | 1192.0929 | 0.025337
|  |  | 28.5696 | 0 | 0 |  |  |  | 799.4708 | 0.042145 | 28.5696 | 78.355 | 63.4675 | 346.1731 |  |  | 1195.2718 | 0.033134
|  |  | 28.6464 | 0 | 0 |  |  |  | 801.5971 | 0.053341 | 28.6464 | 78.495 | 63.9232 | 354.1871 |  |  | 1198.4507 | 0.028465
|  |  | 28.7232 | 0 | 0 |  |  |  | 803.7233 | 0.034164 | 28.7232 | 78.535 | 63.6058 | 361.0914 |  |  | 1201.6296 | 0.057578
|  |  | 28.8 | 0 | 0 |  |  |  | 805.8496 | 0.01589 | 28.8 | 78.33 | 63.3907 | 365.1601 |  |  | 1204.8086 | 0.069301
|  |  | 28.8768 | 0 | 0 |  |  |  | 807.9758 | 0.055156 | 28.8768 | 78.13 | 62.7046 | 365.4683 |  |  | 1207.9875 | 0.037987
|  |  | 28.9536 | 0 | 0 |  |  |  | 810.1021 | 0.078232 | 28.9536 | 77.96 | 62.2438 | 362.0778 |  |  | 1211.1664 | 0.024643
|  |  | 29.0304 | 0 | 0 |  |  |  | 812.2283 | 0.072836 | 29.0304 | 77.62 | 62.8378 | 356.0056 |  |  | 1214.3453 | 0.054572
|  |  | 29.1072 | 0 | 0 |  |  |  | 814.3546 | 0.045322 | 29.1072 | 77.48 | 62.7558 | 348.978 |  |  | 1217.5242 | 0.034152
|  |  | 29.184 | 0 | 0 |  |  |  | 816.4808 | 0.01026 | 29.184 | 77.545 | 61.609 | 342.5668 |  |  | 1220.7031 | 0.023228
|  |  | 29.2608 | 0 | 0 |  |  |  | 818.6071 | 0.024711 | 29.2608 | 77.26 | 61.4195 | 337.5735 |  |  | 1223.882 | 0.066633
|  |  | 29.3376 | 0 | 0 |  |  |  | 820.7333 | 0.062222 | 29.3376 | 76.905 | 60.416 | 333.8747 |  |  | 1227.061 | 0.052786
|  |  | 29.4144 | 0 | 0 |  |  |  | 822.8596 | 0.10088 | 29.4144 | 76.28 | 59.351 | 330.6075 |  |  | 1230.2399 | 0.014447
|  |  | 29.4912 | 0 | 0 |  |  |  | 824.9858 | 0.12512 | 29.4912 | 75.265 | 59.4586 | 326.3539 |  |  | 1233.4188 | 0.07633
|  |  | 29.568 | 0 | 0 |  |  |  | 827.1121 | 0.11768 | 29.568 | 74.695 | 57.7894 | 320.3743 |  |  | 1236.5977 | 0.074214
|  |  | 29.6448 | 0 | 0 |  |  |  | 829.2383 | 0.073799 | 29.6448 | 74.685 | 55.1219 | 313.4083 |  |  | 1239.7766 | 0.01021
|  |  | 29.7216 | 0 | 0 |  |  |  | 831.3646 | 0.0061523 | 29.7216 | 74.7 | 55.6749 | 305.6717 |  |  | 1242.9555 | 0.10386
|  |  | 29.7984 | 0 | 0 |  |  |  | 833.4908 | 0.061652 | 29.7984 | 74.93 | 57.2672 | 298.0892 |  |  | 1246.1344 | 0.12814
|  |  | 29.8752 | 0 | 0 |  |  |  | 835.6171 | 0.10793 | 29.8752 | 75.245 | 58.8646 | 291.8322 |  |  | 1249.3134 | 0.026367
|  |  | 29.952 | 0 | 0 |  |  |  | 837.7433 | 0.12575 | 29.952 | 75.815 | 59.3306 | 287.0854 |  |  | 1252.4923 | 0.19536
|  |  | 30.0288 | 0 | 0 |  |  |  | 839.8696 | 0.12472 | 30.0288 | 76.235 | 59.7504 | 283.7565 |  |  | 1255.6712 | 0.41139
|  |  | 30.1056 | 0 | 0 |  |  |  | 841.9959 | 0.11644 | 30.1056 | 76.385 | 60.1293 | 281.6298 |  |  | 1258.8501 | 0.51612
|  |  | 30.1824 | 0 | 0 |  |  |  | 844.1221 | 0.098122 | 30.1824 | 76.33 | 59.9757 | 280.4276 |  |  | 1262.029 | 0.45664
|  |  | 30.2592 | 0 | 0 |  |  |  | 846.2484 | 0.062177 | 30.2592 | 76.3 | 58.7264 | 279.9345 |  |  | 1265.2079 | 0.27392
|  |  | 30.336 | 0 | 0 |  |  |  | 848.3746 | 0.01511 | 30.336 | 76.005 | 58.1018 | 279.9653 |  |  | 1268.3868 | 0.07188
|  |  | 30.4128 | 0 | 0 |  |  |  | 850.5009 | 0.031411 | 30.4128 | 75.7 | 57.9482 | 280.2427 |  |  | 1271.5658 | 0.057232
|  |  | 30.4896 | 0 | 0 |  |  |  | 852.6271 | 0.050508 | 30.4896 | 75.36 | 58.7008 | 280.6434 |  |  | 1274.7447 | 0.073076
|  |  | 30.5664 | 0 | 0 |  |  |  | 854.7534 | 0.040796 | 30.5664 | 75.22 | 58.9005 | 281.1058 |  |  | 1277.9236 | 0.01712
|  |  | 30.6432 | 0 | 0 |  |  |  | 856.8796 | 0.021636 | 30.6432 | 75.34 | 57.8202 | 281.7839 |  |  | 1281.1025 | 0.043713
|  |  | 30.72 | 0 | 0 |  |  |  | 859.0059 | 0.033059 | 30.72 | 75.49 | 56.919 | 283.8798 |  |  | 1284.2814 | 0.059474
|  |  | 30.7968 | 0 | 0 |  |  |  | 861.1321 | 0.037305 | 30.7968 | 75.825 | 56.6989 | 289.1814 |  |  | 1287.4603 | 0.029642
|  |  | 30.8736 | 0 | 0 |  |  |  | 863.2584 | 0.02422 | 30.8736 | 75.915 | 59.5251 | 298.2433 |  |  | 1290.6392 | 0.016815
|  |  | 30.9504 | 0 | 0 |  |  |  | 865.3846 | 0.047456 | 30.9504 | 76.285 | 60.9638 | 311.4356 |  |  | 1293.8182 | 0.037843
|  |  | 31.0272 | 0 | 0 |  |  |  | 867.5109 | 0.081679 | 31.0272 | 76.765 | 60.7334 | 327.7718 |  |  | 1296.9971 | 0.028837
|  |  | 31.104 | 0 | 0 |  |  |  | 869.6371 | 0.08868 | 31.104 | 77.125 | 60.5133 | 344.7552 |  |  | 1300.176 | 0.0054373
|  |  | 31.1808 | 0 | 0 |  |  |  | 871.7634 | 0.063827 | 31.1808 | 77.205 | 61.1277 | 359.427 |  |  | 1303.3549 | 0.017218
|  |  | 31.2576 | 0 | 0 |  |  |  | 873.8896 | 0.054051 | 31.2576 | 77.25 | 61.4298 | 369.3212 |  |  | 1306.5338 | 0.02389
|  |  | 31.3344 | 0 | 0 |  |  |  | 876.0159 | 0.093919 | 31.3344 | 77.17 | 62.2797 | 373.6364 |  |  | 1309.7127 | 0.022825
|  |  | 31.4112 | 0 | 0 |  |  |  | 878.1421 | 0.1171 | 31.4112 | 77.285 | 62.848 | 373.2357 |  |  | 1312.8916 | 0.019754
|  |  | 31.488 | 0 | 0 |  |  |  | 880.2684 | 0.097167 | 31.488 | 77.685 | 62.8736 | 369.9993 |  |  | 1316.0706 | 0.024515
|  |  | 31.5648 | 0 | 0 |  |  |  | 882.3946 | 0.040624 | 31.5648 | 77.915 | 64.4352 | 366.2081 |  |  | 1319.2495 | 0.037472
|  |  | 31.6416 | 0 | 0 |  |  |  | 884.5209 | 0.026145 | 31.6416 | 78.3 | 65.0342 | 363.6189 |  |  | 1322.4284 | 0.041136
|  |  | 31.7184 | 0 | 0 |  |  |  | 886.6471 | 0.070729 | 31.7184 | 78.56 | 65.2288 | 362.91 |  |  | 1325.6073 | 0.027792
|  |  | 31.7952 | 0 | 0 |  |  |  | 888.7734 | 0.073687 | 31.7952 | 78.51 | 65.1264 | 363.5881 |  |  | 1328.7862 | 0.0075622
|  |  | 31.872 | 0 | 0 |  |  |  | 890.8997 | 0.036722 | 31.872 | 78.215 | 64.5683 | 364.2662 |  |  | 1331.9651 | 0.018175
|  |  | 31.9488 | 0 | 0 |  |  |  | 893.0259 | 0.021028 | 31.9488 | 77.845 | 63.2986 | 363.4032 |  |  | 1335.144 | 0.020151
|  |  | 32.0256 | 0 | 0 |  |  |  | 895.1522 | 0.069209 | 32.0256 | 77.205 | 62.5101 | 359.7661 |  |  | 1338.323 | 0.0097769
|  |  | 32.1024 | 0 | 0 |  |  |  | 897.2784 | 0.094139 | 32.1024 | 76.68 | 61.2864 | 353.2007 |  |  | 1341.5019 | 0.022375
|  |  | 32.1792 | 0 | 0 |  |  |  | 899.4047 | 0.10144 | 32.1792 | 76.4 | 59.7453 | 344.4162 |  |  | 1344.6808 | 0.032599
|  |  | 32.256 | 0 | 0 |  |  |  | 901.5309 | 0.10909 | 32.256 | 76.155 | 60.7027 | 334.1213 |  |  | 1347.8597 | 0.024272
|  |  | 32.3328 | 0 | 0 |  |  |  | 903.6572 | 0.11553 | 32.3328 | 76.465 | 62.1517 | 324.3813 |  |  | 1351.0386 | 0.0074403
|  |  | 32.4096 | 0 | 0 |  |  |  | 905.7834 | 0.10352 | 32.4096 | 76.885 | 63.8566 | 323.8881 |  |  | 1354.2175 | 0.021162
|  |  | 32.4864 | 0 | 0 |  |  |  | 907.9097 | 0.070416 | 32.4864 | 77.415 | 64.2611 | 326.1998 |  |  | 1357.3964 | 0.025045
|  |  | 32.5632 | 0 | 0 |  |  |  | 910.0359 | 0.041398 | 32.5632 | 77.595 | 64.9011 | 321.2681 |  |  | 1360.5754 | 0.012543
|  |  | 32.64 | 0 | 0 |  |  |  | 912.1622 | 0.045936 | 32.64 | 77.62 | 64.9984 | 315.5659 |  |  | 1363.7543 | 0.010479
|  |  | 32.7168 | 0 | 0 |  |  |  | 914.2884 | 0.043964 | 32.7168 | 77.495 | 64.8192 | 310.6958 |  |  | 1366.9332 | 0.017207
|  |  | 32.7936 | 38.605 | 0.54272 |  |  |  | 916.4147 | 0.016633 | 32.7936 | 77.305 | 65.1162 | 306.8738 |  |  | 1370.1121 | 0.010048
|  |  | 32.8704 | 38.605 | 0.94208 |  |  |  | 918.5409 | 0.029488 | 32.8704 | 77.54 | 64.297 | 304.2538 |  |  | 1373.291 | 0.016999
|  |  | 32.9472 | 40.06 | 1.1469 |  |  |  | 920.6672 | 0.063974 | 32.9472 | 77.96 | 63.9949 | 302.5894 |  |  | 1376.4699 | 0.029156
|  |  | 33.024 | 41.735 | 1.2954 |  |  |  | 922.7934 | 0.074289 | 33.024 | 78.295 | 65.0803 | 301.6339 |  |  | 1379.6488 | 0.026246
|  |  | 33.1008 | 43.82 | 1.3261 |  |  |  | 924.9197 | 0.059812 | 33.1008 | 78.685 | 66.1965 | 301.1099 |  |  | 1382.8278 | 0.01179
|  |  | 33.1776 | 45.45 | 1.3978 |  |  |  | 927.0459 | 0.041409 | 33.1776 | 79.165 | 66.432 | 300.9558 |  |  | 1386.0067 | 0.0072568
|  |  | 33.2544 | 46.165 | 1.6384 |  |  |  | 929.1722 | 0.043588 | 33.2544 | 79.5 | 66.2272 | 301.1099 |  |  | 1389.1856 | 0.012201
|  |  | 33.3312 | 47.205 | 1.7613 |  |  |  | 931.2984 | 0.048375 | 33.3312 | 79.605 | 65.8278 | 301.264 |  |  | 1392.3645 | 0.014833
|  |  | 33.408 | 48.115 | 1.8586 |  |  |  | 933.4247 | 0.047599 | 33.408 | 79.4 | 65.6384 | 301.3257 |  |  | 1395.5434 | 0.016308
|  |  | 33.4848 | 48.54 | 2.048 |  |  |  | 935.5509 | 0.053947 | 33.4848 | 79 | 65.6896 | 304.747 |  |  | 1398.7223 | 0.0083706
|  |  | 33.5616 | 49.175 | 2.1248 |  |  |  | 937.6772 | 0.059368 | 33.5616 | 78.595 | 66.0582 | 319.1105 |  |  | 1401.9012 | 0.010779
|  |  | 33.6384 | 49.715 | 2.1709 |  |  |  | 939.8035 | 0.048813 | 33.6384 | 78.43 | 65.8842 | 338.2824 |  |  | 1405.0802 | 0.024214
|  |  | 33.7152 | 50.175 | 2.1965 |  |  |  | 941.9297 | 0.038209 | 33.7152 | 78.6 | 65.5565 | 353.1391 |  |  | 1408.2591 | 0.024921
|  |  | 33.792 | 50.695 | 2.1811 |  |  |  | 944.056 | 0.062538 | 33.792 | 78.87 | 66.601 | 362.5401 |  |  | 1411.438 | 0.017296
|  |  | 33.8688 | 50.825 | 2.2784 |  |  |  | 946.1822 | 0.086777 | 33.8688 | 79.48 | 68.0397 | 367.8725 |  |  | 1414.6169 | 0.011987
|  |  | 33.9456 | 51.04 | 2.3398 |  |  |  | 948.3085 | 0.085756 | 33.9456 | 80.3 | 69.7446 | 371.0781 |  |  | 1417.7958 | 0.0082137
|  |  | 34.0224 | 51.95 | 2.2323 |  |  |  | 950.4347 | 0.063515 | 34.0224 | 81.16 | 71.2806 | 374.0988 |  |  | 1420.9747 | 0.035393
|  |  | 34.0992 | 52.035 | 2.3091 |  |  |  | 952.561 | 0.050407 | 34.0992 | 81.91 | 72.2586 | 378.1674 |  |  | 1424.1536 | 0.058132
|  |  | 34.176 | 52.23 | 2.3501 |  |  |  | 954.6872 | 0.0579 | 34.176 | 82.365 | 72.8781 | 383.4689 |  |  | 1427.3326 | 0.049179
|  |  | 34.2528 | 52.575 | 2.3552 |  |  |  | 956.8135 | 0.052567 | 34.2528 | 82.395 | 73.4054 | 388.863 |  |  | 1430.5115 | 0.023729
|  |  | 34.3296 | 52.225 | 2.519 |  |  |  | 958.9397 | 0.03795 | 34.3296 | 82.3 | 73.0214 | 392.8391 |  |  | 1433.6904 | 0.088031
|  |  | 34.4064 | 52.18 | 2.5754 |  |  |  | 961.066 | 0.062751 | 34.4064 | 82.12 | 72.1459 | 394.0412 |  |  | 1436.8693 | 0.14794
|  |  | 34.4832 | 52.76 | 2.4371 |  |  |  | 963.1922 | 0.095629 | 34.4832 | 81.66 | 72.6323 | 391.6678 |  |  | 1440.0482 | 0.16089
|  |  | 34.56 | 53.105 | 2.3706 |  |  |  | 965.3185 | 0.10093 | 34.56 | 81.095 | 73.8099 | 385.9656 |  |  | 1443.2271 | 0.12122
|  |  | 34.6368 | 52.64 | 2.5139 |  |  |  | 967.4447 | 0.07467 | 34.6368 | 81.005 | 72.5811 | 377.9516 |  |  | 1446.406 | 0.054159
|  |  | 34.7136 | 52.745 | 2.5088 |  |  |  | 969.571 | 0.044237 | 34.7136 | 80.99 | 71.4701 | 369.1054 |  |  | 1449.585 | 0.0084448
|  |  | 34.7904 | 52.725 | 2.519 |  |  |  | 971.6972 | 0.058625 | 34.7904 | 80.905 | 70.2464 | 360.8757 |  |  | 1452.7639 | 0.027764
|  |  | 34.8672 | 52.45 | 2.5805 |  |  |  | 973.8235 | 0.078933 | 34.8672 | 80.605 | 69.1251 | 353.9097 |  |  | 1455.9428 | 0.016635
|  |  | 34.944 | 52.52 | 2.5088 |  |  |  | 975.9497 | 0.081142 | 34.944 | 80.02 | 68.1523 | 348.0533 |  |  | 1459.1217 | 0.009103
|  |  | 35.0208 | 52.385 | 2.5139 |  |  |  | 978.076 | 0.078002 | 35.0208 | 79.295 | 66.4934 | 342.5052 |  |  | 1462.3006 | 0.02577
|  |  | 35.0976 | 51.96 | 2.5805 |  |  |  | 980.2022 | 0.077414 | 35.0976 | 78.51 | 64.9216 | 336.2789 |  |  | 1465.4795 | 0.027249
|  |  | 35.1744 | 51.685 | 2.5958 |  |  |  | 982.3285 | 0.06449 | 35.1744 | 77.895 | 63.6467 | 328.8814 |  |  | 1468.6584 | 0.019941
|  |  | 35.2512 | 51.585 | 2.5395 |  |  |  | 984.4547 | 0.036405 | 35.2512 | 77.45 | 63.4112 | 321.114 |  |  | 1471.8374 | 0.01207
|  |  | 35.328 | 51.41 | 2.4934 |  |  |  | 986.581 | 0.053413 | 35.328 | 77.375 | 64.8653 | 312.8843 |  |  | 1475.0163 | 0.0068908
|  |  | 35.4048 | 51.095 | 2.4781 |  |  |  | 988.7073 | 0.097546 | 35.4048 | 78.045 | 64.7424 | 304.6237 |  |  | 1478.1952 | 0.010457
|  |  | 35.4816 | 50.785 | 2.4218 |  |  |  | 990.8335 | 0.11661 | 35.4816 | 78.84 | 64.3686 | 297.442 |  |  | 1481.3741 | 0.012048
|  |  | 35.5584 | 50.16 | 2.3654 |  |  |  | 992.9598 | 0.097825 | 35.5584 | 79.335 | 64.9267 | 291.4315 |  |  | 1484.553 | 0.011152
|  |  | 35.6352 | 49.595 | 2.1965 |  |  |  | 995.086 | 0.050987 | 35.6352 | 79.39 | 66.3808 | 286.469 |  |  | 1487.7319 | 0.013159
|  |  | 35.712 | 48.185 | 2.1658 |  |  |  | 997.2123 | 0.025726 | 35.712 | 79.26 | 66.9798 | 282.4312 |  |  | 1490.9108 | 0.0094174
|  |  | 35.7888 | 46.81 | 2.1094 |  |  |  | 999.3385 | 0.057868 | 35.7888 | 78.99 | 66.6317 | 279.2872 |  |  | 1494.0898 | 0.012046
|  |  | 35.8656 | 45.81 | 1.9149 |  |  |  | 1001.4648 | 0.068245 | 35.8656 | 78.575 | 65.2134 | 277.0371 |  |  | 1497.2687 | 0.031173
|  |  | 35.9424 | 44.715 | 1.7357 |  |  |  | 1003.591 | 0.053076 | 35.9424 | 78.18 | 63.6211 | 275.5576 |  |  | 1500.4476 | 0.042222
|  |  | 36.0192 | 43.885 | 1.5514 |  |  |  | 1005.7173 | 0.032546 | 36.0192 | 77.675 | 62.9555 | 274.6638 |  |  | 1503.6265 | 0.036424
|  |  | 36.096 | 42.77 | 1.408 |  |  |  | 1007.8435 | 0.038368 | 36.096 | 77.35 | 61.9674 | 274.0473 |  |  | 1506.8054 | 0.017789
|  |  | 36.1728 | 41.53 | 1.3056 |  |  |  | 1009.9698 | 0.064119 | 36.1728 | 76.865 | 62.2234 | 273.585 |  |  | 1509.9843 | 0.010761
|  |  | 36.2496 | 40.655 | 1.1469 |  |  |  | 1012.096 | 0.10125 | 36.2496 | 76.65 | 61.6192 | 273.4 |  |  | 1513.1632 | 0.020064
|  |  | 36.3264 | 40.69 | 0.79872 |  |  |  | 1014.2223 | 0.14276 | 36.3264 | 76.64 | 60.3648 | 274.8179 |  |  | 1516.3422 | 0.017976
|  |  | 36.4032 | 39.575 | 0.6656 |  |  |  | 1016.3485 | 0.17126 | 36.4032 | 76.555 | 60.8154 | 279.5954 |  |  | 1519.5211 | 0.0081576
|  |  | 36.48 | 38.58 | 0.55808 |  |  |  | 1018.4748 | 0.17098 | 36.48 | 76.875 | 60.9331 | 288.4724 |  |  | 1522.7 | 0.0054946
|  |  | 36.5568 | 37.955 | 0.42496 |  |  |  | 1020.601 | 0.13814 | 36.5568 | 76.93 | 61.7114 | 301.8496 |  |  | 1525.8789 | 0.0091247
|  |  | 36.6336 | 0 | 0 |  |  |  | 1022.7273 | 0.086968 | 36.6336 | 76.725 | 62.3616 | 318.3399 |  |  | 1529.0578 | 0.008646
|  |  | 36.7104 | 0 | 0 |  |  |  | 1024.8535 | 0.059769 | 36.7104 | 76.78 | 61.9622 | 335.1385 |  |  | 1532.2367 | 0.007189
|  |  | 36.7872 | 0 | 0 |  |  |  | 1026.9798 | 0.076953 | 36.7872 | 76.89 | 61.9827 | 349.1321 |  |  | 1535.4156 | 0.007713
|  |  | 36.864 | 0 | 0 |  |  |  | 1029.106 | 0.078174 | 36.864 | 76.88 | 61.9776 | 358.0708 |  |  | 1538.5946 | 0.010459
|  |  | 36.9408 | 0 | 0 |  |  |  | 1031.2323 | 0.039685 | 36.9408 | 76.855 | 62.1005 | 361.7079 |  |  | 1541.7735 | 0.011729
|  |  | 37.0176 | 0 | 0 |  |  |  | 1033.3585 | 0.039562 | 37.0176 | 76.715 | 63.5699 | 361.2764 |  |  | 1544.9524 | 0.0079327
|  |  | 37.0944 | 0 | 0 |  |  |  | 1035.4848 | 0.12371 | 37.0944 | 76.97 | 63.9027 | 358.6873 |  |  | 1548.1313 | 0.0019973
|  |  | 37.1712 | 0 | 0 |  |  |  | 1037.611 | 0.19127 | 37.1712 | 77.455 | 63.3139 | 356.1598 |  |  | 1551.3102 | 0.011448
|  |  | 37.248 | 0 | 0 |  |  |  | 1039.7373 | 0.21774 | 37.248 | 77.68 | 63.3856 | 355.081 |  |  | 1554.4891 | 0.017695
|  |  | 37.3248 | 0 | 0 |  |  |  | 1041.8636 | 0.19618 | 37.3248 | 77.575 | 63.9795 | 355.5741 |  |  | 1557.6681 | 0.017705
|  |  | 37.4016 | 0 | 0 |  |  |  | 1043.9898 | 0.13904 | 37.4016 | 77.215 | 64.5325 | 356.9303 |  |  | 1560.847 | 0.013964
|  |  | 37.4784 | 0 | 0 |  |  |  | 1046.1161 | 0.070919 | 37.4784 | 76.905 | 63.2115 | 357.3002 |  |  | 1564.0259 | 0.011
|  |  | 37.5552 | 0 | 0 |  |  |  | 1048.2423 | 0.018817 | 37.5552 | 76.525 | 61.0509 | 355.0501 |  |  | 1567.2048 | 0.0086163
|  |  | 37.632 | 0 | 0 |  |  |  | 1050.3686 | 0.042114 | 37.632 | 75.94 | 59.5251 | 349.7486 |  |  | 1570.3837 | 0.011049
|  |  | 37.7088 | 0 | 0 |  |  |  | 1052.4948 | 0.082328 | 37.7088 | 75.4 | 57.7638 | 341.7654 |  |  | 1573.5626 | 0.019824
|  |  | 37.7856 | 0 | 0 |  |  |  | 1054.6211 | 0.11395 | 37.7856 | 74.71 | 57.4771 | 332.0562 |  |  | 1576.7415 | 0.024447
|  |  | 37.8624 | 0 | 0 |  |  |  | 1056.7473 | 0.12149 | 37.8624 | 74.27 | 58.8902 | 321.638 |  |  | 1579.9205 | 0.019592
|  |  | 37.9392 | 0 | 0 |  |  |  | 1058.8736 | 0.10293 | 37.9392 | 74.63 | 59.1872 | 314.5487 |  |  | 1583.0994 | 0.010469
|  |  | 38.016 | 0 | 0 |  |  |  | 1060.9998 | 0.092361 | 38.016 | 75.055 | 59.6326 | 315.9049 |  |  | 1586.2783 | 0.019054
|  |  | 38.0928 | 0 | 0 |  |  |  | 1063.1261 | 0.11899 | 38.0928 | 75.15 | 60.759 | 313.4391 |  |  | 1589.4572 | 0.031714
|  |  | 38.1696 | 0 | 0 |  |  |  | 1065.2523 | 0.13842 | 38.1696 | 75.065 | 60.7488 | 306.9971 |  |  | 1592.6361 | 0.036612
|  |  | 38.2464 | 0 | 0 |  |  |  | 1067.3786 | 0.11888 | 38.2464 | 74.81 | 59.5302 | 301.3565 |  |  | 1595.815 | 0.030177
|  |  | 38.3232 | 0 | 0 |  |  |  | 1069.5048 | 0.084407 | 38.3232 | 74.455 | 58.711 | 296.9488 |  |  | 1598.9939 | 0.012886
|  |  | 38.4 | 0 | 0 |  |  |  | 1071.6311 | 0.11694 | 38.4 | 74.46 | 56.7654 | 293.7124 |  |  | 1602.1729 | 0.014278
|  |  | 38.4768 | 0 | 0 |  |  |  | 1073.7573 | 0.17782 | 38.4768 | 74.465 | 56.2534 | 291.5239 |  |  | 1605.3518 | 0.028998
|  |  | 38.5536 | 0 | 0 |  |  |  | 1075.8836 | 0.20866 | 38.5536 | 74.415 | 57.4208 | 290.1369 |  |  | 1608.5307 | 0.02175
|  |  | 38.6304 | 0 | 0 |  |  |  | 1078.0098 | 0.20882 | 38.6304 | 74.215 | 60.0832 | 289.3663 |  |  | 1611.7096 | 0.012561
|  |  | 38.7072 | 0 | 0 |  |  |  | 1080.1361 | 0.20617 | 38.7072 | 74.69 | 59.4842 | 289.0889 |  |  | 1614.8885 | 0.061845
|  |  | 38.784 | 0 | 0 |  |  |  | 1082.2623 | 0.20826 | 38.784 | 75.22 | 57.7946 | 289.3663 |  |  | 1618.0674 | 0.10171
|  |  | 38.8608 | 0 | 0 |  |  |  | 1084.3886 | 0.18622 | 38.8608 | 75.19 | 57.8406 | 290.0444 |  |  | 1621.2463 | 0.10963
|  |  | 38.9376 | 0 | 0 |  |  |  | 1086.5148 | 0.12241 | 38.9376 | 74.965 | 57.5488 | 290.9999 |  |  | 1624.4253 | 0.079897
|  |  | 39.0144 | 0 | 0 | [5 empty] | 39.0144 | 74.72 | 56.361 | 295.5617 |  |  | 1627.6042 | 0.028836
|  |  | 39.0912 | 0 | 0 | [5 empty] | 39.0912 | 74.25 | 56.0282 | 311.4972 |  |  | 1630.7831 | 0.019648
|  |  | 39.168 | 0 | 0 | [5 empty] | 39.168 | 73.92 | 55.895 | 332.3952 |  |  | 1633.962 | 0.036134
|  |  | 39.2448 | 0 | 0 | [5 empty] | 39.2448 | 74 | 55.6288 | 348.1766 |  |  | 1637.1409 | 0.022425
|  |  | 39.3216 | 0 | 0 | [5 empty] | 39.3216 | 74.265 | 56.2893 | 357.4543 |  |  | 1640.3198 | 0.0077797
|  |  | 39.3984 | 0 | 0 | [5 empty] | 39.3984 | 74.535 | 58.6035 | 361.8004 |  |  | 1643.4987 | 0.025773
|  |  | 39.4752 | 0 | 0 | [5 empty] | 39.4752 | 75.405 | 59.3101 | 363.6806 |  |  | 1646.6777 | 0.02272
|  |  | 39.552 | 0 | 0 | [5 empty] | 39.552 | 76.33 | 60.3341 | 365.9923 |  |  | 1649.8566 | 0.0045776
|  |  | 39.6288 | 0 | 0 | [5 empty] | 39.6288 | 77.02 | 62.0698 | 370.3692 |  |  | 1653.0355 | 0.019819
|  |  | 39.7056 | 0 | 0 | [5 empty] | 39.7056 | 77.305 | 64.3379 | 376.9345 |  |  | 1656.2144 | 0.026514
|  |  | 39.7824 | 0 | 0 | [5 empty] | 39.7824 | 77.695 | 64.343 | 384.2395 |  |  | 1659.3933 | 0.017384
|  |  | 39.8592 | 0 | 0 | [5 empty] | 39.8592 | 77.81 | 64.2202 | 390.0651 |  |  | 1662.5722 | 0.0091108
|  |  | 39.936 | 0 | 0 | [5 empty] | 39.936 | 77.55 | 64.809 | 392.5926 |  |  | 1665.7511 | 0.014596
|  |  | 40.0128 | 0 | 0 | [5 empty] | 40.0128 | 77.23 | 65.3568 | 390.9897 |  |  | 1668.9301 | 0.012333
|  |  | 40.0896 | 0 | 0 | [5 empty] | 40.0896 | 77.24 | 64.297 | 385.3183 |  |  | 1672.109 | 0.0085634
|  |  | 40.1664 | 0 | 0 | [5 empty] | 40.1664 | 77.315 | 63.5597 | 377.0269 |  |  | 1675.2879 | 0.014177
|  |  | 40.2432 | 0 | 0 | [5 empty] | 40.2432 | 77.27 | 63.6723 | 368.1499 |  |  | 1678.4668 | 0.015929
|  |  | 40.32 | 0 | 0 | [5 empty] | 40.32 | 77.245 | 63.5034 | 360.5058 |  |  | 1681.6457 | 0.010856
|  |  | 40.3968 | 0 | 0 | [5 empty] | 40.3968 | 77.26 | 62.3053 | 355.081 |  |  | 1684.8246 | 0.0015217
|  |  | 40.4736 | 0 | 0 | [5 empty] | 40.4736 | 76.965 | 61.5066 | 351.3822 |  |  | 1688.0035 | 0.010348
|  |  | 40.5504 | 0 | 0 | [5 empty] | 40.5504 | 76.475 | 60.8512 | 348.2074 |  |  | 1691.1825 | 0.017702
|  |  | 40.6272 | 0 | 0 | [5 empty] | 40.6272 | 75.95 | 60.1344 | 343.9539 |  |  | 1694.3614 | 0.015745
|  |  | 40.704 | 0 | 0 | [5 empty] | 40.704 | 75.665 | 59.1206 | 338.2516 |  |  | 1697.5403 | 0.0059673
|  |  | 40.7808 | 0 | 0 | [5 empty] | 40.7808 | 75.665 | 59.8938 | 331.7171 |  |  | 1700.7192 | 0.011237
|  |  | 40.8576 | 0 | 0 | [5 empty] | 40.8576 | 76.19 | 61.609 | 324.1347 |  |  | 1703.8981 | 0.01611
|  |  | 40.9344 | 0 | 0 | [5 empty] | 40.9344 | 77.075 | 62.8429 | 316.9529 |  |  | 1707.077 | 0.011781
|  |  | 41.0112 | 0 | 0 | [5 empty] | 41.0112 | 78.01 | 63.9437 | 311.0349 |  |  | 1710.2559 | 0.015539
|  |  | 41.088 | 0 | 0 | [5 empty] | 41.088 | 78.735 | 65.1674 | 306.4423 |  |  | 1713.4349 | 0.022501
|  |  | 41.1648 | 0 | 0 | [5 empty] | 41.1648 | 79.185 | 66.2426 | 302.9593 |  |  | 1716.6138 | 0.016432
|  |  | 41.2416 | 0 | 0 | [5 empty] | 41.2416 | 79.44 | 66.9952 | 300.401 |  |  | 1719.7927 | 0.0018329
|  |  | 41.3184 | 0 | 0 | [5 empty] | 41.3184 | 79.74 | 66.2579 | 298.5824 |  |  | 1722.9716 | 0.020951
|  |  | 41.3952 | 0 | 0 | [5 empty] | 41.3952 | 79.875 | 65.2339 | 297.3495 |  |  | 1726.1505 | 0.02918
|  |  | 41.472 | 0 | 0 | [5 empty] | 41.472 | 79.77 | 65.2442 | 296.4556 |  |  | 1729.3294 | 0.023454
|  |  | 41.5488 | 0 | 0 | [5 empty] | 41.5488 | 79.635 | 65.5718 | 295.6542 |  |  | 1732.5083 | 0.011945
|  |  | 41.6256 | 0 | 0 | [5 empty] | 41.6256 | 79.61 | 65.664 | 294.8528 |  |  | 1735.6873 | 0.0086759
|  |  | 41.7024 | 0 | 0 | [5 empty] | 41.7024 | 79.66 | 66.2733 | 293.9898 |  |  | 1738.8662 | 0.008293
|  |  | 41.7792 | 0 | 0 | [5 empty] | 41.7792 | 79.605 | 67.3741 | 293.5583 |  |  | 1742.0451 | 0.0050132
|  |  | 41.856 | 0 | 0 | [5 empty] | 41.856 | 79.705 | 68.009 | 295.0994 |  |  | 1745.224 | 0.0038032
|  |  | 41.9328 | 0 | 0 | [5 empty] | 41.9328 | 80.225 | 67.7786 | 300.1852 |  |  | 1748.4029 | 0.0042924
|  |  | 42.0096 | 0 | 0 | [5 empty] | 42.0096 | 80.68 | 69.2992 | 309.4629 |  |  | 1751.5818 | 0.0079164
|  |  | 42.0864 | 0 | 0 | [5 empty] | 42.0864 | 81.005 | 70.5587 | 323.0251 |  |  | 1754.7607 | 0.0096022
|  |  | 42.1632 | 0 | 0 | [5 empty] | 42.1632 | 81.35 | 70.2362 | 339.3921 |  |  | 1757.9397 | 0.011584
|  |  | 42.24 | 0 | 0 | [5 empty] | 42.24 | 81.835 | 69.8266 | 355.6049 |  |  | 1761.1186 | 0.018789
|  |  | 42.3168 | 0 | 0 | [5 empty] | 42.3168 | 82.185 | 70.0979 | 368.5506 |  |  | 1764.2975 | 0.022228
|  |  | 42.3936 | 0 | 0 | [5 empty] | 42.3936 | 82.45 | 70.2771 | 375.9789 |  |  | 1767.4764 | 0.016331
|  |  | 42.4704 | 0 | 0 | [5 empty] | 42.4704 | 82.59 | 70.5741 | 377.6434 |  |  | 1770.6553 | 0.0088759
|  |  | 42.5472 | 0 | 0 | [5 empty] | 42.5472 | 82.65 | 71.5725 | 375.0234 |  |  | 1773.8342 | 0.01412
|  |  | 42.624 | 0 | 0 | [5 empty] | 42.624 | 82.695 | 73.0214 | 370.3075 |  |  | 1777.0131 | 0.015632
|  |  | 42.7008 | 0 | 0 | [5 empty] | 42.7008 | 82.895 | 73.4976 | 365.9307 |  |  | 1780.1921 | 0.012823
|  |  | 42.7776 | 0 | 0 | [5 empty] | 42.7776 | 83.145 | 73.0317 | 363.2182 |  |  | 1783.371 | 0.014591
|  |  | 42.8544 | 0 | 0 | [5 empty] | 42.8544 | 83.085 | 73.0266 | 362.2935 |  |  | 1786.5499 | 0.013268
|  |  | 42.9312 | 0 | 0 | [5 empty] | 42.9312 | 82.805 | 73.0266 | 361.9853 |  |  | 1789.7288 | 0.0022695
|  |  | 43.008 | 0 | 0 | [5 empty] | 43.008 | 82.395 | 71.977 | 360.7832 |  |  | 1792.9077 | 0.014532
|  |  | 43.0848 | 0 | 0 | [5 empty] | 43.0848 | 81.745 | 70.7021 | 357.2077 |  |  | 1796.0866 | 0.028808
|  |  | 43.1616 | 0 | 0 | [5 empty] | 43.1616 | 81.1 | 69.079 | 350.3651 |  |  | 1799.2655 | 0.034008
|  |  | 43.2384 | 0 | 0 | [5 empty] | 43.2384 | 80.605 | 66.729 | 340.7791 |  |  | 1802.4445 | 0.02828
|  |  | 43.3152 | 0 | 0 | [5 empty] | 43.3152 | 80.03 | 65.9149 | 329.5287 |  |  | 1805.6234 | 0.015563
|  |  | 43.392 | 0 | 0 | [5 empty] | 43.392 | 79.665 | 67.0822 | 317.5077 |  |  | 1808.8023 | 0.011776
|  |  | 43.4688 | 0 | 0 | [5 empty] | 43.4688 | 80 | 67.2512 | 306.6888 |  |  | 1811.9812 | 0.018104
|  |  | 43.5456 | 0 | 0 | [5 empty] | 43.5456 | 80.5 | 66.7392 | 304.6237 |  |  | 1815.1601 | 0.014777
|  |  | 43.6224 | 0 | 0 | [5 empty] | 43.6224 | 80.595 | 66.9952 | 304.2538 |  |  | 1818.339 | 0.0021596
|  |  | 43.6992 | 0 | 0 | [5 empty] | 43.6992 | 80.46 | 66.8723 | 297.0104 |  |  | 1821.5179 | 0.013535
|  |  | 43.776 | 0 | 0 | [5 empty] | 43.776 | 80.065 | 66.7238 | 289.2739 |  |  | 1824.6969 | 0.019444
|  |  | 43.8528 | 0 | 0 | [5 empty] | 43.8528 | 79.61 | 66.0224 | 282.462 |  |  | 1827.8758 | 0.013394
|  |  | 43.9296 | 0 | 0 | [5 empty] | 43.9296 | 79.46 | 64.2304 | 276.6981 |  |  | 1831.0547 | 0.0037625
|  |  | 44.0064 | 0 | 0 | [5 empty] | 44.0064 | 79.3 | 63.4778 | 272.1363 |  |  | 1834.2336 | 0.010985
|  |  | 44.0832 | 0 | 0 | [5 empty] | 44.0832 | 79.11 | 63.9334 | 268.7149 |  |  | 1837.4125 | 0.014405
|  |  | 44.16 | 0 | 0 | [5 empty] | 44.16 | 78.89 | 65.1571 | 266.2491 |  |  | 1840.5914 | 0.01661
|  |  | 44.2368 | 0 | 0 | [5 empty] | 44.2368 | 78.76 | 65.5872 | 264.6154 |  |  | 1843.7703 | 0.020226
|  |  | 44.3136 | 0 | 0 | [5 empty] | 44.3136 | 78.77 | 64.6707 | 263.7524 |  |  | 1846.9493 | 0.019954
|  |  | 44.3904 | 0 | 0 | [5 empty] | 44.3904 | 78.8 | 62.9965 | 263.5367 |  |  | 1850.1282 | 0.015271
|  |  | 44.4672 | 0 | 0 | [5 empty] | 44.4672 | 78.31 | 62.6278 | 263.6599 |  |  | 1853.3071 | 0.011231
|  |  | 44.544 | 0 | 0 | [5 empty] | 44.544 | 77.525 | 62.3514 | 264.8312 |  |  | 1856.486 | 0.012428
|  |  | 44.6208 | 0 | 0 | [5 empty] | 44.6208 | 76.93 | 61.0765 | 273.6774 |  |  | 1859.6649 | 0.015487
|  |  | 44.6976 | 0 | 0 | [5 empty] | 44.6976 | 76.67 | 58.7878 | 291.8013 |  |  | 1862.8438 | 0.013094
|  |  | 44.7744 | 0 | 0 | [5 empty] | 44.7744 | 76.26 | 58.3885 | 308.3841 |  |  | 1866.0227 | 0.0023219
|  |  | 44.8512 | 0 | 0 | [5 empty] | 44.8512 | 76.135 | 58.3936 | 319.2338 |  |  | 1869.2017 | 0.011894
|  |  | 44.928 | 0 | 0 | [5 empty] | 44.928 | 76.18 | 59.5302 | 325.4601 |  |  | 1872.3806 | 0.019157
|  |  | 45.0048 | 0 | 0 | [5 empty] | 45.0048 | 76.675 | 59.991 | 329.3438 |  |  | 1875.5595 | 0.013088
|  |  | 45.0816 | 0 | 0 | [5 empty] | 45.0816 | 77.37 | 60.6822 | 333.1966 |  |  | 1878.7384 | 0.0061337
|  |  | 45.1584 | 0 | 0 | [5 empty] | 45.1584 | 78.025 | 61.5526 | 338.529 |  |  | 1881.9173 | 0.018723
|  |  | 45.2352 | 0 | 0 | [5 empty] | 45.2352 | 78.435 | 62.3258 | 345.6491 |  |  | 1885.0962 | 0.018749
|  |  | 45.312 | 0 | 0 | [5 empty] | 45.312 | 78.54 | 63.0016 | 353.6015 |  |  | 1888.2751 | 0.0038442
|  |  | 45.3888 | 0 | 0 | [5 empty] | 45.3888 | 78.405 | 63.5699 | 360.6291 |  |  | 1891.4541 | 0.016262
|  |  | 45.4656 | 0 | 0 | [5 empty] | 45.4656 | 78.345 | 62.6278 | 365.0676 |  |  | 1894.633 | 0.027529
|  |  | 45.5424 | 0 | 0 | [5 empty] | 45.5424 | 78.26 | 61.4195 | 365.8382 |  |  | 1897.8119 | 0.02482
|  |  | 45.6192 | 0 | 0 | [5 empty] | 45.6192 | 77.965 | 61.5424 | 362.9716 |  |  | 1900.9908 | 0.015439
|  |  | 45.696 | 0 | 0 | [5 empty] | 45.696 | 77.505 | 62.807 | 357.3002 |  |  | 1904.1697 | 0.0098954
|  |  | 45.7728 | 0 | 0 | [5 empty] | 45.7728 | 77.385 | 62.7354 | 350.3651 |  |  | 1907.3486 | 0.0029281
|  |  | 45.8496 | 0 | 0 | [5 empty] | 45.8496 | 77.445 | 61.7062 | 343.7689 |  |  | 1910.5275 | 0.0087692
|  |  | 45.9264 | 0 | 0 | [5 empty] | 45.9264 | 77.22 | 61.4195 | 338.529 |  |  | 1913.7065 | 0.016047
|  |  | 46.0032 | 0 | 0 | [5 empty] | 46.0032 | 76.985 | 60.0678 | 334.6761 |  |  | 1916.8854 | 0.01464
|  |  | 46.08 | 0 | 0 | [5 empty] | 46.08 | 76.33 | 59.2896 | 331.4089 |  |  | 1920.0643 | 0.016022
|  |  | 46.1568 | 0 | 0 | [5 empty] | 46.1568 | 75.425 | 58.9824 | 327.4327 |  |  | 1923.2432 | 0.021693
|  |  | 46.2336 | 0 | 0 | [5 empty] | 46.2336 | 74.62 | 58.2042 | 321.8229 |  |  | 1926.4221 | 0.016704
|  |  | 46.3104 | 0 | 0 | [5 empty] | 46.3104 | 74.365 | 56.3456 | 315.1035 |  |  | 1929.601 | 0.005817
|  |  | 46.3872 | 0 | 0 | [5 empty] | 46.3872 | 74.545 | 56.0384 | 307.4902 |  |  | 1932.7799 | 0.019221
|  |  | 46.464 | 0 | 0 | [5 empty] | 46.464 | 74.87 | 57.3696 | 299.6304 |  |  | 1935.9589 | 0.024232
|  |  | 46.5408 | 0 | 0 | [5 empty] | 46.5408 | 75.34 | 58.409 | 292.9418 |  |  | 1939.1378 | 0.01517
|  |  | 46.6176 | 0 | 0 | [5 empty] | 46.6176 | 75.775 | 59.3254 | 287.8252 |  |  | 1942.3167 | 0.0021929
|  |  | 46.6944 | 0 | 0 | [5 empty] | 46.6944 | 76.155 | 59.9091 | 284.2497 |  |  | 1945.4956 | 0.0081445
|  |  | 46.7712 | 0 | 0 | [5 empty] | 46.7712 | 76.325 | 60.2675 | 282.0613 |  |  | 1948.6745 | 0.0087393
|  |  | 46.848 | 0 | 0 | [5 empty] | 46.848 | 76.54 | 59.095 | 280.9208 |  |  | 1951.8534 | 0.012036
|  |  | 46.9248 | 0 | 0 | [5 empty] | 46.9248 | 76.505 | 58.0762 | 280.4893 |  |  | 1955.0323 | 0.012384
|  |  | 47.0016 | 0 | 0 | [5 empty] | 47.0016 | 76.065 | 58.153 | 280.4276 |  |  | 1958.2113 | 0.0068107
|  |  | 47.0784 | 0 | 0 | [5 empty] | 47.0784 | 75.735 | 57.513 | 280.5818 |  |  | 1961.3902 | 0.016518
|  |  | 47.1552 | 0 | 0 | [5 empty] | 47.1552 | 75.64 | 56.4941 | 280.8283 |  |  | 1964.5691 | 0.021848
|  |  | 47.232 | 0 | 0 | [5 empty] | 47.232 | 75.345 | 57.4003 | 281.2291 |  |  | 1967.748 | 0.011536
|  |  | 47.3088 | 0 | 0 | [5 empty] | 47.3088 | 75.225 | 57.8714 | 281.938 |  |  | 1970.9269 | 0.015973
|  |  | 47.3856 | 0 | 0 | [5 empty] | 47.3856 | 75.285 | 57.4464 | 283.9415 |  |  | 1974.1058 | 0.03925
|  |  | 47.4624 | 0 | 0 | [5 empty] | 47.4624 | 75.5 | 57.8714 | 288.9965 |  |  | 1977.2847 | 0.052734
|  |  | 47.5392 | 0 | 0 | [5 empty] | 47.5392 | 76.15 | 58.2963 | 297.6885 |  |  | 1980.4637 | 0.054295
|  |  | 47.616 | 0 | 0 | [5 empty] | 47.616 | 76.445 | 60.3187 | 310.2951 |  |  | 1983.6426 | 0.046656
|  |  | 47.6928 | 0 | 0 | [5 empty] | 47.6928 | 76.655 | 61.2454 | 326.0765 |  |  | 1986.8215 | 0.030638
|  |  | 47.7696 | 0 | 0 | [5 empty] | 47.7696 | 76.95 | 61.5526 | 342.8442 |  |  | 1990.0004 | 0.011461
|  |  | 47.8464 | 0 | 0 | [5 empty] | 47.8464 | 77.245 | 61.7933 | 357.7934 |  |  | 1993.1793 | 0.013979
|  |  | 47.9232 | 0 | 0 | [5 empty] | 47.9232 | 77.32 | 61.9418 | 368.4581 |  |  | 1996.3582 | 0.016793
|  |  | 48 | 0 | 0 | [5 empty] | 48 | 77.16 | 62.9402 | 373.7289 |  |  | 1999.5372 | 0.0098063
|  |  | 48.0768 | 0 | 0 | [5 empty] | 48.0768 | 77.18 | 63.8464 | 374.0063 |  |  | 2002.7161 | 0.015272
|  |  | 48.1536 | 0 | 0 | [5 empty] | 48.1536 | 77.485 | 64.3328 | 371.0781 |  |  | 2005.895 | 0.018295
|  |  | 48.2304 | 0 | 0 | [5 empty] | 48.2304 | 78 | 64.4762 | 367.1328 |  |  | 2009.0739 | 0.0094974
|  |  | 48.3072 | 0 | 0 | [5 empty] | 48.3072 | 78.36 | 65.0342 | 364.1121 |  |  | 2012.2528 | 0.016268
|  |  | 48.384 | 0 | 0 | [5 empty] | 48.384 | 78.425 | 66.2426 | 362.9716 |  |  | 2015.4317 | 0.025394
|  |  | 48.4608 | 0 | 0 | [5 empty] | 48.4608 | 78.34 | 66.5651 | 363.4956 |  |  | 2018.6106 | 0.020863
|  |  | 48.5376 | 0 | 0 | [5 empty] | 48.5376 | 78.105 | 66.0941 | 364.3895 |  |  | 2021.7896 | 0.018983
|  |  | 48.6144 | 0 | 0 | [5 empty] | 48.6144 | 77.575 | 65.5206 | 363.9272 |  |  | 2024.9685 | 0.028473
|  |  | 48.6912 | 0 | 0 | [5 empty] | 48.6912 | 77.19 | 63.4266 | 360.814 |  |  | 2028.1474 | 0.026497
|  |  | 48.768 | 0 | 0 | [5 empty] | 48.768 | 76.54 | 62.3923 | 354.6494 |  |  | 2031.3263 | 0.011216
|  |  | 48.8448 | 0 | 0 | [5 empty] | 48.8448 | 76.085 | 61.3325 | 345.9574 |  |  | 2034.5052 | 0.010114
|  |  | 48.9216 | 0 | 0 | [5 empty] | 48.9216 | 76.205 | 60.1754 | 335.6317 |  |  | 2037.6841 | 0.015791
|  |  | 48.9984 | 0 | 0 | [5 empty] | 48.9984 | 76.55 | 60.9587 | 325.2751 |  |  | 2040.863 | 0.011926
|  |  | 49.0752 | 0 | 0 | [5 empty] | 49.0752 | 77.13 | 61.9776 | 323.1483 |  |  | 2044.042 | 0.018289
|  |  | 49.152 | 0 | 0 | [5 empty] | 49.152 | 77.655 | 63.0938 | 326.2923 |  |  | 2047.2209 | 0.02341
|  |  | 49.2288 | 0 | 0 | [5 empty] | 49.2288 | 77.84 | 64.0512 | 321.9154 |  |  | 2050.3998 | 0.014675
|  |  | 49.3056 | 35.345 | 0.68096 | [5 empty] | 49.3056 | 77.855 | 63.9027 | 315.9974 |  |  | 2053.5787 | 0.0029678
|  |  | 49.3824 | 39.24 | 0.6144 | [5 empty] | 49.3824 | 77.83 | 63.1347 | 310.8191 |  |  | 2056.7576 | 0.016967
|  |  | 49.4592 | 40.615 | 0.87552 | [5 empty] | 49.4592 | 77.49 | 64.1638 | 306.6888 |  |  | 2059.9365 | 0.018607
|  |  | 49.536 | 41.865 | 1.152 | [5 empty] | 49.536 | 77.435 | 64.425 | 303.7607 |  |  | 2063.1154 | 0.0097402
|  |  | 49.6128 | 43.79 | 1.28 | [5 empty] | 49.6128 | 77.84 | 63.6928 | 301.9113 |  |  | 2066.2944 | 0.0041479
|  |  | 49.6896 | 45.455 | 1.4336 | [5 empty] | 49.6896 | 78.4 | 63.7798 | 300.8017 |  |  | 2069.4733 | 0.0069025
|  |  | 49.7664 | 46.395 | 1.7254 | [5 empty] | 49.7664 | 78.705 | 65.408 | 300.1236 |  |  | 2072.6522 | 0.0053358
|  |  | 49.8432 | 47.78 | 1.8893 | [5 empty] | 49.8432 | 78.96 | 66.9338 | 299.8461 |  |  | 2075.8311 | 0.0071884
|  |  | 49.92 | 49.745 | 1.8688 | [5 empty] | 49.92 | 79.32 | 67.0413 | 299.9078 |  |  | 2079.01 | 0.0077638
|  |  | 49.9968 | 51.015 | 1.9661 | [5 empty] | 49.9968 | 79.29 | 67.5277 | 300.0002 |  |  | 2082.1889 | 0.0057289
|  |  | 50.0736 | 51.735 | 2.1146 | [5 empty] | 50.0736 | 78.985 | 67.6915 | 300.0002 |  |  | 2085.3678 | 0.0081193
|  |  | 50.1504 | 52.715 | 2.1453 | [5 empty] | 50.1504 | 78.805 | 66.2323 | 302.5894 |  |  | 2088.5468 | 0.01051
|  |  | 50.2272 | 53.485 | 2.1709 | [5 empty] | 50.2272 | 78.71 | 64.7885 | 315.6275 |  |  | 2091.7257 | 0.012243
|  |  | 50.304 | 53.58 | 2.3091 | [5 empty] | 50.304 | 78.48 | 64.809 | 334.9227 |  |  | 2094.9046 | 0.014775
|  |  | 50.3808 | 53.96 | 2.3501 | [5 empty] | 50.3808 | 78.61 | 64.4864 | 350.4575 |  |  | 2098.0835 | 0.013885
|  |  | 50.4576 | 54.725 | 2.2835 | [5 empty] | 50.4576 | 79.14 | 64.1638 | 360.4133 |  |  | 2101.2624 | 0.0078786
|  |  | 50.5344 | 55 | 2.3194 | [5 empty] | 50.5344 | 79.625 | 65.9661 | 366.2389 |  |  | 2104.4413 | 0.00040554
|  |  | 50.6112 | 54.86 | 2.4525 | [5 empty] | 50.6112 | 80.165 | 69.0227 | 369.6602 |  |  | 2107.6202 | 0.0042132
|  |  | 50.688 | 55.175 | 2.4986 | [5 empty] | 50.688 | 81.005 | 70.9888 | 372.7117 |  |  | 2110.7992 | 0.0043317
|  |  | 50.7648 | 55.64 | 2.5037 | [5 empty] | 50.7648 | 81.745 | 72.2893 | 376.6879 |  |  | 2113.9781 | 0.0040194
|  |  | 50.8416 | 55.805 | 2.5805 | [5 empty] | 50.8416 | 82.13 | 73.472 | 381.8353 |  |  | 2117.157 | 0.0075232
|  |  | 50.9184 | 56.28 | 2.5293 | [5 empty] | 50.9184 | 82.28 | 73.6768 | 387.2602 |  |  | 2120.3359 | 0.010755
|  |  | 50.9952 | 56.58 | 2.5139 | [5 empty] | 50.9952 | 82.09 | 74.1325 | 391.5754 |  |  | 2123.5148 | 0.00923
|  |  | 51.072 | 56.205 | 2.6778 | [5 empty] | 51.072 | 82 | 73.1085 | 393.3323 |  |  | 2126.6937 | 0.0010407
|  |  | 51.1488 | 55.89 | 2.8058 | [5 empty] | 51.1488 | 81.805 | 72.0998 | 391.6678 |  |  | 2129.8726 | 0.010003
|  |  | 51.2256 | 56.21 | 2.729 | [5 empty] | 51.2256 | 81.675 | 70.912 | 386.4896 |  |  | 2133.0516 | 0.015701
|  |  | 51.3024 | 56.23 | 2.7238 | [5 empty] | 51.3024 | 81.52 | 70.103 | 378.6297 |  |  | 2136.2305 | 0.010579
|  |  | 51.3792 | 56.12 | 2.7546 | [5 empty] | 51.3792 | 81.21 | 70.4256 | 369.6294 |  |  | 2139.4094 | 0.011024
|  |  | 51.456 | 56.015 | 2.7904 | [5 empty] | 51.456 | 80.945 | 70.3539 | 360.9681 |  |  | 2142.5883 | 0.025252
|  |  | 51.5328 | 56.21 | 2.7187 | [5 empty] | 51.5328 | 80.715 | 69.2634 | 353.5706 |  |  | 2145.7672 | 0.02841
|  |  | 51.6096 | 55.935 | 2.7443 | [5 empty] | 51.6096 | 80.13 | 68.5978 | 347.4985 |  |  | 2148.9461 | 0.015241
|  |  | 51.6864 | 55.545 | 2.816 | [5 empty] | 51.6864 | 79.28 | 67.5226 | 342.0736 |  |  | 2152.125 | 0.0071075
|  |  | 51.7632 | 55.955 | 2.6624 | [5 empty] | 51.7632 | 78.795 | 64.2867 | 336.2173 |  |  | 2155.304 | 0.024522
|  |  | 51.84 | 56.005 | 2.5958 | [5 empty] | 51.84 | 78.05 | 63.0016 | 329.1896 |  |  | 2158.4829 | 0.026593
|  |  | 51.9168 | 55.495 | 2.647 | [5 empty] | 51.9168 | 77.575 | 62.3514 | 321.6072 |  |  | 2161.6618 | 0.0135
|  |  | 51.9936 | 55.655 | 2.5395 | [5 empty] | 51.9936 | 77.73 | 62.2182 | 313.3774 |  |  | 2164.8407 | 0.0077217
|  |  | 52.0704 | 55.64 | 2.4525 | [5 empty] | 52.0704 | 78.15 | 62.6534 | 304.8703 |  |  | 2168.0196 | 0.020113
|  |  | 52.1472 | 55.19 | 2.4166 | [5 empty] | 52.1472 | 78.69 | 63.4368 | 297.257 |  |  | 2171.1985 | 0.021141
|  |  | 52.224 | 54.2 | 2.4781 | [5 empty] | 52.224 | 79.13 | 64.6963 | 290.9383 |  |  | 2174.3774 | 0.012286
|  |  | 52.3008 | 53.695 | 2.3347 | [5 empty] | 52.3008 | 79.28 | 66.1146 | 285.76 |  |  | 2177.5564 | 0.0022582
|  |  | 52.3776 | 52.645 | 2.1862 | [5 empty] | 52.3776 | 79.205 | 66.8058 | 281.6606 |  |  | 2180.7353 | 0.0095481
|  |  | 52.4544 | 50.65 | 2.2016 | [5 empty] | 52.4544 | 79.11 | 65.6077 | 278.5783 |  |  | 2183.9142 | 0.012693
|  |  | 52.5312 | 49.23 | 2.0736 | [5 empty] | 52.5312 | 78.86 | 64.169 | 276.359 |  |  | 2187.0931 | 0.010423
|  |  | 52.608 | 48.83 | 1.7459 | [5 empty] | 52.608 | 78.505 | 62.551 | 274.8487 |  |  | 2190.272 | 0.0054787
|  |  | 52.6848 | 48.15 | 1.5104 | [5 empty] | 52.6848 | 78.11 | 61.097 | 273.8007 |  |  | 2193.4509 | 0.0039165
|  |  | 52.7616 | 46.55 | 1.4592 | [5 empty] | 52.7616 | 77.575 | 61.1379 | 272.9993 |  |  | 2196.6298 | 0.010671
|  |  | 52.8384 | 45.175 | 1.3466 | [5 empty] | 52.8384 | 77.285 | 60.8768 | 272.352 |  |  | 2199.8088 | 0.01817
|  |  | 52.9152 | 44.6 | 1.0854 | [5 empty] | 52.9152 | 76.975 | 60.841 | 272.0438 |  |  | 2202.9877 | 0.020765
|  |  | 52.992 | 43.795 | 0.90624 | [5 empty] | 52.992 | 76.545 | 61.3837 | 273.4 |  |  | 2206.1666 | 0.014905
|  |  | 53.0688 | 42.32 | 0.83968 | [5 empty] | 53.0688 | 76.29 | 62.3821 | 278.1468 |  |  | 2209.3455 | 0.0030568
|  |  | 53.1456 | 41.21 | 0.7168 | [5 empty] | 53.1456 | 76.57 | 62.761 | 286.9005 |  |  | 2212.5244 | 0.0073883
|  |  | 53.2224 | 40.765 | 0.52736 | [5 empty] | 53.2224 | 76.805 | 62.8122 | 299.9694 |  |  | 2215.7033 | 0.010458
|  |  | 53.2992 | 40.63 | 0.34816 | [5 empty] | 53.2992 | 76.97 | 61.6806 | 316.1207 |  |  | 2218.8822 | 0.011747
|  |  | 53.376 | 0 | 0 | [5 empty] | 53.376 | 77.065 | 61.0918 | 332.7651 |  |  | 2222.0612 | 0.019337
|  |  | 53.4528 | 0 | 0 | [5 empty] | 53.4528 | 77.085 | 61.5373 | 346.8204 |  |  | 2225.2401 | 0.02317
|  |  | 53.5296 | 0 | 0 | [5 empty] | 53.5296 | 77.105 | 61.5475 | 356.0365 |  |  | 2228.419 | 0.018034
|  |  | 53.6064 | 0 | 0 | [5 empty] | 53.6064 | 77.18 | 61.3888 | 360.0435 |  |  | 2231.5979 | 0.0086534
|  |  | 53.6832 | 0 | 0 | [5 empty] | 53.6832 | 77.105 | 62.423 | 359.8585 |  |  | 2234.7768 | 0.0047441
|  |  | 53.76 | 0 | 0 | [5 empty] | 53.76 | 77.175 | 63.5597 | 357.4235 |  |  | 2237.9557 | 0.0054777
|  |  | 53.8368 | 0 | 0 | [5 empty] | 53.8368 | 77.57 | 63.2627 | 355.0193 |  |  | 2241.1346 | 0.010297
|  |  | 53.9136 | 0 | 0 | [5 empty] | 53.9136 | 77.705 | 63.7184 | 353.8481 |  |  | 2244.3136 | 0.01376
|  |  | 53.9904 | 0 | 0 | [5 empty] | 53.9904 | 77.71 | 63.5802 | 354.033 |  |  | 2247.4925 | 0.01
|  |  | 54.0672 | 0 | 0 | [5 empty] | 54.0672 | 77.455 | 63.3344 | 355.081 |  |  | 2250.6714 | 0.0012269
|  |  | 54.144 | 0 | 0 | [5 empty] | 54.144 | 76.915 | 63.2064 | 355.4508 |  |  | 2253.8503 | 0.0052258
|  |  | 54.2208 | 0 | 0 | [5 empty] | 54.2208 | 76.475 | 61.6499 | 353.5706 |  |  | 2257.0292 | 0.0041425
|  |  | 54.2976 | 0 | 0 | [5 empty] | 54.2976 | 76.105 | 59.3101 | 348.6698 |  |  | 2260.2081 | 0.0047332
|  |  | 54.3744 | 0 | 0 | [5 empty] | 54.3744 | 75.585 | 57.7229 | 340.7791 |  |  | 2263.387 | 0.010992
|  |  | 54.4512 | 0 | 0 | [5 empty] | 54.4512 | 74.8 | 57.8253 | 331.039 |  |  | 2266.566 | 0.01319
|  |  | 54.528 | 0 | 0 | [5 empty] | 54.528 | 74.61 | 57.7843 | 320.5284 |  |  | 2269.7449 | 0.016807
|  |  | 54.6048 | 0 | 0 | [5 empty] | 54.6048 | 74.685 | 59.1206 | 312.576 |  |  | 2272.9238 | 0.021299
|  |  | 54.6816 | 0 | 0 | [5 empty] | 54.6816 | 74.98 | 60.0832 | 313.624 |  |  | 2276.1027 | 0.018293
|  |  | 54.7584 | 0 | 0 | [5 empty] | 54.7584 | 75.5 | 59.3152 | 312.1754 |  |  | 2279.2816 | 0.0098326
|  |  | 54.8352 | 0 | 0 | [5 empty] | 54.8352 | 75.56 | 58.6086 | 305.9491 |  |  | 2282.4605 | 0.011361
|  |  | 54.912 | 0 | 0 | [5 empty] | 54.912 | 75.335 | 57.7536 | 300.1852 |  |  | 2285.6394 | 0.01078
|  |  | 54.9888 | 0 | 0 | [5 empty] | 54.9888 | 75.135 | 56.2995 | 295.5001 |  |  | 2288.8184 | 0.0044866
|  |  | 55.0656 | 0 | 0 | [5 empty] | 55.0656 | 74.855 | 55.7568 | 291.863 |  |  | 2291.9973 | 0.017141
|  |  | 55.1424 | 0 | 0 | [5 empty] | 55.1424 | 74.23 | 57.7075 | 289.5204 |  |  | 2295.1762 | 0.021446
|  |  | 55.2192 | 0 | 0 | [5 empty] | 55.2192 | 74.27 | 57.9226 | 288.195 |  |  | 2298.3551 | 0.0096786
|  |  | 55.296 | 0 | 0 | [5 empty] | 55.296 | 74.835 | 57.088 | 287.6094 |  |  | 2301.534 | 0.0124
|  |  | 55.3728 | 0 | 0 | [5 empty] | 55.3728 | 75.34 | 56.617 | 287.5786 |  |  | 2304.7129 | 0.029452
|  |  | 55.4496 | 0 | 0 | [5 empty] | 55.4496 | 75.27 | 58.025 | 287.9485 |  |  | 2307.8918 | 0.031481
|  |  | 55.5264 | 0 | 0 | [5 empty] | 55.5264 | 75.065 | 58.9773 | 288.4416 |  |  | 2311.0708 | 0.020104
|  |  | 55.6032 | 0 | 0 | [5 empty] | 55.6032 | 74.865 | 58.4499 | 289.0273 |  |  | 2314.2497 | 0.0097465
|  |  | 55.68 | 0 | 0 | [5 empty] | 55.68 | 74.335 | 58.5933 | 293.0034 |  |  | 2317.4286 | 0.0092513
|  |  | 55.7568 | 0 | 0 | [5 empty] | 55.7568 | 74.065 | 57.2467 | 308.2917 |  |  | 2320.6075 | 0.0046841
|  |  | 55.8336 | 0 | 0 | [5 empty] | 55.8336 | 74.165 | 55.2141 | 329.2205 |  |  | 2323.7864 | 0.0042744
|  |  | 55.9104 | 0 | 0 | [5 empty] | 55.9104 | 74.11 | 55.5213 | 345.3717 |  |  | 2326.9653 | 0.00265
|  |  | 55.9872 | 0 | 0 | [5 empty] | 55.9872 | 74.185 | 56.832 | 354.9885 |  |  | 2330.1442 | 0.013163
|  |  | 56.064 | 0 | 0 | [5 empty] | 56.064 | 74.675 | 58.1069 | 359.5811 |  |  | 2333.3232 | 0.03023
|  |  | 56.1408 | 0 | 0 | [5 empty] | 56.1408 | 75.63 | 58.9773 | 361.6463 |  |  | 2336.5021 | 0.03811
|  |  | 56.2176 | 0 | 0 | [5 empty] | 56.2176 | 76.475 | 60.7078 | 363.8655 |  |  | 2339.681 | 0.030336
|  |  | 56.2944 | 0 | 0 | [5 empty] | 56.2944 | 77.275 | 61.8035 | 368.1191 |  |  | 2342.8599 | 0.012328
|  |  | 56.3712 | 0 | 0 | [5 empty] | 56.3712 | 77.885 | 62.2746 | 374.4995 |  |  | 2346.0388 | 0.0029907
|  |  | 56.448 | 0 | 0 | [5 empty] | 56.448 | 78.2 | 62.4691 | 381.6196 |  |  | 2349.2177 | 0.006365
|  |  | 56.5248 | 0 | 0 | [5 empty] | 56.5248 | 78.075 | 63.1654 | 387.4759 |  |  | 2352.3966 | 0.010837
|  |  | 56.6016 | 0 | 0 | [5 empty] | 56.6016 | 77.64 | 64.5274 | 390.2192 |  |  | 2355.5756 | 0.024326
|  |  | 56.6784 | 0 | 0 | [5 empty] | 56.6784 | 77.46 | 64.384 | 388.863 |  |  | 2358.7545 | 0.0303
|  |  | 56.7552 | 0 | 0 | [5 empty] | 56.7552 | 77.575 | 63.0016 | 383.5614 |  |  | 2361.9334 | 0.023514
|  |  | 56.832 | 0 | 0 | [5 empty] | 56.832 | 77.575 | 62.7405 | 375.5166 |  |  | 2365.1123 | 0.0088036
|  |  | 56.9088 | 0 | 0 | [5 empty] | 56.9088 | 77.32 | 63.9078 | 366.7629 |  |  | 2368.2912 | 0.0048856
|  |  | 56.9856 | 0 | 0 | [5 empty] | 56.9856 | 77.305 | 63.6518 | 359.088 |  |  | 2371.4701 | 0.010967
|  |  | 57.0624 | 0 | 0 | [5 empty] | 57.0624 | 77.125 | 63.6058 | 353.324 |  |  | 2374.649 | 0.010216
|  |  | 57.1392 | 0 | 0 | [5 empty] | 57.1392 | 76.76 | 63.1757 | 349.4095 |  |  | 2377.828 | 0.0070242
|  |  | 57.216 | 0 | 0 | [5 empty] | 57.216 | 76.49 | 61.4246 | 346.1731 |  |  | 2381.0069 | 0.0050557
|  |  | 57.2928 | 0 | 0 | [5 empty] | 57.2928 | 76.05 | 60.503 | 342.0428 |  |  | 2384.1858 | 0.0062779
|  |  | 57.3696 | 0 | 0 | [5 empty] | 57.3696 | 75.7 | 59.7658 | 336.3714 |  |  | 2387.3647 | 0.0085879
|  |  | 57.4464 | 0 | 0 | [5 empty] | 57.4464 | 75.675 | 59.9347 | 329.9294 |  |  | 2390.5436 | 0.0094857
|  |  | 57.5232 | 0 | 0 | [5 empty] | 57.5232 | 76.155 | 61.4246 | 322.4702 |  |  | 2393.7225 | 0.0095056
|  |  | 57.6 | 0 | 0 | [5 empty] | 57.6 | 77.045 | 62.5715 | 315.196 |  |  | 2396.9014 | 0.010474
|  |  | 57.6768 | 0 | 0 | [5 empty] | 57.6768 | 78.12 | 63.0835 | 309.1239 |  |  | 2400.0804 | 0.011965
|  |  | 57.7536 | 0 | 0 | [5 empty] | 57.7536 | 78.945 | 64.3942 | 304.3463 |  |  | 2403.2593 | 0.011189
|  |  | 57.8304 | 0 | 0 | [5 empty] | 57.8304 | 79.42 | 65.9046 | 300.74 |  |  | 2406.4382 | 0.00685
|  |  | 57.9072 | 0 | 0 | [5 empty] | 57.9072 | 79.705 | 66.5498 | 298.0584 |  |  | 2409.6171 | 0.0049554
|  |  | 57.984 | 0 | 0 | [5 empty] | 57.984 | 79.845 | 66.4883 | 296.1165 |  |  | 2412.796 | 0.0083019
|  |  | 58.0608 | 0 | 0 | [5 empty] | 58.0608 | 79.77 | 66.6624 | 294.7912 |  |  | 2415.9749 | 0.0084046
|  |  | 58.1376 | 0 | 0 | [5 empty] | 58.1376 | 79.625 | 67.0362 | 293.8973 |  |  | 2419.1538 | 0.012027
|  |  | 58.2144 | 0 | 0 | [5 empty] | 58.2144 | 79.78 | 66.0992 | 293.1884 |  |  | 2422.3328 | 0.019817
|  |  | 58.2912 | 0 | 0 | [5 empty] | 58.2912 | 79.615 | 66.944 | 292.4178 |  |  | 2425.5117 | 0.022229
|  |  | 58.368 | 0 | 0 | [5 empty] | 58.368 | 79.825 | 66.6317 | 291.4931 |  |  | 2428.6906 | 0.016526
|  |  | 58.4448 | 0 | 0 | [5 empty] | 58.4448 | 80.11 | 65.6845 | 290.9075 |  |  | 2431.8695 | 0.0085382
|  |  | 58.5216 | 0 | 0 | [5 empty] | 58.5216 | 80.265 | 65.6486 | 292.2637 |  |  | 2435.0484 | 0.0083858
|  |  | 58.5984 | 0 | 0 | [5 empty] | 58.5984 | 80.325 | 67.6198 | 297.1646 |  |  | 2438.2273 | 0.011304
|  |  | 58.6752 | 0 | 0 | [5 empty] | 58.6752 | 80.855 | 68.7462 | 306.319 |  |  | 2441.4062 | 0.01603
|  |  | 58.752 | 0 | 0 | [5 empty] | 58.752 | 81.345 | 69.4682 | 319.7886 |  |  | 2444.5852 | 0.019177
|  |  | 58.8288 | 0 | 0 | [5 empty] | 58.8288 | 81.585 | 69.7907 | 336.1556 |  |  | 2447.7641 | 0.014717
|  |  | 58.9056 | 0 | 0 | [5 empty] | 58.9056 | 81.845 | 70.4666 | 352.4302 |  |  | 2450.943 | 0.0036934
|  |  | 58.9824 | 0 | 0 | [5 empty] | 58.9824 | 82.02 | 71.7926 | 365.4683 |  |  | 2454.1219 | 0.010557
|  |  | 59.0592 | 0 | 0 | [5 empty] | 59.0592 | 82.25 | 72.0179 | 372.9583 |  |  | 2457.3008 | 0.01717
|  |  | 59.136 | 0 | 0 | [5 empty] | 59.136 | 82.405 | 72.1766 | 374.746 |  |  | 2460.4797 | 0.015932
|  |  | 59.2128 | 0 | 0 | [5 empty] | 59.2128 | 82.515 | 72.6579 | 372.311 |  |  | 2463.6587 | 0.012079
|  |  | 59.2896 | 0 | 0 | [5 empty] | 59.2896 | 82.895 | 72.2688 | 367.965 |  |  | 2466.8376 | 0.013373
|  |  | 59.3664 | 0 | 0 | [5 empty] | 59.3664 | 83.225 | 72.0998 | 363.8963 |  |  | 2470.0165 | 0.015739
|  |  | 59.4432 | 0 | 0 | [5 empty] | 59.4432 | 83.34 | 72.576 | 361.2764 |  |  | 2473.1954 | 0.013241
|  |  | 59.52 | 0 | 0 | [5 empty] | 59.52 | 83.265 | 72.8678 | 360.2592 |  |  | 2476.3743 | 0.005345
|  |  | 59.5968 | 0 | 0 | [5 empty] | 59.5968 | 83.045 | 72.2688 | 359.9818 |  |  | 2479.5532 | 0.0040129
|  |  | 59.6736 | 0 | 0 | [5 empty] | 59.6736 | 82.54 | 71.6339 | 358.6564 |  |  | 2482.7321 | 0.0084745
|  |  | 59.7504 | 0 | 0 | [5 empty] | 59.7504 | 81.98 | 70.4205 | 355.0501 |  |  | 2485.9111 | 0.0063893
|  |  | 59.8272 | 0 | 0 | [5 empty] | 59.8272 | 81.53 | 68.0653 | 348.4232 |  |  | 2489.09 | 0.012502
|  |  | 59.904 | 0 | 0 | [5 empty] | 59.904 | 80.925 | 66.2886 | 339.2071 |  |  | 2492.2689 | 0.021044
|  |  | 59.9808 | 0 | 0 | [5 empty] | 59.9808 | 80.18 | 66.2477 | 328.3266 |  |  | 2495.4478 | 0.019373
|  |  | 60.0576 | 0 | 0 | [5 empty] | 60.0576 | 80.07 | 65.8278 | 316.4906 |  |  | 2498.6267 | 0.0068935
|  |  | 60.1344 | 0 | 0 | [5 empty] | 60.1344 | 80.37 | 65.961 | 305.0552 |  |  | 2501.8056 | 0.011387
|  |  | 60.2112 | 0 | 0 | [5 empty] | 60.2112 | 80.715 | 66.2886 | 301.3257 |  |  | 2504.9845 | 0.018752
|  |  | 60.288 | 0 | 0 | [5 empty] | 60.288 | 80.89 | 66.391 | 301.8496 |  |  | 2508.1635 | 0.012698
|  |  | 60.3648 | 0 | 0 | [5 empty] | 60.3648 | 80.58 | 67.1027 | 295.4076 |  |  | 2511.3424 | 0.010244
|  |  | 60.4416 | 0 | 0 | [5 empty] | 60.4416 | 80.035 | 67.3638 | 287.8252 |  |  | 2514.5213 | 0.022313
|  |  | 60.5184 | 0 | 0 | [5 empty] | 60.5184 | 79.74 | 65.9046 | 280.89 |  |  | 2517.7002 | 0.024306
|  |  | 60.5952 | 0 | 0 | [5 empty] | 60.5952 | 79.48 | 64.4659 | 274.8487 |  |  | 2520.8791 | 0.01591
|  |  | 60.672 | 0 | 0 | [5 empty] | 60.672 | 79.255 | 64.1946 | 270.0403 |  |  | 2524.058 | 0.0098448
|  |  | 60.7488 | 0 | 0 | [5 empty] | 60.7488 | 79.155 | 64.0512 | 266.4957 |  |  | 2527.2369 | 0.0087098
|  |  | 60.8256 | 0 | 0 | [5 empty] | 60.8256 | 79.26 | 63.3958 | 264.0606 |  |  | 2530.4159 | 0.0034363
|  |  | 60.9024 | 0 | 0 | [5 empty] | 60.9024 | 79.195 | 63.7491 | 262.5811 |  |  | 2533.5948 | 0.012754
|  |  | 60.9792 | 0 | 0 | [5 empty] | 60.9792 | 78.925 | 64.4045 | 261.903 |  |  | 2536.7737 | 0.017846
|  |  | 61.056 | 0 | 0 | [5 empty] | 61.056 | 78.795 | 63.6877 | 261.8105 |  |  | 2539.9526 | 0.014024
|  |  | 61.1328 | 0 | 0 | [5 empty] | 61.1328 | 78.42 | 62.8275 | 261.9647 |  |  | 2543.1315 | 0.015977
|  |  | 61.2096 | 0 | 0 | [5 empty] | 61.2096 | 77.78 | 62.3462 | 263.0127 |  |  | 2546.3104 | 0.023832
|  |  | 61.2864 | 0 | 0 | [5 empty] | 61.2864 | 77.04 | 61.8342 | 271.6431 |  |  | 2549.4893 | 0.022236
|  |  | 61.3632 | 0 | 0 | [5 empty] | 61.3632 | 76.56 | 60.3443 | 289.6746 |  |  | 2552.6683 | 0.010513
|  |  | 61.44 | 0 | 0 | [5 empty] | 61.44 | 76.435 | 58.5472 | 306.3498 |  |  | 2555.8472 | 0.0028673
|  |  | 61.5168 | 0 | 0 | [5 empty] | 61.5168 | 76.335 | 58.2195 | 317.3536 |  |  | 2559.0261 | 0.0088401
|  |  | 61.5936 | 0 | 0 | [5 empty] | 61.5936 | 76.525 | 58.8032 | 323.7032 |  |  | 2562.205 | 0.0059742
|  |  | 61.6704 | 0 | 0 | [5 empty] | 61.6704 | 76.89 | 60.2726 | 327.6177 |  |  | 2565.3839 | 0.0068283
|  |  | 61.7472 | 0 | 0 | [5 empty] | 61.7472 | 77.485 | 61.824 | 331.4089 |  |  | 2568.5628 | 0.011105
|  |  | 61.824 | 0 | 0 | [5 empty] | 61.824 | 78.14 | 62.8173 | 336.6796 |  |  | 2571.7417 | 0.011147
|  |  | 61.9008 | 0 | 0 | [5 empty] | 61.9008 | 78.495 | 63.5546 | 343.7381 |  |  | 2574.9207 | 0.014588
|  |  | 61.9776 | 0 | 0 | [5 empty] | 61.9776 | 78.625 | 64.0922 | 351.6904 |  |  | 2578.0996 | 0.020954
|  |  | 62.0544 | 0 | 0 | [5 empty] | 62.0544 | 78.745 | 63.3344 | 358.8105 |  |  | 2581.2785 | 0.022396
|  |  | 62.1312 | 0 | 0 | [5 empty] | 62.1312 | 78.765 | 62.1107 | 363.3107 |  |  | 2584.4574 | 0.017482
|  |  | 62.208 | 0 | 0 | [5 empty] | 62.208 | 78.585 | 61.4963 | 364.1121 |  |  | 2587.6363 | 0.0097162
|  |  | 62.2848 | 0 | 0 | [5 empty] | 62.2848 | 78.26 | 61.6653 | 361.1839 |  |  | 2590.8152 | 0.005631
|  |  | 62.3616 | 0 | 0 | [5 empty] | 62.3616 | 77.94 | 62.1056 | 355.42 |  |  | 2593.9941 | 0.0080586
|  |  | 62.4384 | 0 | 0 | [5 empty] | 62.4384 | 77.815 | 61.8342 | 348.4232 |  |  | 2597.1731 | 0.0085916
|  |  | 62.5152 | 0 | 0 | [5 empty] | 62.5152 | 77.605 | 62.2234 | 341.9195 |  |  | 2600.352 | 0.0051112
|  |  | 62.592 | 0 | 0 | [5 empty] | 62.592 | 77.3 | 62.592 | 336.7721 |  |  | 2603.5309 | 0.0020607
|  |  | 62.6688 | 0 | 0 | [5 empty] | 62.6688 | 77.2 | 60.8102 | 333.0425 |  |  | 2606.7098 | 0.0062236
|  |  | 62.7456 | 0 | 0 | [5 empty] | 62.7456 | 76.6 | 59.7709 | 329.8369 |  |  | 2609.8887 | 0.007454
|  |  | 62.8224 | 0 | 0 | [5 empty] | 62.8224 | 75.725 | 59.4534 | 325.7066 |  |  | 2613.0676 | 0.0057453
|  |  | 62.8992 | 0 | 0 | [5 empty] | 62.8992 | 75.265 | 57.1187 | 319.727 |  |  | 2616.2465 | 0.0040769
|  |  | 62.976 | 0 | 0 | [5 empty] | 62.976 | 74.92 | 55.4547 | 312.5452 |  |  | 2619.4255 | 0.0052458
|  |  | 63.0528 | 0 | 0 | [5 empty] | 63.0528 | 74.86 | 55.9411 | 304.7162 |  |  | 2622.6044 | 0.0084185
|  |  | 63.1296 | 0 | 0 | [5 empty] | 63.1296 | 75.245 | 56.9344 | 296.8255 |  |  | 2625.7833 | 0.010503
|  |  | 63.2064 | 0 | 0 | [5 empty] | 63.2064 | 75.69 | 58.0813 | 290.2602 |  |  | 2628.9622 | 0.0091355
|  |  | 63.2832 | 0 | 0 | [5 empty] | 63.2832 | 76.19 | 58.9466 | 285.3593 |  |  | 2632.1411 | 0.0076628
|  |  | 63.36 | 0 | 0 | [5 empty] | 63.36 | 76.655 | 58.9978 | 281.938 |  |  | 2635.32 | 0.012646
|  |  | 63.4368 | 0 | 0 | [5 empty] | 63.4368 | 76.64 | 59.904 | 279.6879 |  |  | 2638.4989 | 0.017416
|  |  | 63.5136 | 0 | 0 | [5 empty] | 63.5136 | 76.36 | 61.0714 | 278.3317 |  |  | 2641.6779 | 0.017685
|  |  | 63.5904 | 0 | 0 | [5 empty] | 63.5904 | 76.43 | 59.8323 | 277.6228 |  |  | 2644.8568 | 0.014475
|  |  | 63.6672 | 0 | 0 | [5 empty] | 63.6672 | 76.26 | 58.9824 | 277.4995 |  |  | 2648.0357 | 0.010716
|  |  | 63.744 | 0 | 0 | [5 empty] | 63.744 | 76.06 | 58.0301 | 277.7769 |  |  | 2651.2146 | 0.0081469
|  |  | 63.8208 | 0 | 0 | [5 empty] | 63.8208 | 75.92 | 57.5642 | 278.3317 |  |  | 2654.3935 | 0.0085684
|  |  | 63.8976 | 0 | 0 | [5 empty] | 63.8976 | 75.58 | 58.4755 | 279.0406 |  |  | 2657.5724 | 0.010287
|  |  | 63.9744 | 0 | 0 | [5 empty] | 63.9744 | 75.29 | 59.6685 | 279.8728 |  |  | 2660.7513 | 0.0088427
|  |  | 64.0512 | 0 | 0 | [5 empty] | 64.0512 | 75.455 | 58.9312 | 281.7222 |  |  | 2663.9303 | 0.0051699
|  |  | 64.128 | 0 | 0 | [5 empty] | 64.128 | 75.69 | 58.8032 | 286.4381 |  |  | 2667.1092 | 0.0069765
|  |  | 64.2048 | 0 | 0 | [5 empty] | 64.2048 | 76.125 | 59.8784 | 294.7912 |  |  | 2670.2881 | 0.0083911
|  |  | 64.2816 | 0 | 0 | [5 empty] | 64.2816 | 76.54 | 61.1072 | 307.2745 |  |  | 2673.467 | 0.0058014
|  |  | 64.3584 | 0 | 0 | [5 empty] | 64.3584 | 77.185 | 59.8784 | 323.2408 |  |  | 2676.6459 | 0.0069685
|  |  | 64.4352 | 0 | 0 | [5 empty] | 64.4352 | 77.46 | 60.4928 | 340.3476 |  |  | 2679.8248 | 0.0099974
|  |  | 64.512 | 0 | 0 | [5 empty] | 64.512 | 77.64 | 60.8563 | 355.6049 |  |  | 2683.0037 | 0.0084972
|  |  | 64.5888 | 0 | 0 | [5 empty] | 64.5888 | 77.605 | 61.5066 | 366.3622 |  |  | 2686.1827 | 0.0043869
|  |  | 64.6656 | 0 | 0 | [5 empty] | 64.6656 | 77.185 | 63.913 | 371.4171 |  |  | 2689.3616 | 0.0018615
|  |  | 64.7424 | 0 | 0 | [5 empty] | 64.7424 | 77.335 | 64.4096 | 371.4171 |  |  | 2692.5405 | 0.012344
|  |  | 64.8192 | 0 | 0 | [5 empty] | 64.8192 | 77.815 | 64.0768 | 368.3657 |  |  | 2695.7194 | 0.028056
|  |  | 64.896 | 0 | 0 | [5 empty] | 64.896 | 78.475 | 63.488 | 364.5436 |  |  | 2698.8983 | 0.040109
|  |  | 64.9728 | 0 | 0 | [5 empty] | 64.9728 | 78.725 | 64.512 | 361.8004 |  |  | 2702.0772 | 0.04042
|  |  | 65.0496 | 0 | 0 | [5 empty] | 65.0496 | 78.77 | 65.7459 | 360.9373 |  |  | 2705.2561 | 0.029178
|  |  | 65.1264 | 0 | 0 | [5 empty] | 65.1264 | 78.86 | 65.3773 | 361.5538 |  |  | 2708.4351 | 0.01743
|  |  | 65.2032 | 0 | 0 | [5 empty] | 65.2032 | 78.645 | 64.7322 | 362.2935 |  |  | 2711.614 | 0.0146
|  |  | 65.28 | 0 | 0 | [5 empty] | 65.28 | 78.135 | 64.0256 | 361.523 |  |  | 2714.7929 | 0.010245
|  |  | 65.3568 | 0 | 0 | [5 empty] | 65.3568 | 77.605 | 62.5715 | 358.1016 |  |  | 2717.9718 | 0.0093429
|  |  | 65.4336 | 0 | 0 | [5 empty] | 65.4336 | 77.04 | 61.5373 | 351.7521 |  |  | 2721.1507 | 0.014031
|  |  | 65.5104 | 0 | 0 | [5 empty] | 65.5104 | 76.61 | 60.5491 | 343.1525 |  |  | 2724.3296 | 0.0096548
|  |  | 65.5872 | 0 | 0 | [5 empty] | 65.5872 | 76.56 | 60.0422 | 333.1042 |  |  | 2727.5085 | 0.0082239
|  |  | 65.664 | 0 | 0 | [5 empty] | 65.664 | 76.83 | 61.2301 | 322.9634 |  |  | 2730.6875 | 0.018548
|  |  | 65.7408 | 0 | 0 | [5 empty] | 65.7408 | 77.28 | 62.9658 | 320.4976 |  |  | 2733.8664 | 0.01764
|  |  | 65.8176 | 0 | 0 | [5 empty] | 65.8176 | 77.78 | 63.9283 | 323.3333 |  |  | 2737.0453 | 0.0056079
|  |  | 65.8944 | 0 | 0 | [5 empty] | 65.8944 | 78.115 | 63.8157 | 318.9564 |  |  | 2740.2242 | 0.016389
|  |  | 65.9712 | 37.97 | 0.76288 | [5 empty] | 65.9712 | 77.965 | 64.2918 | 313.0384 |  |  | 2743.4031 | 0.02544
|  |  | 66.048 | 39.845 | 0.96256 | [5 empty] | 66.048 | 77.725 | 64.4352 | 307.9834 |  |  | 2746.582 | 0.020808
|  |  | 66.1248 | 42.18 | 1.0854 | [5 empty] | 66.1248 | 77.745 | 63.7542 | 303.9148 |  |  | 2749.7609 | 0.0082593
|  |  | 66.2016 | 44.245 | 1.2288 | [5 empty] | 66.2016 | 77.845 | 63.5085 | 301.0791 |  |  | 2752.9399 | 0.0054985
|  |  | 66.2784 | 45.33 | 1.5667 | [5 empty] | 66.2784 | 78.035 | 63.9795 | 299.2913 |  |  | 2756.1188 | 0.013687
|  |  | 66.3552 | 46.775 | 1.8022 | [5 empty] | 66.3552 | 78.405 | 64.8397 | 298.2125 |  |  | 2759.2977 | 0.024543
|  |  | 66.432 | 49.005 | 1.7869 | [5 empty] | 66.432 | 78.83 | 65.6333 | 297.6269 |  |  | 2762.4766 | 0.031231
|  |  | 66.5088 | 51.29 | 1.751 | [5 empty] | 66.5088 | 79.2 | 66.3808 | 297.442 |  |  | 2765.6555 | 0.025054
|  |  | 66.5856 | 52.77 | 1.8534 | [5 empty] | 66.5856 | 79.415 | 66.7802 | 297.5961 |  |  | 2768.8344 | 0.0070815
|  |  | 66.6624 | 53.495 | 2.0531 | [5 empty] | 66.6624 | 79.31 | 67.7581 | 297.8427 |  |  | 2772.0133 | 0.013368
|  |  | 66.7392 | 54.42 | 2.1658 | [5 empty] | 66.7392 | 79.215 | 67.5328 | 297.9659 |  |  | 2775.1923 | 0.023601
|  |  | 66.816 | 55.395 | 2.2118 | [5 empty] | 66.816 | 79.255 | 65.5616 | 300.5859 |  |  | 2778.3712 | 0.020589
|  |  | 66.8928 | 55.385 | 2.3962 | [5 empty] | 66.8928 | 79.12 | 64.2611 | 313.5316 |  |  | 2781.5501 | 0.010445
|  |  | 66.9696 | 55.685 | 2.4934 | [5 empty] | 66.9696 | 78.86 | 64.0256 | 332.7959 |  |  | 2784.729 | 0.0055458
|  |  | 67.0464 | 56.675 | 2.391 | [5 empty] | 67.0464 | 78.885 | 64.0307 | 348.3615 |  |  | 2787.9079 | 0.0063835
|  |  | 67.1232 | 56.975 | 2.4166 | [5 empty] | 67.1232 | 79.085 | 65.3824 | 358.3482 |  |  | 2791.0868 | 0.0049853
|  |  | 67.2 | 57.325 | 2.4474 | [5 empty] | 67.2 | 79.565 | 66.8314 | 364.1121 |  |  | 2794.2657 | 0.0046289
|  |  | 67.2768 | 57.645 | 2.473 | [5 empty] | 67.2768 | 80.4 | 68.2854 | 367.4718 |  |  | 2797.4447 | 0.0070385
|  |  | 67.3536 | 57.795 | 2.5293 | [5 empty] | 67.3536 | 81.235 | 70.1542 | 370.5541 |  |  | 2800.6236 | 0.008762
|  |  | 67.4304 | 58.1 | 2.5754 | [5 empty] | 67.4304 | 81.955 | 71.5571 | 374.6844 |  |  | 2803.8025 | 0.0049305
|  |  | 67.5072 | 58.155 | 2.6522 | [5 empty] | 67.5072 | 82.315 | 72.8115 | 380.0476 |  |  | 2806.9814 | 0.0052417
|  |  | 67.584 | 58.795 | 2.5907 | [5 empty] | 67.584 | 82.33 | 74.0096 | 385.6882 |  |  | 2810.1603 | 0.015097
|  |  | 67.6608 | 59.28 | 2.5498 | [5 empty] | 67.6608 | 82.315 | 73.2672 | 390.0651 |  |  | 2813.3392 | 0.017936
|  |  | 67.7376 | 59.04 | 2.6573 | [5 empty] | 67.7376 | 82.08 | 72.7603 | 391.6678 |  |  | 2816.5181 | 0.012751
|  |  | 67.8144 | 58.825 | 2.7443 | [5 empty] | 67.8144 | 81.67 | 72.8678 | 389.6644 |  |  | 2819.6971 | 0.009426
|  |  | 67.8912 | 58.955 | 2.729 | [5 empty] | 67.8912 | 81.45 | 72.1152 | 384.2087 |  |  | 2822.876 | 0.010685
|  |  | 67.968 | 59.015 | 2.729 | [5 empty] | 67.968 | 81.33 | 71.6288 | 376.318 |  |  | 2826.0549 | 0.0079219
|  |  | 68.0448 | 59.24 | 2.688 | [5 empty] | 68.0448 | 81.16 | 71.3984 | 367.5951 |  |  | 2829.2338 | 0.010634
|  |  | 68.1216 | 58.71 | 2.7904 | [5 empty] | 68.1216 | 80.925 | 71.1168 | 359.4886 |  |  | 2832.4127 | 0.013281
|  |  | 68.1984 | 58.755 | 2.7699 | [5 empty] | 68.1984 | 80.59 | 70.4563 | 352.6459 |  |  | 2835.5916 | 0.0074526
|  |  | 68.2752 | 59.12 | 2.6726 | [5 empty] | 68.2752 | 80.11 | 69.033 | 346.9129 |  |  | 2838.7705 | 0.010752
|  |  | 68.352 | 58.83 | 2.7034 | [5 empty] | 68.352 | 79.495 | 66.9645 | 341.3647 |  |  | 2841.9495 | 0.017741
|  |  | 68.4288 | 58.235 | 2.7904 | [5 empty] | 68.4288 | 78.72 | 65.2698 | 335.046 |  |  | 2845.1284 | 0.011852
|  |  | 68.5056 | 57.28 | 2.9747 | [5 empty] | 68.5056 | 78 | 64.1997 | 327.4636 |  |  | 2848.3073 | 0.0047255
|  |  | 68.5824 | 57.3 | 2.903 | [5 empty] | 68.5824 | 77.675 | 63.2371 | 319.3571 |  |  | 2851.4862 | 0.019596
|  |  | 68.6592 | 57.32 | 2.8058 | [5 empty] | 68.6592 | 77.915 | 62.935 | 310.8808 |  |  | 2854.6651 | 0.022733
|  |  | 68.736 | 57.235 | 2.7136 | [5 empty] | 68.736 | 78.37 | 63.6109 | 302.4661 |  |  | 2857.844 | 0.018161
|  |  | 68.8128 | 57.19 | 2.6061 | [5 empty] | 68.8128 | 78.765 | 64.9626 | 295.2227 |  |  | 2861.0229 | 0.021049
|  |  | 68.8896 | 56.64 | 2.5549 | [5 empty] | 68.8896 | 79.03 | 66.647 | 289.243 |  |  | 2864.2019 | 0.020466
|  |  | 68.9664 | 55.855 | 2.4525 | [5 empty] | 68.9664 | 79.39 | 66.6675 | 284.373 |  |  | 2867.3808 | 0.0067286
|  |  | 69.0432 | 54.785 | 2.3398 | [5 empty] | 69.0432 | 79.475 | 66.4986 | 280.366 |  |  | 2870.5597 | 0.014905
|  |  | 69.12 | 53.29 | 2.2682 | [5 empty] | 69.12 | 79.315 | 65.6538 | 277.0988 |  |  | 2873.7386 | 0.025742
|  |  | 69.1968 | 51.935 | 2.1248 | [5 empty] | 69.1968 | 78.92 | 64.7322 | 274.6946 |  |  | 2876.9175 | 0.019436
|  |  | 69.2736 | 50.625 | 1.9507 | [5 empty] | 69.2736 | 78.43 | 64.0256 | 273.0609 |  |  | 2880.0964 | 0.002009
|  |  | 69.3504 | 49.475 | 1.7818 | [5 empty] | 69.3504 | 77.675 | 64.5478 | 272.1054 |  |  | 2883.2753 | 0.012245
|  |  | 69.4272 | 48.235 | 1.6486 | [5 empty] | 69.4272 | 77.235 | 63.8669 | 271.6431 |  |  | 2886.4543 | 0.011424
|  |  | 69.504 | 47.09 | 1.5002 | [5 empty] | 69.504 | 77.16 | 62.1568 | 271.3657 |  |  | 2889.6332 | 0.0030476
|  |  | 69.5808 | 45.895 | 1.3568 | [5 empty] | 69.5808 | 76.925 | 61.6192 | 271.2732 |  |  | 2892.8121 | 0.015705
|  |  | 69.6576 | 44.9 | 1.1674 | [5 empty] | 69.6576 | 76.48 | 62.2541 | 272.4137 |  |  | 2895.991 | 0.015676
|  |  | 69.7344 | 43.815 | 0.99328 | [5 empty] | 69.7344 | 76.59 | 61.8342 | 276.6981 |  |  | 2899.1699 | 0.00188
|  |  | 69.8112 | 42.07 | 0.94208 | [5 empty] | 69.8112 | 77.005 | 61.5782 | 284.9278 |  |  | 2902.3488 | 0.016568
|  |  | 69.888 | 40.345 | 0.87552 | [5 empty] | 69.888 | 77.23 | 61.952 | 297.6269 |  |  | 2905.5278 | 0.023644
|  |  | 69.9648 | 39.335 | 0.67584 | [5 empty] | 69.9648 | 77.11 | 62.1261 | 313.809 |  |  | 2908.7067 | 0.016511
|  |  | 70.0416 | 37.63 | 0.58368 | [5 empty] | 70.0416 | 76.91 | 62.3565 | 330.8541 |  |  | 2911.8856 | 0.011416
|  |  | 70.1184 | 0 | 0 | [5 empty] | 70.1184 | 77.25 | 61.2045 | 345.3717 |  |  | 2915.0645 | 0.017569
|  |  | 70.1952 | 0 | 0 | [5 empty] | 70.1952 | 77.265 | 61.3018 | 355.1118 |  |  | 2918.2434 | 0.014758
|  |  | 70.272 | 0 | 0 | [5 empty] | 70.272 | 77.275 | 61.3376 | 359.3037 |  |  | 2921.4223 | 0.0081897
|  |  | 70.3488 | 0 | 0 | [5 empty] | 70.3488 | 77.455 | 61.0202 | 359.0263 |  |  | 2924.6012 | 0.010929
|  |  | 70.4256 | 0 | 0 | [5 empty] | 70.4256 | 77.51 | 62.167 | 356.4372 |  |  | 2927.7802 | 0.0085455
|  |  | 70.5024 | 0 | 0 | [5 empty] | 70.5024 | 77.655 | 63.4368 | 353.9713 |  |  | 2930.9591 | 0.0078981
|  |  | 70.5792 | 0 | 0 | [5 empty] | 70.5792 | 77.75 | 64.4659 | 353.0158 |  |  | 2934.138 | 0.015491
|  |  | 70.656 | 0 | 0 | [5 empty] | 70.656 | 77.6 | 65.2902 | 353.6323 |  |  | 2937.3169 | 0.014965
|  |  | 70.7328 | 0 | 0 | [5 empty] | 70.7328 | 77.35 | 65.1213 | 355.0193 |  |  | 2940.4958 | 0.0067904
|  |  | 70.8096 | 0 | 0 | [5 empty] | 70.8096 | 77.235 | 62.7558 | 355.5741 |  |  | 2943.6747 | 0.0041706
|  |  | 70.8864 | 0 | 0 | [5 empty] | 70.8864 | 76.865 | 60.7386 | 353.6939 |  |  | 2946.8536 | 0.0038171
|  |  | 70.9632 | 0 | 0 | [5 empty] | 70.9632 | 76.175 | 59.9296 | 348.6698 |  |  | 2950.0326 | 0.0028557
|  |  | 71.04 | 0 | 0 | [5 empty] | 71.04 | 75.5 | 59.2896 | 340.8099 |  |  | 2953.2115 | 0.008297
|  |  | 71.1168 | 0 | 0 | [5 empty] | 71.1168 | 75.08 | 57.9635 | 331.224 |  |  | 2956.3904 | 0.0056576
|  |  | 71.1936 | 0 | 0 | [5 empty] | 71.1936 | 74.785 | 58.2963 | 320.9291 |  |  | 2959.5693 | 0.0077027
|  |  | 71.2704 | 0 | 0 | [5 empty] | 71.2704 | 74.81 | 60.0064 | 313.0076 |  |  | 2962.7482 | 0.016905
|  |  | 71.3472 | 0 | 0 | [5 empty] | 71.3472 | 75.335 | 59.7965 | 313.809 |  |  | 2965.9271 | 0.016525
|  |  | 71.424 | 0 | 0 | [5 empty] | 71.424 | 75.71 | 59.392 | 312.1445 |  |  | 2969.106 | 0.0097146
|  |  | 71.5008 | 0 | 0 | [5 empty] | 71.5008 | 75.71 | 58.8698 | 305.4868 |  |  | 2972.285 | 0.016657
|  |  | 71.5776 | 0 | 0 | [5 empty] | 71.5776 | 75.275 | 58.8237 | 299.4763 |  |  | 2975.4639 | 0.021409
|  |  | 71.6544 | 0 | 0 | [5 empty] | 71.6544 | 74.915 | 57.9482 | 294.7603 |  |  | 2978.6428 | 0.015285
|  |  | 71.7312 | 0 | 0 | [5 empty] | 71.7312 | 74.68 | 57.0317 | 291.4623 |  |  | 2981.8217 | 0.010258
|  |  | 71.808 | 0 | 0 | [5 empty] | 71.808 | 74.63 | 56.6835 | 289.3355 |  |  | 2985.0006 | 0.013994
|  |  | 71.8848 | 0 | 0 | [5 empty] | 71.8848 | 74.63 | 57.4925 | 288.0101 |  |  | 2988.1795 | 0.011802
|  |  | 71.9616 | 0 | 0 | [5 empty] | 71.9616 | 74.77 | 58.4909 | 287.2087 |  |  | 2991.3584 | 0.011176
|  |  | 72.0384 | 0 | 0 | [5 empty] | 72.0384 | 75.04 | 58.9107 | 286.7464 |  |  | 2994.5374 | 0.015631
|  |  | 72.1152 | 0 | 0 | [5 empty] | 72.1152 | 75.195 | 59.1667 | 286.7772 |  |  | 2997.7163 | 0.013488
|  |  | 72.192 | 0 | 0 | [5 empty] | 72.192 | 75.25 | 58.7725 | 287.332 |  |  | 3000.8952 | 0.0077503
|  |  | 72.2688 | 0 | 0 | [5 empty] | 72.2688 | 75.035 | 58.6701 | 288.2567 |  |  | 3004.0741 | 0.005921
|  |  | 72.3456 | 0 | 0 | [5 empty] | 72.3456 | 74.79 | 57.6 | 291.9863 |  |  | 3007.253 | 0.0018817
|  |  | 72.4224 | 0 | 0 | [5 empty] | 72.4224 | 74.325 | 57.1853 | 306.5039 |  |  | 3010.4319 | 0.014804
|  |  | 72.4992 | 0 | 0 | [5 empty] | 72.4992 | 73.98 | 56.9446 | 327.4944 |  |  | 3013.6108 | 0.02403
|  |  | 72.576 | 0 | 0 | [5 empty] | 72.576 | 73.9 | 56.9805 | 344.1388 |  |  | 3016.7898 | 0.020039
|  |  | 72.6528 | 0 | 0 | [5 empty] | 72.6528 | 74.18 | 57.4566 | 354.1871 |  |  | 3019.9687 | 0.0049433
|  |  | 72.7296 | 0 | 0 | [5 empty] | 72.7296 | 74.545 | 59.0797 | 359.0263 |  |  | 3023.1476 | 0.016403
|  |  | 72.8064 | 0 | 0 | [5 empty] | 72.8064 | 75.245 | 60.5082 | 361.0914 |  |  | 3026.3265 | 0.025347
|  |  | 72.8832 | 0 | 0 | [5 empty] | 72.8832 | 76.18 | 61.4042 | 363.2799 |  |  | 3029.5054 | 0.021066
|  |  | 72.96 | 0 | 0 | [5 empty] | 72.96 | 76.98 | 62.4845 | 367.3793 |  |  | 3032.6843 | 0.0099269
|  |  | 73.0368 | 0 | 0 | [5 empty] | 73.0368 | 77.575 | 63.5238 | 373.7905 |  |  | 3035.8632 | 0.0040388
|  |  | 73.1136 | 0 | 0 | [5 empty] | 73.1136 | 77.93 | 64.1178 | 381.0956 |  |  | 3039.0422 | 0.0079487
|  |  | 73.1904 | 0 | 0 | [5 empty] | 73.1904 | 77.94 | 64.599 | 387.3526 |  |  | 3042.2211 | 0.012554
|  |  | 73.2672 | 0 | 0 | [5 empty] | 73.2672 | 77.88 | 64.3686 | 390.6199 |  |  | 3045.4 | 0.012917
|  |  | 73.344 | 0 | 0 | [5 empty] | 73.344 | 77.715 | 64.1382 | 389.6952 |  |  | 3048.5789 | 0.0066713
|  |  | 73.4208 | 0 | 0 | [5 empty] | 73.4208 | 77.66 | 63.2064 | 384.5169 |  |  | 3051.7578 | 0.0071559
|  |  | 73.4976 | 0 | 0 | [5 empty] | 73.4976 | 77.7 | 62.3923 | 376.4413 |  |  | 3054.9367 | 0.015313
|  |  | 73.5744 | 0 | 0 | [5 empty] | 73.5744 | 77.76 | 61.9059 | 367.5643 |  |  | 3058.1156 | 0.022544
|  |  | 73.6512 | 0 | 0 | [5 empty] | 73.6512 | 77.715 | 62.1005 | 359.7352 |  |  | 3061.2946 | 0.028555
|  |  | 73.728 | 0 | 0 | [5 empty] | 73.728 | 77.405 | 62.8122 | 353.9713 |  |  | 3064.4735 | 0.026598
|  |  | 73.8048 | 0 | 0 | [5 empty] | 73.8048 | 77.14 | 62.1926 | 350.1801 |  |  | 3067.6524 | 0.014132
|  |  | 73.8816 | 0 | 0 | [5 empty] | 73.8816 | 76.68 | 61.5168 | 347.2211 |  |  | 3070.8313 | 0.014424
|  |  | 73.9584 | 0 | 0 | [5 empty] | 73.9584 | 76.12 | 60.7334 | 343.3066 |  |  | 3074.0102 | 0.025013
|  |  | 74.0352 | 0 | 0 | [5 empty] | 74.0352 | 75.885 | 59.1155 | 337.7892 |  |  | 3077.1891 | 0.020681
|  |  | 74.112 | 0 | 0 | [5 empty] | 74.112 | 76 | 58.8902 | 331.4089 |  |  | 3080.368 | 0.0084007
|  |  | 74.1888 | 0 | 0 | [5 empty] | 74.1888 | 76.52 | 60.5235 | 323.8264 |  |  | 3083.547 | 0.020373
|  |  | 74.2656 | 0 | 0 | [5 empty] | 74.2656 | 77.2 | 62.7251 | 316.3981 |  |  | 3086.7259 | 0.025649
|  |  | 74.3424 | 0 | 0 | [5 empty] | 74.3424 | 77.92 | 64.4762 | 310.1102 |  |  | 3089.9048 | 0.015693
|  |  | 74.4192 | 0 | 0 | [5 empty] | 74.4192 | 78.71 | 65.5514 | 305.1785 |  |  | 3093.0837 | 0.00026579
|  |  | 74.496 | 0 | 0 | [5 empty] | 74.496 | 79.475 | 65.7152 | 301.4181 |  |  | 3096.2626 | 0.0089243
|  |  | 74.5728 | 0 | 0 | [5 empty] | 74.5728 | 79.62 | 67.0566 | 298.6749 |  |  | 3099.4415 | 0.0044706
|  |  | 74.6496 | 0 | 0 | [5 empty] | 74.6496 | 79.9 | 66.5498 | 296.7022 |  |  | 3102.6204 | 0.0084983
|  |  | 74.7264 | 0 | 0 | [5 empty] | 74.7264 | 80.015 | 65.751 | 295.346 |  |  | 3105.7994 | 0.016987
|  |  | 74.8032 | 0 | 0 | [5 empty] | 74.8032 | 79.745 | 66.3552 | 294.3905 |  |  | 3108.9783 | 0.014654
|  |  | 74.88 | 0 | 0 | [5 empty] | 74.88 | 79.69 | 66.2528 | 293.6199 |  |  | 3112.1572 | 0.0043813
|  |  | 74.9568 | 0 | 0 | [5 empty] | 74.9568 | 79.815 | 65.8278 | 292.8185 |  |  | 3115.3361 | 0.0044768
|  |  | 75.0336 | 0 | 0 | [5 empty] | 75.0336 | 79.795 | 66.4576 | 291.9554 |  |  | 3118.515 | 0.0049351
|  |  | 75.1104 | 0 | 0 | [5 empty] | 75.1104 | 79.57 | 68.3622 | 291.4006 |  |  | 3121.6939 | 0.0014524
|  |  | 75.1872 | 0 | 0 | [5 empty] | 75.1872 | 79.72 | 68.5158 | 292.6336 |  |  | 3124.8728 | 0.0057382
|  |  | 75.264 | 0 | 0 | [5 empty] | 75.264 | 80.345 | 67.5584 | 297.257 |  |  | 3128.0518 | 0.0018724
|  |  | 75.3408 | 0 | 0 | [5 empty] | 75.3408 | 80.85 | 68.6285 | 305.9799 |  |  | 3131.2307 | 0.0090398
|  |  | 75.4176 | 0 | 0 | [5 empty] | 75.4176 | 81.225 | 69.8112 | 319.018 |  |  | 3134.4096 | 0.01856
|  |  | 75.4944 | 0 | 0 | [5 empty] | 75.4944 | 81.575 | 69.5296 | 335.0768 |  |  | 3137.5885 | 0.019291
|  |  | 75.5712 | 0 | 0 | [5 empty] | 75.5712 | 81.73 | 70.3437 | 351.5363 |  |  | 3140.7674 | 0.012613
|  |  | 75.648 | 0 | 0 | [5 empty] | 75.648 | 82.02 | 71.168 | 365.1293 |  |  | 3143.9463 | 0.012101
|  |  | 75.7248 | 0 | 0 | [5 empty] | 75.7248 | 82.14 | 72.105 | 373.5748 |  |  | 3147.1252 | 0.01613
|  |  | 75.8016 | 0 | 0 | [5 empty] | 75.8016 | 82.345 | 72.0026 | 376.1639 |  |  | 3150.3042 | 0.015288
|  |  | 75.8784 | 0 | 0 | [5 empty] | 75.8784 | 82.525 | 72.3405 | 374.0371 |  |  | 3153.4831 | 0.012806
|  |  | 75.9552 | 0 | 0 | [5 empty] | 75.9552 | 82.84 | 72.4019 | 369.537 |  |  | 3156.662 | 0.011281
|  |  | 76.032 | 0 | 0 | [5 empty] | 76.032 | 83.07 | 72.8269 | 364.9751 |  |  | 3159.8409 | 0.012069
|  |  | 76.1088 | 0 | 0 | [5 empty] | 76.1088 | 83.075 | 73.7894 | 361.9853 |  |  | 3163.0198 | 0.016366
|  |  | 76.1856 | 0 | 0 | [5 empty] | 76.1856 | 83 | 74.071 | 360.9373 |  |  | 3166.1987 | 0.017294
|  |  | 76.2624 | 0 | 0 | [5 empty] | 76.2624 | 82.865 | 72.9139 | 360.8448 |  |  | 3169.3776 | 0.012914
|  |  | 76.3392 | 0 | 0 | [5 empty] | 76.3392 | 82.725 | 70.6867 | 360.0126 |  |  | 3172.5566 | 0.011513
|  |  | 76.416 | 0 | 0 | [5 empty] | 76.416 | 82.255 | 69.0176 | 356.9303 |  |  | 3175.7355 | 0.013384
|  |  | 76.4928 | 0 | 0 | [5 empty] | 76.4928 | 81.55 | 67.8349 | 350.7349 |  |  | 3178.9144 | 0.011013
|  |  | 76.5696 | 0 | 0 | [5 empty] | 76.5696 | 80.87 | 66.6163 | 341.5188 |  |  | 3182.0933 | 0.010342
|  |  | 76.6464 | 0 | 0 | [5 empty] | 76.6464 | 80.3 | 65.4131 | 330.2376 |  |  | 3185.2722 | 0.013892
|  |  | 76.7232 | 0 | 0 | [5 empty] | 76.7232 | 80.12 | 65.065 | 318.0009 |  |  | 3188.4511 | 0.013944
|  |  | 76.8 | 0 | 0 | [5 empty] | 76.8 | 80.3 | 65.751 | 306.5039 |  |  | 3191.63 | 0.012137
|  |  | 76.8768 | 0 | 0 | [5 empty] | 76.8768 | 80.545 | 66.7238 | 303.5449 |  |  | 3194.809 | 0.011484
|  |  | 76.9536 | 0 | 0 | [5 empty] | 76.9536 | 80.715 | 67.2154 | 304.408 |  |  | 3197.9879 | 0.0086528
|  |  | 77.0304 | 0 | 0 | [5 empty] | 77.0304 | 80.74 | 66.6982 | 297.5344 |  |  | 3201.1668 | 0.0050914
|  |  | 77.1072 | 0 | 0 | [5 empty] | 77.1072 | 80.435 | 65.8278 | 289.428 |  |  | 3204.3457 | 0.0067501
|  |  | 77.184 | 0 | 0 | [5 empty] | 77.184 | 80.125 | 64.1178 | 282.1846 |  |  | 3207.5246 | 0.010509
|  |  | 77.2608 | 0 | 0 | [5 empty] | 77.2608 | 79.81 | 62.8429 | 276.1124 |  |  | 3210.7035 | 0.016368
|  |  | 77.3376 | 0 | 0 | [5 empty] | 77.3376 | 79.505 | 62.9965 | 271.3965 |  |  | 3213.8824 | 0.019549
|  |  | 77.4144 | 0 | 0 | [5 empty] | 77.4144 | 79.16 | 64.1792 | 267.9752 |  |  | 3217.0614 | 0.015743
|  |  | 77.4912 | 0 | 0 | [5 empty] | 77.4912 | 78.965 | 65.2237 | 265.5401 |  |  | 3220.2403 | 0.0086604
|  |  | 77.568 | 0 | 0 | [5 empty] | 77.568 | 78.915 | 65.449 | 263.7832 |  |  | 3223.4192 | 0.006326
|  |  | 77.6448 | 0 | 0 | [5 empty] | 77.6448 | 78.95 | 64.8346 | 262.7353 |  |  | 3226.5981 | 0.0022843
|  |  | 77.7216 | 0 | 0 | [5 empty] | 77.7216 | 78.92 | 63.5853 | 262.3037 |  |  | 3229.777 | 0.011853
|  |  | 77.7984 | 0 | 0 | [5 empty] | 77.7984 | 78.76 | 61.7165 | 262.2729 |  |  | 3232.9559 | 0.023535
|  |  | 77.8752 | 0 | 0 | [5 empty] | 77.8752 | 78.26 | 60.3853 | 263.0435 |  |  | 3236.1348 | 0.026613
|  |  | 77.952 | 0 | 0 | [5 empty] | 77.952 | 77.605 | 59.3408 | 270.6876 |  |  | 3239.3138 | 0.021011
|  |  | 78.0288 | 0 | 0 | [5 empty] | 78.0288 | 76.945 | 58.665 | 288.3492 |  |  | 3242.4927 | 0.017385
|  |  | 78.1056 | 0 | 0 | [5 empty] | 78.1056 | 76.69 | 57.7229 | 305.5792 |  |  | 3245.6716 | 0.019042
|  |  | 78.1824 | 0 | 0 | [5 empty] | 78.1824 | 76.625 | 57.472 | 317.107 |  |  | 3248.8505 | 0.017814
|  |  | 78.2592 | 0 | 0 | [5 empty] | 78.2592 | 76.645 | 58.1478 | 323.734 |  |  | 3252.0294 | 0.017594
|  |  | 78.336 | 0 | 0 | [5 empty] | 78.336 | 77.075 | 58.5216 | 327.7718 |  |  | 3255.2083 | 0.020134
|  |  | 78.4128 | 0 | 0 | [5 empty] | 78.4128 | 77.61 | 59.5712 | 331.4706 |  |  | 3258.3872 | 0.017902
|  |  | 78.4896 | 0 | 0 | [5 empty] | 78.4896 | 78.05 | 61.6806 | 336.618 |  |  | 3261.5662 | 0.010315
|  |  | 78.5664 | 0 | 0 | [5 empty] | 78.5664 | 78.6 | 62.2848 | 343.6456 |  |  | 3264.7451 | 0.0069642
|  |  | 78.6432 | 0 | 0 | [5 empty] | 78.6432 | 78.665 | 63.3702 | 351.5671 |  |  | 3267.924 | 0.010768
|  |  | 78.72 | 0 | 0 | [5 empty] | 78.72 | 78.685 | 63.2832 | 358.7489 |  |  | 3271.1029 | 0.01017
|  |  | 78.7968 | 0 | 0 | [5 empty] | 78.7968 | 78.655 | 62.4896 | 363.4648 |  |  | 3274.2818 | 0.0031572
|  |  | 78.8736 | 0 | 0 | [5 empty] | 78.8736 | 78.675 | 61.097 | 364.6053 |  |  | 3277.4607 | 0.0097563
|  |  | 78.9504 | 0 | 0 | [5 empty] | 78.9504 | 78.505 | 60.6208 | 362.0161 |  |  | 3280.6396 | 0.019402
|  |  | 79.0272 | 0 | 0 | [5 empty] | 79.0272 | 78.335 | 60.2061 | 356.5296 |  |  | 3283.8186 | 0.019263
|  |  | 79.104 | 0 | 0 | [5 empty] | 79.104 | 78.18 | 60.2982 | 349.6869 |  |  | 3286.9975 | 0.0077022
|  |  | 79.1808 | 0 | 0 | [5 empty] | 79.1808 | 78.02 | 60.4672 | 343.0292 |  |  | 3290.1764 | 0.009539
|  |  | 79.2576 | 0 | 0 | [5 empty] | 79.2576 | 77.74 | 60.5389 | 337.6351 |  |  | 3293.3553 | 0.018871
|  |  | 79.3344 | 0 | 0 | [5 empty] | 79.3344 | 77.255 | 60.4262 | 333.6281 |  |  | 3296.5342 | 0.014824
|  |  | 79.4112 | 0 | 0 | [5 empty] | 79.4112 | 76.72 | 59.3562 | 330.2684 |  |  | 3299.7131 | 0.0014511
|  |  | 79.488 | 0 | 0 | [5 empty] | 79.488 | 75.99 | 58.3936 | 326.3231 |  |  | 3302.892 | 0.0093392
|  |  | 79.5648 | 0 | 0 | [5 empty] | 79.5648 | 75.35 | 56.8218 | 320.8058 |  |  | 3306.071 | 0.0075829
|  |  | 79.6416 | 0 | 0 | [5 empty] | 79.6416 | 74.835 | 56.0282 | 314.0864 |  |  | 3309.2499 | 0.0081364
|  |  | 79.7184 | 0 | 0 | [5 empty] | 79.7184 | 74.85 | 56.0947 | 306.6272 |  |  | 3312.4288 | 0.023149
|  |  | 79.7952 | 0 | 0 | [5 empty] | 79.7952 | 75.19 | 56.9907 | 298.7982 |  |  | 3315.6077 | 0.027612
|  |  | 79.872 | 0 | 0 | [5 empty] | 79.872 | 75.59 | 58.1376 | 291.9246 |  |  | 3318.7866 | 0.018827
|  |  | 79.9488 | 0 | 0 | [5 empty] | 79.9488 | 76.09 | 58.9261 | 286.5306 |  |  | 3321.9655 | 0.01095
|  |  | 80.0256 | 0 | 0 | [5 empty] | 80.0256 | 76.44 | 59.6019 | 282.6469 |  |  | 3325.1444 | 0.014755
|  |  | 80.1024 | 0 | 0 | [5 empty] | 80.1024 | 76.685 | 59.6122 | 280.1502 |  |  | 3328.3234 | 0.013339
|  |  | 80.1792 | 0 | 0 | [5 empty] | 80.1792 | 76.72 | 59.3357 | 278.8249 |  |  | 3331.5023 | 0.014571
|  |  | 80.256 | 0 | 0 | [5 empty] | 80.256 | 76.525 | 59.1309 | 278.3009 |  |  | 3334.6812 | 0.01819
|  |  | 80.3328 | 0 | 0 | [5 empty] | 80.3328 | 76.175 | 59.2589 | 278.2701 |  |  | 3337.8601 | 0.012147
|  |  | 80.4096 | 0 | 0 | [5 empty] | 80.4096 | 75.87 | 59.1309 | 278.4242 |  |  | 3341.039 | 0.0041631
|  |  | 80.4864 | 0 | 0 | [5 empty] | 80.4864 | 75.8 | 58.3526 | 278.6091 |  |  | 3344.2179 | 0.01298
|  |  | 80.5632 | 0 | 0 | [5 empty] | 80.5632 | 75.68 | 58.1786 | 278.8557 |  |  | 3347.3969 | 0.01289
|  |  | 80.64 | 0 | 0 | [5 empty] | 80.64 | 75.655 | 57.8202 | 279.3488 |  |  | 3350.5758 | 0.015248
|  |  | 80.7168 | 0 | 0 | [5 empty] | 80.7168 | 75.6 | 58.0813 | 281.0133 |  |  | 3353.7547 | 0.024455
|  |  | 80.7936 | 0 | 0 | [5 empty] | 80.7936 | 75.715 | 58.3526 | 285.6984 |  |  | 3356.9336 | 0.022795
|  |  | 80.8704 | 0 | 0 | [5 empty] | 80.8704 | 75.945 | 60.073 | 294.0514 |  |  | 3360.1125 | 0.0073328
|  |  | 80.9472 | 0 | 0 | [5 empty] | 80.9472 | 76.36 | 61.6397 | 306.4114 |  |  | 3363.2914 | 0.012873
|  |  | 81.024 | 0 | 0 | [5 empty] | 81.024 | 76.765 | 62.0442 | 322.0695 |  |  | 3366.4703 | 0.022896
|  |  | 81.1008 | 0 | 0 | [5 empty] | 81.1008 | 77.17 | 62.0339 | 338.8989 |  |  | 3369.6493 | 0.01783
|  |  | 81.1776 | 0 | 0 | [5 empty] | 81.1776 | 77.545 | 61.783 | 353.9713 |  |  | 3372.8282 | 0.003851
|  |  | 81.2544 | 0 | 0 | [5 empty] | 81.2544 | 77.645 | 61.6141 | 364.8827 |  |  | 3376.0071 | 0.0068749
|  |  | 81.3312 | 0 | 0 | [5 empty] | 81.3312 | 77.47 | 62.423 | 370.4616 |  |  | 3379.186 | 0.0067435
|  |  | 81.408 | 0 | 0 | [5 empty] | 81.408 | 77.39 | 63.2678 | 371.1089 |  |  | 3382.3649 | 0.0041964
|  |  | 81.4848 | 0 | 0 | [5 empty] | 81.4848 | 77.75 | 63.4624 | 368.5506 |  |  | 3385.5438 | 0.012041
|  |  | 81.5616 | 0 | 0 | [5 empty] | 81.5616 | 78.15 | 64.256 | 364.7902 |  |  | 3388.7227 | 0.013069
|  |  | 81.6384 | 0 | 0 | [5 empty] | 81.6384 | 78.765 | 63.7645 | 361.6771 |  |  | 3391.9017 | 0.0074825
|  |  | 81.7152 | 0 | 0 | [5 empty] | 81.7152 | 78.92 | 64.4762 | 360.2284 |  |  | 3395.0806 | 0.0030137
|  |  | 81.792 | 0 | 0 | [5 empty] | 81.792 | 78.66 | 65.9354 | 360.4133 |  |  | 3398.2595 | 0.0032905
|  |  | 81.8688 | 0 | 0 | [5 empty] | 81.8688 | 78.565 | 64.855 | 361.1223 |  |  | 3401.4384 | 0.0058579
|  |  | 81.9456 | 0 | 0 | [5 empty] | 81.9456 | 78.11 | 64.1638 | 360.814 |  |  | 3404.6173 | 0.0071082
|  |  | 82.0224 | 0 | 0 | [5 empty] | 82.0224 | 77.47 | 63.7133 | 358.0708 |  |  | 3407.7962 | 0.0022655
|  |  | 82.0992 | 0 | 0 | [5 empty] | 82.0992 | 76.93 | 62.6432 | 352.3377 |  |  | 3410.9751 | 0.0092337
|  |  | 82.176 | 0 | 0 | [5 empty] | 82.176 | 76.545 | 61.312 | 343.9847 |  |  | 3414.1541 | 0.016453
|  |  | 82.2528 | 0 | 0 | [5 empty] | 82.2528 | 76.555 | 60.4877 | 333.7823 |  |  | 3417.333 | 0.015683
|  |  | 82.3296 | 0 | 0 | [5 empty] | 82.3296 | 76.845 | 61.527 | 322.9942 |  |  | 3420.5119 | 0.014493
|  |  | 82.4064 | 0 | 0 | [5 empty] | 82.4064 | 77.325 | 62.7507 | 319.1105 |  |  | 3423.6908 | 0.020417
|  |  | 82.4832 | 0 | 0 | [5 empty] | 82.4832 | 77.855 | 63.5187 | 322.7784 |  |  | 3426.8697 | 0.020475
|  |  | 82.56 | 0 | 0 | [5 empty] | 82.56 | 78.16 | 64.0461 | 319.2338 |  |  | 3430.0486 | 0.010504
|  |  | 82.6368 | 39.805 | 0.75776 | [5 empty] | 82.6368 | 78.075 | 64.4454 | 313.2233 |  |  | 3433.2275 | 0.0062573
|  |  | 82.7136 | 41.62 | 0.94208 | [5 empty] | 82.7136 | 77.855 | 64.7322 | 307.891 |  |  | 3436.4065 | 0.012992
|  |  | 82.7904 | 43.565 | 1.111 | [5 empty] | 82.7904 | 77.88 | 63.9642 | 303.5449 |  |  | 3439.5854 | 0.011254
|  |  | 82.8672 | 45.76 | 1.239 | [5 empty] | 82.8672 | 77.93 | 63.9437 | 300.401 |  |  | 3442.7643 | 0.0095358
|  |  | 82.944 | 47.775 | 1.3824 | [5 empty] | 82.944 | 78.12 | 64.2099 | 298.3975 |  |  | 3445.9432 | 0.012572
|  |  | 83.0208 | 48.82 | 1.6794 | [5 empty] | 83.0208 | 78.55 | 64.3021 | 297.1954 |  |  | 3449.1221 | 0.011134
|  |  | 83.0976 | 50.39 | 1.833 | [5 empty] | 83.0976 | 78.995 | 64.7424 | 296.4865 |  |  | 3452.301 | 0.0090158
|  |  | 83.1744 | 52.605 | 1.7971 | [5 empty] | 83.1744 | 79.3 | 65.6333 | 296.209 |  |  | 3455.4799 | 0.0093471
|  |  | 83.2512 | 54.14 | 1.8842 | [5 empty] | 83.2512 | 79.48 | 66.1555 | 296.3323 |  |  | 3458.6589 | 0.0057273
|  |  | 83.328 | 55.475 | 1.9917 | [5 empty] | 83.328 | 79.565 | 65.7562 | 296.4556 |  |  | 3461.8378 | 0.010168
|  |  | 83.4048 | 56.54 | 2.0838 | [5 empty] | 83.4048 | 79.5 | 64.8397 | 296.4556 |  |  | 3465.0167 | 0.01657
|  |  | 83.4816 | 57.045 | 2.217 | [5 empty] | 83.4816 | 79.105 | 65.367 | 298.644 |  |  | 3468.1956 | 0.013734
|  |  | 83.5584 | 57.955 | 2.217 | [5 empty] | 83.5584 | 78.89 | 64.9984 | 310.9732 |  |  | 3471.3745 | 0.012124
|  |  | 83.6352 | 58.355 | 2.2835 | [5 empty] | 83.6352 | 78.84 | 64.0666 | 330.0835 |  |  | 3474.5534 | 0.024694
|  |  | 83.712 | 58.135 | 2.4525 | [5 empty] | 83.712 | 78.87 | 64.1075 | 345.8032 |  |  | 3477.7323 | 0.031136
|  |  | 83.7888 | 58.185 | 2.5498 | [5 empty] | 83.7888 | 79.04 | 65.4541 | 356.0056 |  |  | 3480.9113 | 0.024044
|  |  | 83.8656 | 58.42 | 2.5958 | [5 empty] | 83.8656 | 79.46 | 67.159 | 362.047 |  |  | 3484.0902 | 0.0068185
|  |  | 83.9424 | 58.985 | 2.5549 | [5 empty] | 83.9424 | 80.295 | 68.2906 | 365.7457 |  |  | 3487.2691 | 0.012021
|  |  | 84.0192 | 59.88 | 2.4627 | [5 empty] | 84.0192 | 81.24 | 69.888 | 369.0438 |  |  | 3490.448 | 0.023319
|  |  | 84.096 | 60.525 | 2.4474 | [5 empty] | 84.096 | 82.02 | 71.3779 | 373.0816 |  |  | 3493.6269 | 0.023344
|  |  | 84.1728 | 60.245 | 2.5958 | [5 empty] | 84.1728 | 82.645 | 71.4035 | 378.1982 |  |  | 3496.8058 | 0.013788
|  |  | 84.2496 | 60.16 | 2.6931 | [5 empty] | 84.2496 | 82.705 | 72.1254 | 383.5614 |  |  | 3499.9847 | 0.0011422
|  |  | 84.3264 | 60.775 | 2.6266 | [5 empty] | 84.3264 | 82.585 | 72.1357 | 387.815 |  |  | 3503.1637 | 0.0088235
|  |  | 84.4032 | 61.05 | 2.601 | [5 empty] | 84.4032 | 82.31 | 71.7261 | 389.6335 |  |  | 3506.3426 | 0.010604
|  |  | 84.48 | 60.45 | 2.7443 | [5 empty] | 84.48 | 81.965 | 71.4138 | 388.1849 |  |  | 3509.5215 | 0.0067571
|  |  | 84.5568 | 59.905 | 2.8979 | [5 empty] | 84.5568 | 81.68 | 71.3318 | 383.4073 |  |  | 3512.7004 | 0.0064733
|  |  | 84.6336 | 60.245 | 2.8365 | [5 empty] | 84.6336 | 81.525 | 71.04 | 376.0406 |  |  | 3515.8793 | 0.010385
|  |  | 84.7104 | 60.825 | 2.7136 | [5 empty] | 84.7104 | 81.49 | 70.3693 | 367.4102 |  |  | 3519.0582 | 0.01171
|  |  | 84.7872 | 60.945 | 2.6778 | [5 empty] | 84.7872 | 81.355 | 69.5245 | 358.8722 |  |  | 3522.2371 | 0.00718
|  |  | 84.864 | 60.895 | 2.6982 | [5 empty] | 84.864 | 81.085 | 68.4288 | 351.3206 |  |  | 3525.4161 | 0.0052178
|  |  | 84.9408 | 60.68 | 2.729 | [5 empty] | 84.9408 | 80.61 | 67.415 | 344.9094 |  |  | 3528.595 | 0.021034
|  |  | 85.0176 | 60.355 | 2.7546 | [5 empty] | 85.0176 | 79.875 | 66.2886 | 339.1763 |  |  | 3531.7739 | 0.031132
|  |  | 85.0944 | 60.405 | 2.6931 | [5 empty] | 85.0944 | 79.135 | 64.5888 | 333.1658 |  |  | 3534.9528 | 0.027377
|  |  | 85.1712 | 60.115 | 2.7034 | [5 empty] | 85.1712 | 78.475 | 62.9197 | 326.2306 |  |  | 3538.1317 | 0.010897
|  |  | 85.248 | 59.595 | 2.7597 | [5 empty] | 85.248 | 78.045 | 61.993 | 318.7098 |  |  | 3541.3106 | 0.010165
|  |  | 85.3248 | 59.145 | 2.7802 | [5 empty] | 85.3248 | 78.03 | 62.3565 | 310.7575 |  |  | 3544.4895 | 0.018407
|  |  | 85.4016 | 58.85 | 2.7597 | [5 empty] | 85.4016 | 78.4 | 63.5546 | 302.3428 |  |  | 3547.6685 | 0.012115
|  |  | 85.4784 | 58.855 | 2.6522 | [5 empty] | 85.4784 | 79.155 | 63.4112 | 294.6371 |  |  | 3550.8474 | 0.014525
|  |  | 85.5552 | 58.01 | 2.6829 | [5 empty] | 85.5552 | 79.61 | 64.215 | 288.1026 |  |  | 3554.0263 | 0.027629
|  |  | 85.632 | 57.01 | 2.6522 | [5 empty] | 85.632 | 79.805 | 65.2134 | 282.6161 |  |  | 3557.2052 | 0.027954
|  |  | 85.7088 | 55.99 | 2.5037 | [5 empty] | 85.7088 | 79.92 | 64.983 | 278.2392 |  |  | 3560.3841 | 0.015397
|  |  | 85.7856 | 54.885 | 2.3245 | [5 empty] | 85.7856 | 79.685 | 64.7834 | 274.9103 |  |  | 3563.563 | 0.0058515
|  |  | 85.8624 | 54.16 | 2.0531 | [5 empty] | 85.8624 | 78.945 | 65.6538 | 272.6294 |  |  | 3566.7419 | 0.0098121
|  |  | 85.9392 | 52.625 | 1.9354 | [5 empty] | 85.9392 | 78.33 | 65.1622 | 271.1499 |  |  | 3569.9209 | 0.0068663
|  |  | 86.016 | 50.7 | 1.92 | [5 empty] | 86.016 | 77.865 | 64.1946 | 270.2561 |  |  | 3573.0998 | 0.011345
|  |  | 86.0928 | 49.41 | 1.7715 | [5 empty] | 86.0928 | 77.685 | 62.4435 | 269.6088 |  |  | 3576.2787 | 0.015773
|  |  | 86.1696 | 48.4 | 1.5872 | [5 empty] | 86.1696 | 77.525 | 61.1738 | 268.9615 |  |  | 3579.4576 | 0.010482
|  |  | 86.2464 | 47.46 | 1.4029 | [5 empty] | 86.2464 | 77.33 | 60.3341 | 268.5608 |  |  | 3582.6365 | 0.0010483
|  |  | 86.3232 | 46.16 | 1.2544 | [5 empty] | 86.3232 | 77.085 | 60.119 | 269.6396 |  |  | 3585.8154 | 0.0094149
|  |  | 86.4 | 44.83 | 1.1264 | [5 empty] | 86.4 | 76.91 | 61.0406 | 274.1089 |  |  | 3588.9943 | 0.010964
|  |  | 86.4768 | 43.17 | 1.0394 | [5 empty] | 86.4768 | 77.05 | 62.0032 | 282.6161 |  |  | 3592.1733 | 0.012247
|  |  | 86.5536 | 41.995 | 0.86016 | [5 empty] | 86.5536 | 77.195 | 62.5101 | 295.4693 |  |  | 3595.3522 | 0.017634
|  |  | 86.6304 | 40.505 | 0.75264 | [5 empty] | 86.6304 | 77.28 | 61.7472 | 311.6822 |  |  | 3598.5311 | 0.019726
|  |  | 86.7072 | 37.51 | 0.78848 | [5 empty] | 86.7072 | 77.265 | 61.3274 | 328.4807 |  |  | 3601.71 | 0.016753
|  |  | 86.784 | 0 | 0 | [5 empty] | 86.784 | 77.275 | 61.737 | 342.6901 |  |  | 3604.8889 | 0.01174
|  |  | 86.8608 | 0 | 0 | [5 empty] | 86.8608 | 77.285 | 61.8086 | 352.1219 |  |  | 3608.0678 | 0.0077055
|  |  | 86.9376 | 0 | 0 | [5 empty] | 86.9376 | 77.225 | 62.1261 | 356.3447 |  |  | 3611.2467 | 0.0069318
|  |  | 87.0144 | 0 | 0 | [5 empty] | 87.0144 | 77.38 | 62.4589 | 356.3755 |  |  | 3614.4257 | 0.0071241
|  |  | 87.0912 | 0 | 0 | [5 empty] | 87.0912 | 77.615 | 62.7712 | 354.3104 |  |  | 3617.6046 | 0.0068448
|  |  | 87.168 | 0 | 0 | [5 empty] | 87.168 | 77.97 | 62.807 | 352.1219 |  |  | 3620.7835 | 0.0078019
|  |  | 87.2448 | 0 | 0 | [5 empty] | 87.2448 | 78.525 | 61.5219 | 350.9199 |  |  | 3623.9624 | 0.0051478
|  |  | 87.3216 | 0 | 0 | [5 empty] | 87.3216 | 78.345 | 62.4435 | 351.1973 |  |  | 3627.1413 | 0.0063837
|  |  | 87.3984 | 0 | 0 | [5 empty] | 87.3984 | 78.005 | 62.9043 | 352.2452 |  |  | 3630.3202 | 0.018778
|  |  | 87.4752 | 0 | 0 | [5 empty] | 87.4752 | 77.81 | 61.3837 | 352.5843 |  |  | 3633.4991 | 0.025667
|  |  | 87.552 | 0 | 0 | [5 empty] | 87.552 | 77.28 | 60.329 | 350.8582 |  |  | 3636.6781 | 0.021362
|  |  | 87.6288 | 0 | 0 | [5 empty] | 87.6288 | 76.585 | 59.6736 | 346.2348 |  |  | 3639.857 | 0.0084646
|  |  | 87.7056 | 0 | 0 | [5 empty] | 87.7056 | 75.89 | 58.9978 | 338.7448 |  |  | 3643.0359 | 0.0044561
|  |  | 87.7824 | 0 | 0 | [5 empty] | 87.7824 | 75.34 | 58.2042 | 329.3129 |  |  | 3646.2148 | 0.010002
|  |  | 87.8592 | 0 | 0 | [5 empty] | 87.8592 | 75.115 | 57.9891 | 318.8331 |  |  | 3649.3937 | 0.0081867
|  |  | 87.936 | 0 | 0 | [5 empty] | 87.936 | 75.34 | 58.4602 | 310.0486 |  |  | 3652.5726 | 0.0038886
|  |  | 88.0128 | 0 | 0 | [5 empty] | 88.0128 | 75.615 | 59.5712 | 309.9561 |  |  | 3655.7515 | 0.0032772
|  |  | 88.0896 | 0 | 0 | [5 empty] | 88.0896 | 76.055 | 59.1104 | 309.1239 |  |  | 3658.9305 | 0.013776
|  |  | 88.1664 | 0 | 0 | [5 empty] | 88.1664 | 76.19 | 58.1274 | 302.9593 |  |  | 3662.1094 | 0.022185
|  |  | 88.2432 | 0 | 0 | [5 empty] | 88.2432 | 75.88 | 57.7792 | 297.1029 |  |  | 3665.2883 | 0.02127
|  |  | 88.32 | 0 | 0 | [5 empty] | 88.32 | 75.375 | 57.5949 | 292.3253 |  |  | 3668.4672 | 0.011088
|  |  | 88.3968 | 0 | 0 | [5 empty] | 88.3968 | 74.92 | 57.6461 | 288.5958 |  |  | 3671.6461 | 0.0032967
|  |  | 88.4736 | 0 | 0 | [5 empty] | 88.4736 | 74.92 | 56.8269 | 286.0683 |  |  | 3674.825 | 0.0043856
|  |  | 88.5504 | 0 | 0 | [5 empty] | 88.5504 | 74.84 | 57.641 | 284.4655 |  |  | 3678.0039 | 0.0047506
|  |  | 88.6272 | 0 | 0 | [5 empty] | 88.6272 | 74.92 | 58.9107 | 283.6949 |  |  | 3681.1829 | 0.016613
|  |  | 88.704 | 0 | 0 | [5 empty] | 88.704 | 75.1 | 59.5763 | 283.6332 |  |  | 3684.3618 | 0.019643
|  |  | 88.7808 | 0 | 0 | [5 empty] | 88.7808 | 75.12 | 60.201 | 284.0956 |  |  | 3687.5407 | 0.0079655
|  |  | 88.8576 | 0 | 0 | [5 empty] | 88.8576 | 75.125 | 60.032 | 284.8354 |  |  | 3690.7196 | 0.01265
|  |  | 88.9344 | 0 | 0 | [5 empty] | 88.9344 | 75.195 | 58.7776 | 285.6367 |  |  | 3693.8985 | 0.027616
|  |  | 89.0112 | 0 | 0 | [5 empty] | 89.0112 | 75.12 | 57.2109 | 288.904 |  |  | 3697.0774 | 0.026645
|  |  | 89.088 | 0 | 0 | [5 empty] | 89.088 | 74.76 | 56.5146 | 302.8668 |  |  | 3700.2563 | 0.010992
|  |  | 89.1648 | 0 | 0 | [5 empty] | 89.1648 | 74.39 | 56.3814 | 323.6723 |  |  | 3703.4353 | 0.0095475
|  |  | 89.2416 | 0 | 0 | [5 empty] | 89.2416 | 74.225 | 56.832 | 340.4092 |  |  | 3706.6142 | 0.019724
|  |  | 89.3184 | 0 | 0 | [5 empty] | 89.3184 | 74.33 | 58.0352 | 350.7041 |  |  | 3709.7931 | 0.01882
|  |  | 89.3952 | 0 | 0 | [5 empty] | 89.3952 | 74.74 | 59.7862 | 355.8207 |  |  | 3712.972 | 0.015988
|  |  | 89.472 | 0 | 0 | [5 empty] | 89.472 | 75.76 | 59.991 | 358.1016 |  |  | 3716.1509 | 0.015659
|  |  | 89.5488 | 0 | 0 | [5 empty] | 89.5488 | 76.65 | 61.2045 | 360.4442 |  |  | 3719.3298 | 0.010729
|  |  | 89.6256 | 0 | 0 | [5 empty] | 89.6256 | 77.24 | 63.0323 | 364.6361 |  |  | 3722.5087 | 0.0056174
|  |  | 89.7024 | 0 | 0 | [5 empty] | 89.7024 | 77.685 | 64.0563 | 370.924 |  |  | 3725.6877 | 0.0086333
|  |  | 89.7792 | 0 | 0 | [5 empty] | 89.7792 | 78.005 | 64.5274 | 378.0749 |  |  | 3728.8666 | 0.0070756
|  |  | 89.856 | 0 | 0 | [5 empty] | 89.856 | 78.215 | 63.7798 | 384.2087 |  |  | 3732.0455 | 0.005599
|  |  | 89.9328 | 0 | 0 | [5 empty] | 89.9328 | 78.2 | 63.3395 | 387.4143 |  |  | 3735.2244 | 0.0111
|  |  | 90.0096 | 0 | 0 | [5 empty] | 90.0096 | 77.67 | 64.9318 | 386.5204 |  |  | 3738.4033 | 0.014828
|  |  | 90.0864 | 0 | 0 | [5 empty] | 90.0864 | 77.265 | 66.0531 | 381.6196 |  |  | 3741.5822 | 0.015925
|  |  | 90.1632 | 0 | 0 | [5 empty] | 90.1632 | 77.44 | 64.9216 | 373.9138 |  |  | 3744.7611 | 0.012754
|  |  | 90.24 | 0 | 0 | [5 empty] | 90.24 | 77.615 | 64.2406 | 365.3142 |  |  | 3747.9401 | 0.0050071
|  |  | 90.3168 | 0 | 0 | [5 empty] | 90.3168 | 77.695 | 63.6365 | 357.6085 |  |  | 3751.119 | 0.0096172
|  |  | 90.3936 | 0 | 0 | [5 empty] | 90.3936 | 77.58 | 63.2627 | 351.7829 |  |  | 3754.2979 | 0.015179
|  |  | 90.4704 | 0 | 0 | [5 empty] | 90.4704 | 77.22 | 62.7098 | 347.8376 |  |  | 3757.4768 | 0.011907
|  |  | 90.5472 | 0 | 0 | [5 empty] | 90.5472 | 76.74 | 62.0698 | 344.5395 |  |  | 3760.6557 | 0.0037762
|  |  | 90.624 | 0 | 0 | [5 empty] | 90.624 | 76.27 | 61.1891 | 340.4092 |  |  | 3763.8346 | 0.0044687
|  |  | 90.7008 | 0 | 0 | [5 empty] | 90.7008 | 76.13 | 59.392 | 334.7994 |  |  | 3767.0135 | 0.0032851
|  |  | 90.7776 | 0 | 0 | [5 empty] | 90.7776 | 76.125 | 59.3971 | 328.3882 |  |  | 3770.1925 | 0.0047479
|  |  | 90.8544 | 0 | 0 | [5 empty] | 90.8544 | 76.38 | 61.6602 | 320.9291 |  |  | 3773.3714 | 0.0084329
|  |  | 90.9312 | 0 | 0 | [5 empty] | 90.9312 | 76.97 | 63.8566 | 313.4699 |  |  | 3776.5503 | 0.0046702
|  |  | 91.008 | 0 | 0 | [5 empty] | 91.008 | 77.885 | 65.0803 | 307.2437 |  |  | 3779.7292 | 0.0068853
|  |  | 91.0848 | 0 | 0 | [5 empty] | 91.0848 | 78.925 | 65.0394 | 302.312 |  |  | 3782.9081 | 0.014013
|  |  | 91.1616 | 0 | 0 | [5 empty] | 91.1616 | 79.56 | 65.6947 | 298.5516 |  |  | 3786.087 | 0.013061
|  |  | 91.2384 | 0 | 0 | [5 empty] | 91.2384 | 79.765 | 66.8467 | 295.7775 |  |  | 3789.266 | 0.011355
|  |  | 91.3152 | 0 | 0 | [5 empty] | 91.3152 | 79.795 | 67.4611 | 293.8048 |  |  | 3792.4449 | 0.018516
|  |  | 91.392 | 0 | 0 | [5 empty] | 91.392 | 79.915 | 66.6061 | 292.4486 |  |  | 3795.6238 | 0.022555
|  |  | 91.4688 | 0 | 0 | [5 empty] | 91.4688 | 79.965 | 65.8022 | 291.4931 |  |  | 3798.8027 | 0.022208
|  |  | 91.5456 | 0 | 0 | [5 empty] | 91.5456 | 79.925 | 65.6845 | 290.7842 |  |  | 3801.9816 | 0.020509
|  |  | 91.6224 | 0 | 0 | [5 empty] | 91.6224 | 80.1 | 65.2288 | 290.1061 |  |  | 3805.1605 | 0.013473
|  |  | 91.6992 | 0 | 0 | [5 empty] | 91.6992 | 80.145 | 65.6128 | 289.3663 |  |  | 3808.3394 | 0.01101
|  |  | 91.776 | 0 | 0 | [5 empty] | 91.776 | 80.125 | 66.1197 | 288.8115 |  |  | 3811.5184 | 0.024524
|  |  | 91.8528 | 0 | 0 | [5 empty] | 91.8528 | 80.075 | 66.8314 | 289.8595 |  |  | 3814.6973 | 0.029075
|  |  | 91.9296 | 0 | 0 | [5 empty] | 91.9296 | 80.535 | 66.6214 | 294.2672 |  |  | 3817.8762 | 0.02022
|  |  | 92.0064 | 0 | 0 | [5 empty] | 92.0064 | 81.025 | 67.8042 | 302.7435 |  |  | 3821.0551 | 0.0083086
|  |  | 92.0832 | 0 | 0 | [5 empty] | 92.0832 | 81.28 | 69.4579 | 315.6584 |  |  | 3824.234 | 0.0043383
|  |  | 92.16 | 0 | 0 | [5 empty] | 92.16 | 81.435 | 70.2208 | 331.7788 |  |  | 3827.4129 | 0.0069622
|  |  | 92.2368 | 0 | 0 | [5 empty] | 92.2368 | 81.975 | 69.335 | 348.2999 |  |  | 3830.5918 | 0.023878
|  |  | 92.3136 | 0 | 0 | [5 empty] | 92.3136 | 82.33 | 69.9238 | 362.0778 |  |  | 3833.7708 | 0.032176
|  |  | 92.3904 | 0 | 0 | [5 empty] | 92.3904 | 82.555 | 70.144 | 370.6158 |  |  | 3836.9497 | 0.022321
|  |  | 92.4672 | 0 | 0 | [5 empty] | 92.4672 | 82.65 | 70.6765 | 373.2049 |  |  | 3840.1286 | 0.0058858
|  |  | 92.544 | 0 | 0 | [5 empty] | 92.544 | 82.77 | 71.1782 | 371.0781 |  |  | 3843.3075 | 0.023319
|  |  | 92.6208 | 0 | 0 | [5 empty] | 92.6208 | 82.995 | 71.6083 | 366.7629 |  |  | 3846.4864 | 0.026968
|  |  | 92.6976 | 0 | 0 | [5 empty] | 92.6976 | 83.19 | 72.0384 | 362.571 |  |  | 3849.6653 | 0.013786
|  |  | 92.7744 | 0 | 0 | [5 empty] | 92.7744 | 83.51 | 71.2294 | 359.951 |  |  | 3852.8442 | 0.0034864
|  |  | 92.8512 | 0 | 0 | [5 empty] | 92.8512 | 83.495 | 71.1834 | 358.9955 |  |  | 3856.0232 | 0.0070852
|  |  | 92.928 | 0 | 0 | [5 empty] | 92.928 | 83.405 | 70.4512 | 358.8414 |  |  | 3859.2021 | 0.0078364
|  |  | 93.0048 | 0 | 0 | [5 empty] | 93.0048 | 83.075 | 69.6986 | 357.8242 |  |  | 3862.381 | 0.026001
|  |  | 93.0816 | 0 | 0 | [5 empty] | 93.0816 | 82.4 | 69.2787 | 354.4953 |  |  | 3865.5599 | 0.033655
|  |  | 93.1584 | 0 | 0 | [5 empty] | 93.1584 | 81.69 | 68.311 | 348.2074 |  |  | 3868.7388 | 0.022419
|  |  | 93.2352 | 0 | 0 | [5 empty] | 93.2352 | 81.085 | 66.601 | 339.1763 |  |  | 3871.9177 | 0.0025667
|  |  | 93.312 | 0 | 0 | [5 empty] | 93.312 | 80.67 | 64.9882 | 328.4191 |  |  | 3875.0966 | 0.01786
|  |  | 93.3888 | 0 | 0 | [5 empty] | 93.3888 | 80.415 | 64.9933 | 316.8913 |  |  | 3878.2756 | 0.018627
|  |  | 93.4656 | 0 | 0 | [5 empty] | 93.4656 | 80.625 | 65.4694 | 305.4868 |  |  | 3881.4545 | 0.0060179
|  |  | 93.5424 | 0 | 0 | [5 empty] | 93.5424 | 80.81 | 66.3194 | 300.401 |  |  | 3884.6334 | 0.0067813
|  |  | 93.6192 | 0 | 0 | [5 empty] | 93.6192 | 81.015 | 66.048 | 300.6784 |  |  | 3887.8123 | 0.0074956
|  |  | 93.696 | 0 | 0 | [5 empty] | 93.696 | 81.055 | 65.1725 | 294.5446 |  |  | 3890.9912 | 0.0047481
|  |  | 93.7728 | 0 | 0 | [5 empty] | 93.7728 | 80.81 | 64.3942 | 286.9005 |  |  | 3894.1701 | 0.013695
|  |  | 93.8496 | 0 | 0 | [5 empty] | 93.8496 | 80.355 | 64.0102 | 280.0886 |  |  | 3897.349 | 0.011392
|  |  | 93.9264 | 0 | 0 | [5 empty] | 93.9264 | 79.845 | 63.9846 | 274.2014 |  |  | 3900.528 | 0.011068
|  |  | 94.0032 | 0 | 0 | [5 empty] | 94.0032 | 79.655 | 63.5187 | 269.3622 |  |  | 3903.7069 | 0.026633
|  |  | 94.08 | 0 | 0 | [5 empty] | 94.08 | 79.57 | 63.0221 | 265.571 |  |  | 3906.8858 | 0.030267
|  |  | 94.1568 | 0 | 0 | [5 empty] | 94.1568 | 79.51 | 63.0733 | 262.8586 |  |  | 3910.0647 | 0.016869
|  |  | 94.2336 | 0 | 0 | [5 empty] | 94.2336 | 79.385 | 63.4624 | 261.0708 |  |  | 3913.2436 | 0.0085818
|  |  | 94.3104 | 0 | 0 | [5 empty] | 94.3104 | 79.27 | 63.7696 | 260.1769 |  |  | 3916.4225 | 0.019692
|  |  | 94.3872 | 0 | 0 | [5 empty] | 94.3872 | 78.995 | 63.8771 | 260.0536 |  |  | 3919.6014 | 0.014965
|  |  | 94.464 | 0 | 0 | [5 empty] | 94.464 | 78.605 | 63.2883 | 260.3002 |  |  | 3922.7804 | 0.0019805
|  |  | 94.5408 | 0 | 0 | [5 empty] | 94.5408 | 78.18 | 61.7523 | 261.1324 |  |  | 3925.9593 | 0.017371
|  |  | 94.6176 | 0 | 0 | [5 empty] | 94.6176 | 77.625 | 60.4365 | 268.0676 |  |  | 3929.1382 | 0.021482
|  |  | 94.6944 | 0 | 0 | [5 empty] | 94.6944 | 77.005 | 59.3818 | 285.2052 |  |  | 3932.3171 | 0.018606
|  |  | 94.7712 | 0 | 0 | [5 empty] | 94.7712 | 76.76 | 57.9584 | 302.5277 |  |  | 3935.496 | 0.02007
|  |  | 94.848 | 0 | 0 | [5 empty] | 94.848 | 76.82 | 57.047 | 314.2713 |  |  | 3938.6749 | 0.019108
|  |  | 94.9248 | 0 | 0 | [5 empty] | 94.9248 | 77.085 | 56.8832 | 321.1757 |  |  | 3941.8538 | 0.010184
|  |  | 95.0016 | 0 | 0 | [5 empty] | 95.0016 | 77.24 | 58.7878 | 325.3984 |  |  | 3945.0328 | 0.0053903
|  |  | 95.0784 | 0 | 0 | [5 empty] | 95.0784 | 77.81 | 60.2522 | 329.1896 |  |  | 3948.2117 | 0.0087363
|  |  | 95.1552 | 0 | 0 | [5 empty] | 95.1552 | 78.415 | 61.5475 | 334.2446 |  |  | 3951.3906 | 0.005446
|  |  | 95.232 | 0 | 0 | [5 empty] | 95.232 | 78.75 | 62.9402 | 341.149 |  |  | 3954.5695 | 0.008052
|  |  | 95.3088 | 0 | 0 | [5 empty] | 95.3088 | 78.925 | 63.36 | 349.1321 |  |  | 3957.7484 | 0.013209
|  |  | 95.3856 | 0 | 0 | [5 empty] | 95.3856 | 78.915 | 63.3344 | 356.4988 |  |  | 3960.9273 | 0.013895
|  |  | 95.4624 | 0 | 0 | [5 empty] | 95.4624 | 78.65 | 63.4214 | 361.4922 |  |  | 3964.1062 | 0.012324
|  |  | 95.5392 | 0 | 0 | [5 empty] | 95.5392 | 78.54 | 62.6381 | 363.0025 |  |  | 3967.2852 | 0.0088377
|  |  | 95.616 | 0 | 0 | [5 empty] | 95.616 | 78.28 | 62.5613 | 360.6599 |  |  | 3970.4641 | 0.008801
|  |  | 95.6928 | 0 | 0 | [5 empty] | 95.6928 | 78.05 | 62.6483 | 355.2043 |  |  | 3973.643 | 0.015908
|  |  | 95.7696 | 0 | 0 | [5 empty] | 95.7696 | 78.13 | 61.6653 | 348.2691 |  |  | 3976.8219 | 0.01784
|  |  | 95.8464 | 0 | 0 | [5 empty] | 95.8464 | 78.19 | 60.7846 | 341.488 |  |  | 3980.0008 | 0.011054
|  |  | 95.9232 | 0 | 0 | [5 empty] | 95.9232 | 78.03 | 60.2163 | 336.1248 |  |  | 3983.1797 | 0.0034221
|  |  | 96 | 0 | 0 | [5 empty] | 96 | 77.45 | 60.585 | 332.2719 |  |  | 3986.3586 | 0.0087208
|  |  | 96.0768 | 0 | 0 | [5 empty] | 96.0768 | 76.72 | 60.5338 | 329.1588 |  |  | 3989.5376 | 0.0082478
|  |  | 96.1536 | 0 | 0 | [5 empty] | 96.1536 | 76.01 | 59.4381 | 325.4292 |  |  | 3992.7165 | 0.0042459
|  |  | 96.2304 | 0 | 0 | [5 empty] | 96.2304 | 75.475 | 57.687 | 319.8811 |  |  | 3995.8954 | 0.0056164
|  |  | 96.3072 | 0 | 0 | [5 empty] | 96.3072 | 75.105 | 56.3917 | 312.8226 |  |  | 3999.0743 | 0.0056544
|  |  | 96.384 | 0 | 0 | [5 empty] | 96.384 | 75.12 | 56.279 | 304.9628
|  |  | 96.4608 | 0 | 0 | [5 empty] | 96.4608 | 75.36 | 57.4771 | 296.733
|  |  | 96.5376 | 0 | 0 | [5 empty] | 96.5376 | 75.8 | 58.2093 | 289.7054
|  |  | 96.6144 | 0 | 0 | [5 empty] | 96.6144 | 76.2 | 59.049 | 284.4346
|  |  | 96.6912 | 0 | 0 | [5 empty] | 96.6912 | 76.485 | 60.0832 | 280.8283
|  |  | 96.768 | 0 | 0 | [5 empty] | 96.768 | 76.7 | 60.5235 | 278.5475
|  |  | 96.8448 | 0 | 0 | [5 empty] | 96.8448 | 76.8 | 60.3802 | 277.1296
|  |  | 96.9216 | 0 | 0 | [5 empty] | 96.9216 | 76.665 | 60.2061 | 276.3282
|  |  | 96.9984 | 0 | 0 | [5 empty] | 96.9984 | 76.32 | 60.1498 | 275.9275
|  |  | 97.0752 | 0 | 0 | [5 empty] | 97.0752 | 75.98 | 59.8528 | 275.9275
|  |  | 97.152 | 0 | 0 | [5 empty] | 97.152 | 75.725 | 59.7299 | 276.2665
|  |  | 97.2288 | 0 | 0 | [5 empty] | 97.2288 | 75.56 | 59.561 | 276.9138
|  |  | 97.3056 | 0 | 0 | [5 empty] | 97.3056 | 75.615 | 58.9722 | 277.8385
|  |  | 97.3824 | 0 | 0 | [5 empty] | 97.3824 | 75.92 | 57.5437 | 279.6263
|  |  | 97.4592 | 0 | 0 | [5 empty] | 97.4592 | 76.16 | 57.5283 | 284.0031
|  |  | 97.536 | 0 | 0 | [5 empty] | 97.536 | 76.455 | 59.0285 | 291.8013
|  |  | 97.6128 | 0 | 0 | [5 empty] | 97.6128 | 76.725 | 60.9075 | 303.5757
|  |  | 97.6896 | 0 | 0 | [5 empty] | 97.6896 | 76.89 | 62.0032 | 318.9872
|  |  | 97.7664 | 0 | 0 | [5 empty] | 97.7664 | 77.225 | 61.993 | 336.0015
|  |  | 97.8432 | 0 | 0 | [5 empty] | 97.8432 | 77.7 | 61.4451 | 351.6596
|  |  | 97.92 | 0 | 0 | [5 empty] | 97.92 | 77.755 | 61.5987 | 363.2491
|  |  | 97.9968 | 0 | 0 | [5 empty] | 97.9968 | 77.54 | 62.6381 | 369.2287
|  |  | 98.0736 | 0 | 0 | [5 empty] | 98.0736 | 77.67 | 62.8173 | 369.9377
|  |  | 98.1504 | 0 | 0 | [5 empty] | 98.1504 | 77.88 | 63.5392 | 367.0403
|  |  | 98.2272 | 0 | 0 | [5 empty] | 98.2272 | 78.27 | 64.2765 | 363.0025
|  |  | 98.304 | 0 | 0 | [5 empty] | 98.304 | 78.685 | 64.7322 | 359.8585
|  |  | 98.3808 | 0 | 0 | [5 empty] | 98.3808 | 78.99 | 64.9216 | 358.7489
|  |  | 98.4576 | 0 | 0 | [5 empty] | 98.4576 | 79.03 | 64.727 | 359.427
|  |  | 98.5344 | 0 | 0 | [5 empty] | 98.5344 | 78.77 | 64.3686 | 360.5366
|  |  | 98.6112 | 0 | 0 | [5 empty] | 98.6112 | 78.28 | 63.9232 | 360.2592
|  |  | 98.688 | 0 | 0 | [5 empty] | 98.688 | 77.625 | 63.3139 | 357.2386
|  |  | 98.7648 | 0 | 0 | [5 empty] | 98.7648 | 77.11 | 61.7933 | 351.074
|  |  | 98.8416 | 0 | 0 | [5 empty] | 98.8416 | 76.7 | 60.759 | 342.4744
|  |  | 98.9184 | 0 | 0 | [5 empty] | 98.9184 | 76.52 | 60.8051 | 332.4261
|  |  | 98.9952 | 0 | 0 | [5 empty] | 98.9952 | 76.965 | 60.9843 | 322.0387
|  |  | 99.072 | 0 | 0 | [5 empty] | 99.072 | 77.445 | 62.4538 | 318.4016
|  |  | 99.1488 | 0 | 0 | [5 empty] | 99.1488 | 77.885 | 63.7645 | 321.6688
|  |  | 99.2256 | 0 | 0 | [5 empty] | 99.2256 | 78.185 | 64.087 | 317.8159
|  |  | 99.3024 | 39.735 | 0.8448 | [5 empty] | 99.3024 | 78.305 | 63.4317 | 311.7438
|  |  | 99.3792 | 42.915 | 0.90112 | [5 empty] | 99.3792 | 78.185 | 63.2525 | 306.4114
|  |  | 99.456 | 44.845 | 1.0803 | [5 empty] | 99.456 | 77.865 | 64.0922 | 302.1579
|  |  | 99.5328 | 46.68 | 1.2646 | [5 empty] | 99.5328 | 77.645 | 65.1315 | 299.1064
|  |  | 99.6096 | 48.46 | 1.4336 | [5 empty] | 99.6096 | 77.845 | 65.6179 | 297.1954
|  |  | 99.6864 | 50.525 | 1.5104 | [5 empty] | 99.6864 | 78.405 | 65.2493 | 296.0549
|  |  | 99.7632 | 52.255 | 1.6435 | [5 empty] | 99.7632 | 78.745 | 66.5702 | 295.4076
|  |  | 99.84 | 53.27 | 1.8637 | [5 empty] | 99.84 | 79.16 | 67.1488 | 295.1302
|  |  | 99.9168 | 54.93 | 1.9251 | [5 empty] | 99.9168 | 79.555 | 66.8467 | 295.2535
|  |  | 99.9936 | 56.195 | 2.0275 | [5 empty] | 99.9936 | 79.67 | 66.6931 | 295.5309
|  |  | 100.0704 | 57.32 | 2.1146 | [5 empty] | 100.0704 | 79.445 | 66.8058 | 295.6542
|  |  | 100.1472 | 57.91 | 2.2579 | [5 empty] | 100.1472 | 79.245 | 65.9405 | 297.5652
|  |  | 100.224 | 58.285 | 2.3706 | [5 empty] | 100.224 | 79.035 | 65.0445 | 309.0314
|  |  | 100.3008 | 58.535 | 2.4627 | [5 empty] | 100.3008 | 78.715 | 65.5565 | 328.1725
|  |  | 100.3776 | 58.97 | 2.5037 | [5 empty] | 100.3776 | 78.875 | 64.9216 | 344.4778
|  |  | 100.4544 | 59.605 | 2.4678 | [5 empty] | 100.4544 | 79.425 | 64.2099 | 355.2043
|  |  | 100.5312 | 59.78 | 2.5139 | [5 empty] | 100.5312 | 79.875 | 65.751 | 361.4305
|  |  | 100.608 | 60.03 | 2.5446 | [5 empty] | 100.608 | 80.435 | 68.393 | 364.9751
|  |  | 100.6848 | 59.955 | 2.647 | [5 empty] | 100.6848 | 81.25 | 70.5229 | 367.8417
|  |  | 100.7616 | 60.585 | 2.6214 | [5 empty] | 100.7616 | 81.95 | 72.1408 | 371.6329
|  |  | 100.8384 | 60.52 | 2.7238 | [5 empty] | 100.8384 | 82.325 | 73.3696 | 376.8112
|  |  | 100.9152 | 60.57 | 2.7955 | [5 empty] | 100.9152 | 82.45 | 74.1427 | 382.6367
|  |  | 100.992 | 60.99 | 2.7802 | [5 empty] | 100.992 | 82.56 | 73.303 | 387.4759
|  |  | 101.0688 | 61.06 | 2.7955 | [5 empty] | 101.0688 | 82.31 | 72.8422 | 389.6952
|  |  | 101.1456 | 61.595 | 2.7034 | [5 empty] | 101.1456 | 81.975 | 72.535 | 388.339
|  |  | 101.2224 | 61.295 | 2.8006 | [5 empty] | 101.2224 | 81.705 | 72.105 | 383.2224
|  |  | 101.2992 | 60.905 | 2.903 | [5 empty] | 101.2992 | 81.625 | 71.0349 | 375.3625
|  |  | 101.376 | 61.525 | 2.7904 | [5 empty] | 101.376 | 81.56 | 70.3642 | 366.3005
|  |  | 101.4528 | 61.82 | 2.7494 | [5 empty] | 101.4528 | 81.33 | 70.2157 | 357.7934
|  |  | 101.5296 | 61.365 | 2.8262 | [5 empty] | 101.5296 | 81.015 | 69.6166 | 350.7966
|  |  | 101.6064 | 60.795 | 2.9082 | [5 empty] | 101.6064 | 80.665 | 67.8707 | 345.2176
|  |  | 101.6832 | 60.59 | 2.9235 | [5 empty] | 101.6832 | 79.98 | 66.6112 | 340.1318
|  |  | 101.76 | 60.67 | 2.8723 | [5 empty] | 101.76 | 79.14 | 65.449 | 334.3062
|  |  | 101.8368 | 60.355 | 2.8979 | [5 empty] | 101.8368 | 78.445 | 63.8106 | 327.032
|  |  | 101.9136 | 60.375 | 2.8262 | [5 empty] | 101.9136 | 78.09 | 62.6227 | 318.9256
|  |  | 101.9904 | 60.365 | 2.7494 | [5 empty] | 101.9904 | 77.975 | 63.3549 | 310.2027
|  |  | 102.0672 | 59.79 | 2.7802 | [5 empty] | 102.0672 | 78.42 | 63.5853 | 301.3257
|  |  | 102.144 | 59.795 | 2.6675 | [5 empty] | 102.144 | 78.905 | 64.512 | 293.5891
|  |  | 102.2208 | 59.17 | 2.6522 | [5 empty] | 102.2208 | 79.105 | 66.5907 | 287.2704
|  |  | 102.2976 | 58.2 | 2.6368 | [5 empty] | 102.2976 | 79.34 | 67.8605 | 282.2462
|  |  | 102.3744 | 57.34 | 2.4832 | [5 empty] | 102.3744 | 79.605 | 67.241 | 278.2084
|  |  | 102.4512 | 56.19 | 2.2938 | [5 empty] | 102.4512 | 79.475 | 66.5907 | 275.0028
|  |  | 102.528 | 54.435 | 2.2221 | [5 empty] | 102.528 | 78.99 | 66.3962 | 272.5061
|  |  | 102.6048 | 52.83 | 2.1146 | [5 empty] | 102.6048 | 78.62 | 64.7834 | 270.7492
|  |  | 102.6816 | 51.81 | 1.92 | [5 empty] | 102.6816 | 78.505 | 62.0032 | 269.5471
|  |  | 102.7584 | 51.03 | 1.705 | [5 empty] | 102.7584 | 77.78 | 62.5152 | 268.8074
|  |  | 102.8352 | 49.69 | 1.5872 | [5 empty] | 102.8352 | 77.45 | 62.0083 | 268.4375
|  |  | 102.912 | 48.14 | 1.4746 | [5 empty] | 102.912 | 77.385 | 60.7334 | 268.3759
|  |  | 102.9888 | 46.74 | 1.3568 | [5 empty] | 102.9888 | 76.965 | 61.0509 | 269.5471
|  |  | 103.0656 | 44.915 | 1.3005 | [5 empty] | 103.0656 | 77.08 | 60.3648 | 273.7391
|  |  | 103.1424 | 44.48 | 1.0086 | [5 empty] | 103.1424 | 77.61 | 59.4176 | 281.6914
|  |  | 103.2192 | 44.225 | 0.76288 | [5 empty] | 103.2192 | 77.76 | 59.9142 | 293.959
|  |  | 103.296 | 42.185 | 0.71168 | [5 empty] | 103.296 | 77.64 | 60.3597 | 309.802
|  |  | 103.3728 | 39.85 | 0.6656 | [5 empty] | 103.3728 | 77.325 | 61.3632 | 326.7546
|  |  | 103.4496 | 0 | 0 | [5 empty] | 103.4496 | 77.195 | 62.4845 | 341.5188
|  |  | 103.5264 | 0 | 0 | [5 empty] | 103.5264 | 77.195 | 62.6842 | 351.8137
|  |  | 103.6032 | 0 | 0 | [5 empty] | 103.6032 | 77.235 | 62.5766 | 356.6221
|  |  | 103.68 | 0 | 0 | [5 empty] | 103.68 | 77.355 | 62.72 | 356.7762
|  |  | 103.7568 | 0 | 0 | [5 empty] | 103.7568 | 77.605 | 62.72 | 354.372
|  |  | 103.8336 | 0 | 0 | [5 empty] | 103.8336 | 78.04 | 62.2746 | 351.6288
|  |  | 103.9104 | 0 | 0 | [5 empty] | 103.9104 | 78.135 | 63.2115 | 350.2726
|  |  | 103.9872 | 0 | 0 | [5 empty] | 103.9872 | 78.175 | 63.2576 | 350.7041
|  |  | 104.064 | 0 | 0 | [5 empty] | 104.064 | 78.115 | 62.249 | 352.0911
|  |  | 104.1408 | 0 | 0 | [5 empty] | 104.1408 | 77.545 | 62.4589 | 352.8925
|  |  | 104.2176 | 0 | 0 | [5 empty] | 104.2176 | 77.07 | 61.271 | 351.5055
|  |  | 104.2944 | 0 | 0 | [5 empty] | 104.2944 | 76.515 | 59.8579 | 346.9129
|  |  | 104.3712 | 0 | 0 | [5 empty] | 104.3712 | 75.97 | 58.4141 | 339.2379
|  |  | 104.448 | 0 | 0 | [5 empty] | 104.448 | 75.49 | 57.175 | 329.6212
|  |  | 104.5248 | 0 | 0 | [5 empty] | 104.5248 | 75.275 | 57.0522 | 319.2955
|  |  | 104.6016 | 0 | 0 | [5 empty] | 104.6016 | 75.385 | 58.0864 | 310.7575
|  |  | 104.6784 | 0 | 0 | [5 empty] | 104.6784 | 75.63 | 59.4995 | 311.0349
|  |  | 104.7552 | 0 | 0 | [5 empty] | 104.7552 | 75.83 | 60.2573 | 310.5725
|  |  | 104.832 | 0 | 0 | [5 empty] | 104.832 | 76.045 | 59.1206 | 304.0997
|  |  | 104.9088 | 0 | 0 | [5 empty] | 104.9088 | 75.695 | 58.9926 | 297.7502
|  |  | 104.9856 | 0 | 0 | [5 empty] | 104.9856 | 75.105 | 59.2794 | 292.6028
|  |  | 105.0624 | 0 | 0 | [5 empty] | 105.0624 | 74.94 | 57.7946 | 288.8423
|  |  | 105.1392 | 0 | 0 | [5 empty] | 105.1392 | 74.81 | 57.431 | 286.5306
|  |  | 105.216 | 0 | 0 | [5 empty] | 105.216 | 74.785 | 58.0045 | 285.2669
|  |  | 105.2928 | 0 | 0 | [5 empty] | 105.2928 | 75.045 | 58.3578 | 284.6504
|  |  | 105.3696 | 0 | 0 | [5 empty] | 105.3696 | 75.305 | 59.1974 | 284.3113
|  |  | 105.4464 | 0 | 0 | [5 empty] | 105.4464 | 75.33 | 60.2214 | 284.3113
|  |  | 105.5232 | 0 | 0 | [5 empty] | 105.5232 | 75.535 | 59.2333 | 284.6504
|  |  | 105.6 | 0 | 0 | [5 empty] | 105.6 | 75.31 | 58.7366 | 285.2669
|  |  | 105.6768 | 0 | 0 | [5 empty] | 105.6768 | 74.84 | 58.5933 | 288.2259
|  |  | 105.7536 | 0 | 0 | [5 empty] | 105.7536 | 74.39 | 58.3014 | 301.5106
|  |  | 105.8304 | 0 | 0 | [5 empty] | 105.8304 | 74.07 | 58.1325 | 322.4702
|  |  | 105.9072 | 0 | 0 | [5 empty] | 105.9072 | 73.875 | 58.5472 | 339.8544
|  |  | 105.984 | 0 | 0 | [5 empty] | 105.984 | 74.145 | 58.4909 | 350.6116
|  |  | 106.0608 | 0 | 0 | [5 empty] | 106.0608 | 74.91 | 58.2758 | 355.9748
|  |  | 106.1376 | 0 | 0 | [5 empty] | 106.1376 | 75.925 | 58.1325 | 358.3174
|  |  | 106.2144 | 0 | 0 | [5 empty] | 106.2144 | 76.545 | 60.4262 | 360.475
|  |  | 106.2912 | 0 | 0 | [5 empty] | 106.2912 | 77.12 | 62.8429 | 364.3587
|  |  | 106.368 | 0 | 0 | [5 empty] | 106.368 | 77.605 | 64.4096 | 370.3692
|  |  | 106.4448 | 0 | 0 | [5 empty] | 106.4448 | 77.965 | 64.7987 | 377.6434
|  |  | 106.5216 | 0 | 0 | [5 empty] | 106.5216 | 77.985 | 65.1878 | 384.0854
|  |  | 106.5984 | 0 | 0 | [5 empty] | 106.5984 | 77.665 | 66.4218 | 387.7533
|  |  | 106.6752 | 0 | 0 | [5 empty] | 106.6752 | 77.795 | 64.512 | 387.4143
|  |  | 106.752 | 0 | 0 | [5 empty] | 106.752 | 78.065 | 61.9162 | 382.9141
|  |  | 106.8288 | 0 | 0 | [5 empty] | 106.8288 | 77.71 | 62.8941 | 375.2392
|  |  | 106.9056 | 0 | 0 | [5 empty] | 106.9056 | 77.305 | 64.6707 | 366.4855
|  |  | 106.9824 | 0 | 0 | [5 empty] | 106.9824 | 76.925 | 66.391 | 358.4715
|  |  | 107.0592 | 0 | 0 | [5 empty] | 107.0592 | 76.875 | 65.9251 | 352.3994
|  |  | 107.136 | 0 | 0 | [5 empty] | 107.136 | 76.905 | 64.1946 | 348.3615
|  |  | 107.2128 | 0 | 0 | [5 empty] | 107.2128 | 76.72 | 62.2182 | 345.3409
|  |  | 107.2896 | 0 | 0 | [5 empty] | 107.2896 | 76.11 | 61.9059 | 341.7038
|  |  | 107.3664 | 0 | 0 | [5 empty] | 107.3664 | 75.695 | 61.2045 | 336.4947
|  |  | 107.4432 | 0 | 0 | [5 empty] | 107.4432 | 75.81 | 59.8118 | 330.2993
|  |  | 107.52 | 0 | 0 | [5 empty] | 107.52 | 76.565 | 59.351 | 322.9017
|  |  | 107.5968 | 0 | 0 | [5 empty] | 107.5968 | 77.495 | 60.3699 | 315.3193
|  |  | 107.6736 | 0 | 0 | [5 empty] | 107.6736 | 78.13 | 62.8326 | 308.8773
|  |  | 107.7504 | 0 | 0 | [5 empty] | 107.7504 | 78.63 | 65.2646 | 303.699
|  |  | 107.8272 | 0 | 0 | [5 empty] | 107.8272 | 79.415 | 65.4643 | 299.7228
|  |  | 107.904 | 0 | 0 | [5 empty] | 107.904 | 79.855 | 65.9507 | 296.7639
|  |  | 107.9808 | 0 | 0 | [5 empty] | 107.9808 | 79.885 | 66.5088 | 294.6679
|  |  | 108.0576 | 0 | 0 | [5 empty] | 108.0576 | 79.81 | 66.4832 | 293.25
|  |  | 108.1344 | 0 | 0 | [5 empty] | 108.1344 | 79.775 | 66.1197 | 292.2945
|  |  | 108.2112 | 0 | 0 | [5 empty] | 108.2112 | 79.68 | 66.3603 | 291.5239
|  |  | 108.288 | 0 | 0 | [5 empty] | 108.288 | 79.69 | 66.4371 | 290.7842
|  |  | 108.3648 | 0 | 0 | [5 empty] | 108.3648 | 80.05 | 65.2237 | 289.9211
|  |  | 108.4416 | 0 | 0 | [5 empty] | 108.4416 | 80.2 | 64.9984 | 289.243
|  |  | 108.5184 | 0 | 0 | [5 empty] | 108.5184 | 80.25 | 65.3773 | 290.1061
|  |  | 108.5952 | 0 | 0 | [5 empty] | 108.5952 | 80.28 | 67.0362 | 294.2364
|  |  | 108.672 | 0 | 0 | [5 empty] | 108.672 | 80.575 | 69.207 | 302.3736
|  |  | 108.7488 | 0 | 0 | [5 empty] | 108.7488 | 81.1 | 69.8112 | 314.9186
|  |  | 108.8256 | 0 | 0 | [5 empty] | 108.8256 | 81.325 | 70.2208 | 330.7308
|  |  | 108.9024 | 0 | 0 | [5 empty] | 108.9024 | 81.66 | 70.1184 | 347.2519
|  |  | 108.9792 | 0 | 0 | [5 empty] | 108.9792 | 81.98 | 71.3062 | 361.4613
|  |  | 109.056 | 0 | 0 | [5 empty] | 109.056 | 82.15 | 71.9872 | 370.739
|  |  | 109.1328 | 0 | 0 | [5 empty] | 109.1328 | 82.51 | 71.0605 | 374.1912
|  |  | 109.2096 | 0 | 0 | [5 empty] | 109.2096 | 82.805 | 70.5946 | 372.7426
|  |  | 109.2864 | 0 | 0 | [5 empty] | 109.2864 | 83.08 | 70.6253 | 368.5814
|  |  | 109.3632 | 0 | 0 | [5 empty] | 109.3632 | 83.315 | 71.1936 | 363.9888
|  |  | 109.44 | 0 | 0 | [5 empty] | 109.44 | 83.325 | 72.0742 | 360.6599
|  |  | 109.5168 | 0 | 0 | [5 empty] | 109.5168 | 83.335 | 71.721 | 359.1804
|  |  | 109.5936 | 0 | 0 | [5 empty] | 109.5936 | 83.22 | 70.8966 | 358.7797
|  |  | 109.6704 | 0 | 0 | [5 empty] | 109.6704 | 82.88 | 70.3795 | 358.1016
|  |  | 109.7472 | 0 | 0 | [5 empty] | 109.7472 | 82.455 | 69.2429 | 355.5741
|  |  | 109.824 | 0 | 0 | [5 empty] | 109.824 | 81.835 | 67.84 | 350.1801
|  |  | 109.9008 | 0 | 0 | [5 empty] | 109.9008 | 81.015 | 66.857 | 341.5805
|  |  | 109.9776 | 0 | 0 | [5 empty] | 109.9776 | 80.32 | 66.0787 | 330.7
|  |  | 110.0544 | 0 | 0 | [5 empty] | 110.0544 | 80.165 | 65.2493 | 318.5865
|  |  | 110.1312 | 0 | 0 | [5 empty] | 110.1312 | 80.32 | 65.9149 | 306.2573
|  |  | 110.208 | 0 | 0 | [5 empty] | 110.208 | 80.74 | 66.1197 | 300.5551
|  |  | 110.2848 | 0 | 0 | [5 empty] | 110.2848 | 81.055 | 66.3552 | 302.6819
|  |  | 110.3616 | 0 | 0 | [5 empty] | 110.3616 | 81.01 | 66.519 | 297.4728
|  |  | 110.4384 | 0 | 0 | [5 empty] | 110.4384 | 80.775 | 65.8637 | 289.2739
|  |  | 110.5152 | 0 | 0 | [5 empty] | 110.5152 | 80.365 | 64.8909 | 281.6914
|  |  | 110.592 | 0 | 0 | [5 empty] | 110.592 | 79.97 | 63.959 | 275.0336
|  |  | 110.6688 | 0 | 0 | [5 empty] | 110.6688 | 79.69 | 63.3139 | 269.7321
|  |  | 110.7456 | 0 | 0 | [5 empty] | 110.7456 | 79.415 | 63.6109 | 265.8792
|  |  | 110.8224 | 0 | 0 | [5 empty] | 110.8224 | 79.28 | 64.0154 | 263.2901
|  |  | 110.8992 | 0 | 0 | [5 empty] | 110.8992 | 79.335 | 63.7952 | 261.5948
|  |  | 110.976 | 0 | 0 | [5 empty] | 110.976 | 79.28 | 63.959 | 260.6085
|  |  | 111.0528 | 0 | 0 | [5 empty] | 111.0528 | 79.025 | 63.8771 | 260.2078
|  |  | 111.1296 | 0 | 0 | [5 empty] | 111.1296 | 78.675 | 62.807 | 260.1153
|  |  | 111.2064 | 0 | 0 | [5 empty] | 111.2064 | 78 | 62.1773 | 260.6085
|  |  | 111.2832 | 0 | 0 | [5 empty] | 111.2832 | 77.24 | 61.4502 | 267.1121
|  |  | 111.36 | 0 | 0 | [5 empty] | 111.36 | 77.015 | 58.6752 | 284.1264
|  |  | 111.4368 | 0 | 0 | [5 empty] | 111.4368 | 76.705 | 57.4157 | 301.8496
|  |  | 111.5136 | 0 | 0 | [5 empty] | 111.5136 | 76.72 | 56.5658 | 314.0555
|  |  | 111.5904 | 0 | 0 | [5 empty] | 111.5904 | 76.855 | 57.3082 | 321.1757
|  |  | 111.6672 | 0 | 0 | [5 empty] | 111.6672 | 76.95 | 59.7606 | 325.4601
|  |  | 111.744 | 0 | 0 | [5 empty] | 111.744 | 77.665 | 60.4262 | 329.128
|  |  | 111.8208 | 0 | 0 | [5 empty] | 111.8208 | 78.37 | 60.8666 | 334.0905
|  |  | 111.8976 | 0 | 0 | [5 empty] | 111.8976 | 78.64 | 62.4998 | 340.8407
|  |  | 111.9744 | 0 | 0 | [5 empty] | 111.9744 | 78.91 | 62.72 | 348.7622
|  |  | 112.0512 | 0 | 0 | [5 empty] | 112.0512 | 78.88 | 63.0272 | 356.1598
|  |  | 112.128 | 0 | 0 | [5 empty] | 112.128 | 78.525 | 63.8106 | 361.1839
|  |  | 112.2048 | 0 | 0 | [5 empty] | 112.2048 | 78.47 | 62.9606 | 362.6942
|  |  | 112.2816 | 0 | 0 | [5 empty] | 112.2816 | 78.33 | 62.3206 | 360.4442
|  |  | 112.3584 | 0 | 0 | [5 empty] | 112.3584 | 78.145 | 62.0134 | 355.1734
|  |  | 112.4352 | 0 | 0 | [5 empty] | 112.4352 | 77.92 | 62.4026 | 348.454
|  |  | 112.512 | 0 | 0 | [5 empty] | 112.512 | 77.835 | 62.2029 | 341.8579
|  |  | 112.5888 | 0 | 0 | [5 empty] | 112.5888 | 77.83 | 60.8358 | 336.3714
|  |  | 112.6656 | 0 | 0 | [5 empty] | 112.6656 | 77.615 | 59.6531 | 332.2719
|  |  | 112.7424 | 0 | 0 | [5 empty] | 112.7424 | 77.165 | 58.6342 | 328.8198
|  |  | 112.8192 | 0 | 0 | [5 empty] | 112.8192 | 76.605 | 57.047 | 324.8128
|  |  | 112.896 | 0 | 0 | [5 empty] | 112.896 | 75.755 | 56.3302 | 319.2646
|  |  | 112.9728 | 0 | 0 | [5 empty] | 112.9728 | 75.295 | 55.0349 | 312.4836
|  |  | 113.0496 | 0 | 0 | [5 empty] | 113.0496 | 75.33 | 54.4205 | 305.1477
|  |  | 113.1264 | 0 | 0 | [5 empty] | 113.1264 | 75.71 | 54.9939 | 297.3187
|  |  | 113.2032 | 0 | 0 | [5 empty] | 113.2032 | 76.08 | 56.6682 | 290.3835
|  |  | 113.28 | 0 | 0 | [5 empty] | 113.28 | 76.475 | 58.2451 | 284.8662
|  |  | 113.3568 | 0 | 0 | [5 empty] | 113.3568 | 76.89 | 58.9926 | 280.7667
|  |  | 113.4336 | 0 | 0 | [5 empty] | 113.4336 | 76.86 | 60.3648 | 277.931
|  |  | 113.5104 | 0 | 0 | [5 empty] | 113.5104 | 76.695 | 61.1174 | 276.2357
|  |  | 113.5872 | 0 | 0 | [5 empty] | 113.5872 | 76.785 | 59.6378 | 275.4343
|  |  | 113.664 | 0 | 0 | [5 empty] | 113.664 | 76.58 | 58.9773 | 275.3419
|  |  | 113.7408 | 0 | 0 | [5 empty] | 113.7408 | 76.345 | 58.2912 | 275.6193
|  |  | 113.8176 | 0 | 0 | [5 empty] | 113.8176 | 76.045 | 58.1786 | 276.0508
|  |  | 113.8944 | 0 | 0 | [5 empty] | 113.8944 | 75.95 | 57.8611 | 276.4515
|  |  | 113.9712 | 0 | 0 | [5 empty] | 113.9712 | 75.82 | 58.1478 | 276.883
|  |  | 114.048 | 0 | 0 | [5 empty] | 114.048 | 75.735 | 58.3782 | 278.2084
|  |  | 114.1248 | 0 | 0 | [5 empty] | 114.1248 | 75.71 | 59.3203 | 282.3079
|  |  | 114.2016 | 0 | 0 | [5 empty] | 114.2016 | 76.305 | 59.2384 | 290.1369
|  |  | 114.2784 | 0 | 0 | [5 empty] | 114.2784 | 76.65 | 60.713 | 302.0346
|  |  | 114.3552 | 0 | 0 | [5 empty] | 114.3552 | 76.945 | 61.4349 | 317.5385
|  |  | 114.432 | 0 | 0 | [5 empty] | 114.432 | 77.37 | 60.8358 | 334.522
|  |  | 114.5088 | 0 | 0 | [5 empty] | 114.5088 | 77.67 | 60.5542 | 350.0568
|  |  | 114.5856 | 0 | 0 | [5 empty] | 114.5856 | 77.75 | 60.5491 | 361.4613
|  |  | 114.6624 | 0 | 0 | [5 empty] | 114.6624 | 77.75 | 60.9536 | 367.5334
|  |  | 114.7392 | 0 | 0 | [5 empty] | 114.7392 | 77.87 | 61.5424 | 368.5814
|  |  | 114.816 | 0 | 0 | [5 empty] | 114.816 | 78.135 | 62.1261 | 366.2697
|  |  | 114.8928 | 0 | 0 | [5 empty] | 114.8928 | 78.445 | 62.8531 | 362.7559
|  |  | 114.9696 | 0 | 0 | [5 empty] | 114.9696 | 78.71 | 64.0307 | 359.7969
|  |  | 115.0464 | 0 | 0 | [5 empty] | 115.0464 | 78.855 | 64.9574 | 358.379
|  |  | 115.1232 | 0 | 0 | [5 empty] | 115.1232 | 78.985 | 64.4045 | 358.4407
|  |  | 115.2 | 0 | 0 | [5 empty] | 115.2 | 78.955 | 63.2474 | 358.9955
|  |  | 115.2768 | 0 | 0 | [5 empty] | 115.2768 | 78.53 | 62.6278 | 358.5948
|  |  | 115.3536 | 0 | 0 | [5 empty] | 115.3536 | 77.86 | 62.5613 | 355.8824
|  |  | 115.4304 | 0 | 0 | [5 empty] | 115.4304 | 77.17 | 62.3002 | 350.3959
|  |  | 115.5072 | 0 | 0 | [5 empty] | 115.5072 | 76.485 | 62.0544 | 342.4744
|  |  | 115.584 | 0 | 0 | [5 empty] | 115.584 | 76.455 | 61.0048 | 332.7035
|  |  | 115.6608 | 0 | 0 | [5 empty] | 115.6608 | 76.825 | 61.0867 | 322.0079
|  |  | 115.7376 | 0 | 0 | [5 empty] | 115.7376 | 77.36 | 62.1773 | 316.4597
|  |  | 115.8144 | 0 | 0 | [5 empty] | 115.8144 | 77.865 | 63.4522 | 319.7578
|  |  | 115.8912 | 0 | 0 | [5 empty] | 115.8912 | 78.14 | 64.3789 | 317.0454
|  |  | 115.968 | 40.19 | 0.82432 | [5 empty] | 115.968 | 78.41 | 63.4317 | 311.0041
|  |  | 116.0448 | 43.08 | 0.90112 | [5 empty] | 116.0448 | 78.39 | 62.7302 | 305.61
|  |  | 116.1216 | 44.725 | 1.1315 | [5 empty] | 116.1216 | 77.8 | 64.5837 | 301.2024
|  |  | 116.1984 | 46.665 | 1.3005 | [5 empty] | 116.1984 | 77.85 | 64.4454 | 297.9659
|  |  | 116.2752 | 48.84 | 1.4029 | [5 empty] | 116.2752 | 78.29 | 63.616 | 295.87
|  |  | 116.352 | 51.09 | 1.4746 | [5 empty] | 116.352 | 78.73 | 63.6621 | 294.6062
|  |  | 116.4288 | 52.565 | 1.6589 | [5 empty] | 116.4288 | 78.945 | 65.239 | 293.8665
|  |  | 116.5056 | 53.66 | 1.879 | [5 empty] | 116.5056 | 79.18 | 66.8211 | 293.5891
|  |  | 116.5824 | 54.955 | 2.0429 | [5 empty] | 116.5824 | 79.525 | 67.2154 | 293.6815
|  |  | 116.6592 | 56.355 | 2.1197 | [5 empty] | 116.6592 | 79.925 | 65.7766 | 293.959
|  |  | 116.736 | 57.85 | 2.1197 | [5 empty] | 116.736 | 79.855 | 65.3722 | 294.0206
|  |  | 116.8128 | 58.205 | 2.2784 | [5 empty] | 116.8128 | 79.575 | 65.0854 | 295.5926
|  |  | 116.8896 | 58.52 | 2.4115 | [5 empty] | 116.8896 | 78.97 | 66.1402 | 306.4114
|  |  | 116.9664 | 59.025 | 2.473 | [5 empty] | 116.9664 | 78.7 | 66.1248 | 325.3059
|  |  | 117.0432 | 60 | 2.4064 | [5 empty] | 117.0432 | 78.86 | 65.4541 | 341.6729
|  |  | 117.12 | 60 | 2.4934 | [5 empty] | 117.12 | 79.15 | 66.0224 | 352.461
|  |  | 117.1968 | 60.265 | 2.5139 | [5 empty] | 117.1968 | 79.58 | 67.5379 | 358.8414
|  |  | 117.2736 | 60.32 | 2.5856 | [5 empty] | 117.2736 | 80.5 | 68.0704 | 362.571
|  |  | 117.3504 | 60.74 | 2.6061 | [5 empty] | 117.3504 | 81.26 | 70.1184 | 365.7766
|  |  | 117.4272 | 60.955 | 2.6522 | [5 empty] | 117.4272 | 81.95 | 71.7414 | 369.7835
|  |  | 117.504 | 61.615 | 2.6163 | [5 empty] | 117.504 | 82.425 | 72.6784 | 374.931
|  |  | 117.5808 | 62.155 | 2.5856 | [5 empty] | 117.5808 | 82.75 | 72.5402 | 380.3558
|  |  | 117.6576 | 61.165 | 2.8928 | [5 empty] | 117.6576 | 82.88 | 71.552 | 384.7019
|  |  | 117.7344 | 61.68 | 2.7802 | [5 empty] | 117.7344 | 82.7 | 70.8557 | 386.6129
|  |  | 117.8112 | 62.295 | 2.7136 | [5 empty] | 117.8112 | 82.235 | 71.168 | 385.3492
|  |  | 117.888 | 61.175 | 3.0464 | [5 empty] | 117.888 | 81.92 | 70.825 | 380.7874
|  |  | 117.9648 | 61.84 | 2.8211 | [5 empty] | 117.9648 | 81.855 | 69.7446 | 373.6981
|  |  | 118.0416 | 62.25 | 2.7546 | [5 empty] | 118.0416 | 81.885 | 68.5926 | 365.4067
|  |  | 118.1184 | 60.295 | 3.3741 | [5 empty] | 118.1184 | 81.575 | 68.7206 | 357.1769
|  |  | 118.1952 | 62.39 | 2.729 | [5 empty] | 118.1952 | 81.035 | 69.2531 | 349.8102
|  |  | 118.272 | 59.92 | 3.4202 | [5 empty] | 118.272 | 80.645 | 68.0499 | 343.399
|  |  | 118.3488 | 61.45 | 2.8672 | [5 empty] | 118.3488 | 80.145 | 66.1197 | 337.5118
|  |  | 118.4256 | 61.525 | 2.8211 | [5 empty] | 118.4256 | 79.5 | 64.0666 | 331.224
|  |  | 118.5024 | 61.71 | 2.7238 | [5 empty] | 118.5024 | 78.93 | 62.0032 | 324.0422
|  |  | 118.5792 | 61.59 | 2.6931 | [5 empty] | 118.5792 | 78.385 | 61.225 | 316.2132
|  |  | 118.656 | 60.95 | 2.7648 | [5 empty] | 118.656 | 78.325 | 61.3683 | 308.2608
|  |  | 118.7328 | 60.51 | 2.7699 | [5 empty] | 118.7328 | 78.705 | 62.3821 | 299.9386
|  |  | 118.8096 | 60.35 | 2.6778 | [5 empty] | 118.8096 | 79.38 | 62.6586 | 292.2945
|  |  | 118.8864 | 59.935 | 2.6266 | [5 empty] | 118.8864 | 79.745 | 64.1331 | 285.6984
|  |  | 118.9632 | 59.735 | 2.4678 | [5 empty] | 118.9632 | 79.8 | 65.751 | 280.1194
|  |  | 119.04 | 58.135 | 2.473 | [5 empty] | 119.04 | 79.97 | 65.6179 | 275.496
|  |  | 119.1168 | 56.36 | 2.4474 | [5 empty] | 119.1168 | 79.905 | 64.8294 | 271.7972
|  |  | 119.1936 | 55.01 | 2.2784 | [5 empty] | 119.1936 | 79.445 | 64.4045 | 269.1464
|  |  | 119.2704 | 53.9 | 2.0736 | [5 empty] | 119.2704 | 78.865 | 63.9181 | 267.482
|  |  | 119.3472 | 52.83 | 1.8893 | [5 empty] | 119.3472 | 78.375 | 62.8736 | 266.5573
|  |  | 119.424 | 51.975 | 1.664 | [5 empty] | 119.424 | 78.01 | 61.7728 | 266.0641
|  |  | 119.5008 | 50.685 | 1.5258 | [5 empty] | 119.5008 | 77.6 | 61.609 | 265.6635
|  |  | 119.5776 | 49.215 | 1.4387 | [5 empty] | 119.5776 | 77.44 | 60.6464 | 265.3244
|  |  | 119.6544 | 48.435 | 1.2083 | [5 empty] | 119.6544 | 77.23 | 60.2163 | 266.095
|  |  | 119.7312 | 46.74 | 1.1315 | [5 empty] | 119.7312 | 77.08 | 60.7232 | 269.9787
|  |  | 119.808 | 45.04 | 1.0445 | [5 empty] | 119.808 | 77.35 | 61.1277 | 277.931
|  |  | 119.8848 | 44.08 | 0.85504 | [5 empty] | 119.8848 | 77.67 | 60.8819 | 290.3527
|  |  | 119.9616 | 42.115 | 0.79872 | [5 empty] | 119.9616 | 77.37 | 61.7165 | 306.3806
|  |  | 120.0384 | 40.225 | 0.71168 | [5 empty] | 120.0384 | 77.3 | 61.1584 | 323.4566
|  |  | 120.1152 | 0 | 0 | [5 empty] | 120.1152 | 77.48 | 60.9638 | 338.2516
|  |  | 120.192 | 0 | 0 | [5 empty] | 120.192 | 77.665 | 60.6822 | 348.3307
|  |  | 120.2688 | 0 | 0 | [5 empty] | 120.2688 | 77.67 | 60.9894 | 353.0158
|  |  | 120.3456 | 0 | 0 | [5 empty] | 120.3456 | 77.63 | 61.9469 | 353.3549
|  |  | 120.4224 | 0 | 0 | [5 empty] | 120.4224 | 77.65 | 63.2883 | 351.3514
|  |  | 120.4992 | 0 | 0 | [5 empty] | 120.4992 | 77.975 | 63.6518 | 349.0396
|  |  | 120.576 | 0 | 0 | [5 empty] | 120.576 | 78.135 | 64.2765 | 348.0533
|  |  | 120.6528 | 0 | 0 | [5 empty] | 120.6528 | 78.17 | 64.4813 | 348.4848
|  |  | 120.7296 | 0 | 0 | [5 empty] | 120.7296 | 78.06 | 63.8669 | 349.7177
|  |  | 120.8064 | 0 | 0 | [5 empty] | 120.8064 | 77.735 | 62.8378 | 350.1801
|  |  | 120.8832 | 0 | 0 | [5 empty] | 120.8832 | 77.145 | 61.8496 | 348.5773
|  |  | 120.96 | 0 | 0 | [5 empty] | 120.96 | 76.61 | 60.3392 | 344.108
|  |  | 121.0368 | 0 | 0 | [5 empty] | 121.0368 | 75.89 | 59.6326 | 336.9262
|  |  | 121.1136 | 0 | 0 | [5 empty] | 121.1136 | 75.565 | 57.8253 | 327.8334
|  |  | 121.1904 | 0 | 0 | [5 empty] | 121.1904 | 75.6 | 56.4019 | 317.7851
|  |  | 121.2672 | 0 | 0 | [5 empty] | 121.2672 | 75.665 | 57.6256 | 308.6615
|  |  | 121.344 | 0 | 0 | [5 empty] | 121.344 | 76.065 | 58.3219 | 307.4902
|  |  | 121.4208 | 0 | 0 | [5 empty] | 121.4208 | 76.3 | 58.9158 | 307.1512
|  |  | 121.4976 | 0 | 0 | [5 empty] | 121.4976 | 76.19 | 59.1616 | 301.0791
|  |  | 121.5744 | 0 | 0 | [5 empty] | 121.5744 | 75.915 | 58.8646 | 295.1302
|  |  | 121.6512 | 0 | 0 | [5 empty] | 121.6512 | 75.555 | 58.153 | 290.3527
|  |  | 121.728 | 0 | 0 | [5 empty] | 121.728 | 74.975 | 58.711 | 286.7464
|  |  | 121.8048 | 0 | 0 | [5 empty] | 121.8048 | 74.665 | 59.1309 | 284.2497
|  |  | 121.8816 | 0 | 0 | [5 empty] | 121.8816 | 74.925 | 58.2963 | 282.5853
|  |  | 121.9584 | 0 | 0 | [5 empty] | 121.9584 | 75.325 | 57.9584 | 281.5989
|  |  | 122.0352 | 0 | 0 | [5 empty] | 122.0352 | 75.435 | 58.8646 | 281.2291
|  |  | 122.112 | 0 | 0 | [5 empty] | 122.112 | 75.5 | 59.6173 | 281.5373
|  |  | 122.1888 | 0 | 0 | [5 empty] | 122.1888 | 75.46 | 59.945 | 282.3079
|  |  | 122.2656 | 0 | 0 | [5 empty] | 122.2656 | 75.415 | 59.2077 | 283.325
|  |  | 122.3424 | 0 | 0 | [5 empty] | 122.3424 | 75.16 | 58.4294 | 286.0374
|  |  | 122.4192 | 0 | 0 | [5 empty] | 122.4192 | 74.69 | 57.9277 | 298.3358
|  |  | 122.496 | 0 | 0 | [5 empty] | 122.496 | 74.34 | 57.6 | 318.9872
|  |  | 122.5728 | 0 | 0 | [5 empty] | 122.5728 | 74 | 58.5984 | 336.4947
|  |  | 122.6496 | 0 | 0 | [5 empty] | 122.6496 | 74.165 | 59.1258 | 347.591
|  |  | 122.7264 | 0 | 0 | [5 empty] | 122.7264 | 74.97 | 58.6138 | 353.2316
|  |  | 122.8032 | 0 | 0 | [5 empty] | 122.8032 | 75.76 | 59.8682 | 355.8207
|  |  | 122.88 | 0 | 0 | [5 empty] | 122.88 | 76.625 | 61.2096 | 358.1324
|  |  | 122.9568 | 0 | 0 | [5 empty] | 122.9568 | 77.19 | 63.6262 | 362.047
|  |  | 123.0336 | 0 | 0 | [5 empty] | 123.0336 | 77.91 | 64.1075 | 368.0883
|  |  | 123.1104 | 0 | 0 | [5 empty] | 123.1104 | 78.39 | 63.9795 | 375.27
|  |  | 123.1872 | 0 | 0 | [5 empty] | 123.1872 | 78.5 | 63.6774 | 381.6812
|  |  | 123.264 | 0 | 0 | [5 empty] | 123.264 | 78.27 | 64.0666 | 385.3183
|  |  | 123.3408 | 0 | 0 | [5 empty] | 123.3408 | 78.075 | 63.9642 | 384.9793
|  |  | 123.4176 | 0 | 0 | [5 empty] | 123.4176 | 78.01 | 63.3088 | 380.4791
|  |  | 123.4944 | 0 | 0 | [5 empty] | 123.4944 | 77.985 | 62.848 | 373.02
|  |  | 123.5712 | 0 | 0 | [5 empty] | 123.5712 | 77.75 | 63.7594 | 364.482
|  |  | 123.648 | 0 | 0 | [5 empty] | 123.648 | 77.64 | 64.2048 | 356.6837
|  |  | 123.7248 | 0 | 0 | [5 empty] | 123.7248 | 77.76 | 62.9299 | 350.7349
|  |  | 123.8016 | 0 | 0 | [5 empty] | 123.8016 | 77.525 | 62.1107 | 346.6355
|  |  | 123.8784 | 0 | 0 | [5 empty] | 123.8784 | 77.08 | 61.5117 | 343.399
|  |  | 123.9552 | 0 | 0 | [5 empty] | 123.9552 | 76.86 | 59.3613 | 339.5154
|  |  | 124.032 | 0 | 0 | [5 empty] | 124.032 | 76.22 | 59.2384 | 333.998
|  |  | 124.1088 | 0 | 0 | [5 empty] | 124.1088 | 75.885 | 60.1498 | 327.5252
|  |  | 124.1856 | 0 | 0 | [5 empty] | 124.1856 | 76.47 | 60.6669 | 320.0352
|  |  | 124.2624 | 0 | 0 | [5 empty] | 124.2624 | 77.445 | 60.9434 | 312.3603
|  |  | 124.3392 | 0 | 0 | [5 empty] | 124.3392 | 78.31 | 62.2592 | 305.8566
|  |  | 124.416 | 0 | 0 | [5 empty] | 124.416 | 78.845 | 64.4301 | 300.74
|  |  | 124.4928 | 0 | 0 | [5 empty] | 124.4928 | 79.405 | 65.664 | 296.7947
|  |  | 124.5696 | 0 | 0 | [5 empty] | 124.5696 | 79.925 | 65.4848 | 293.8973
|  |  | 124.6464 | 0 | 0 | [5 empty] | 124.6464 | 79.925 | 66.4422 | 291.863
|  |  | 124.7232 | 0 | 0 | [5 empty] | 124.7232 | 79.895 | 66.4525 | 290.476
|  |  | 124.8 | 0 | 0 | [5 empty] | 124.8 | 79.995 | 65.3824 | 289.5513
|  |  | 124.8768 | 0 | 0 | [5 empty] | 124.8768 | 79.985 | 64.8653 | 288.904
|  |  | 124.9536 | 0 | 0 | [5 empty] | 124.9536 | 80.09 | 64.4608 | 288.2567
|  |  | 125.0304 | 0 | 0 | [5 empty] | 125.0304 | 80.035 | 65.5309 | 287.5786
|  |  | 125.1072 | 0 | 0 | [5 empty] | 125.1072 | 80.08 | 66.1043 | 287.0238
|  |  | 125.184 | 0 | 0 | [5 empty] | 125.184 | 80.28 | 65.7357 | 287.7635
|  |  | 125.2608 | 0 | 0 | [5 empty] | 125.2608 | 80.445 | 66.647 | 291.5856
|  |  | 125.3376 | 0 | 0 | [5 empty] | 125.3376 | 80.855 | 67.968 | 299.353
|  |  | 125.4144 | 0 | 0 | [5 empty] | 125.4144 | 81.315 | 68.905 | 311.5897
|  |  | 125.4912 | 0 | 0 | [5 empty] | 125.4912 | 81.525 | 69.6781 | 327.3403
|  |  | 125.568 | 0 | 0 | [5 empty] | 125.568 | 81.895 | 69.7702 | 344.0463
|  |  | 125.6448 | 0 | 0 | [5 empty] | 125.6448 | 82.42 | 69.7344 | 358.5023
|  |  | 125.7216 | 0 | 0 | [5 empty] | 125.7216 | 82.775 | 69.3248 | 368.0883
|  |  | 125.7984 | 0 | 0 | [5 empty] | 125.7984 | 82.775 | 69.9955 | 371.6329
|  |  | 125.8752 | 0 | 0 | [5 empty] | 125.8752 | 82.945 | 70.0621 | 370.1226
|  |  | 125.952 | 0 | 0 | [5 empty] | 125.952 | 83.105 | 70.6304 | 365.8998
|  |  | 126.0288 | 0 | 0 | [5 empty] | 126.0288 | 83.285 | 71.511 | 361.3997
|  |  | 126.1056 | 0 | 0 | [5 empty] | 126.1056 | 83.425 | 72.2483 | 358.379
|  |  | 126.1824 | 0 | 0 | [5 empty] | 126.1824 | 83.41 | 72.8371 | 357.0844
|  |  | 126.2592 | 0 | 0 | [5 empty] | 126.2592 | 83.29 | 72.361 | 356.9612
|  |  | 126.336 | 0 | 0 | [5 empty] | 126.336 | 82.895 | 71.465 | 356.4063
|  |  | 126.4128 | 0 | 0 | [5 empty] | 126.4128 | 82.41 | 70.1133 | 353.7556
|  |  | 126.4896 | 0 | 0 | [5 empty] | 126.4896 | 81.96 | 68.0346 | 347.8992
|  |  | 126.5664 | 0 | 0 | [5 empty] | 126.5664 | 81.405 | 65.8995 | 339.0222
|  |  | 126.6432 | 0 | 0 | [5 empty] | 126.6432 | 80.97 | 64.0154 | 328.265
|  |  | 126.72 | 0 | 0 | [5 empty] | 126.72 | 80.625 | 64.2202 | 316.583
|  |  | 126.7968 | 0 | 0 | [5 empty] | 126.7968 | 80.615 | 65.4746 | 305.0552
|  |  | 126.8736 | 0 | 0 | [5 empty] | 126.8736 | 80.71 | 67.1334 | 299.4763
|  |  | 126.9504 | 0 | 0 | [5 empty] | 126.9504 | 80.835 | 67.5277 | 300.1236
|  |  | 127.0272 | 0 | 0 | [5 empty] | 127.0272 | 80.865 | 66.8774 | 294.1747
|  |  | 127.104 | 0 | 0 | [5 empty] | 127.104 | 80.55 | 66.3603 | 286.1916
|  |  | 127.1808 | 0 | 0 | [5 empty] | 127.1808 | 80.255 | 64.8704 | 279.1023
|  |  | 127.2576 | 0 | 0 | [5 empty] | 127.2576 | 80.05 | 63.4112 | 273.1226
|  |  | 127.3344 | 0 | 0 | [5 empty] | 127.3344 | 79.62 | 63.8771 | 268.2526
|  |  | 127.4112 | 0 | 0 | [5 empty] | 127.4112 | 79.405 | 64.2355 | 264.4305
|  |  | 127.488 | 0 | 0 | [5 empty] | 127.488 | 79.44 | 63.831 | 261.564
|  |  | 127.5648 | 0 | 0 | [5 empty] | 127.5648 | 79.465 | 63.6416 | 259.5605
|  |  | 127.6416 | 0 | 0 | [5 empty] | 127.6416 | 79.145 | 64.599 | 258.42
|  |  | 127.7184 | 0 | 0 | [5 empty] | 127.7184 | 78.825 | 64.5376 | 258.081
|  |  | 127.7952 | 0 | 0 | [5 empty] | 127.7952 | 78.5 | 63.6774 | 258.2659
|  |  | 127.872 | 0 | 0 | [5 empty] | 127.872 | 78.2 | 61.783 | 258.944
|  |  | 127.9488 | 0 | 0 | [5 empty] | 127.9488 | 77.715 | 60.0576 | 264.8004
|  |  | 128.0256 | 0 | 0 | [5 empty] | 128.0256 | 76.965 | 59.8272 | 281.1982
|  |  | 128.1024 | 0 | 0 | [5 empty] | 128.1024 | 76.53 | 59.264 | 299.168
|  |  | 128.1792 | 0 | 0 | [5 empty] | 128.1792 | 76.475 | 58.4294 | 311.7438
|  |  | 128.256 | 0 | 0 | [5 empty] | 128.256 | 76.665 | 58.4448 | 319.1105
|  |  | 128.3328 | 0 | 0 | [5 empty] | 128.3328 | 77.09 | 58.8595 | 323.5182
|  |  | 128.4096 | 0 | 0 | [5 empty] | 128.4096 | 77.46 | 60.887 | 327.1862
|  |  | 128.4864 | 0 | 0 | [5 empty] | 128.4864 | 78 | 62.1978 | 332.0562
|  |  | 128.5632 | 0 | 0 | [5 empty] | 128.5632 | 78.53 | 62.7507 | 338.714
|  |  | 128.64 | 0 | 0 | [5 empty] | 128.64 | 78.975 | 62.4896 | 346.6046
|  |  | 128.7168 | 0 | 0 | [5 empty] | 128.7168 | 78.915 | 62.9811 | 354.1255
|  |  | 128.7936 | 0 | 0 | [5 empty] | 128.7936 | 78.665 | 63.2115 | 359.5503
|  |  | 128.8704 | 0 | 0 | [5 empty] | 128.8704 | 78.6 | 61.8854 | 361.523
|  |  | 128.9472 | 0 | 0 | [5 empty] | 128.9472 | 78.165 | 62.3821 | 359.6428
|  |  | 129.024 | 0 | 0 | [5 empty] | 129.024 | 77.925 | 62.5101 | 354.5878
|  |  | 129.1008 | 0 | 0 | [5 empty] | 129.1008 | 78.18 | 60.9843 | 347.8067
|  |  | 129.1776 | 0 | 0 | [5 empty] | 129.1776 | 78.11 | 60.8666 | 340.964
|  |  | 129.2544 | 0 | 0 | [5 empty] | 129.2544 | 77.805 | 61.2966 | 335.3234
|  |  | 129.3312 | 0 | 0 | [5 empty] | 129.3312 | 77.7 | 59.6173 | 331.1931
|  |  | 129.408 | 0 | 0 | [5 empty] | 129.408 | 77.295 | 58.112 | 328.0492
|  |  | 129.4848 | 0 | 0 | [5 empty] | 129.4848 | 76.515 | 57.4003 | 324.5354
|  |  | 129.5616 | 0 | 0 | [5 empty] | 129.5616 | 75.84 | 56.3661 | 319.3879
|  |  | 129.6384 | 0 | 0 | [5 empty] | 129.6384 | 75.085 | 56.4326 | 312.6993
|  |  | 129.7152 | 0 | 0 | [5 empty] | 129.7152 | 74.85 | 57.1034 | 305.1169
|  |  | 129.792 | 0 | 0 | [5 empty] | 129.792 | 75.21 | 57.7075 | 296.8563
|  |  | 129.8688 | 0 | 0 | [5 empty] | 129.8688 | 75.795 | 57.9994 | 289.5204
|  |  | 129.9456 | 0 | 0 | [5 empty] | 129.9456 | 76.27 | 58.8288 | 283.8182
|  |  | 130.0224 | 0 | 0 | [5 empty] | 130.0224 | 76.75 | 59.1514 | 279.842
|  |  | 130.0992 | 0 | 0 | [5 empty] | 130.0992 | 77.21 | 58.3117 | 277.3145
|  |  | 130.176 | 0 | 0 | [5 empty] | 130.176 | 77.335 | 57.897 | 275.8658
|  |  | 130.2528 | 0 | 0 | [5 empty] | 130.2528 | 76.78 | 59.5968 | 275.1261
|  |  | 130.3296 | 0 | 0 | [5 empty] | 130.3296 | 76.35 | 60.0474 | 274.7562
|  |  | 130.4064 | 0 | 0 | [5 empty] | 130.4064 | 76.225 | 59.0848 | 274.6638
|  |  | 130.4832 | 0 | 0 | [5 empty] | 130.4832 | 76.12 | 58.2246 | 274.787
|  |  | 130.56 | 0 | 0 | [5 empty] | 130.56 | 76.03 | 57.8611 | 275.2186
|  |  | 130.6368 | 0 | 0 | [5 empty] | 130.6368 | 75.92 | 57.9072 | 275.9583
|  |  | 130.7136 | 0 | 0 | [5 empty] | 130.7136 | 76.015 | 57.1699 | 277.5919
|  |  | 130.7904 | 0 | 0 | [5 empty] | 130.7904 | 76.175 | 57.0266 | 281.7222
|  |  | 130.8672 | 0 | 0 | [5 empty] | 130.8672 | 76.31 | 59.177 | 289.243
|  |  | 130.944 | 0 | 0 | [5 empty] | 130.944 | 76.655 | 61.015 | 300.5551
|  |  | 131.0208 | 0 | 0 | [5 empty] | 131.0208 | 77.1 | 61.015 | 315.5042
|  |  | 131.0976 | 0 | 0 | [5 empty] | 131.0976 | 77.445 | 61.1379 | 332.3028
|  |  | 131.1744 | 0 | 0 | [5 empty] | 131.1744 | 77.83 | 61.0253 | 348.1458
|  |  | 131.2512 | 0 | 0 | [5 empty] | 131.2512 | 77.885 | 61.4502 | 360.3209
|  |  | 131.328 | 0 | 0 | [5 empty] | 131.328 | 77.6 | 62.889 | 367.1636
|  |  | 131.4048 | 0 | 0 | [5 empty] | 131.4048 | 77.615 | 63.3139 | 368.6739
|  |  | 131.4816 | 0 | 0 | [5 empty] | 131.4816 | 77.865 | 63.6467 | 366.3314
|  |  | 131.5584 | 0 | 0 | [5 empty] | 131.5584 | 78.445 | 63.4061 | 362.4168
|  |  | 131.6352 | 0 | 0 | [5 empty] | 131.6352 | 78.94 | 63.4778 | 358.9647
|  |  | 131.712 | 0 | 0 | [5 empty] | 131.712 | 79.11 | 64.3994 | 357.3618
|  |  | 131.7888 | 0 | 0 | [5 empty] | 131.7888 | 79.155 | 64.7885 | 357.6701
|  |  | 131.8656 | 0 | 0 | [5 empty] | 131.8656 | 78.91 | 64.8909 | 358.7797
|  |  | 131.9424 | 0 | 0 | [5 empty] | 131.9424 | 78.38 | 64.4403 | 358.9338
|  |  | 132.0192 | 0 | 0 | [5 empty] | 132.0192 | 77.88 | 63.0067 | 356.5605
|  |  | 132.096 | 0 | 0 | [5 empty] | 132.096 | 77.44 | 61.3939 | 350.9507
|  |  | 132.1728 | 0 | 0 | [5 empty] | 132.1728 | 76.95 | 60.5798 | 342.6593
|  |  | 132.2496 | 0 | 0 | [5 empty] | 132.2496 | 76.795 | 60.0371 | 332.611
|  |  | 132.3264 | 0 | 0 | [5 empty] | 132.3264 | 77.08 | 60.6362 | 321.8846
|  |  | 132.4032 | 0 | 0 | [5 empty] | 132.4032 | 77.65 | 61.5834 | 316.5214
|  |  | 132.48 | 0 | 0 | [5 empty] | 132.48 | 77.92 | 63.8003 | 320.0352
|  |  | 132.5568 | 0 | 0 | [5 empty] | 132.5568 | 78 | 65.3824 | 317.2611
|  |  | 132.6336 | 41.275 | 0.70144 | [5 empty] | 132.6336 | 78.07 | 65.0394 | 311.0349
|  |  | 132.7104 | 42.495 | 0.98304 | [5 empty] | 132.7104 | 78.095 | 64.0614 | 305.4559
|  |  | 132.7872 | 44.04 | 1.2851 | [5 empty] | 132.7872 | 78.085 | 63.2832 | 300.925
|  |  | 132.864 | 46.555 | 1.3414 | [5 empty] | 132.864 | 77.925 | 64.1178 | 297.5961
|  |  | 132.9408 | 48.755 | 1.408 | [5 empty] | 132.9408 | 78.085 | 64.4966 | 295.4693
|  |  | 133.0176 | 50.15 | 1.6486 | [5 empty] | 133.0176 | 78.565 | 64.4301 | 294.2055
|  |  | 133.0944 | 51.98 | 1.7869 | [5 empty] | 133.0944 | 79.02 | 65.0291 | 293.4966
|  |  | 133.1712 | 53.865 | 1.8534 | [5 empty] | 133.1712 | 79.365 | 65.6845 | 293.1576
|  |  | 133.248 | 55 | 2.0224 | [5 empty] | 133.248 | 79.835 | 65.0189 | 293.2192
|  |  | 133.3248 | 56.395 | 2.1146 | [5 empty] | 133.3248 | 80.04 | 64.6758 | 293.435
|  |  | 133.4016 | 57.655 | 2.1914 | [5 empty] | 133.4016 | 79.83 | 65.239 | 293.5583
|  |  | 133.4784 | 58.76 | 2.2374 | [5 empty] | 133.4784 | 79.505 | 65.5002 | 294.9145
|  |  | 133.5552 | 59.32 | 2.3296 | [5 empty] | 133.5552 | 79.195 | 65.0086 | 304.8703
|  |  | 133.632 | 59.54 | 2.4422 | [5 empty] | 133.632 | 78.765 | 65.6384 | 323.6415
|  |  | 133.7088 | 60.215 | 2.4422 | [5 empty] | 133.7088 | 78.68 | 66.007 | 340.5325
|  |  | 133.7856 | 60.5 | 2.473 | [5 empty] | 133.7856 | 79.04 | 66.2989 | 351.937
|  |  | 133.8624 | 60.555 | 2.5344 | [5 empty] | 133.8624 | 79.615 | 67.0771 | 358.7181
|  |  | 133.9392 | 61.4 | 2.4474 | [5 empty] | 133.9392 | 80.58 | 67.2819 | 362.6018
|  |  | 134.016 | 61.675 | 2.4781 | [5 empty] | 134.016 | 81.345 | 69.4374 | 365.5608
|  |  | 134.0928 | 61.22 | 2.6675 | [5 empty] | 134.0928 | 82.02 | 71.6032 | 369.1363
|  |  | 134.1696 | 61.5 | 2.7034 | [5 empty] | 134.1696 | 82.6 | 72.361 | 373.9446
|  |  | 134.2464 | 61.795 | 2.7238 | [5 empty] | 134.2464 | 83.03 | 71.8182 | 379.4928
|  |  | 134.3232 | 61.865 | 2.7648 | [5 empty] | 134.3232 | 82.975 | 71.8182 | 384.3936
|  |  | 134.4 | 62.265 | 2.7392 | [5 empty] | 134.4 | 82.635 | 72.023 | 387.106
|  |  | 134.4768 | 61.455 | 3.0259 | [5 empty] | 134.4768 | 82.29 | 71.424 | 386.4896
|  |  | 134.5536 | 62.56 | 2.7341 | [5 empty] | 134.5536 | 82.15 | 70.231 | 382.1127
|  |  | 134.6304 | 61.965 | 2.903 | [5 empty] | 134.6304 | 82.045 | 69.7242 | 374.6844
|  |  | 134.7072 | 62.455 | 2.7648 | [5 empty] | 134.7072 | 81.735 | 70.2618 | 365.6841
|  |  | 134.784 | 61.98 | 2.944 | [5 empty] | 134.784 | 81.43 | 70.4512 | 356.807
|  |  | 134.8608 | 62.695 | 2.729 | [5 empty] | 134.8608 | 81.19 | 69.6371 | 349.1629
|  |  | 134.9376 | 61.86 | 2.944 | [5 empty] | 134.9376 | 80.8 | 68.2598 | 343.1216
|  |  | 135.0144 | 62.53 | 2.7546 | [5 empty] | 135.0144 | 80.25 | 66.3347 | 338.0358
|  |  | 135.0912 | 60.995 | 3.1334 | [5 empty] | 135.0912 | 79.68 | 63.7184 | 332.6418
|  |  | 135.168 | 61.825 | 2.7904 | [5 empty] | 135.168 | 78.83 | 62.633 | 325.9532
|  |  | 135.2448 | 61.81 | 2.7392 | [5 empty] | 135.2448 | 77.96 | 63.2781 | 318.3091
|  |  | 135.3216 | 61.68 | 2.6829 | [5 empty] | 135.3216 | 77.94 | 63.3088 | 309.9561
|  |  | 135.3984 | 61.08 | 2.7187 | [5 empty] | 135.3984 | 78.59 | 62.8582 | 300.9558
|  |  | 135.4752 | 60.45 | 2.7546 | [5 empty] | 135.4752 | 79.26 | 63.2371 | 292.7877
|  |  | 135.552 | 60.075 | 2.6829 | [5 empty] | 135.552 | 79.645 | 64.4506 | 285.945
|  |  | 135.6288 | 59.46 | 2.6061 | [5 empty] | 135.6288 | 79.865 | 65.577 | 280.4585
|  |  | 135.7056 | 58.34 | 2.5395 | [5 empty] | 135.7056 | 79.705 | 67.113 | 276.1432
|  |  | 135.7824 | 56.945 | 2.432 | [5 empty] | 135.7824 | 79.39 | 67.3178 | 272.7835
|  |  | 135.8592 | 56.04 | 2.176 | [5 empty] | 135.8592 | 78.995 | 66.5651 | 270.3177
|  |  | 135.936 | 54.755 | 2.0122 | [5 empty] | 135.936 | 78.725 | 64.6144 | 268.5608
|  |  | 136.0128 | 53.08 | 1.9456 | [5 empty] | 136.0128 | 78.34 | 63.1757 | 267.3279
|  |  | 136.0896 | 51.855 | 1.8074 | [5 empty] | 136.0896 | 78.015 | 62.1824 | 266.4648
|  |  | 136.1664 | 50.31 | 1.7254 | [5 empty] | 136.1664 | 77.545 | 62.2848 | 265.8484
|  |  | 136.2432 | 48.775 | 1.6435 | [5 empty] | 136.2432 | 77.345 | 61.6704 | 265.6018
|  |  | 136.32 | 46.92 | 1.5923 | [5 empty] | 136.32 | 77.145 | 61.2659 | 266.5573
|  |  | 136.3968 | 45.85 | 1.3875 | [5 empty] | 136.3968 | 77.13 | 61.056 | 270.5027
|  |  | 136.4736 | 44.725 | 1.1981 | [5 empty] | 136.4736 | 77.33 | 61.3683 | 278.3009
|  |  | 136.5504 | 43.9 | 0.96768 | [5 empty] | 136.5504 | 77.4 | 62.7456 | 290.291
|  |  | 136.6272 | 42.43 | 0.83456 | [5 empty] | 136.6272 | 77.32 | 63.2218 | 305.8258
|  |  | 136.704 | 40.5 | 0.75264 | [5 empty] | 136.704 | 77.315 | 62.5869 | 322.686
|  |  | 136.7808 | 0 | 0 | [5 empty] | 136.7808 | 77.635 | 61.3325 | 337.6968
|  |  | 136.8576 | 0 | 0 | [5 empty] | 136.8576 | 77.725 | 61.2813 | 348.2691
|  |  | 136.9344 | 0 | 0 | [5 empty] | 136.9344 | 77.515 | 62.4435 | 353.5706
|  |  | 137.0112 | 0 | 0 | [5 empty] | 137.0112 | 77.4 | 63.4522 | 354.372
|  |  | 137.088 | 0 | 0 | [5 empty] | 137.088 | 77.68 | 63.2986 | 352.461
|  |  | 137.1648 | 0 | 0 | [5 empty] | 137.1648 | 77.93 | 63.8618 | 349.6561
|  |  | 137.2416 | 0 | 0 | [5 empty] | 137.2416 | 78.155 | 64.2099 | 348.0225
|  |  | 137.3184 | 0 | 0 | [5 empty] | 137.3184 | 78.3 | 63.7389 | 348.1458
|  |  | 137.3952 | 0 | 0 | [5 empty] | 137.3952 | 78.295 | 62.7507 | 349.4403
|  |  | 137.472 | 0 | 0 | [5 empty] | 137.472 | 77.89 | 62.3462 | 350.3342
|  |  | 137.5488 | 0 | 0 | [5 empty] | 137.5488 | 77.255 | 62.0646 | 349.1938
|  |  | 137.6256 | 0 | 0 | [5 empty] | 137.6256 | 76.765 | 60.5286 | 345.1559
|  |  | 137.7024 | 0 | 0 | [5 empty] | 137.7024 | 76.19 | 59.1616 | 338.0975
|  |  | 137.7792 | 0 | 0 | [5 empty] | 137.7792 | 75.535 | 58.4806 | 328.7273
|  |  | 137.856 | 0 | 0 | [5 empty] | 137.856 | 75.29 | 57.7178 | 318.3708
|  |  | 137.9328 | 0 | 0 | [5 empty] | 137.9328 | 75.735 | 56.8166 | 309.0006
|  |  | 138.0096 | 0 | 0 | [5 empty] | 138.0096 | 76.02 | 57.8458 | 307.8601
|  |  | 138.0864 | 0 | 0 | [5 empty] | 138.0864 | 76.025 | 59.6941 | 308.7232
|  |  | 138.1632 | 0 | 0 | [5 empty] | 138.1632 | 75.84 | 60.8512 | 302.9284
|  |  | 138.24 | 0 | 0 | [5 empty] | 138.24 | 75.56 | 60.6925 | 296.6406
|  |  | 138.3168 | 0 | 0 | [5 empty] | 138.3168 | 75.225 | 59.8528 | 291.4006
|  |  | 138.3936 | 0 | 0 | [5 empty] | 138.3936 | 74.995 | 58.7571 | 287.332
|  |  | 138.4704 | 0 | 0 | [5 empty] | 138.4704 | 74.895 | 58.0301 | 284.6812
|  |  | 138.5472 | 0 | 0 | [5 empty] | 138.5472 | 74.815 | 58.4806 | 283.2017
|  |  | 138.624 | 0 | 0 | [5 empty] | 138.624 | 75.125 | 58.0608 | 282.5544
|  |  | 138.7008 | 0 | 0 | [5 empty] | 138.7008 | 75.365 | 58.5779 | 282.3387
|  |  | 138.7776 | 0 | 0 | [5 empty] | 138.7776 | 75.42 | 59.945 | 282.462
|  |  | 138.8544 | 0 | 0 | [5 empty] | 138.8544 | 75.475 | 59.9654 | 282.8935
|  |  | 138.9312 | 0 | 0 | [5 empty] | 138.9312 | 75.505 | 58.6342 | 283.5099
|  |  | 139.008 | 0 | 0 | [5 empty] | 139.008 | 75.435 | 56.7091 | 285.6676
|  |  | 139.0848 | 0 | 0 | [5 empty] | 139.0848 | 75.09 | 55.7517 | 297.1646
|  |  | 139.1616 | 0 | 0 | [5 empty] | 139.1616 | 74.57 | 56.0742 | 317.7543
|  |  | 139.2384 | 0 | 0 | [5 empty] | 139.2384 | 74.545 | 55.6339 | 335.8474
|  |  | 139.3152 | 0 | 0 | [5 empty] | 139.3152 | 74.81 | 55.8899 | 347.4985
|  |  | 139.392 | 0 | 0 | [5 empty] | 139.392 | 75.225 | 57.2518 | 353.6323
|  |  | 139.4688 | 0 | 0 | [5 empty] | 139.4688 | 75.84 | 59.049 | 356.3447
|  |  | 139.5456 | 0 | 0 | [5 empty] | 139.5456 | 76.74 | 59.9296 | 358.5023
|  |  | 139.6224 | 0 | 0 | [5 empty] | 139.6224 | 77.475 | 61.2659 | 362.1086
|  |  | 139.6992 | 0 | 0 | [5 empty] | 139.6992 | 77.9 | 62.9299 | 367.9033
|  |  | 139.776 | 0 | 0 | [5 empty] | 139.776 | 78.22 | 63.7286 | 374.9926
|  |  | 139.8528 | 0 | 0 | [5 empty] | 139.8528 | 78.37 | 63.8976 | 381.5887
|  |  | 139.9296 | 0 | 0 | [5 empty] | 139.9296 | 78.345 | 63.5085 | 385.6266
|  |  | 140.0064 | 0 | 0 | [5 empty] | 140.0064 | 78.15 | 63.319 | 385.7807
|  |  | 140.0832 | 0 | 0 | [5 empty] | 140.0832 | 77.955 | 63.1654 | 381.8353
|  |  | 140.16 | 0 | 0 | [5 empty] | 140.16 | 78.005 | 62.4845 | 374.6227
|  |  | 140.2368 | 0 | 0 | [5 empty] | 140.2368 | 77.935 | 62.72 | 365.9307
|  |  | 140.3136 | 0 | 0 | [5 empty] | 140.3136 | 77.62 | 63.657 | 357.7934
|  |  | 140.3904 | 0 | 0 | [5 empty] | 140.3904 | 77.545 | 63.4368 | 351.4438
|  |  | 140.4672 | 0 | 0 | [5 empty] | 140.4672 | 77.63 | 61.5834 | 347.1286
|  |  | 140.544 | 0 | 0 | [5 empty] | 140.544 | 77.135 | 61.1584 | 343.9847
|  |  | 140.6208 | 0 | 0 | [5 empty] | 140.6208 | 76.64 | 60.1139 | 340.4092
|  |  | 140.6976 | 0 | 0 | [5 empty] | 140.6976 | 76.035 | 59.5354 | 335.2926
|  |  | 140.7744 | 0 | 0 | [5 empty] | 140.7744 | 75.77 | 59.1053 | 329.128
|  |  | 140.8512 | 0 | 0 | [5 empty] | 140.8512 | 76.43 | 58.8646 | 321.8846
|  |  | 140.928 | 0 | 0 | [5 empty] | 140.928 | 77.435 | 59.6429 | 314.148
|  |  | 141.0048 | 0 | 0 | [5 empty] | 141.0048 | 78.245 | 61.4195 | 307.4594
|  |  | 141.0816 | 0 | 0 | [5 empty] | 141.0816 | 78.835 | 63.5546 | 302.0654
|  |  | 141.1584 | 0 | 0 | [5 empty] | 141.1584 | 79.415 | 64.8141 | 297.8735
|  |  | 141.2352 | 0 | 0 | [5 empty] | 141.2352 | 79.945 | 65.1366 | 294.7603
|  |  | 141.312 | 0 | 0 | [5 empty] | 141.312 | 80.28 | 64.8192 | 292.5103
|  |  | 141.3888 | 0 | 0 | [5 empty] | 141.3888 | 80.275 | 64.7834 | 290.9691
|  |  | 141.4656 | 0 | 0 | [5 empty] | 141.4656 | 80.155 | 64.9062 | 289.9211
|  |  | 141.5424 | 0 | 0 | [5 empty] | 141.5424 | 80.04 | 64.8755 | 289.2122
|  |  | 141.6192 | 0 | 0 | [5 empty] | 141.6192 | 80.17 | 64.2714 | 288.5341
|  |  | 141.696 | 0 | 0 | [5 empty] | 141.696 | 80.12 | 64.9011 | 287.7327
|  |  | 141.7728 | 0 | 0 | [5 empty] | 141.7728 | 80.26 | 64.983 | 286.9313
|  |  | 141.8496 | 0 | 0 | [5 empty] | 141.8496 | 80.365 | 65.2646 | 287.3628
|  |  | 141.9264 | 0 | 0 | [5 empty] | 141.9264 | 80.615 | 65.6845 | 290.9075
|  |  | 142.0032 | 0 | 0 | [5 empty] | 142.0032 | 81.11 | 66.5446 | 298.3666
|  |  | 142.08 | 0 | 0 | [5 empty] | 142.08 | 81.63 | 67.2051 | 310.2951
|  |  | 142.1568 | 0 | 0 | [5 empty] | 142.1568 | 81.705 | 68.7462 | 325.7683
|  |  | 142.2336 | 0 | 0 | [5 empty] | 142.2336 | 81.695 | 70.1952 | 342.3511
|  |  | 142.3104 | 0 | 0 | [5 empty] | 142.3104 | 82.16 | 70.3795 | 356.992
|  |  | 142.3872 | 0 | 0 | [5 empty] | 142.3872 | 82.475 | 70.3181 | 367.0403
|  |  | 142.464 | 0 | 0 | [5 empty] | 142.464 | 82.765 | 69.673 | 371.263
|  |  | 142.5408 | 0 | 0 | [5 empty] | 142.5408 | 82.835 | 70.1952 | 370.4308
|  |  | 142.6176 | 0 | 0 | [5 empty] | 142.6176 | 82.98 | 70.7379 | 366.6704
|  |  | 142.6944 | 0 | 0 | [5 empty] | 142.6944 | 83.195 | 71.2294 | 362.2011
|  |  | 142.7712 | 0 | 0 | [5 empty] | 142.7712 | 83.205 | 72.5606 | 358.8414
|  |  | 142.848 | 0 | 0 | [5 empty] | 142.848 | 83.395 | 72.0282 | 357.1153
|  |  | 142.9248 | 0 | 0 | [5 empty] | 142.9248 | 83.335 | 71.3216 | 356.5605
|  |  | 143.0016 | 0 | 0 | [5 empty] | 143.0016 | 82.975 | 70.7021 | 355.8824
|  |  | 143.0784 | 0 | 0 | [5 empty] | 143.0784 | 82.49 | 69.7242 | 353.4165
|  |  | 143.1552 | 0 | 0 | [5 empty] | 143.1552 | 81.885 | 68.2035 | 348.2383
|  |  | 143.232 | 0 | 0 | [5 empty] | 143.232 | 81.305 | 66.2067 | 340.3167
|  |  | 143.3088 | 0 | 0 | [5 empty] | 143.3088 | 80.71 | 64.8397 | 330.1451
|  |  | 143.3856 | 0 | 0 | [5 empty] | 143.3856 | 80.27 | 64.9267 | 318.5249
|  |  | 143.4624 | 0 | 0 | [5 empty] | 143.4624 | 80.405 | 65.4029 | 306.2265
|  |  | 143.5392 | 0 | 0 | [5 empty] | 143.5392 | 80.75 | 65.8227 | 297.8735
|  |  | 143.616 | 0 | 0 | [5 empty] | 143.616 | 80.785 | 67.0515 | 299.1064
|  |  | 143.6928 | 0 | 0 | [5 empty] | 143.6928 | 80.68 | 67.5072 | 295.8391
|  |  | 143.7696 | 0 | 0 | [5 empty] | 143.7696 | 80.61 | 66.1965 | 288.1026
|  |  | 143.8464 | 0 | 0 | [5 empty] | 143.8464 | 80.27 | 64.8909 | 280.6434
|  |  | 143.9232 | 0 | 0 | [5 empty] | 143.9232 | 79.735 | 64.768 | 273.9857
|  |  | 144 | 0 | 0 | [5 empty] | 144 | 79.45 | 64.4096 | 268.4067
|  |  | 144.0768 | 0 | 0 | [5 empty] | 144.0768 | 79.425 | 63.4317 | 264.2456
|  |  | 144.1536 | 0 | 0 | [5 empty] | 144.1536 | 79.47 | 62.9146 | 261.3174
|  |  | 144.2304 | 0 | 0 | [5 empty] | 144.2304 | 79.395 | 63.2678 | 259.4064
|  |  | 144.3072 | 0 | 0 | [5 empty] | 144.3072 | 79.31 | 63.3958 | 258.3892
|  |  | 144.384 | 0 | 0 | [5 empty] | 144.384 | 79.04 | 63.7133 | 258.081
|  |  | 144.4608 | 0 | 0 | [5 empty] | 144.4608 | 78.76 | 62.8531 | 258.1426
|  |  | 144.5376 | 0 | 0 | [5 empty] | 144.5376 | 78.545 | 60.8 | 258.5125
|  |  | 144.6144 | 0 | 0 | [5 empty] | 144.6144 | 77.985 | 59.4125 | 263.6908
|  |  | 144.6912 | 0 | 0 | [5 empty] | 144.6912 | 77.345 | 58.2502 | 279.6571
|  |  | 144.768 | 0 | 0 | [5 empty] | 144.768 | 76.735 | 58.1581 | 297.7194
|  |  | 144.8448 | 0 | 0 | [5 empty] | 144.8448 | 76.535 | 58.0301 | 310.5417
|  |  | 144.9216 | 0 | 0 | [5 empty] | 144.9216 | 76.735 | 57.8816 | 318.2166
|  |  | 144.9984 | 0 | 0 | [5 empty] | 144.9984 | 77.095 | 58.8083 | 322.7784
|  |  | 145.0752 | 0 | 0 | [5 empty] | 145.0752 | 77.64 | 60.4672 | 326.5388
|  |  | 145.152 | 0 | 0 | [5 empty] | 145.152 | 78.335 | 61.2966 | 331.4397
|  |  | 145.2288 | 0 | 0 | [5 empty] | 145.2288 | 78.945 | 61.7114 | 338.0667
|  |  | 145.3056 | 0 | 0 | [5 empty] | 145.3056 | 79.32 | 61.7728 | 345.9574
|  |  | 145.3824 | 0 | 0 | [5 empty] | 145.3824 | 79.26 | 62.0237 | 353.509
|  |  | 145.4592 | 0 | 0 | [5 empty] | 145.4592 | 78.925 | 62.4947 | 358.8722
|  |  | 145.536 | 0 | 0 | [5 empty] | 145.536 | 78.71 | 62.0646 | 360.7832
|  |  | 145.6128 | 0 | 0 | [5 empty] | 145.6128 | 78.475 | 61.5731 | 358.9338
|  |  | 145.6896 | 0 | 0 | [5 empty] | 145.6896 | 78.225 | 61.2352 | 353.9713
|  |  | 145.7664 | 0 | 0 | [5 empty] | 145.7664 | 78.1 | 60.631 | 347.3444
|  |  | 145.8432 | 0 | 0 | [5 empty] | 145.8432 | 78.06 | 60.5133 | 340.6558
|  |  | 145.92 | 0 | 0 | [5 empty] | 145.92 | 77.975 | 60.3392 | 335.0768
|  |  | 145.9968 | 0 | 0 | [5 empty] | 145.9968 | 77.54 | 60.4211 | 330.8233
|  |  | 146.0736 | 0 | 0 | [5 empty] | 146.0736 | 76.985 | 59.6634 | 327.4019
|  |  | 146.1504 | 0 | 0 | [5 empty] | 146.1504 | 76.44 | 57.9584 | 323.5799
|  |  | 146.2272 | 0 | 0 | [5 empty] | 146.2272 | 75.805 | 56.617 | 318.155
|  |  | 146.304 | 0 | 0 | [5 empty] | 146.304 | 75.37 | 55.2602 | 311.2506
|  |  | 146.3808 | 0 | 0 | [5 empty] | 146.3808 | 75.37 | 54.4717 | 303.7298
|  |  | 146.4576 | 0 | 0 | [5 empty] | 146.4576 | 75.54 | 55.6749 | 295.8391
|  |  | 146.5344 | 0 | 0 | [5 empty] | 146.5344 | 75.645 | 58.3782 | 288.8115
|  |  | 146.6112 | 0 | 0 | [5 empty] | 146.6112 | 76.155 | 59.2845 | 283.2942
|  |  | 146.688 | 0 | 0 | [5 empty] | 146.688 | 76.6 | 59.6736 | 279.2564
|  |  | 146.7648 | 0 | 0 | [5 empty] | 146.7648 | 76.875 | 59.8835 | 276.4515
|  |  | 146.8416 | 0 | 0 | [5 empty] | 146.8416 | 77.11 | 59.2333 | 274.6638
|  |  | 146.9184 | 0 | 0 | [5 empty] | 146.9184 | 77.05 | 58.6035 | 273.6774
|  |  | 146.9952 | 0 | 0 | [5 empty] | 146.9952 | 76.73 | 58.3987 | 273.3692
|  |  | 147.072 | 0 | 0 | [5 empty] | 147.072 | 76.35 | 58.4806 | 273.4617
|  |  | 147.1488 | 0 | 0 | [5 empty] | 147.1488 | 76.04 | 58.7213 | 273.8932
|  |  | 147.2256 | 0 | 0 | [5 empty] | 147.2256 | 76.11 | 57.6819 | 274.4788
|  |  | 147.3024 | 0 | 0 | [5 empty] | 147.3024 | 76.075 | 57.6205 | 275.1261
|  |  | 147.3792 | 0 | 0 | [5 empty] | 147.3792 | 75.83 | 58.4858 | 276.359
|  |  | 147.456 | 0 | 0 | [5 empty] | 147.456 | 76.085 | 57.7792 | 279.9961
|  |  | 147.5328 | 0 | 0 | [5 empty] | 147.5328 | 76.315 | 58.9363 | 287.2087
|  |  | 147.6096 | 0 | 0 | [5 empty] | 147.6096 | 76.65 | 60.6669 | 298.3975
|  |  | 147.6864 | 0 | 0 | [5 empty] | 147.6864 | 77.055 | 61.2198 | 313.4699
|  |  | 147.7632 | 0 | 0 | [5 empty] | 147.7632 | 77.64 | 59.991 | 330.4534
|  |  | 147.84 | 0 | 0 | [5 empty] | 147.84 | 77.79 | 60.841 | 346.4197
|  |  | 147.9168 | 0 | 0 | [5 empty] | 147.9168 | 77.765 | 61.911 | 358.564
|  |  | 147.9936 | 0 | 0 | [5 empty] | 147.9936 | 77.875 | 61.7882 | 365.3142
|  |  | 148.0704 | 0 | 0 | [5 empty] | 148.0704 | 77.915 | 62.3616 | 366.7937
|  |  | 148.1472 | 0 | 0 | [5 empty] | 148.1472 | 78.12 | 63.0835 | 364.5744
|  |  | 148.224 | 0 | 0 | [5 empty] | 148.224 | 78.32 | 64.4659 | 360.8757
|  |  | 148.3008 | 0 | 0 | [5 empty] | 148.3008 | 78.76 | 64.9523 | 357.7009
|  |  | 148.3776 | 0 | 0 | [5 empty] | 148.3776 | 79.04 | 65.4387 | 356.2214
|  |  | 148.4544 | 0 | 0 | [5 empty] | 148.4544 | 79.24 | 64.7578 | 356.3755
|  |  | 148.5312 | 0 | 0 | [5 empty] | 148.5312 | 78.945 | 64.5683 | 357.0536
|  |  | 148.608 | 0 | 0 | [5 empty] | 148.608 | 78.31 | 64.7322 | 356.7762
|  |  | 148.6848 | 0 | 0 | [5 empty] | 148.6848 | 77.97 | 62.8531 | 354.1871
|  |  | 148.7616 | 0 | 0 | [5 empty] | 148.7616 | 77.465 | 61.8342 | 348.7314
|  |  | 148.8384 | 0 | 0 | [5 empty] | 148.8384 | 76.84 | 61.5936 | 340.8407
|  |  | 148.9152 | 0 | 0 | [5 empty] | 148.9152 | 76.78 | 60.2624 | 331.224
|  |  | 148.992 | 0 | 0 | [5 empty] | 148.992 | 77.235 | 59.776 | 320.6825
|  |  | 149.0688 | 0 | 0 | [5 empty] | 149.0688 | 77.735 | 60.9894 | 314.1172
|  |  | 149.1456 | 0 | 0 | [5 empty] | 149.1456 | 78.15 | 62.2387 | 316.9221
|  |  | 149.2224 | 0 | 0 | [5 empty] | 149.2224 | 78.4 | 63.145 | 314.9186
|  |  | 149.2992 | 39.345 | 0.95744 | [5 empty] | 149.2992 | 78.435 | 63.4112 | 308.8465
|  |  | 149.376 | 42.945 | 0.91648 | [5 empty] | 149.376 | 78.415 | 62.7507 | 303.3291
|  |  | 149.4528 | 45.33 | 1.0445 | [5 empty] | 149.4528 | 78.355 | 62.2387 | 298.7673
|  |  | 149.5296 | 47.32 | 1.1981 | [5 empty] | 149.5296 | 78.15 | 62.8224 | 295.4076
|  |  | 149.6064 | 49.265 | 1.3517 | [5 empty] | 149.6064 | 78.255 | 63.2986 | 293.25
|  |  | 149.6832 | 50.705 | 1.5667 | [5 empty] | 149.6832 | 78.57 | 63.9386 | 291.8938
|  |  | 149.76 | 52.565 | 1.6742 | [5 empty] | 149.76 | 78.955 | 64.7987 | 291.1541
|  |  | 149.8368 | 54 | 1.833 | [5 empty] | 149.8368 | 79.31 | 65.7203 | 290.8458
|  |  | 149.9136 | 55.47 | 1.961 | [5 empty] | 149.9136 | 79.745 | 65.792 | 290.9075
|  |  | 149.9904 | 56.99 | 2.0326 | [5 empty] | 149.9904 | 79.82 | 66.2016 | 291.1849
|  |  | 150.0672 | 57.815 | 2.1862 | [5 empty] | 150.0672 | 79.82 | 65.5514 | 291.339
|  |  | 150.144 | 58.675 | 2.2835 | [5 empty] | 150.144 | 79.65 | 64.7731 | 292.4795
|  |  | 150.2208 | 59.375 | 2.3552 | [5 empty] | 150.2208 | 79.45 | 64.0666 | 301.8805
|  |  | 150.2976 | 59.905 | 2.4064 | [5 empty] | 150.2976 | 79.225 | 63.6518 | 320.4667
|  |  | 150.3744 | 60.185 | 2.4576 | [5 empty] | 150.3744 | 79.155 | 63.8362 | 337.481
|  |  | 150.4512 | 60.445 | 2.5037 | [5 empty] | 150.4512 | 79.445 | 64.4045 | 348.978
|  |  | 150.528 | 60.845 | 2.5088 | [5 empty] | 150.528 | 80.02 | 65.239 | 355.7591
|  |  | 150.6048 | 61.055 | 2.5446 | [5 empty] | 150.6048 | 80.49 | 68.0192 | 359.6428
|  |  | 150.6816 | 61.435 | 2.5651 | [5 empty] | 150.6816 | 81.245 | 70.2618 | 362.7559
|  |  | 150.7584 | 61.725 | 2.5958 | [5 empty] | 150.7584 | 81.99 | 71.5827 | 366.6396
|  |  | 150.8352 | 61.925 | 2.6419 | [5 empty] | 150.8352 | 82.635 | 71.8336 | 371.7562
|  |  | 150.912 | 62.14 | 2.6829 | [5 empty] | 150.912 | 82.765 | 72.8474 | 377.3968
|  |  | 150.9888 | 61.28 | 3.0259 | [5 empty] | 150.9888 | 82.65 | 73.3952 | 382.1744
|  |  | 151.0656 | 62.74 | 2.688 | [5 empty] | 151.0656 | 82.6 | 72.3917 | 384.5478
|  |  | 151.1424 | 61.345 | 3.113 | [5 empty] | 151.1424 | 82.54 | 70.6662 | 383.5306
|  |  | 151.2192 | 62.79 | 2.7341 | [5 empty] | 151.2192 | 82.42 | 69.3094 | 379.0304
|  |  | 151.296 | 61.365 | 3.1437 | [5 empty] | 151.296 | 82.23 | 68.7718 | 371.8178
|  |  | 151.3728 | 62.66 | 2.775 | [5 empty] | 151.3728 | 81.995 | 68.9971 | 363.3415
|  |  | 151.4496 | 61.515 | 3.1078 | [5 empty] | 151.4496 | 81.645 | 69.5654 | 355.0501
|  |  | 151.5264 | 62.83 | 2.7494 | [5 empty] | 151.5264 | 81.17 | 69.9802 | 347.8376
|  |  | 151.6032 | 61.42 | 3.1232 | [5 empty] | 151.6032 | 80.65 | 69.2582 | 341.7654
|  |  | 151.68 | 62.35 | 2.8006 | [5 empty] | 151.68 | 80.275 | 66.6163 | 336.1556
|  |  | 151.7568 | 60.775 | 3.241 | [5 empty] | 151.7568 | 79.42 | 65.6589 | 330.0835
|  |  | 151.8336 | 62.435 | 2.7136 | [5 empty] | 151.8336 | 78.72 | 64.1075 | 322.8093
|  |  | 151.9104 | 60.38 | 3.2563 | [5 empty] | 151.9104 | 78.39 | 62.3411 | 314.672
|  |  | 151.9872 | 62.17 | 2.6368 | [5 empty] | 151.9872 | 78.33 | 62.1158 | 306.3806
|  |  | 152.064 | 60.645 | 2.9235 | [5 empty] | 152.064 | 78.54 | 63.4266 | 297.781
|  |  | 152.1408 | 60.265 | 2.8262 | [5 empty] | 152.1408 | 78.94 | 64.599 | 290.0136
|  |  | 152.2176 | 59.9 | 2.7699 | [5 empty] | 152.2176 | 79.505 | 64.9677 | 283.4791
|  |  | 152.2944 | 59.4 | 2.7034 | [5 empty] | 152.2944 | 79.875 | 65.4746 | 278.0543
|  |  | 152.3712 | 58.605 | 2.5907 | [5 empty] | 152.3712 | 79.985 | 65.8381 | 273.5541
|  |  | 152.448 | 57.485 | 2.4422 | [5 empty] | 152.448 | 79.86 | 65.5974 | 269.8862
|  |  | 152.5248 | 56.66 | 2.1914 | [5 empty] | 152.5248 | 79.425 | 65.1469 | 267.1121
|  |  | 152.6016 | 55.31 | 2.0122 | [5 empty] | 152.6016 | 78.995 | 63.7798 | 265.2319
|  |  | 152.6784 | 53.255 | 2.007 | [5 empty] | 152.6784 | 78.635 | 62.4538 | 264.1223
|  |  | 152.7552 | 51.515 | 1.9763 | [5 empty] | 152.7552 | 78.29 | 61.2659 | 263.5675
|  |  | 152.832 | 50.46 | 1.7869 | [5 empty] | 152.832 | 77.985 | 60.3187 | 263.2901
|  |  | 152.9088 | 49.4 | 1.5974 | [5 empty] | 152.9088 | 77.69 | 59.9194 | 263.1976
|  |  | 152.9856 | 48.43 | 1.3926 | [5 empty] | 152.9856 | 77.235 | 60.4416 | 263.999
|  |  | 153.0624 | 47.095 | 1.2493 | [5 empty] | 153.0624 | 77.235 | 60.3341 | 267.6053
|  |  | 153.1392 | 45.6 | 1.1162 | [5 empty] | 153.1392 | 77.395 | 61.0765 | 275.0028
|  |  | 153.216 | 43.53 | 1.0906 | [5 empty] | 153.216 | 77.51 | 62.08 | 286.7772
|  |  | 153.2928 | 41.275 | 1.0752 | [5 empty] | 153.2928 | 77.67 | 61.312 | 302.4044
|  |  | 153.3696 | 39.97 | 0.88064 | [5 empty] | 153.3696 | 77.51 | 61.4042 | 319.4496
|  |  | 153.4464 | 0 | 0 | [5 empty] | 153.4464 | 77.4 | 62.2182 | 334.7686
|  |  | 153.5232 | 0 | 0 | [5 empty] | 153.5232 | 77.485 | 61.9776 | 345.7108
|  |  | 153.6 | 0 | 0 | [5 empty] | 153.6 | 77.695 | 61.184 | 351.1356
|  |  | 153.6768 | 0 | 0 | [5 empty] | 153.6768 | 77.65 | 62.0339 | 351.9678
|  |  | 153.7536 | 0 | 0 | [5 empty] | 153.7536 | 77.75 | 63.1808 | 349.9644
|  |  | 153.8304 | 0 | 0 | [5 empty] | 153.8304 | 78.025 | 63.7594 | 347.3444
|  |  | 153.9072 | 0 | 0 | [5 empty] | 153.9072 | 78.21 | 64.4454 | 346.0807
|  |  | 153.984 | 0 | 0 | [5 empty] | 153.984 | 78.325 | 64.3328 | 346.4505
|  |  | 154.0608 | 0 | 0 | [5 empty] | 154.0608 | 78.235 | 63.8003 | 347.6526
|  |  | 154.1376 | 0 | 0 | [5 empty] | 154.1376 | 77.91 | 62.7917 | 348.454
|  |  | 154.2144 | 0 | 0 | [5 empty] | 154.2144 | 77.27 | 62.3411 | 347.2827
|  |  | 154.2912 | 0 | 0 | [5 empty] | 154.2912 | 76.68 | 61.2147 | 343.06
|  |  | 154.368 | 0 | 0 | [5 empty] | 154.368 | 76.285 | 59.1821 | 335.9091
|  |  | 154.4448 | 0 | 0 | [5 empty] | 154.4448 | 75.885 | 57.5744 | 326.7854
|  |  | 154.5216 | 0 | 0 | [5 empty] | 154.5216 | 75.425 | 57.7382 | 316.7988
|  |  | 154.5984 | 0 | 0 | [5 empty] | 154.5984 | 75.545 | 58.3987 | 307.3978
|  |  | 154.6752 | 0 | 0 | [5 empty] | 154.6752 | 76.065 | 58.5421 | 305.5792
|  |  | 154.752 | 0 | 0 | [5 empty] | 154.752 | 76.445 | 58.6035 | 305.9183
|  |  | 154.8288 | 0 | 0 | [5 empty] | 154.8288 | 76.195 | 59.6992 | 299.9386
|  |  | 154.9056 | 0 | 0 | [5 empty] | 154.9056 | 75.905 | 59.2282 | 293.7124
|  |  | 154.9824 | 0 | 0 | [5 empty] | 154.9824 | 75.59 | 58.3629 | 288.6574
|  |  | 155.0592 | 0 | 0 | [5 empty] | 155.0592 | 75.45 | 56.8013 | 284.9278
|  |  | 155.136 | 0 | 0 | [5 empty] | 155.136 | 75.29 | 56.1562 | 282.4928
|  |  | 155.2128 | 0 | 0 | [5 empty] | 155.2128 | 75.295 | 56.2893 | 280.9825
|  |  | 155.2896 | 0 | 0 | [5 empty] | 155.2896 | 75.55 | 56.5658 | 280.0886
|  |  | 155.3664 | 0 | 0 | [5 empty] | 155.3664 | 75.82 | 57.2365 | 279.6263
|  |  | 155.4432 | 0 | 0 | [5 empty] | 155.4432 | 76.165 | 56.8781 | 279.6879
|  |  | 155.52 | 0 | 0 | [5 empty] | 155.52 | 76.205 | 56.9907 | 280.2119
|  |  | 155.5968 | 0 | 0 | [5 empty] | 155.5968 | 75.695 | 58.1427 | 281.1366
|  |  | 155.6736 | 0 | 0 | [5 empty] | 155.6736 | 75.16 | 58.537 | 283.4791
|  |  | 155.7504 | 0 | 0 | [5 empty] | 155.7504 | 74.71 | 58.0659 | 294.5446
|  |  | 155.8272 | 0 | 0 | [5 empty] | 155.8272 | 74.355 | 57.7178 | 314.8878
|  |  | 155.904 | 0 | 0 | [5 empty] | 155.904 | 74.595 | 56.0486 | 333.0425
|  |  | 155.9808 | 0 | 0 | [5 empty] | 155.9808 | 75.015 | 55.0707 | 344.8169
|  |  | 156.0576 | 0 | 0 | [5 empty] | 156.0576 | 75.315 | 56.3763 | 351.0123
|  |  | 156.1344 | 0 | 0 | [5 empty] | 156.1344 | 75.745 | 59.0387 | 353.8481
|  |  | 156.2112 | 0 | 0 | [5 empty] | 156.2112 | 76.41 | 61.3581 | 356.0673
|  |  | 156.288 | 0 | 0 | [5 empty] | 156.288 | 77.215 | 62.6842 | 359.7969
|  |  | 156.3648 | 0 | 0 | [5 empty] | 156.3648 | 77.81 | 63.9846 | 365.7149
|  |  | 156.4416 | 0 | 0 | [5 empty] | 156.4416 | 78.31 | 63.9078 | 372.9891
|  |  | 156.5184 | 0 | 0 | [5 empty] | 156.5184 | 78.48 | 63.8106 | 379.6777
|  |  | 156.5952 | 0 | 0 | [5 empty] | 156.5952 | 78.19 | 64.5427 | 383.9005
|  |  | 156.672 | 0 | 0 | [5 empty] | 156.672 | 78.04 | 64.2918 | 384.1471
|  |  | 156.7488 | 0 | 0 | [5 empty] | 156.7488 | 77.895 | 63.9846 | 380.14
|  |  | 156.8256 | 0 | 0 | [5 empty] | 156.8256 | 77.72 | 64.3482 | 372.8658
|  |  | 156.9024 | 0 | 0 | [5 empty] | 156.9024 | 77.75 | 64.1894 | 364.2354
|  |  | 156.9792 | 0 | 0 | [5 empty] | 156.9792 | 77.81 | 63.6467 | 356.2214
|  |  | 157.056 | 0 | 0 | [5 empty] | 157.056 | 77.795 | 62.8275 | 349.9335
|  |  | 157.1328 | 0 | 0 | [5 empty] | 157.1328 | 77.63 | 62.2285 | 345.6799
|  |  | 157.2096 | 0 | 0 | [5 empty] | 157.2096 | 77 | 62.7712 | 342.5977
|  |  | 157.2864 | 0 | 0 | [5 empty] | 157.2864 | 76.64 | 61.225 | 339.0222
|  |  | 157.3632 | 0 | 0 | [5 empty] | 157.3632 | 76.24 | 60.288 | 333.8747
|  |  | 157.44 | 0 | 0 | [5 empty] | 157.44 | 75.82 | 61.2557 | 327.5869
|  |  | 157.5168 | 0 | 0 | [5 empty] | 157.5168 | 76.315 | 61.5066 | 320.1893
|  |  | 157.5936 | 0 | 0 | [5 empty] | 157.5936 | 77.115 | 62.4538 | 312.3603
|  |  | 157.6704 | 0 | 0 | [5 empty] | 157.6704 | 78.05 | 63.2115 | 305.5176
|  |  | 157.7472 | 0 | 0 | [5 empty] | 157.7472 | 78.955 | 63.7747 | 300.0002
|  |  | 157.824 | 0 | 0 | [5 empty] | 157.824 | 79.585 | 64.9114 | 295.8391
|  |  | 157.9008 | 0 | 0 | [5 empty] | 157.9008 | 79.935 | 66.0736 | 292.6952
|  |  | 157.9776 | 0 | 0 | [5 empty] | 157.9776 | 80.135 | 66.1965 | 290.4451
|  |  | 158.0544 | 0 | 0 | [5 empty] | 158.0544 | 80.3 | 65.3414 | 288.9348
|  |  | 158.1312 | 0 | 0 | [5 empty] | 158.1312 | 80.185 | 65.2134 | 287.8868
|  |  | 158.208 | 0 | 0 | [5 empty] | 158.208 | 80.115 | 65.0138 | 287.1471
|  |  | 158.2848 | 0 | 0 | [5 empty] | 158.2848 | 80.085 | 65.408 | 286.469
|  |  | 158.3616 | 0 | 0 | [5 empty] | 158.3616 | 80.145 | 65.6691 | 285.76
|  |  | 158.4384 | 0 | 0 | [5 empty] | 158.4384 | 80.22 | 65.792 | 285.1128
|  |  | 158.5152 | 0 | 0 | [5 empty] | 158.5152 | 80.3 | 66.0787 | 285.5135
|  |  | 158.592 | 0 | 0 | [5 empty] | 158.592 | 80.235 | 67.8144 | 288.8732
|  |  | 158.6688 | 0 | 0 | [5 empty] | 158.6688 | 80.78 | 68.2957 | 296.0241
|  |  | 158.7456 | 0 | 0 | [5 empty] | 158.7456 | 81.48 | 68.2138 | 307.5211
|  |  | 158.8224 | 0 | 0 | [5 empty] | 158.8224 | 81.735 | 68.7718 | 322.7168
|  |  | 158.8992 | 0 | 0 | [5 empty] | 158.8992 | 81.89 | 69.4272 | 339.4229
|  |  | 158.976 | 0 | 0 | [5 empty] | 158.976 | 82.035 | 70.9171 | 354.4645
|  |  | 159.0528 | 0 | 0 | [5 empty] | 159.0528 | 82.44 | 70.5536 | 364.9751
|  |  | 159.1296 | 0 | 0 | [5 empty] | 159.1296 | 82.725 | 70.0877 | 369.6602
|  |  | 159.2064 | 0 | 0 | [5 empty] | 159.2064 | 82.715 | 70.9888 | 369.013
|  |  | 159.2832 | 0 | 0 | [5 empty] | 159.2832 | 82.925 | 71.5059 | 365.1909
|  |  | 159.36 | 0 | 0 | [5 empty] | 159.36 | 83.195 | 71.9309 | 360.5674
|  |  | 159.4368 | 0 | 0 | [5 empty] | 159.4368 | 83.245 | 72.832 | 356.9612
|  |  | 159.5136 | 0 | 0 | [5 empty] | 159.5136 | 83.315 | 72.8064 | 355.3275
|  |  | 159.5904 | 0 | 0 | [5 empty] | 159.5904 | 83.285 | 71.8643 | 355.0501
|  |  | 159.6672 | 0 | 0 | [5 empty] | 159.6672 | 82.905 | 71.2704 | 354.7419
|  |  | 159.744 | 0 | 0 | [5 empty] | 159.744 | 82.34 | 70.4666 | 352.6459
|  |  | 159.8208 | 0 | 0 | [5 empty] | 159.8208 | 81.715 | 69.0637 | 347.4369
|  |  | 159.8976 | 0 | 0 | [5 empty] | 159.8976 | 81.06 | 67.5482 | 339.1763
|  |  | 159.9744 | 0 | 0 | [5 empty] | 159.9744 | 80.5 | 65.8688 | 328.6965
|  |  | 160.0512 | 0 | 0 | [5 empty] | 160.0512 | 80 | 66.0992 | 316.9837
|  |  | 160.128 | 0 | 0 | [5 empty] | 160.128 | 80.21 | 66.2272 | 305.0244
|  |  | 160.2048 | 0 | 0 | [5 empty] | 160.2048 | 80.735 | 65.792 | 297.8118
|  |  | 160.2816 | 0 | 0 | [5 empty] | 160.2816 | 80.855 | 66.6675 | 299.2297
|  |  | 160.3584 | 0 | 0 | [5 empty] | 160.3584 | 80.955 | 66.3142 | 294.8528
|  |  | 160.4352 | 0 | 0 | [5 empty] | 160.4352 | 80.87 | 65.0394 | 286.6847
|  |  | 160.512 | 0 | 0 | [5 empty] | 160.512 | 80.43 | 64.2611 | 279.2256
|  |  | 160.5888 | 0 | 0 | [5 empty] | 160.5888 | 80.025 | 63.5546 | 272.8144
|  |  | 160.6656 | 0 | 0 | [5 empty] | 160.6656 | 79.775 | 62.6227 | 267.6669
|  |  | 160.7424 | 0 | 0 | [5 empty] | 160.7424 | 79.49 | 63.1603 | 263.7832
|  |  | 160.8192 | 0 | 0 | [5 empty] | 160.8192 | 79.34 | 63.6006 | 260.9475
|  |  | 160.896 | 0 | 0 | [5 empty] | 160.896 | 79.31 | 63.6262 | 258.944
|  |  | 160.9728 | 0 | 0 | [5 empty] | 160.9728 | 79.425 | 62.9146 | 257.6803
|  |  | 161.0496 | 0 | 0 | [5 empty] | 161.0496 | 79.3 | 62.5152 | 257.1255
|  |  | 161.1264 | 0 | 0 | [5 empty] | 161.1264 | 78.87 | 62.2848 | 257.033
|  |  | 161.2032 | 0 | 0 | [5 empty] | 161.2032 | 78.355 | 61.2915 | 257.4029
|  |  | 161.28 | 0 | 0 | [5 empty] | 161.28 | 77.84 | 59.479 | 262.1496
|  |  | 161.3568 | 0 | 0 | [5 empty] | 161.3568 | 77.355 | 57.9328 | 277.6844
|  |  | 161.4336 | 0 | 0 | [5 empty] | 161.4336 | 76.72 | 58.0659 | 296.0241
|  |  | 161.5104 | 0 | 0 | [5 empty] | 161.5104 | 76.38 | 58.5574 | 309.3088
|  |  | 161.5872 | 0 | 0 | [5 empty] | 161.5872 | 76.7 | 58.0506 | 317.1995
|  |  | 161.664 | 0 | 0 | [5 empty] | 161.664 | 77.115 | 58.5318 | 321.8538
|  |  | 161.7408 | 0 | 0 | [5 empty] | 161.7408 | 77.58 | 60.16 | 325.5217
|  |  | 161.8176 | 0 | 0 | [5 empty] | 161.8176 | 78.34 | 60.6106 | 330.2068
|  |  | 161.8944 | 0 | 0 | [5 empty] | 161.8944 | 78.88 | 61.225 | 336.7104
|  |  | 161.9712 | 0 | 0 | [5 empty] | 161.9712 | 79.115 | 61.7779 | 344.5703
|  |  | 162.048 | 0 | 0 | [5 empty] | 162.048 | 78.94 | 62.9453 | 352.2761
|  |  | 162.1248 | 0 | 0 | [5 empty] | 162.1248 | 78.79 | 62.9043 | 358.0092
|  |  | 162.2016 | 0 | 0 | [5 empty] | 162.2016 | 78.705 | 62.0186 | 360.4133
|  |  | 162.2784 | 0 | 0 | [5 empty] | 162.2784 | 78.6 | 61.2557 | 358.9955
|  |  | 162.3552 | 0 | 0 | [5 empty] | 162.3552 | 78.25 | 62.0083 | 354.2487
|  |  | 162.432 | 0 | 0 | [5 empty] | 162.432 | 77.95 | 62.7917 | 347.591
|  |  | 162.5088 | 0 | 0 | [5 empty] | 162.5088 | 77.875 | 62.2336 | 340.6866
|  |  | 162.5856 | 0 | 0 | [5 empty] | 162.5856 | 77.765 | 61.527 | 334.7994
|  |  | 162.6624 | 0 | 0 | [5 empty] | 162.6624 | 77.54 | 60.6362 | 330.4225
|  |  | 162.7392 | 0 | 0 | [5 empty] | 162.7392 | 77.205 | 59.1206 | 327.0937
|  |  | 162.816 | 0 | 0 | [5 empty] | 162.816 | 76.465 | 58.4602 | 323.5799
|  |  | 162.8928 | 0 | 0 | [5 empty] | 162.8928 | 75.79 | 57.1699 | 318.6482
|  |  | 162.9696 | 0 | 0 | [5 empty] | 162.9696 | 75.375 | 55.7363 | 312.1137
|  |  | 163.0464 | 0 | 0 | [5 empty] | 163.0464 | 75.17 | 55.6083 | 304.7778
|  |  | 163.1232 | 0 | 0 | [5 empty] | 163.1232 | 75.33 | 56.7194 | 296.733
|  |  | 163.2 | 0 | 0 | [5 empty] | 163.2 | 75.87 | 57.2877 | 289.3355
|  |  | 163.2768 | 0 | 0 | [5 empty] | 163.2768 | 76.38 | 58.112 | 283.4483
|  |  | 163.3536 | 0 | 0 | [5 empty] | 163.3536 | 76.76 | 58.9466 | 279.1947
|  |  | 163.4304 | 0 | 0 | [5 empty] | 163.4304 | 77 | 59.7197 | 276.4515
|  |  | 163.5072 | 0 | 0 | [5 empty] | 163.5072 | 76.915 | 60.4774 | 274.8487
|  |  | 163.584 | 0 | 0 | [5 empty] | 163.584 | 76.765 | 60.2982 | 274.0781
|  |  | 163.6608 | 0 | 0 | [5 empty] | 163.6608 | 76.57 | 59.7504 | 273.8624
|  |  | 163.7376 | 0 | 0 | [5 empty] | 163.7376 | 76.205 | 59.8426 | 273.8624
|  |  | 163.8144 | 0 | 0 | [5 empty] | 163.8144 | 75.93 | 59.6941 | 274.0165
|  |  | 163.8912 | 0 | 0 | [5 empty] | 163.8912 | 75.81 | 59.0746 | 274.2939
|  |  | 163.968 | 0 | 0 | [5 empty] | 163.968 | 75.66 | 59.2333 | 274.787
|  |  | 164.0448 | 0 | 0 | [5 empty] | 164.0448 | 75.63 | 59.3254 | 276.02
|  |  | 164.1216 | 0 | 0 | [5 empty] | 164.1216 | 75.9 | 58.6291 | 279.6879
|  |  | 164.1984 | 0 | 0 | [5 empty] | 164.1984 | 76.165 | 60.0986 | 286.8388
|  |  | 164.2752 | 0 | 0 | [5 empty] | 164.2752 | 76.46 | 62.0237 | 297.8118
|  |  | 164.352 | 0 | 0 | [5 empty] | 164.352 | 76.945 | 61.7779 | 312.4836
|  |  | 164.4288 | 0 | 0 | [5 empty] | 164.4288 | 77.43 | 60.8563 | 329.1896
|  |  | 164.5056 | 0 | 0 | [5 empty] | 164.5056 | 77.48 | 62.3155 | 345.1868
|  |  | 164.5824 | 0 | 0 | [5 empty] | 164.5824 | 77.56 | 62.8582 | 357.7009
|  |  | 164.6592 | 0 | 0 | [5 empty] | 164.6592 | 77.82 | 61.9213 | 365.006
|  |  | 164.736 | 0 | 0 | [5 empty] | 164.736 | 77.99 | 61.7472 | 367.1019
|  |  | 164.8128 | 0 | 0 | [5 empty] | 164.8128 | 78.19 | 62.4282 | 365.2834
|  |  | 164.8896 | 0 | 0 | [5 empty] | 164.8896 | 78.33 | 64.0922 | 361.6463
|  |  | 164.9664 | 0 | 0 | [5 empty] | 164.9664 | 78.625 | 65.1418 | 358.1633
|  |  | 165.0432 | 0 | 0 | [5 empty] | 165.0432 | 79 | 65.1981 | 356.2831
|  |  | 165.12 | 0 | 0 | [5 empty] | 165.12 | 79.18 | 64.9574 | 356.1289
|  |  | 165.1968 | 0 | 0 | [5 empty] | 165.1968 | 78.895 | 65.193 | 356.9303
|  |  | 165.2736 | 0 | 0 | [5 empty] | 165.2736 | 78.355 | 65.193 | 357.0536
|  |  | 165.3504 | 0 | 0 | [5 empty] | 165.3504 | 77.915 | 63.9283 | 355.0193
|  |  | 165.4272 | 0 | 0 | [5 empty] | 165.4272 | 77.44 | 62.336 | 350.0568
|  |  | 165.504 | 0 | 0 | [5 empty] | 165.504 | 76.845 | 61.7984 | 342.3203
|  |  | 165.5808 | 0 | 0 | [5 empty] | 165.5808 | 76.47 | 61.8854 | 332.611
|  |  | 165.6576 | 0 | 0 | [5 empty] | 165.6576 | 76.78 | 61.7779 | 321.7921
|  |  | 165.7344 | 0 | 0 | [5 empty] | 165.7344 | 77.37 | 62.6637 | 314.3638
|  |  | 165.8112 | 0 | 0 | [5 empty] | 165.8112 | 77.91 | 63.872 | 317.1687
|  |  | 165.888 | 0 | 0 | [5 empty] | 165.888 | 78.34 | 63.8157 | 316.0899
|  |  | 165.9648 | 39.3 | 0.90624 | [5 empty] | 165.9648 | 78.32 | 64.3635 | 309.9869
|  |  | 166.0416 | 41.92 | 1.0394 | [5 empty] | 166.0416 | 78.185 | 64.3379 | 304.2847
|  |  | 166.1184 | 44.73 | 1.111 | [5 empty] | 166.1184 | 78.175 | 63.447 | 299.5379
|  |  | 166.1952 | 46.665 | 1.2851 | [5 empty] | 166.1952 | 78.3 | 62.4384 | 295.9624
|  |  | 166.272 | 48.78 | 1.4131 | [5 empty] | 166.272 | 78.505 | 62.4282 | 293.6507
|  |  | 166.3488 | 50.58 | 1.5565 | [5 empty] | 166.3488 | 78.52 | 64.3533 | 292.2637
|  |  | 166.4256 | 52.365 | 1.6896 | [5 empty] | 166.4256 | 78.92 | 64.8294 | 291.4931
|  |  | 166.5024 | 54.135 | 1.7869 | [5 empty] | 166.5024 | 79.415 | 65.0086 | 291.1232
|  |  | 166.5792 | 55.775 | 1.8842 | [5 empty] | 166.5792 | 79.515 | 66.8723 | 291.1541
|  |  | 166.656 | 57.03 | 2.0173 | [5 empty] | 166.656 | 79.71 | 66.8672 | 291.4006
|  |  | 166.7328 | 57.875 | 2.1555 | [5 empty] | 166.7328 | 79.88 | 65.3619 | 291.5239
|  |  | 166.8096 | 58.63 | 2.2835 | [5 empty] | 166.8096 | 79.84 | 64.3328 | 292.4486
|  |  | 166.8864 | 59.16 | 2.4013 | [5 empty] | 166.8864 | 79.465 | 64.3072 | 300.9558
|  |  | 166.9632 | 59.71 | 2.4627 | [5 empty] | 166.9632 | 78.945 | 65.3005 | 319.1105
|  |  | 167.04 | 60.025 | 2.5293 | [5 empty] | 167.04 | 78.9 | 65.2749 | 336.4022
|  |  | 167.1168 | 60.725 | 2.4781 | [5 empty] | 167.1168 | 79.4 | 64.2048 | 348.3307
|  |  | 167.1936 | 61.415 | 2.4269 | [5 empty] | 167.1936 | 80.02 | 64.3789 | 355.5433
|  |  | 167.2704 | 61.665 | 2.4576 | [5 empty] | 167.2704 | 80.68 | 66.1402 | 359.8585
|  |  | 167.3472 | 61.585 | 2.5651 | [5 empty] | 167.3472 | 81.38 | 68.9613 | 363.0641
|  |  | 167.424 | 61.895 | 2.5958 | [5 empty] | 167.424 | 82.1 | 71.1578 | 366.7321
|  |  | 167.5008 | 62.11 | 2.6214 | [5 empty] | 167.5008 | 82.645 | 72.3149 | 371.5404
|  |  | 167.5776 | 61.335 | 2.9286 | [5 empty] | 167.5776 | 82.88 | 72.7194 | 376.9961
|  |  | 167.6544 | 62.575 | 2.6624 | [5 empty] | 167.6544 | 82.905 | 72.2432 | 381.8353
|  |  | 167.7312 | 62.275 | 2.8262 | [5 empty] | 167.7312 | 82.75 | 71.7926 | 384.6402
|  |  | 167.808 | 62.265 | 2.8723 | [5 empty] | 167.808 | 82.59 | 70.8813 | 384.332
|  |  | 167.8848 | 62.02 | 2.9645 | [5 empty] | 167.8848 | 82.39 | 69.8829 | 380.5099
|  |  | 167.9616 | 61.41 | 3.1846 | [5 empty] | 167.9616 | 82.05 | 69.6986 | 373.6981
|  |  | 168.0384 | 63.18 | 2.729 | [5 empty] | 168.0384 | 81.75 | 70.1798 | 365.1601
|  |  | 168.1152 | 61.805 | 3.0925 | [5 empty] | 168.1152 | 81.455 | 70.5485 | 356.4063
|  |  | 168.192 | 62.72 | 2.7955 | [5 empty] | 168.192 | 81.305 | 69.4221 | 348.5465
|  |  | 168.2688 | 61.495 | 3.113 | [5 empty] | 168.2688 | 81.015 | 67.8349 | 342.012
|  |  | 168.3456 | 62.075 | 2.903 | [5 empty] | 168.3456 | 80.32 | 66.7034 | 336.4022
|  |  | 168.4224 | 61.625 | 3.0157 | [5 empty] | 168.4224 | 79.695 | 64.2662 | 330.7924
|  |  | 168.4992 | 62.06 | 2.8314 | [5 empty] | 168.4992 | 79.025 | 62.3667 | 324.2271
|  |  | 168.576 | 61.035 | 3.072 | [5 empty] | 168.576 | 78.335 | 62.1363 | 316.6447
|  |  | 168.6528 | 61.745 | 2.7648 | [5 empty] | 168.6528 | 77.985 | 63.2474 | 308.6924
|  |  | 168.7296 | 61.14 | 2.7955 | [5 empty] | 168.7296 | 78.245 | 64.3789 | 300.0002
|  |  | 168.8064 | 60.9 | 2.775 | [5 empty] | 168.8064 | 78.9 | 64.7219 | 291.8322
|  |  | 168.8832 | 60.92 | 2.647 | [5 empty] | 168.8832 | 79.495 | 65.0854 | 284.8662
|  |  | 168.96 | 60.47 | 2.5498 | [5 empty] | 168.96 | 79.7 | 66.3808 | 279.1639
|  |  | 169.0368 | 59.295 | 2.4934 | [5 empty] | 169.0368 | 79.835 | 66.1862 | 274.5405
|  |  | 169.1136 | 57.985 | 2.3757 | [5 empty] | 169.1136 | 79.83 | 65.0189 | 270.9958
|  |  | 169.1904 | 56.575 | 2.2374 | [5 empty] | 169.1904 | 79.425 | 64.5939 | 268.4067
|  |  | 169.2672 | 54.955 | 2.1402 | [5 empty] | 169.2672 | 78.95 | 63.7133 | 266.6498
|  |  | 169.344 | 53.68 | 1.9866 | [5 empty] | 169.344 | 78.515 | 62.6534 | 265.5401
|  |  | 169.4208 | 52.21 | 1.8842 | [5 empty] | 169.4208 | 78.005 | 62.4845 | 264.8004
|  |  | 169.4976 | 50.98 | 1.7254 | [5 empty] | 169.4976 | 77.71 | 61.7114 | 264.2148
|  |  | 169.5744 | 49.74 | 1.5872 | [5 empty] | 169.5744 | 77.385 | 61.6141 | 263.8449
|  |  | 169.6512 | 48.57 | 1.4234 | [5 empty] | 169.6512 | 76.94 | 61.8957 | 264.4305
|  |  | 169.728 | 47.36 | 1.2646 | [5 empty] | 169.728 | 76.835 | 61.8598 | 267.9135
|  |  | 169.8048 | 45.685 | 1.193 | [5 empty] | 169.8048 | 77.265 | 61.2813 | 275.2494
|  |  | 169.8816 | 43.82 | 1.1213 | [5 empty] | 169.8816 | 77.54 | 61.4861 | 286.8697
|  |  | 169.9584 | 42.425 | 0.96256 | [5 empty] | 169.9584 | 77.54 | 61.6397 | 302.1887
|  |  | 170.0352 | 41.375 | 0.74752 | [5 empty] | 170.0352 | 77.475 | 61.1174 | 319.0797
|  |  | 170.112 | 39.58 | 0.64 | [5 empty] | 170.112 | 77.735 | 60.2368 | 334.4604
|  |  | 170.1888 | 0 | 0 | [5 empty] | 170.1888 | 77.815 | 60.2368 | 345.5875
|  |  | 170.2656 | 0 | 0 | [5 empty] | 170.2656 | 77.765 | 60.713 | 351.3206
|  |  | 170.3424 | 0 | 0 | [5 empty] | 170.3424 | 77.795 | 61.056 | 352.4302
|  |  | 170.4192 | 0 | 0 | [5 empty] | 170.4192 | 77.94 | 61.5936 | 350.7349
|  |  | 170.496 | 0 | 0 | [5 empty] | 170.496 | 78.165 | 62.2797 | 348.4848
|  |  | 170.5728 | 0 | 0 | [5 empty] | 170.5728 | 78.38 | 62.7405 | 346.9745
|  |  | 170.6496 | 0 | 0 | [5 empty] | 170.6496 | 78.34 | 63.6058 | 347.0053
|  |  | 170.7264 | 0 | 0 | [5 empty] | 170.7264 | 78.255 | 63.2166 | 348.1458
|  |  | 170.8032 | 0 | 0 | [5 empty] | 170.8032 | 78.1 | 61.737 | 348.8855
|  |  | 170.88 | 0 | 0 | [5 empty] | 170.88 | 77.645 | 60.5133 | 347.8992
|  |  | 170.9568 | 0 | 0 | [5 empty] | 170.9568 | 76.86 | 60.3443 | 344.2004
|  |  | 171.0336 | 0 | 0 | [5 empty] | 171.0336 | 76.345 | 59.0541 | 337.4194
|  |  | 171.1104 | 0 | 0 | [5 empty] | 171.1104 | 75.88 | 57.3901 | 328.4191
|  |  | 171.1872 | 0 | 0 | [5 empty] | 171.1872 | 75.57 | 56.873 | 318.3091
|  |  | 171.264 | 0 | 0 | [5 empty] | 171.264 | 75.485 | 58.0096 | 308.4458
|  |  | 171.3408 | 0 | 0 | [5 empty] | 171.3408 | 75.75 | 59.0746 | 305.5792
|  |  | 171.4176 | 0 | 0 | [5 empty] | 171.4176 | 76.2 | 59.1718 | 306.8121
|  |  | 171.4944 | 0 | 0 | [5 empty] | 171.4944 | 76.22 | 59.4534 | 301.5106
|  |  | 171.5712 | 0 | 0 | [5 empty] | 171.5712 | 75.905 | 59.4278 | 295.2227
|  |  | 171.648 | 0 | 0 | [5 empty] | 171.648 | 75.675 | 58.1171 | 289.952
|  |  | 171.7248 | 0 | 0 | [5 empty] | 171.7248 | 75.425 | 57.0573 | 285.9142
|  |  | 171.8016 | 0 | 0 | [5 empty] | 171.8016 | 75.045 | 57.2314 | 283.1401
|  |  | 171.8784 | 0 | 0 | [5 empty] | 171.8784 | 74.97 | 57.2467 | 281.5989
|  |  | 171.9552 | 0 | 0 | [5 empty] | 171.9552 | 75.245 | 56.96 | 280.7975
|  |  | 172.032 | 0 | 0 | [5 empty] | 172.032 | 75.69 | 56.9344 | 280.5509
|  |  | 172.1088 | 0 | 0 | [5 empty] | 172.1088 | 75.95 | 57.472 | 280.7975
|  |  | 172.1856 | 0 | 0 | [5 empty] | 172.1856 | 76.06 | 57.4515 | 281.3524
|  |  | 172.2624 | 0 | 0 | [5 empty] | 172.2624 | 75.845 | 57.7126 | 282.0613
|  |  | 172.3392 | 0 | 0 | [5 empty] | 172.3392 | 75.55 | 57.344 | 283.7565
|  |  | 172.416 | 0 | 0 | [5 empty] | 172.416 | 75.205 | 56.5043 | 293.3425
|  |  | 172.4928 | 0 | 0 | [5 empty] | 172.4928 | 74.705 | 56.5965 | 313.2233
|  |  | 172.5696 | 0 | 0 | [5 empty] | 172.5696 | 74.555 | 56.4992 | 331.9945
|  |  | 172.6464 | 0 | 0 | [5 empty] | 172.6464 | 75 | 55.7056 | 344.5087
|  |  | 172.7232 | 0 | 0 | [5 empty] | 172.7232 | 75.39 | 56.7552 | 351.2897
|  |  | 172.8 | 0 | 0 | [5 empty] | 172.8 | 76 | 58.3066 | 354.3412
|  |  | 172.8768 | 0 | 0 | [5 empty] | 172.8768 | 76.73 | 60.2214 | 356.468
|  |  | 172.9536 | 0 | 0 | [5 empty] | 172.9536 | 77.265 | 62.7251 | 359.8893
|  |  | 173.0304 | 0 | 0 | [5 empty] | 173.0304 | 78.095 | 62.7251 | 365.4067
|  |  | 173.1072 | 0 | 0 | [5 empty] | 173.1072 | 78.395 | 63.7338 | 372.4651
|  |  | 173.184 | 0 | 0 | [5 empty] | 173.184 | 78.245 | 65.3107 | 379.2154
|  |  | 173.2608 | 0 | 0 | [5 empty] | 173.2608 | 78.2 | 65.2493 | 383.6847
|  |  | 173.3376 | 0 | 0 | [5 empty] | 173.3376 | 78.155 | 64.5018 | 384.3628
|  |  | 173.4144 | 0 | 0 | [5 empty] | 173.4144 | 78.03 | 63.744 | 380.9106
|  |  | 173.4912 | 0 | 0 | [5 empty] | 173.4912 | 77.855 | 63.7082 | 374.0371
|  |  | 173.568 | 0 | 0 | [5 empty] | 173.568 | 77.795 | 63.575 | 365.4991
|  |  | 173.6448 | 0 | 0 | [5 empty] | 173.6448 | 77.875 | 62.8275 | 357.331
|  |  | 173.7216 | 0 | 0 | [5 empty] | 173.7216 | 77.92 | 61.9776 | 350.7349
|  |  | 173.7984 | 0 | 0 | [5 empty] | 173.7984 | 77.75 | 61.3632 | 346.0498
|  |  | 173.8752 | 0 | 0 | [5 empty] | 173.8752 | 77.36 | 60.8768 | 342.7209
|  |  | 173.952 | 0 | 0 | [5 empty] | 173.952 | 76.955 | 59.7555 | 339.1763
|  |  | 174.0288 | 0 | 0 | [5 empty] | 174.0288 | 76.525 | 58.5933 | 334.1521
|  |  | 174.1056 | 0 | 0 | [5 empty] | 174.1056 | 76.315 | 58.0352 | 327.9567
|  |  | 174.1824 | 0 | 0 | [5 empty] | 174.1824 | 76.595 | 58.7827 | 320.8058
|  |  | 174.2592 | 0 | 0 | [5 empty] | 174.2592 | 77.31 | 60.0678 | 313.0384
|  |  | 174.336 | 0 | 0 | [5 empty] | 174.336 | 78.18 | 61.1277 | 306.1032
|  |  | 174.4128 | 0 | 0 | [5 empty] | 174.4128 | 78.855 | 63.0733 | 300.5243
|  |  | 174.4896 | 0 | 0 | [5 empty] | 174.4896 | 79.45 | 64.4864 | 296.1782
|  |  | 174.5664 | 0 | 0 | [5 empty] | 174.5664 | 79.895 | 65.3824 | 292.9418
|  |  | 174.6432 | 0 | 0 | [5 empty] | 174.6432 | 80.24 | 65.1622 | 290.5684
|  |  | 174.72 | 0 | 0 | [5 empty] | 174.72 | 80.5 | 64.1997 | 288.9348
|  |  | 174.7968 | 0 | 0 | [5 empty] | 174.7968 | 80.36 | 64.5478 | 287.8252
|  |  | 174.8736 | 0 | 0 | [5 empty] | 174.8736 | 80.39 | 63.5085 | 287.1162
|  |  | 174.9504 | 0 | 0 | [5 empty] | 174.9504 | 80.275 | 63.6314 | 286.469
|  |  | 175.0272 | 0 | 0 | [5 empty] | 175.0272 | 80.11 | 64.7936 | 285.7909
|  |  | 175.104 | 0 | 0 | [5 empty] | 175.104 | 79.97 | 66.1658 | 285.0511
|  |  | 175.1808 | 0 | 0 | [5 empty] | 175.1808 | 80.03 | 66.263 | 285.2052
|  |  | 175.2576 | 0 | 0 | [5 empty] | 175.2576 | 80.445 | 65.6333 | 288.1642
|  |  | 175.3344 | 0 | 0 | [5 empty] | 175.3344 | 80.785 | 67.1181 | 294.9761
|  |  | 175.4112 | 0 | 0 | [5 empty] | 175.4112 | 81.255 | 68.4698 | 306.1032
|  |  | 175.488 | 0 | 0 | [5 empty] | 175.488 | 81.715 | 68.5414 | 321.114
|  |  | 175.5648 | 0 | 0 | [5 empty] | 175.5648 | 81.89 | 68.8947 | 337.7276
|  |  | 175.6416 | 0 | 0 | [5 empty] | 175.6416 | 82.21 | 69.4221 | 352.8925
|  |  | 175.7184 | 0 | 0 | [5 empty] | 175.7184 | 82.58 | 69.2838 | 363.773
|  |  | 175.7952 | 0 | 0 | [5 empty] | 175.7952 | 82.88 | 68.992 | 368.8897
|  |  | 175.872 | 0 | 0 | [5 empty] | 175.872 | 82.89 | 69.9648 | 368.7355
|  |  | 175.9488 | 0 | 0 | [5 empty] | 175.9488 | 82.985 | 70.9069 | 365.2834
|  |  | 176.0256 | 0 | 0 | [5 empty] | 176.0256 | 83.255 | 71.0451 | 360.8757
|  |  | 176.1024 | 0 | 0 | [5 empty] | 176.1024 | 83.465 | 71.1782 | 357.4235
|  |  | 176.1792 | 0 | 0 | [5 empty] | 176.1792 | 83.465 | 71.721 | 355.7282
|  |  | 176.256 | 0 | 0 | [5 empty] | 176.256 | 83.435 | 71.2294 | 355.2043
|  |  | 176.3328 | 0 | 0 | [5 empty] | 176.3328 | 83.12 | 70.6509 | 354.6494
|  |  | 176.4096 | 0 | 0 | [5 empty] | 176.4096 | 82.49 | 70.3181 | 352.5535
|  |  | 176.4864 | 0 | 0 | [5 empty] | 176.4864 | 81.89 | 69.0074 | 347.7451
|  |  | 176.5632 | 0 | 0 | [5 empty] | 176.5632 | 81.335 | 67.0618 | 339.9777
|  |  | 176.64 | 0 | 0 | [5 empty] | 176.64 | 80.775 | 65.3978 | 329.9294
|  |  | 176.7168 | 0 | 0 | [5 empty] | 176.7168 | 80.415 | 64.6195 | 318.5249
|  |  | 176.7936 | 0 | 0 | [5 empty] | 176.7936 | 80.62 | 63.8515 | 306.5039
|  |  | 176.8704 | 0 | 0 | [5 empty] | 176.8704 | 80.915 | 63.9846 | 297.442
|  |  | 176.9472 | 0 | 0 | [5 empty] | 176.9472 | 81.065 | 64.5376 | 297.7194
|  |  | 177.024 | 0 | 0 | [5 empty] | 177.024 | 80.825 | 66.0429 | 294.8836
|  |  | 177.1008 | 0 | 0 | [5 empty] | 177.1008 | 80.575 | 66.2067 | 287.0854
|  |  | 177.1776 | 0 | 0 | [5 empty] | 177.1776 | 80.4 | 64.8704 | 279.7187
|  |  | 177.2544 | 0 | 0 | [5 empty] | 177.2544 | 80.065 | 63.7645 | 273.1842
|  |  | 177.3312 | 0 | 0 | [5 empty] | 177.3312 | 79.77 | 63.1603 | 267.7286
|  |  | 177.408 | 0 | 0 | [5 empty] | 177.408 | 79.715 | 62.2336 | 263.475
|  |  | 177.4848 | 0 | 0 | [5 empty] | 177.4848 | 79.795 | 61.4349 | 260.3619
|  |  | 177.5616 | 0 | 0 | [5 empty] | 177.5616 | 79.945 | 60.9024 | 258.2968
|  |  | 177.6384 | 0 | 0 | [5 empty] | 177.6384 | 79.56 | 62.6944 | 257.0946
|  |  | 177.7152 | 0 | 0 | [5 empty] | 177.7152 | 79.3 | 62.9965 | 256.6323
|  |  | 177.792 | 0 | 0 | [5 empty] | 177.792 | 79.055 | 62.121 | 256.6323
|  |  | 177.8688 | 0 | 0 | [5 empty] | 177.8688 | 78.545 | 61.4298 | 256.9405
|  |  | 177.9456 | 0 | 0 | [5 empty] | 177.9456 | 78.09 | 59.6531 | 261.0092
|  |  | 178.0224 | 0 | 0 | [5 empty] | 178.0224 | 77.53 | 58.1734 | 275.7426
|  |  | 178.0992 | 0 | 0 | [5 empty] | 178.0992 | 76.955 | 58.0557 | 294.0822
|  |  | 178.176 | 0 | 0 | [5 empty] | 178.176 | 76.58 | 58.3936 | 307.6443
|  |  | 178.2528 | 0 | 0 | [5 empty] | 178.2528 | 76.555 | 58.9466 | 315.8433
|  |  | 178.3296 | 0 | 0 | [5 empty] | 178.3296 | 76.905 | 59.7555 | 320.6825
|  |  | 178.4064 | 0 | 0 | [5 empty] | 178.4064 | 77.385 | 61.1942 | 324.4121
|  |  | 178.4832 | 0 | 0 | [5 empty] | 178.4832 | 78.28 | 60.928 | 329.128
|  |  | 178.56 | 0 | 0 | [5 empty] | 178.56 | 79.035 | 60.6976 | 335.5392
|  |  | 178.6368 | 0 | 0 | [5 empty] | 178.6368 | 79.335 | 61.2813 | 343.399
|  |  | 178.7136 | 0 | 0 | [5 empty] | 178.7136 | 79.195 | 62.3002 | 351.1356
|  |  | 178.7904 | 0 | 0 | [5 empty] | 178.7904 | 78.945 | 62.7354 | 356.9303
|  |  | 178.8672 | 0 | 0 | [5 empty] | 178.8672 | 78.55 | 63.3293 | 359.4578
|  |  | 178.944 | 0 | 0 | [5 empty] | 178.944 | 78.315 | 62.9504 | 358.1941
|  |  | 179.0208 | 0 | 0 | [5 empty] | 179.0208 | 78.26 | 62.0595 | 353.6323
|  |  | 179.0976 | 0 | 0 | [5 empty] | 179.0976 | 78.195 | 61.7984 | 347.0978
|  |  | 179.1744 | 0 | 0 | [5 empty] | 179.1744 | 78.04 | 62.0442 | 340.2859
|  |  | 179.2512 | 0 | 0 | [5 empty] | 179.2512 | 78.02 | 61.2454 | 334.4295
|  |  | 179.328 | 0 | 0 | [5 empty] | 179.328 | 77.765 | 60.6925 | 330.0219
|  |  | 179.4048 | 0 | 0 | [5 empty] | 179.4048 | 77.16 | 60.457 | 326.5697
|  |  | 179.4816 | 0 | 0 | [5 empty] | 179.4816 | 76.605 | 58.8595 | 322.9017
|  |  | 179.5584 | 0 | 0 | [5 empty] | 179.5584 | 76.135 | 56.7808 | 317.7235
|  |  | 179.6352 | 0 | 0 | [5 empty] | 179.6352 | 75.71 | 55.0195 | 310.9116
|  |  | 179.712 | 0 | 0 | [5 empty] | 179.712 | 75.4 | 55.0963 | 303.4216
|  |  | 179.7888 | 0 | 0 | [5 empty] | 179.7888 | 75.735 | 55.4547 | 295.346
|  |  | 179.8656 | 0 | 0 | [5 empty] | 179.8656 | 76 | 57.3082 | 287.9485
|  |  | 179.9424 | 0 | 0 | [5 empty] | 179.9424 | 76.435 | 58.3117 | 282.0921
|  |  | 180.0192 | 0 | 0 | [5 empty] | 180.0192 | 76.855 | 58.839 | 277.8694
|  |  | 180.096 | 0 | 0 | [5 empty] | 180.096 | 77.225 | 58.9005 | 275.0028
|  |  | 180.1728 | 0 | 0 | [5 empty] | 180.1728 | 77.205 | 59.4586 | 273.1534
|  |  | 180.2496 | 0 | 0 | [5 empty] | 180.2496 | 77.065 | 59.2691 | 272.1671
|  |  | 180.3264 | 0 | 0 | [5 empty] | 180.3264 | 76.92 | 58.4243 | 271.7664
|  |  | 180.4032 | 0 | 0 | [5 empty] | 180.4032 | 76.605 | 58.0454 | 271.7972
|  |  | 180.48 | 0 | 0 | [5 empty] | 180.48 | 76.325 | 58.1069 | 272.0746
|  |  | 180.5568 | 0 | 0 | [5 empty] | 180.5568 | 75.975 | 58.7315 | 272.5678
|  |  | 180.6336 | 0 | 0 | [5 empty] | 180.6336 | 75.86 | 58.8698 | 273.1842
|  |  | 180.7104 | 0 | 0 | [5 empty] | 180.7104 | 76.12 | 57.7946 | 274.3247
|  |  | 180.7872 | 0 | 0 | [5 empty] | 180.7872 | 76.33 | 57.2262 | 277.6536
|  |  | 180.864 | 0 | 0 | [5 empty] | 180.864 | 76.635 | 57.897 | 284.4346
|  |  | 180.9408 | 0 | 0 | [5 empty] | 180.9408 | 77.01 | 59.4944 | 295.0069
|  |  | 181.0176 | 0 | 0 | [5 empty] | 181.0176 | 77.39 | 60.3392 | 309.4629
|  |  | 181.0944 | 0 | 0 | [5 empty] | 181.0944 | 77.75 | 60.2266 | 326.1998
|  |  | 181.1712 | 0 | 0 | [5 empty] | 181.1712 | 78.01 | 60.6566 | 342.4435
|  |  | 181.248 | 0 | 0 | [5 empty] | 181.248 | 78.08 | 61.1584 | 355.2967
|  |  | 181.3248 | 0 | 0 | [5 empty] | 181.3248 | 78.085 | 61.3325 | 362.8792
|  |  | 181.4016 | 0 | 0 | [5 empty] | 181.4016 | 78.215 | 61.312 | 365.0676
|  |  | 181.4784 | 0 | 0 | [5 empty] | 181.4784 | 78.4 | 62.1619 | 363.2491
|  |  | 181.5552 | 0 | 0 | [5 empty] | 181.5552 | 78.6 | 63.6006 | 359.5195
|  |  | 181.632 | 0 | 0 | [5 empty] | 181.632 | 78.79 | 65.0957 | 356.0673
|  |  | 181.7088 | 0 | 0 | [5 empty] | 181.7088 | 78.92 | 66.1914 | 354.1871
|  |  | 181.7856 | 0 | 0 | [5 empty] | 181.7856 | 79.105 | 65.9302 | 354.1563
|  |  | 181.8624 | 0 | 0 | [5 empty] | 181.8624 | 79.04 | 65.1059 | 355.0193
|  |  | 181.9392 | 0 | 0 | [5 empty] | 181.9392 | 78.73 | 64.2611 | 355.1118
|  |  | 182.016 | 0 | 0 | [5 empty] | 182.016 | 78.305 | 62.8582 | 352.9542
|  |  | 182.0928 | 0 | 0 | [5 empty] | 182.0928 | 77.845 | 61.3376 | 347.8376
|  |  | 182.1696 | 0 | 0 | [5 empty] | 182.1696 | 77.6 | 59.1821 | 340.0085
|  |  | 182.2464 | 0 | 0 | [5 empty] | 182.2464 | 77.165 | 59.392 | 330.3917
|  |  | 182.3232 | 0 | 0 | [5 empty] | 182.3232 | 77.235 | 60.1088 | 319.6962
|  |  | 182.4 | 0 | 0 | [5 empty] | 182.4 | 77.585 | 61.5066 | 312.1137
|  |  | 182.4768 | 0 | 0 | [5 empty] | 182.4768 | 78.125 | 62.5306 | 314.5795
|  |  | 182.5536 | 0 | 0 | [5 empty] | 182.5536 | 78.55 | 63.3702 | 313.5316
|  |  | 182.6304 | 39.47 | 0.8704 | [5 empty] | 182.6304 | 78.495 | 64.2662 | 307.3978
|  |  | 182.7072 | 42.73 | 0.90112 | [5 empty] | 182.7072 | 78.17 | 64.9779 | 301.6647
|  |  | 182.784 | 44.905 | 1.0342 | [5 empty] | 182.784 | 78.145 | 63.9078 | 296.918
|  |  | 182.8608 | 46.72 | 1.2237 | [5 empty] | 182.8608 | 78.215 | 63.3242 | 293.3733
|  |  | 182.9376 | 48.445 | 1.4131 | [5 empty] | 182.9376 | 78.425 | 63.401 | 291.0616
|  |  | 183.0144 | 50.16 | 1.5821 | [5 empty] | 183.0144 | 78.615 | 64.6758 | 289.6746
|  |  | 183.0912 | 52.28 | 1.6794 | [5 empty] | 183.0912 | 79.04 | 65.321 | 288.8423
|  |  | 183.168 | 54.285 | 1.7613 | [5 empty] | 183.168 | 79.555 | 65.4848 | 288.4724
|  |  | 183.2448 | 55.82 | 1.8688 | [5 empty] | 183.2448 | 79.82 | 66.089 | 288.5341
|  |  | 183.3216 | 56.945 | 2.0173 | [5 empty] | 183.3216 | 79.86 | 66.5805 | 288.8423
|  |  | 183.3984 | 58.095 | 2.1094 | [5 empty] | 183.3984 | 79.965 | 65.5053 | 289.0581
|  |  | 183.4752 | 59.025 | 2.1914 | [5 empty] | 183.4752 | 79.91 | 64.553 | 289.952
|  |  | 183.552 | 59.56 | 2.2989 | [5 empty] | 183.552 | 79.6 | 64.2406 | 298.1817
|  |  | 183.6288 | 60.23 | 2.3552 | [5 empty] | 183.6288 | 79.35 | 64.0358 | 316.2748
|  |  | 183.7056 | 60.805 | 2.391 | [5 empty] | 183.7056 | 79.24 | 64.471 | 333.7206
|  |  | 183.7824 | 61.015 | 2.4422 | [5 empty] | 183.7824 | 79.41 | 65.2902 | 345.7724
|  |  | 183.8592 | 61.62 | 2.3962 | [5 empty] | 183.8592 | 79.81 | 66.8365 | 353.0158
|  |  | 183.936 | 61.84 | 2.4269 | [5 empty] | 183.936 | 80.59 | 67.8298 | 357.1769
|  |  | 184.0128 | 61.845 | 2.519 | [5 empty] | 184.0128 | 81.435 | 69.4835 | 360.29
|  |  | 184.0896 | 61.7 | 2.6317 | [5 empty] | 184.0896 | 82.125 | 71.4035 | 363.9888
|  |  | 184.1664 | 61.42 | 2.7802 | [5 empty] | 184.1664 | 82.825 | 71.5162 | 368.9205
|  |  | 184.2432 | 61.595 | 2.8262 | [5 empty] | 184.2432 | 83.21 | 71.5571 | 374.5611
|  |  | 184.32 | 62.44 | 2.7238 | [5 empty] | 184.32 | 83.135 | 71.7824 | 379.5544
|  |  | 184.3968 | 61.805 | 2.9798 | [5 empty] | 184.3968 | 82.935 | 71.465 | 382.421
|  |  | 184.4736 | 62.715 | 2.775 | [5 empty] | 184.4736 | 82.645 | 71.1117 | 381.9894
|  |  | 184.5504 | 61.1 | 3.2768 | [5 empty] | 184.5504 | 82.395 | 70.441 | 377.9208
|  |  | 184.6272 | 63.305 | 2.6778 | [5 empty] | 184.6272 | 81.995 | 70.8147 | 370.8932
|  |  | 184.704 | 62 | 3.031 | [5 empty] | 184.704 | 81.87 | 70.2566 | 362.386
|  |  | 184.7808 | 62.555 | 2.8621 | [5 empty] | 184.7808 | 81.815 | 69.1456 | 353.8789
|  |  | 184.8576 | 62.33 | 2.9133 | [5 empty] | 184.8576 | 81.67 | 68.0806 | 346.4197
|  |  | 184.9344 | 62.105 | 2.9645 | [5 empty] | 184.9344 | 81.305 | 66.7955 | 340.3167
|  |  | 185.0112 | 61.515 | 3.1078 | [5 empty] | 185.0112 | 80.71 | 65.3824 | 334.9536
|  |  | 185.088 | 62.6 | 2.7648 | [5 empty] | 185.088 | 79.845 | 64.169 | 329.3438
|  |  | 185.1648 | 61.635 | 2.9747 | [5 empty] | 185.1648 | 79.175 | 62.4691 | 322.5627
|  |  | 185.2416 | 61.82 | 2.8672 | [5 empty] | 185.2416 | 78.53 | 62.0237 | 314.6412
|  |  | 185.3184 | 61.32 | 2.9082 | [5 empty] | 185.3184 | 78.305 | 62.505 | 306.4114
|  |  | 185.3952 | 61.27 | 2.8365 | [5 empty] | 185.3952 | 78.505 | 63.913 | 297.5652
|  |  | 185.472 | 61.175 | 2.7494 | [5 empty] | 185.472 | 79.145 | 63.9795 | 289.4588
|  |  | 185.5488 | 61.075 | 2.6214 | [5 empty] | 185.5488 | 79.48 | 65.5872 | 282.6469
|  |  | 185.6256 | 60.295 | 2.6061 | [5 empty] | 185.6256 | 79.8 | 66.2374 | 277.0988
|  |  | 185.7024 | 59.155 | 2.5805 | [5 empty] | 185.7024 | 79.825 | 66.7802 | 272.537
|  |  | 185.7792 | 57.695 | 2.519 | [5 empty] | 185.7792 | 79.74 | 66.3501 | 268.9615
|  |  | 185.856 | 56.635 | 2.304 | [5 empty] | 185.856 | 79.52 | 65.0957 | 266.2491
|  |  | 185.9328 | 55.48 | 2.1146 | [5 empty] | 185.9328 | 79.125 | 63.9949 | 264.338
|  |  | 186.0096 | 54.2 | 1.9456 | [5 empty] | 186.0096 | 78.55 | 63.529 | 263.136
|  |  | 186.0864 | 52.84 | 1.8074 | [5 empty] | 186.0864 | 78.005 | 63.3088 | 262.427
|  |  | 186.1632 | 51.2 | 1.7613 | [5 empty] | 186.1632 | 77.91 | 61.6499 | 262.0263
|  |  | 186.24 | 49.675 | 1.6742 | [5 empty] | 186.24 | 77.625 | 61.0406 | 261.8414
|  |  | 186.3168 | 48.11 | 1.5821 | [5 empty] | 186.3168 | 77.18 | 61.2352 | 262.4887
|  |  | 186.3936 | 46.77 | 1.4234 | [5 empty] | 186.3936 | 77.01 | 61.3683 | 265.8484
|  |  | 186.4704 | 45.345 | 1.2749 | [5 empty] | 186.4704 | 77.375 | 61.2352 | 272.9068
|  |  | 186.5472 | 44.26 | 1.0854 | [5 empty] | 186.5472 | 77.91 | 60.6976 | 284.2189
|  |  | 186.624 | 43.155 | 0.89088 | [5 empty] | 186.624 | 77.89 | 61.0509 | 299.3221
|  |  | 186.7008 | 41.205 | 0.82432 | [5 empty] | 186.7008 | 77.56 | 61.7062 | 316.1515
|  |  | 186.7776 | 40.51 | 0.57344 | [5 empty] | 186.7776 | 77.59 | 61.5782 | 331.563
|  |  | 186.8544 | 0 | 0 | [5 empty] | 186.8544 | 77.675 | 61.6192 | 342.8134
|  |  | 186.9312 | 0 | 0 | [5 empty] | 186.9312 | 77.815 | 61.0918 | 348.9472
|  |  | 187.008 | 0 | 0 | [5 empty] | 187.008 | 77.855 | 61.4144 | 350.3342
|  |  | 187.0848 | 0 | 0 | [5 empty] | 187.0848 | 78.06 | 61.6448 | 348.6389
|  |  | 187.1616 | 0 | 0 | [5 empty] | 187.1616 | 78.445 | 61.5475 | 346.1731
|  |  | 187.2384 | 0 | 0 | [5 empty] | 187.2384 | 78.71 | 61.9622 | 344.7244
|  |  | 187.3152 | 0 | 0 | [5 empty] | 187.3152 | 78.75 | 62.3002 | 344.7861
|  |  | 187.392 | 0 | 0 | [5 empty] | 187.392 | 78.465 | 63.0272 | 346.0498
|  |  | 187.4688 | 0 | 0 | [5 empty] | 187.4688 | 78.075 | 62.7302 | 347.0362
|  |  | 187.5456 | 0 | 0 | [5 empty] | 187.5456 | 77.58 | 61.8086 | 346.2964
|  |  | 187.6224 | 0 | 0 | [5 empty] | 187.6224 | 76.845 | 61.3427 | 342.6285
|  |  | 187.6992 | 0 | 0 | [5 empty] | 187.6992 | 76.26 | 60.032 | 335.8474
|  |  | 187.776 | 0 | 0 | [5 empty] | 187.776 | 75.765 | 58.7418 | 326.8471
|  |  | 187.8528 | 0 | 0 | [5 empty] | 187.8528 | 75.625 | 57.5744 | 316.768
|  |  | 187.9296 | 0 | 0 | [5 empty] | 187.9296 | 75.875 | 57.7024 | 306.9662
|  |  | 188.0064 | 0 | 0 | [5 empty] | 188.0064 | 76.075 | 59.1718 | 303.9148
|  |  | 188.0832 | 0 | 0 | [5 empty] | 188.0832 | 76.295 | 59.7862 | 305.0861
|  |  | 188.16 | 0 | 0 | [5 empty] | 188.16 | 76.23 | 60.3136 | 299.7228
|  |  | 188.2368 | 0 | 0 | [5 empty] | 188.2368 | 75.955 | 59.9859 | 293.3425
|  |  | 188.3136 | 0 | 0 | [5 empty] | 188.3136 | 75.68 | 58.9005 | 288.0101
|  |  | 188.3904 | 0 | 0 | [5 empty] | 188.3904 | 75.655 | 57.0214 | 283.9106
|  |  | 188.4672 | 0 | 0 | [5 empty] | 188.4672 | 75.42 | 56.873 | 281.1674
|  |  | 188.544 | 0 | 0 | [5 empty] | 188.544 | 75.185 | 57.9686 | 279.5646
|  |  | 188.6208 | 0 | 0 | [5 empty] | 188.6208 | 75.2 | 59.1718 | 278.7632
|  |  | 188.6976 | 0 | 0 | [5 empty] | 188.6976 | 75.715 | 58.5728 | 278.4242
|  |  | 188.7744 | 0 | 0 | [5 empty] | 188.7744 | 76.105 | 57.9379 | 278.5475
|  |  | 188.8512 | 0 | 0 | [5 empty] | 188.8512 | 75.975 | 58.6342 | 279.0714
|  |  | 188.928 | 0 | 0 | [5 empty] | 188.928 | 75.77 | 58.8595 | 279.9037
|  |  | 189.0048 | 0 | 0 | [5 empty] | 189.0048 | 75.46 | 58.6035 | 281.5989
|  |  | 189.0816 | 0 | 0 | [5 empty] | 189.0816 | 75.32 | 56.9651 | 290.815
|  |  | 189.1584 | 0 | 0 | [5 empty] | 189.1584 | 75.06 | 56.064 | 310.4492
|  |  | 189.2352 | 0 | 0 | [5 empty] | 189.2352 | 74.71 | 56.832 | 329.3438
|  |  | 189.312 | 0 | 0 | [5 empty] | 189.312 | 74.55 | 58.24 | 342.1045
|  |  | 189.3888 | 0 | 0 | [5 empty] | 189.3888 | 75.105 | 58.3936 | 349.1321
|  |  | 189.4656 | 0 | 0 | [5 empty] | 189.4656 | 75.98 | 58.8544 | 352.3377
|  |  | 189.5424 | 0 | 0 | [5 empty] | 189.5424 | 76.845 | 60.1958 | 354.4953
|  |  | 189.6192 | 0 | 0 | [5 empty] | 189.6192 | 77.67 | 61.1584 | 357.855
|  |  | 189.696 | 0 | 0 | [5 empty] | 189.696 | 78.24 | 62.5101 | 363.3723
|  |  | 189.7728 | 0 | 0 | [5 empty] | 189.7728 | 78.485 | 63.9232 | 370.4
|  |  | 189.8496 | 0 | 0 | [5 empty] | 189.8496 | 78.6 | 64.1024 | 377.1811
|  |  | 189.9264 | 0 | 0 | [5 empty] | 189.9264 | 78.435 | 64.4915 | 381.7429
|  |  | 190.0032 | 0 | 0 | [5 empty] | 190.0032 | 78.395 | 63.8771 | 382.6675
|  |  | 190.08 | 0 | 0 | [5 empty] | 190.08 | 78.135 | 63.9437 | 379.4311
|  |  | 190.1568 | 0 | 0 | [5 empty] | 190.1568 | 77.925 | 63.9283 | 372.6809
|  |  | 190.2336 | 0 | 0 | [5 empty] | 190.2336 | 77.815 | 64.0358 | 364.2046
|  |  | 190.3104 | 0 | 0 | [5 empty] | 190.3104 | 78.05 | 62.7507 | 355.9748
|  |  | 190.3872 | 0 | 0 | [5 empty] | 190.3872 | 78.165 | 61.9264 | 349.3787
|  |  | 190.464 | 0 | 0 | [5 empty] | 190.464 | 77.655 | 63.0733 | 344.7244
|  |  | 190.5408 | 0 | 0 | [5 empty] | 190.5408 | 77.34 | 62.0134 | 341.4572
|  |  | 190.6176 | 0 | 0 | [5 empty] | 190.6176 | 76.885 | 61.0253 | 338.0358
|  |  | 190.6944 | 0 | 0 | [5 empty] | 190.6944 | 76.305 | 60.5235 | 333.135
|  |  | 190.7712 | 0 | 0 | [5 empty] | 190.7712 | 76.285 | 58.752 | 326.9396
|  |  | 190.848 | 0 | 0 | [5 empty] | 190.848 | 76.825 | 57.9686 | 319.8195
|  |  | 190.9248 | 0 | 0 | [5 empty] | 190.9248 | 77.19 | 60.4826 | 311.9288
|  |  | 191.0016 | 0 | 0 | [5 empty] | 191.0016 | 77.595 | 63.6467 | 304.9011
|  |  | 191.0784 | 0 | 0 | [5 empty] | 191.0784 | 78.585 | 64.3021 | 299.168
|  |  | 191.1552 | 0 | 0 | [5 empty] | 191.1552 | 79.19 | 65.5616 | 294.7603
|  |  | 191.232 | 0 | 0 | [5 empty] | 191.232 | 79.63 | 66.3808 | 291.4931
|  |  | 191.3088 | 0 | 0 | [5 empty] | 191.3088 | 79.97 | 65.9814 | 289.0889
|  |  | 191.3856 | 0 | 0 | [5 empty] | 191.3856 | 80.105 | 65.4848 | 287.4553
|  |  | 191.4624 | 0 | 0 | [5 empty] | 191.4624 | 80.23 | 64.5632 | 286.3765
|  |  | 191.5392 | 0 | 0 | [5 empty] | 191.5392 | 80.23 | 64.0256 | 285.6367
|  |  | 191.616 | 0 | 0 | [5 empty] | 191.616 | 80.07 | 64.5683 | 285.0511
|  |  | 191.6928 | 0 | 0 | [5 empty] | 191.6928 | 79.975 | 65.449 | 284.4038
|  |  | 191.7696 | 0 | 0 | [5 empty] | 191.7696 | 80.03 | 65.879 | 283.7257
|  |  | 191.8464 | 0 | 0 | [5 empty] | 191.8464 | 80.08 | 65.9558 | 283.9106
|  |  | 191.9232 | 0 | 0 | [5 empty] | 191.9232 | 80.48 | 65.0906 | 286.7772
|  |  | 192 | 0 | 0 | [5 empty] | 192 | 81.005 | 65.2902 | 293.3733
|  |  | 192.0768 | 0 | 0 | [5 empty] | 192.0768 | 81.27 | 67.3587 | 304.2847
|  |  | 192.1536 | 0 | 0 | [5 empty] | 192.1536 | 81.415 | 68.9766 | 319.0489
|  |  | 192.2304 | 0 | 0 | [5 empty] | 192.2304 | 81.68 | 69.2787 | 335.6933
|  |  | 192.3072 | 0 | 0 | [5 empty] | 192.3072 | 82.125 | 69.3965 | 351.1048
|  |  | 192.384 | 0 | 0 | [5 empty] | 192.384 | 82.34 | 70.2618 | 362.386
|  |  | 192.4608 | 0 | 0 | [5 empty] | 192.4608 | 82.445 | 70.6458 | 367.8725
|  |  | 192.5376 | 0 | 0 | [5 empty] | 192.5376 | 82.59 | 70.8198 | 367.9958
|  |  | 192.6144 | 0 | 0 | [5 empty] | 192.6144 | 82.99 | 70.359 | 364.6361
|  |  | 192.6912 | 0 | 0 | [5 empty] | 192.6912 | 83.34 | 70.2362 | 360.2284
|  |  | 192.768 | 0 | 0 | [5 empty] | 192.768 | 83.535 | 70.7174 | 356.5605
|  |  | 192.8448 | 0 | 0 | [5 empty] | 192.8448 | 83.585 | 71.0042 | 354.6494
|  |  | 192.9216 | 0 | 0 | [5 empty] | 192.9216 | 83.395 | 71.2653 | 354.1563
|  |  | 192.9984 | 0 | 0 | [5 empty] | 192.9984 | 82.96 | 71.1117 | 353.9405
|  |  | 193.0752 | 0 | 0 | [5 empty] | 193.0752 | 82.56 | 69.7293 | 352.1219
|  |  | 193.152 | 0 | 0 | [5 empty] | 193.152 | 81.965 | 68.567 | 347.4985
|  |  | 193.2288 | 0 | 0 | [5 empty] | 193.2288 | 81.225 | 67.5942 | 339.916
|  |  | 193.3056 | 0 | 0 | [5 empty] | 193.3056 | 80.655 | 65.8483 | 329.8369
|  |  | 193.3824 | 0 | 0 | [5 empty] | 193.3824 | 80.32 | 64.8192 | 318.3708
|  |  | 193.4592 | 0 | 0 | [5 empty] | 193.4592 | 80.42 | 64.7424 | 306.2265
|  |  | 193.536 | 0 | 0 | [5 empty] | 193.536 | 80.93 | 64.599 | 297.0721
|  |  | 193.6128 | 0 | 0 | [5 empty] | 193.6128 | 81.045 | 65.7869 | 297.5961
|  |  | 193.6896 | 0 | 0 | [5 empty] | 193.6896 | 81.05 | 65.9558 | 295.161
|  |  | 193.7664 | 0 | 0 | [5 empty] | 193.7664 | 80.91 | 64.9523 | 287.3012
|  |  | 193.8432 | 0 | 0 | [5 empty] | 193.8432 | 80.555 | 63.8106 | 279.7495
|  |  | 193.92 | 0 | 0 | [5 empty] | 193.92 | 80.185 | 63.0426 | 273.0918
|  |  | 193.9968 | 0 | 0 | [5 empty] | 193.9968 | 79.74 | 63.3805 | 267.6053
|  |  | 194.0736 | 0 | 0 | [5 empty] | 194.0736 | 79.41 | 63.913 | 263.4134
|  |  | 194.1504 | 0 | 0 | [5 empty] | 194.1504 | 79.44 | 63.529 | 260.4235
|  |  | 194.2272 | 0 | 0 | [5 empty] | 194.2272 | 79.565 | 62.9402 | 258.3584
|  |  | 194.304 | 0 | 0 | [5 empty] | 194.304 | 79.49 | 63.1603 | 257.1255
|  |  | 194.3808 | 0 | 0 | [5 empty] | 194.3808 | 79.295 | 63.0067 | 256.6323
|  |  | 194.4576 | 0 | 0 | [5 empty] | 194.4576 | 78.89 | 62.9811 | 256.6323
|  |  | 194.5344 | 0 | 0 | [5 empty] | 194.5344 | 78.28 | 62.9094 | 256.8789
|  |  | 194.6112 | 0 | 0 | [5 empty] | 194.6112 | 77.71 | 61.9059 | 260.3619
|  |  | 194.688 | 0 | 0 | [5 empty] | 194.688 | 77.335 | 59.5098 | 274.3555
|  |  | 194.7648 | 0 | 0 | [5 empty] | 194.7648 | 76.915 | 58.1786 | 292.911
|  |  | 194.8416 | 0 | 0 | [5 empty] | 194.8416 | 76.7 | 57.8202 | 306.9354
|  |  | 194.9184 | 0 | 0 | [5 empty] | 194.9184 | 76.745 | 58.071 | 315.5351
|  |  | 194.9952 | 0 | 0 | [5 empty] | 194.9952 | 77.18 | 58.3117 | 320.5592
|  |  | 195.072 | 0 | 0 | [5 empty] | 195.072 | 77.73 | 59.2794 | 324.2888
|  |  | 195.1488 | 0 | 0 | [5 empty] | 195.1488 | 78.205 | 61.015 | 328.7273
|  |  | 195.2256 | 0 | 0 | [5 empty] | 195.2256 | 78.665 | 62.3462 | 334.9227
|  |  | 195.3024 | 0 | 0 | [5 empty] | 195.3024 | 79.035 | 62.5766 | 342.5668
|  |  | 195.3792 | 0 | 0 | [5 empty] | 195.3792 | 78.965 | 63.3702 | 350.3034
|  |  | 195.456 | 0 | 0 | [5 empty] | 195.456 | 78.75 | 63.5187 | 356.3755
|  |  | 195.5328 | 0 | 0 | [5 empty] | 195.5328 | 78.81 | 61.7523 | 359.2729
|  |  | 195.6096 | 0 | 0 | [5 empty] | 195.6096 | 78.54 | 61.5782 | 358.3482
|  |  | 195.6864 | 0 | 0 | [5 empty] | 195.6864 | 78.1 | 62.3053 | 354.033
|  |  | 195.7632 | 0 | 0 | [5 empty] | 195.7632 | 78.065 | 61.6653 | 347.591
|  |  | 195.84 | 0 | 0 | [5 empty] | 195.84 | 77.91 | 62.0749 | 340.6866
|  |  | 195.9168 | 0 | 0 | [5 empty] | 195.9168 | 77.92 | 61.2352 | 334.6761
|  |  | 195.9936 | 0 | 0 | [5 empty] | 195.9936 | 77.7 | 60.6515 | 330.176
|  |  | 196.0704 | 0 | 0 | [5 empty] | 196.0704 | 77.275 | 59.5866 | 326.8471
|  |  | 196.1472 | 0 | 0 | [5 empty] | 196.1472 | 76.765 | 57.9738 | 323.4566
|  |  | 196.224 | 0 | 0 | [5 empty] | 196.224 | 76.23 | 56.2483 | 318.679
|  |  | 196.3008 | 0 | 0 | [5 empty] | 196.3008 | 75.525 | 55.7926 | 312.1754
|  |  | 196.3776 | 0 | 0 | [5 empty] | 196.3776 | 75.14 | 56.1101 | 304.8087
|  |  | 196.4544 | 0 | 0 | [5 empty] | 196.4544 | 75.485 | 56.2432 | 296.7639
|  |  | 196.5312 | 0 | 0 | [5 empty] | 196.5312 | 76.12 | 56.3917 | 289.243
|  |  | 196.608 | 0 | 0 | [5 empty] | 196.608 | 76.635 | 57.0163 | 283.2017
|  |  | 196.6848 | 0 | 0 | [5 empty] | 196.6848 | 76.935 | 58.4346 | 278.8249
|  |  | 196.7616 | 0 | 0 | [5 empty] | 196.7616 | 77.055 | 59.7094 | 275.8967
|  |  | 196.8384 | 0 | 0 | [5 empty] | 196.8384 | 76.885 | 61.1174 | 274.0473
|  |  | 196.9152 | 0 | 0 | [5 empty] | 196.9152 | 76.64 | 61.8496 | 273.0301
|  |  | 196.992 | 0 | 0 | [5 empty] | 196.992 | 76.76 | 59.9142 | 272.6294
|  |  | 197.0688 | 0 | 0 | [5 empty] | 197.0688 | 76.615 | 58.8698 | 272.6294
|  |  | 197.1456 | 0 | 0 | [5 empty] | 197.1456 | 76.255 | 58.9722 | 272.8452
|  |  | 197.2224 | 0 | 0 | [5 empty] | 197.2224 | 76.27 | 57.9686 | 273.2151
|  |  | 197.2992 | 0 | 0 | [5 empty] | 197.2992 | 76.175 | 57.7946 | 273.8007
|  |  | 197.376 | 0 | 0 | [5 empty] | 197.376 | 75.98 | 58.537 | 274.8487
|  |  | 197.4528 | 0 | 0 | [5 empty] | 197.4528 | 75.91 | 59.3152 | 278.0235
|  |  | 197.5296 | 0 | 0 | [5 empty] | 197.5296 | 76.145 | 60.1549 | 284.5271
|  |  | 197.6064 | 0 | 0 | [5 empty] | 197.6064 | 76.695 | 60.7232 | 294.7912
|  |  | 197.6832 | 0 | 0 | [5 empty] | 197.6832 | 77.145 | 61.2198 | 308.9389
|  |  | 197.76 | 0 | 0 | [5 empty] | 197.76 | 77.51 | 61.1635 | 325.4601
|  |  | 197.8368 | 0 | 0 | [5 empty] | 197.8368 | 77.7 | 61.9213 | 341.7654
|  |  | 197.9136 | 0 | 0 | [5 empty] | 197.9136 | 77.845 | 62.167 | 354.896
|  |  | 197.9904 | 0 | 0 | [5 empty] | 197.9904 | 78.045 | 61.6755 | 362.91
|  |  | 198.0672 | 0 | 0 | [5 empty] | 198.0672 | 78.065 | 62.1261 | 365.5608
|  |  | 198.144 | 0 | 0 | [5 empty] | 198.144 | 78.095 | 63.3088 | 364.0504
|  |  | 198.2208 | 0 | 0 | [5 empty] | 198.2208 | 78.465 | 63.6365 | 360.475
|  |  | 198.2976 | 0 | 0 | [5 empty] | 198.2976 | 78.87 | 63.8464 | 356.992
|  |  | 198.3744 | 0 | 0 | [5 empty] | 198.3744 | 79.22 | 64.2253 | 354.9885
|  |  | 198.4512 | 0 | 0 | [5 empty] | 198.4512 | 79.355 | 64.4301 | 354.7727
|  |  | 198.528 | 0 | 0 | [5 empty] | 198.528 | 79.1 | 64.896 | 355.5741
|  |  | 198.6048 | 0 | 0 | [5 empty] | 198.6048 | 78.715 | 64.5581 | 355.8207
|  |  | 198.6816 | 0 | 0 | [5 empty] | 198.6816 | 78.345 | 62.935 | 354.0022
|  |  | 198.7584 | 0 | 0 | [5 empty] | 198.7584 | 77.835 | 61.5014 | 349.317
|  |  | 198.8352 | 0 | 0 | [5 empty] | 198.8352 | 77.355 | 60.5798 | 341.9195
|  |  | 198.912 | 0 | 0 | [5 empty] | 198.912 | 76.985 | 60.3187 | 332.5802
|  |  | 198.9888 | 0 | 0 | [5 empty] | 198.9888 | 77.375 | 59.52 | 322.0079
|  |  | 199.0656 | 0 | 0 | [5 empty] | 199.0656 | 77.525 | 62.0442 | 313.4391
|  |  | 199.1424 | 0 | 0 | [5 empty] | 199.1424 | 77.95 | 63.447 | 315.011
|  |  | 199.2192 | 0 | 0 | [5 empty] | 199.2192 | 78.47 | 63.5085 | 315.1035
|  |  | 199.296 | 38.905 | 0.89088 | [5 empty] | 199.296 | 78.625 | 63.3498 | 309.1239
|  |  | 199.3728 | 41.65 | 1.024 | [5 empty] | 199.3728 | 78.53 | 63.1142 | 303.2983
|  |  | 199.4496 | 43.94 | 1.1827 | [5 empty] | 199.4496 | 78.19 | 63.6928 | 298.3666
|  |  | 199.5264 | 46.625 | 1.239 | [5 empty] | 199.5264 | 78.11 | 63.6416 | 294.6371
|  |  | 199.6032 | 49.415 | 1.239 | [5 empty] | 199.6032 | 78.445 | 62.7302 | 292.1712
|  |  | 199.68 | 51.055 | 1.4131 | [5 empty] | 199.68 | 78.72 | 63.3088 | 290.6609
|  |  | 199.7568 | 52.41 | 1.6282 | [5 empty] | 199.7568 | 79.04 | 64.6554 | 289.767
|  |  | 199.8336 | 54.06 | 1.751 | [5 empty] | 199.8336 | 79.41 | 65.7101 | 289.3355
|  |  | 199.9104 | 55.605 | 1.8586 | [5 empty] | 199.9104 | 79.68 | 66.5293 | 289.3355
|  |  | 199.9872 | 56.63 | 2.0378 | [5 empty] | 199.9872 | 79.91 | 66.6573 | 289.6129
|  |  | 200.064 | 57.365 | 2.2272 | [5 empty] | 200.064 | 80.07 | 65.6845 | 289.767
|  |  | 200.1408 | 58.375 | 2.304 | [5 empty] | 200.1408 | 79.875 | 65.0138 | 290.3218
|  |  | 200.2176 | 59.405 | 2.3142 | [5 empty] | 200.2176 | 79.62 | 64.0307 | 297.1954
|  |  | 200.2944 | 60.165 | 2.3398 | [5 empty] | 200.2944 | 79.365 | 63.6928 | 314.6104
|  |  | 200.3712 | 60.23 | 2.4627 | [5 empty] | 200.3712 | 79.28 | 63.5904 | 332.5802
|  |  | 200.448 | 60.56 | 2.5139 | [5 empty] | 200.448 | 79.495 | 63.7594 | 345.2484
|  |  | 200.5248 | 61.135 | 2.4781 | [5 empty] | 200.5248 | 79.73 | 65.6589 | 352.9542
|  |  | 200.6016 | 61.39 | 2.5037 | [5 empty] | 200.6016 | 80.35 | 67.3587 | 357.4543
|  |  | 200.6784 | 61.585 | 2.5498 | [5 empty] | 200.6784 | 81.13 | 69.4528 | 360.6599
|  |  | 200.7552 | 61.685 | 2.6163 | [5 empty] | 200.7552 | 81.915 | 71.0298 | 364.297
|  |  | 200.832 | 61.92 | 2.6624 | [5 empty] | 200.832 | 82.58 | 71.8285 | 369.1054
|  |  | 200.9088 | 62.515 | 2.647 | [5 empty] | 200.9088 | 82.805 | 72.6272 | 374.6536
|  |  | 200.9856 | 60.46 | 3.3178 | [5 empty] | 200.9856 | 82.945 | 72.0333 | 379.801
|  |  | 201.0624 | 62.655 | 2.7341 | [5 empty] | 201.0624 | 82.895 | 70.9325 | 383.0066
|  |  | 201.1392 | 61.795 | 3.0413 | [5 empty] | 201.1392 | 82.645 | 70.3539 | 383.1299
|  |  | 201.216 | 63.38 | 2.6624 | [5 empty] | 201.216 | 82.335 | 70.1286 | 379.7085
|  |  | 201.2928 | 61.28 | 3.2205 | [5 empty] | 201.2928 | 82.05 | 69.8112 | 373.1432
|  |  | 201.3696 | 62.57 | 2.8518 | [5 empty] | 201.3696 | 81.59 | 70.912 | 364.7594
|  |  | 201.4464 | 61.365 | 3.2256 | [5 empty] | 201.4464 | 81.43 | 70.697 | 356.1289
|  |  | 201.5232 | 63.055 | 2.7597 | [5 empty] | 201.5232 | 81.41 | 69.0381 | 348.3615
|  |  | 201.6 | 61.505 | 3.1744 | [5 empty] | 201.6 | 81.13 | 67.6454 | 341.8579
|  |  | 201.6768 | 62.93 | 2.7443 | [5 empty] | 201.6768 | 80.705 | 65.7715 | 336.3098
|  |  | 201.7536 | 61.665 | 3.0208 | [5 empty] | 201.7536 | 79.88 | 64.4915 | 330.6692
|  |  | 201.8304 | 62.37 | 2.7802 | [5 empty] | 201.8304 | 78.955 | 63.6365 | 324.0114
|  |  | 201.9072 | 61.25 | 3.0669 | [5 empty] | 201.9072 | 78.535 | 61.8598 | 316.244
|  |  | 201.984 | 61.985 | 2.7597 | [5 empty] | 201.984 | 78.215 | 62.0595 | 308.1375
|  |  | 202.0608 | 61.37 | 2.816 | [5 empty] | 202.0608 | 78.39 | 63.1398 | 299.4146
|  |  | 202.1376 | 60.99 | 2.8058 | [5 empty] | 202.1376 | 79.01 | 63.5648 | 291.1232
|  |  | 202.2144 | 60.415 | 2.7955 | [5 empty] | 202.2144 | 79.485 | 64.3738 | 284.0956
|  |  | 202.2912 | 60.04 | 2.6931 | [5 empty] | 202.2912 | 79.775 | 65.2646 | 278.2701
|  |  | 202.368 | 59.755 | 2.4781 | [5 empty] | 202.368 | 79.875 | 65.92 | 273.5233
|  |  | 202.4448 | 58.375 | 2.4064 | [5 empty] | 202.4448 | 79.8 | 65.6589 | 269.7629
|  |  | 202.5216 | 56.76 | 2.3296 | [5 empty] | 202.5216 | 79.42 | 65.3722 | 266.8964
|  |  | 202.5984 | 55.31 | 2.1965 | [5 empty] | 202.5984 | 78.895 | 64.6758 | 264.9545
|  |  | 202.6752 | 54.21 | 2.0019 | [5 empty] | 202.6752 | 78.43 | 63.4522 | 263.6908
|  |  | 202.752 | 53 | 1.8534 | [5 empty] | 202.752 | 78.2 | 61.5731 | 262.9818
|  |  | 202.8288 | 51.74 | 1.7101 | [5 empty] | 202.8288 | 78.005 | 60.4006 | 262.5195
|  |  | 202.9056 | 50.165 | 1.6179 | [5 empty] | 202.9056 | 77.455 | 61.1482 | 262.3037
|  |  | 202.9824 | 48.915 | 1.4643 | [5 empty] | 202.9824 | 77.005 | 61.4861 | 262.7969
|  |  | 203.0592 | 47.77 | 1.2902 | [5 empty] | 203.0592 | 77.095 | 60.0422 | 265.8792
|  |  | 203.136 | 45.955 | 1.2237 | [5 empty] | 203.136 | 77.65 | 58.6701 | 272.6294
|  |  | 203.2128 | 45.285 | 0.96768 | [5 empty] | 203.2128 | 78.07 | 58.8851 | 283.6332
|  |  | 203.2896 | 43.895 | 0.82944 | [5 empty] | 203.2896 | 78 | 59.9654 | 298.6132
|  |  | 203.3664 | 42.285 | 0.7168 | [5 empty] | 203.3664 | 77.785 | 60.0371 | 315.4426
|  |  | 203.4432 | 41.355 | 0.5376 | [5 empty] | 203.4432 | 77.84 | 59.945 | 331.1623
|  |  | 203.52 | 0 | 0 | [5 empty] | 203.52 | 78.09 | 59.4637 | 342.9984
|  |  | 203.5968 | 0 | 0 | [5 empty] | 203.5968 | 78.1 | 59.6685 | 349.4712
|  |  | 203.6736 | 0 | 0 | [5 empty] | 203.6736 | 77.905 | 61.1174 | 350.9507
|  |  | 203.7504 | 0 | 0 | [5 empty] | 203.7504 | 78.18 | 61.1482 | 349.4712
|  |  | 203.8272 | 0 | 0 | [5 empty] | 203.8272 | 78.245 | 62.679 | 347.0053
|  |  | 203.904 | 0 | 0 | [5 empty] | 203.904 | 78.535 | 62.9146 | 345.2792
|  |  | 203.9808 | 0 | 0 | [5 empty] | 203.9808 | 78.715 | 62.8173 | 345.1251
|  |  | 204.0576 | 0 | 0 | [5 empty] | 204.0576 | 78.61 | 62.6125 | 346.2964
|  |  | 204.1344 | 0 | 0 | [5 empty] | 204.1344 | 78.215 | 62.5766 | 347.3752
|  |  | 204.2112 | 0 | 0 | [5 empty] | 204.2112 | 77.8 | 61.3581 | 346.8204
|  |  | 204.288 | 0 | 0 | [5 empty] | 204.288 | 77.175 | 60.3699 | 343.5532
|  |  | 204.3648 | 0 | 0 | [5 empty] | 204.3648 | 76.445 | 59.9194 | 337.0803
|  |  | 204.4416 | 0 | 0 | [5 empty] | 204.4416 | 75.725 | 59.6275 | 328.2033
|  |  | 204.5184 | 0 | 0 | [5 empty] | 204.5184 | 75.425 | 58.6803 | 318.2166
|  |  | 204.5952 | 0 | 0 | [5 empty] | 204.5952 | 75.705 | 58.0762 | 308.2608
|  |  | 204.672 | 0 | 0 | [5 empty] | 204.672 | 76.175 | 58.199 | 304.1922
|  |  | 204.7488 | 0 | 0 | [5 empty] | 204.7488 | 76.34 | 59.3254 | 305.7642
|  |  | 204.8256 | 0 | 0 | [5 empty] | 204.8256 | 76.165 | 60.5338 | 300.925
|  |  | 204.9024 | 0 | 0 | [5 empty] | 204.9024 | 76.015 | 59.8374 | 294.3288
|  |  | 204.9792 | 0 | 0 | [5 empty] | 204.9792 | 75.855 | 58.2554 | 288.7807
|  |  | 205.056 | 0 | 0 | [5 empty] | 205.056 | 75.715 | 56.663 | 284.5579
|  |  | 205.1328 | 0 | 0 | [5 empty] | 205.1328 | 75.6 | 55.977 | 281.6606
|  |  | 205.2096 | 0 | 0 | [5 empty] | 205.2096 | 75.475 | 56.448 | 279.9961
|  |  | 205.2864 | 0 | 0 | [5 empty] | 205.2864 | 75.7 | 56.5299 | 279.1331
|  |  | 205.3632 | 0 | 0 | [5 empty] | 205.3632 | 75.71 | 58.2502 | 278.7632
|  |  | 205.44 | 0 | 0 | [5 empty] | 205.44 | 75.605 | 59.9962 | 278.794
|  |  | 205.5168 | 0 | 0 | [5 empty] | 205.5168 | 75.805 | 59.4074 | 279.2256
|  |  | 205.5936 | 0 | 0 | [5 empty] | 205.5936 | 75.92 | 58.4243 | 279.9653
|  |  | 205.6704 | 0 | 0 | [5 empty] | 205.6704 | 75.71 | 57.8765 | 281.414
|  |  | 205.7472 | 0 | 0 | [5 empty] | 205.7472 | 75.27 | 57.5078 | 289.5821
|  |  | 205.824 | 0 | 0 | [5 empty] | 205.824 | 74.81 | 57.4003 | 308.6924
|  |  | 205.9008 | 0 | 0 | [5 empty] | 205.9008 | 74.72 | 56.9754 | 328.0492
|  |  | 205.9776 | 0 | 0 | [5 empty] | 205.9776 | 75.025 | 56.4122 | 341.3031
|  |  | 206.0544 | 0 | 0 | [5 empty] | 206.0544 | 75.625 | 56.1715 | 348.7314
|  |  | 206.1312 | 0 | 0 | [5 empty] | 206.1312 | 76.185 | 57.5386 | 352.1528
|  |  | 206.208 | 0 | 0 | [5 empty] | 206.208 | 76.81 | 59.479 | 354.2796
|  |  | 206.2848 | 0 | 0 | [5 empty] | 206.2848 | 77.325 | 61.8906 | 357.4235
|  |  | 206.3616 | 0 | 0 | [5 empty] | 206.3616 | 78.105 | 62.6022 | 362.6942
|  |  | 206.4384 | 0 | 0 | [5 empty] | 206.4384 | 78.53 | 63.2576 | 369.5986
|  |  | 206.5152 | 0 | 0 | [5 empty] | 206.5152 | 78.745 | 63.3907 | 376.503
|  |  | 206.592 | 0 | 0 | [5 empty] | 206.592 | 78.56 | 64.0358 | 381.373
|  |  | 206.6688 | 0 | 0 | [5 empty] | 206.6688 | 78.285 | 64.3789 | 382.5751
|  |  | 206.7456 | 0 | 0 | [5 empty] | 206.7456 | 77.995 | 64.5376 | 379.5544
|  |  | 206.8224 | 0 | 0 | [5 empty] | 206.8224 | 77.995 | 63.7594 | 373.0816
|  |  | 206.8992 | 0 | 0 | [5 empty] | 206.8992 | 78.105 | 62.935 | 364.6669
|  |  | 206.976 | 0 | 0 | [5 empty] | 206.976 | 78.145 | 62.6176 | 356.2831
|  |  | 207.0528 | 0 | 0 | [5 empty] | 207.0528 | 77.96 | 63.1296 | 349.5328
|  |  | 207.1296 | 0 | 0 | [5 empty] | 207.1296 | 77.715 | 62.7661 | 344.7552
|  |  | 207.2064 | 0 | 0 | [5 empty] | 207.2064 | 77.39 | 61.8342 | 341.4264
|  |  | 207.2832 | 0 | 0 | [5 empty] | 207.2832 | 76.76 | 61.6141 | 338.1591
|  |  | 207.36 | 0 | 0 | [5 empty] | 207.36 | 76.195 | 61.2096 | 333.5049
|  |  | 207.4368 | 0 | 0 | [5 empty] | 207.4368 | 76.17 | 59.5917 | 327.4019
|  |  | 207.5136 | 0 | 0 | [5 empty] | 207.5136 | 76.565 | 59.2794 | 320.3743
|  |  | 207.5904 | 0 | 0 | [5 empty] | 207.5904 | 77.1 | 61.0816 | 312.5144
|  |  | 207.6672 | 0 | 0 | [5 empty] | 207.6672 | 77.905 | 62.2643 | 305.3943
|  |  | 207.744 | 0 | 0 | [5 empty] | 207.744 | 78.775 | 63.0221 | 299.5687
|  |  | 207.8208 | 0 | 0 | [5 empty] | 207.8208 | 79.665 | 63.0682 | 295.0686
|  |  | 207.8976 | 0 | 0 | [5 empty] | 207.8976 | 80.15 | 63.6314 | 291.6164
|  |  | 207.9744 | 0 | 0 | [5 empty] | 207.9744 | 80.18 | 64.5786 | 289.1198
|  |  | 208.0512 | 0 | 0 | [5 empty] | 208.0512 | 80.215 | 64.384 | 287.3936
|  |  | 208.128 | 0 | 0 | [5 empty] | 208.128 | 80.115 | 64.4301 | 286.3148
|  |  | 208.2048 | 0 | 0 | [5 empty] | 208.2048 | 80.2 | 64.0051 | 285.5443
|  |  | 208.2816 | 0 | 0 | [5 empty] | 208.2816 | 79.98 | 65.2493 | 284.8662
|  |  | 208.3584 | 0 | 0 | [5 empty] | 208.3584 | 80.01 | 65.5309 | 284.188
|  |  | 208.4352 | 0 | 0 | [5 empty] | 208.4352 | 80.105 | 65.6282 | 283.4175
|  |  | 208.512 | 0 | 0 | [5 empty] | 208.512 | 80.145 | 65.7408 | 283.325
|  |  | 208.5888 | 0 | 0 | [5 empty] | 208.5888 | 80.555 | 64.64 | 285.8217
|  |  | 208.6656 | 0 | 0 | [5 empty] | 208.6656 | 81.07 | 64.9933 | 291.9863
|  |  | 208.7424 | 0 | 0 | [5 empty] | 208.7424 | 81.6 | 66.1811 | 302.4661
|  |  | 208.8192 | 0 | 0 | [5 empty] | 208.8192 | 81.815 | 67.4867 | 316.9837
|  |  | 208.896 | 0 | 0 | [5 empty] | 208.896 | 81.95 | 68.137 | 333.5357
|  |  | 208.9728 | 0 | 0 | [5 empty] | 208.9728 | 82.245 | 68.6848 | 349.1013
|  |  | 209.0496 | 0 | 0 | [5 empty] | 209.0496 | 82.5 | 69.3453 | 360.7216
|  |  | 209.1264 | 0 | 0 | [5 empty] | 209.1264 | 82.515 | 70.0621 | 366.7012
|  |  | 209.2032 | 0 | 0 | [5 empty] | 209.2032 | 82.54 | 70.9325 | 367.1944
|  |  | 209.28 | 0 | 0 | [5 empty] | 209.28 | 82.715 | 71.4957 | 364.0504
|  |  | 209.3568 | 0 | 0 | [5 empty] | 209.3568 | 83.085 | 71.2858 | 359.4886
|  |  | 209.4336 | 0 | 0 | [5 empty] | 209.4336 | 83.405 | 71.0861 | 355.6049
|  |  | 209.5104 | 0 | 0 | [5 empty] | 209.5104 | 83.585 | 70.8301 | 353.5398
|  |  | 209.5872 | 0 | 0 | [5 empty] | 209.5872 | 83.535 | 70.528 | 352.9233
|  |  | 209.664 | 0 | 0 | [5 empty] | 209.664 | 83.26 | 69.9699 | 352.5535
|  |  | 209.7408 | 0 | 0 | [5 empty] | 209.7408 | 82.67 | 69.4323 | 350.9815
|  |  | 209.8176 | 0 | 0 | [5 empty] | 209.8176 | 81.965 | 68.3571 | 346.7279
|  |  | 209.8944 | 0 | 0 | [5 empty] | 209.8944 | 81.4 | 66.8109 | 339.4845
|  |  | 209.9712 | 0 | 0 | [5 empty] | 209.9712 | 80.73 | 65.5309 | 329.7136
|  |  | 210.048 | 0 | 0 | [5 empty] | 210.048 | 80.26 | 64.9062 | 318.4324
|  |  | 210.1248 | 0 | 0 | [5 empty] | 210.1248 | 80.475 | 64.3123 | 306.3498
|  |  | 210.2016 | 0 | 0 | [5 empty] | 210.2016 | 80.785 | 64.7526 | 296.4248
|  |  | 210.2784 | 0 | 0 | [5 empty] | 210.2784 | 81.02 | 65.3978 | 296.1782
|  |  | 210.3552 | 0 | 0 | [5 empty] | 210.3552 | 81.1 | 65.6845 | 294.5754
|  |  | 210.432 | 0 | 0 | [5 empty] | 210.432 | 80.795 | 66.1555 | 286.8697
|  |  | 210.5088 | 0 | 0 | [5 empty] | 210.5088 | 80.455 | 65.3056 | 279.2256
|  |  | 210.5856 | 0 | 0 | [5 empty] | 210.5856 | 80.105 | 64.2253 | 272.352
|  |  | 210.6624 | 0 | 0 | [5 empty] | 210.6624 | 79.7 | 63.9642 | 266.6498
|  |  | 210.7392 | 0 | 0 | [5 empty] | 210.7392 | 79.455 | 63.6979 | 262.2729
|  |  | 210.816 | 0 | 0 | [5 empty] | 210.816 | 79.43 | 63.3907 | 259.129
|  |  | 210.8928 | 0 | 0 | [5 empty] | 210.8928 | 79.675 | 62.3616 | 256.9713
|  |  | 210.9696 | 0 | 0 | [5 empty] | 210.9696 | 79.49 | 63.0272 | 255.646
|  |  | 211.0464 | 0 | 0 | [5 empty] | 211.0464 | 79.28 | 63.2269 | 255.0295
|  |  | 211.1232 | 0 | 0 | [5 empty] | 211.1232 | 79.105 | 61.993 | 254.9062
|  |  | 211.2 | 0 | 0 | [5 empty] | 211.2 | 78.69 | 60.8307 | 255.0603
|  |  | 211.2768 | 0 | 0 | [5 empty] | 211.2768 | 77.79 | 61.0355 | 258.3276
|  |  | 211.3536 | 0 | 0 | [5 empty] | 211.3536 | 77.11 | 60.2214 | 272.013
|  |  | 211.4304 | 0 | 0 | [5 empty] | 211.4304 | 76.95 | 58.199 | 290.6609
|  |  | 211.5072 | 0 | 0 | [5 empty] | 211.5072 | 76.82 | 57.4106 | 304.9319
|  |  | 211.584 | 0 | 0 | [5 empty] | 211.584 | 76.965 | 56.9907 | 313.624
|  |  | 211.6608 | 0 | 0 | [5 empty] | 211.6608 | 77.3 | 57.4003 | 318.679
|  |  | 211.7376 | 0 | 0 | [5 empty] | 211.7376 | 77.66 | 59.479 | 322.3469
|  |  | 211.8144 | 0 | 0 | [5 empty] | 211.8144 | 78.215 | 61.1072 | 326.8163
|  |  | 211.8912 | 0 | 0 | [5 empty] | 211.8912 | 78.685 | 62.3411 | 333.0733
|  |  | 211.968 | 0 | 0 | [5 empty] | 211.968 | 79.005 | 62.848 | 340.8099
|  |  | 212.0448 | 0 | 0 | [5 empty] | 212.0448 | 79.33 | 61.9776 | 348.7006
|  |  | 212.1216 | 0 | 0 | [5 empty] | 212.1216 | 79.035 | 62.8531 | 354.8344
|  |  | 212.1984 | 0 | 0 | [5 empty] | 212.1984 | 78.765 | 62.807 | 357.7626
|  |  | 212.2752 | 0 | 0 | [5 empty] | 212.2752 | 78.63 | 62.2438 | 356.8379
|  |  | 212.352 | 0 | 0 | [5 empty] | 212.352 | 78.68 | 60.8461 | 352.5226
|  |  | 212.4288 | 0 | 0 | [5 empty] | 212.4288 | 78.57 | 60.375 | 346.2039
|  |  | 212.5056 | 0 | 0 | [5 empty] | 212.5056 | 78.18 | 61.609 | 339.3921
|  |  | 212.5824 | 0 | 0 | [5 empty] | 212.5824 | 77.89 | 62.2797 | 333.474
|  |  | 212.6592 | 0 | 0 | [5 empty] | 212.6592 | 77.77 | 61.3171 | 329.0047
|  |  | 212.736 | 0 | 0 | [5 empty] | 212.736 | 77.46 | 59.8989 | 325.5217
|  |  | 212.8128 | 0 | 0 | [5 empty] | 212.8128 | 76.825 | 58.9312 | 321.9462
|  |  | 212.8896 | 0 | 0 | [5 empty] | 212.8896 | 76.325 | 56.919 | 317.0146
|  |  | 212.9664 | 0 | 0 | [5 empty] | 212.9664 | 75.745 | 55.6442 | 310.4184
|  |  | 213.0432 | 0 | 0 | [5 empty] | 213.0432 | 75.4 | 55.5264 | 303.0517
|  |  | 213.12 | 0 | 0 | [5 empty] | 213.12 | 75.575 | 56.4122 | 295.0686
|  |  | 213.1968 | 0 | 0 | [5 empty] | 213.1968 | 75.995 | 57.5181 | 287.5478
|  |  | 213.2736 | 0 | 0 | [5 empty] | 213.2736 | 76.59 | 58.1786 | 281.3832
|  |  | 213.3504 | 0 | 0 | [5 empty] | 213.3504 | 76.97 | 59.1206 | 276.8214
|  |  | 213.4272 | 0 | 0 | [5 empty] | 213.4272 | 77.11 | 59.863 | 273.6158
|  |  | 213.504 | 0 | 0 | [5 empty] | 213.504 | 77.08 | 60.4314 | 271.5506
|  |  | 213.5808 | 0 | 0 | [5 empty] | 213.5808 | 76.895 | 60.8819 | 270.4102
|  |  | 213.6576 | 0 | 0 | [5 empty] | 213.6576 | 76.83 | 59.9245 | 269.9787
|  |  | 213.7344 | 0 | 0 | [5 empty] | 213.7344 | 76.595 | 59.2026 | 270.0711
|  |  | 213.8112 | 0 | 0 | [5 empty] | 213.8112 | 76.475 | 58.6445 | 270.3485
|  |  | 213.888 | 0 | 0 | [5 empty] | 213.888 | 76.43 | 58.1376 | 270.7801
|  |  | 213.9648 | 0 | 0 | [5 empty] | 213.9648 | 76.39 | 57.943 | 271.304
|  |  | 214.0416 | 0 | 0 | [5 empty] | 214.0416 | 76.305 | 58.0966 | 272.2904
|  |  | 214.1184 | 0 | 0 | [5 empty] | 214.1184 | 76.13 | 58.9619 | 275.311
|  |  | 214.1952 | 0 | 0 | [5 empty] | 214.1952 | 76.415 | 59.7197 | 281.7222
|  |  | 214.272 | 0 | 0 | [5 empty] | 214.272 | 76.89 | 60.7283 | 291.9246
|  |  | 214.3488 | 0 | 0 | [5 empty] | 214.3488 | 77.37 | 61.056 | 306.0724
|  |  | 214.4256 | 0 | 0 | [5 empty] | 214.4256 | 77.675 | 61.3632 | 322.6552
|  |  | 214.5024 | 0 | 0 | [5 empty] | 214.5024 | 77.925 | 61.8957 | 338.9605
|  |  | 214.5792 | 0 | 0 | [5 empty] | 214.5792 | 78.045 | 62.2848 | 352.1836
|  |  | 214.656 | 0 | 0 | [5 empty] | 214.656 | 77.84 | 63.4573 | 360.29
|  |  | 214.7328 | 0 | 0 | [5 empty] | 214.7328 | 77.9 | 63.7133 | 363.0949
|  |  | 214.8096 | 0 | 0 | [5 empty] | 214.8096 | 78.37 | 62.7507 | 361.8004
|  |  | 214.8864 | 0 | 0 | [5 empty] | 214.8864 | 78.55 | 64.0973 | 358.3482
|  |  | 214.9632 | 0 | 0 | [5 empty] | 214.9632 | 78.78 | 65.495 | 354.8652
|  |  | 215.04 | 0 | 0 | [5 empty] | 215.04 | 79.195 | 65.6486 | 352.7692
|  |  | 215.1168 | 0 | 0 | [5 empty] | 215.1168 | 79.26 | 66.1709 | 352.4302
|  |  | 215.1936 | 0 | 0 | [5 empty] | 215.1936 | 79.03 | 66.2733 | 353.0775
|  |  | 215.2704 | 0 | 0 | [5 empty] | 215.2704 | 78.825 | 64.7578 | 353.2932
|  |  | 215.3472 | 0 | 0 | [5 empty] | 215.3472 | 78.405 | 63.447 | 351.5671
|  |  | 215.424 | 0 | 0 | [5 empty] | 215.424 | 77.905 | 62.0954 | 347.0362
|  |  | 215.5008 | 0 | 0 | [5 empty] | 215.5008 | 77.34 | 61.1738 | 339.7928
|  |  | 215.5776 | 0 | 0 | [5 empty] | 215.5776 | 77.15 | 60.0474 | 330.5458
|  |  | 215.6544 | 0 | 0 | [5 empty] | 215.6544 | 77.305 | 60.2931 | 320.0044
|  |  | 215.7312 | 0 | 0 | [5 empty] | 215.7312 | 77.695 | 61.4605 | 311.1273
|  |  | 215.808 | 0 | 0 | [5 empty] | 215.808 | 78.155 | 62.7712 | 312.2986
|  |  | 215.8848 | 0 | 0 | [5 empty] | 215.8848 | 78.645 | 62.9197 | 312.7918
|  |  | 215.9616 | 38.755 | 0.85504 | [5 empty] | 215.9616 | 78.81 | 62.9043 | 306.9354
|  |  | 216.0384 | 41.71 | 0.92672 | [5 empty] | 216.0384 | 78.715 | 62.6944 | 301.0482
|  |  | 216.1152 | 43.96 | 1.065 | [5 empty] | 216.1152 | 78.445 | 62.8531 | 296.1165
|  |  | 216.192 | 46.435 | 1.152 | [5 empty] | 216.192 | 78.45 | 62.4077 | 292.3561
|  |  | 216.2688 | 48.43 | 1.3158 | [5 empty] | 216.2688 | 78.655 | 62.3258 | 289.8595
|  |  | 216.3456 | 50.495 | 1.4694 | [5 empty] | 216.3456 | 78.885 | 62.9197 | 288.3492
|  |  | 216.4224 | 51.85 | 1.6691 | [5 empty] | 216.4224 | 79.23 | 63.9283 | 287.5169
|  |  | 216.4992 | 53.975 | 1.705 | [5 empty] | 216.4992 | 79.73 | 64.4301 | 287.0854
|  |  | 216.576 | 55.415 | 1.8483 | [5 empty] | 216.576 | 79.955 | 65.3773 | 287.1471
|  |  | 216.6528 | 56.63 | 1.9968 | [5 empty] | 216.6528 | 79.895 | 66.473 | 287.4861
|  |  | 216.7296 | 58.35 | 2.0122 | [5 empty] | 216.7296 | 79.98 | 66.0378 | 287.7635
|  |  | 216.8064 | 59.43 | 2.0838 | [5 empty] | 216.8064 | 79.795 | 65.5667 | 288.3492
|  |  | 216.8832 | 59.74 | 2.2426 | [5 empty] | 216.8832 | 79.35 | 65.6435 | 294.8836
|  |  | 216.96 | 59.945 | 2.3706 | [5 empty] | 216.96 | 79.25 | 64.6912 | 312.0829
|  |  | 217.0368 | 60.515 | 2.4064 | [5 empty] | 217.0368 | 79.345 | 64.0307 | 330.1451
|  |  | 217.1136 | 61.305 | 2.3706 | [5 empty] | 217.1136 | 79.51 | 64.6093 | 343.06
|  |  | 217.1904 | 61.265 | 2.4525 | [5 empty] | 217.1904 | 79.845 | 65.8381 | 350.9815
|  |  | 217.2672 | 61.365 | 2.5242 | [5 empty] | 217.2672 | 80.51 | 67.3946 | 355.6049
|  |  | 217.344 | 61.87 | 2.5139 | [5 empty] | 217.344 | 81.415 | 68.5926 | 358.7797
|  |  | 217.4208 | 62.01 | 2.5651 | [5 empty] | 217.4208 | 82.265 | 69.5501 | 362.3552
|  |  | 217.4976 | 61.27 | 2.8621 | [5 empty] | 217.4976 | 82.81 | 70.5997 | 367.0711
|  |  | 217.5744 | 62.46 | 2.647 | [5 empty] | 217.5744 | 82.93 | 72.2842 | 372.5884
|  |  | 217.6512 | 61.515 | 3.0054 | [5 empty] | 217.6512 | 83.165 | 71.7722 | 377.6742
|  |  | 217.728 | 62.94 | 2.688 | [5 empty] | 217.728 | 83.09 | 70.9786 | 380.8182
|  |  | 217.8048 | 61.57 | 3.1027 | [5 empty] | 217.8048 | 82.955 | 69.8931 | 380.9415
|  |  | 217.8816 | 63.41 | 2.6675 | [5 empty] | 217.8816 | 82.6 | 69.7754 | 377.5201
|  |  | 217.9584 | 61.645 | 3.1437 | [5 empty] | 217.9584 | 82.21 | 69.9187 | 371.0473
|  |  | 218.0352 | 62.595 | 2.8518 | [5 empty] | 218.0352 | 82.08 | 69.2275 | 362.7867
|  |  | 218.112 | 61.72 | 3.1283 | [5 empty] | 218.112 | 82.005 | 68.5517 | 354.2179
|  |  | 218.1888 | 62.785 | 2.816 | [5 empty] | 218.1888 | 81.74 | 67.7632 | 346.4814
|  |  | 218.2656 | 62.19 | 2.985 | [5 empty] | 218.2656 | 81.24 | 67.1181 | 339.9777
|  |  | 218.3424 | 61.515 | 3.1539 | [5 empty] | 218.3424 | 80.75 | 65.5104 | 334.3679
|  |  | 218.4192 | 62.47 | 2.8109 | [5 empty] | 218.4192 | 80.09 | 63.913 | 328.6965
|  |  | 218.496 | 61.58 | 3.0413 | [5 empty] | 218.496 | 79.375 | 62.5306 | 322.0695
|  |  | 218.5728 | 62.245 | 2.7904 | [5 empty] | 218.5728 | 78.715 | 61.7728 | 314.3329
|  |  | 218.6496 | 61.27 | 2.9286 | [5 empty] | 218.6496 | 78.465 | 61.7011 | 306.319
|  |  | 218.7264 | 60.915 | 2.8774 | [5 empty] | 218.7264 | 78.715 | 62.3155 | 297.7194
|  |  | 218.8032 | 60.95 | 2.775 | [5 empty] | 218.8032 | 78.99 | 64.1178 | 289.4588
|  |  | 218.88 | 60.435 | 2.7853 | [5 empty] | 218.88 | 79.485 | 64.7885 | 282.3695
|  |  | 218.9568 | 60.165 | 2.6982 | [5 empty] | 218.9568 | 79.915 | 65.2954 | 276.4207
|  |  | 219.0336 | 59.44 | 2.6163 | [5 empty] | 219.0336 | 80.155 | 65.577 | 271.5506
|  |  | 219.1104 | 58.26 | 2.5139 | [5 empty] | 219.1104 | 80.04 | 65.495 | 267.6978
|  |  | 219.1872 | 57.275 | 2.2886 | [5 empty] | 219.1872 | 79.715 | 64.6656 | 264.8004
|  |  | 219.264 | 55.605 | 2.1862 | [5 empty] | 219.264 | 79.295 | 63.3754 | 262.8586
|  |  | 219.3408 | 54.235 | 2.0224 | [5 empty] | 219.3408 | 78.785 | 62.5613 | 261.6564
|  |  | 219.4176 | 53.035 | 1.8483 | [5 empty] | 219.4176 | 78.17 | 62.7763 | 260.9783
|  |  | 219.4944 | 51.885 | 1.7152 | [5 empty] | 219.4944 | 77.815 | 62.0083 | 260.5468
|  |  | 219.5712 | 50.27 | 1.6589 | [5 empty] | 219.5712 | 77.755 | 60.4006 | 260.3002
|  |  | 219.648 | 48.715 | 1.5667 | [5 empty] | 219.648 | 77.415 | 60.6413 | 260.7009
|  |  | 219.7248 | 47.415 | 1.3875 | [5 empty] | 219.7248 | 77.255 | 60.759 | 263.6599
|  |  | 219.8016 | 45.925 | 1.2646 | [5 empty] | 219.8016 | 77.645 | 60.1242 | 270.2869
|  |  | 219.8784 | 44.585 | 1.1315 | [5 empty] | 219.8784 | 78.02 | 60.0883 | 281.1982
|  |  | 219.9552 | 43.335 | 0.95232 | [5 empty] | 219.9552 | 77.965 | 60.8461 | 296.0857
|  |  | 220.032 | 41.905 | 0.78848 | [5 empty] | 220.032 | 77.8 | 60.7642 | 313.0076
|  |  | 220.1088 | 41.43 | 0.54784 | [5 empty] | 220.1088 | 77.655 | 61.6038 | 328.8814
|  |  | 220.1856 | 0 | 0 | [5 empty] | 220.1856 | 77.62 | 62.1875 | 340.8407
|  |  | 220.2624 | 0 | 0 | [5 empty] | 220.2624 | 77.655 | 62.0339 | 347.4985
|  |  | 220.3392 | 0 | 0 | [5 empty] | 220.3392 | 77.71 | 62.3155 | 349.2862
|  |  | 220.416 | 0 | 0 | [5 empty] | 220.416 | 77.875 | 62.7405 | 347.9608
|  |  | 220.4928 | 0 | 0 | [5 empty] | 220.4928 | 78.22 | 62.9606 | 345.6183
|  |  | 220.5696 | 0 | 0 | [5 empty] | 220.5696 | 78.45 | 63.5904 | 344.1388
|  |  | 220.6464 | 0 | 0 | [5 empty] | 220.6464 | 78.605 | 63.7952 | 344.0463
|  |  | 220.7232 | 0 | 0 | [5 empty] | 220.7232 | 78.645 | 63.2013 | 344.971
|  |  | 220.8 | 0 | 0 | [5 empty] | 220.8 | 78.34 | 62.8275 | 345.8649
|  |  | 220.8768 | 0 | 0 | [5 empty] | 220.8768 | 77.83 | 62.1107 | 345.3409
|  |  | 220.9536 | 0 | 0 | [5 empty] | 220.9536 | 77.46 | 59.8784 | 342.1353
|  |  | 221.0304 | 0 | 0 | [5 empty] | 221.0304 | 76.795 | 58.7059 | 335.8474
|  |  | 221.1072 | 0 | 0 | [5 empty] | 221.1072 | 75.78 | 59.5763 | 327.2786
|  |  | 221.184 | 0 | 0 | [5 empty] | 221.184 | 75.385 | 59.2333 | 317.3844
|  |  | 221.2608 | 0 | 0 | [5 empty] | 221.2608 | 75.72 | 58.2605 | 307.2745
|  |  | 221.3376 | 0 | 0 | [5 empty] | 221.3376 | 76.23 | 58.1888 | 302.651
|  |  | 221.4144 | 0 | 0 | [5 empty] | 221.4144 | 76.39 | 59.479 | 304.1614
|  |  | 221.4912 | 0 | 0 | [5 empty] | 221.4912 | 76.545 | 59.4432 | 299.5995
|  |  | 221.568 | 0 | 0 | [5 empty] | 221.568 | 76.575 | 58.1734 | 293.1267
|  |  | 221.6448 | 0 | 0 | [5 empty] | 221.6448 | 76.125 | 58.0454 | 287.7635
|  |  | 221.7216 | 0 | 0 | [5 empty] | 221.7216 | 75.6 | 58.0813 | 283.6332
|  |  | 221.7984 | 0 | 0 | [5 empty] | 221.7984 | 75.455 | 57.2826 | 280.705
|  |  | 221.8752 | 0 | 0 | [5 empty] | 221.8752 | 75.795 | 55.5059 | 278.8249
|  |  | 221.952 | 0 | 0 | [5 empty] | 221.952 | 75.955 | 55.9821 | 277.7769
|  |  | 222.0288 | 0 | 0 | [5 empty] | 222.0288 | 75.83 | 58.1683 | 277.2529
|  |  | 222.1056 | 0 | 0 | [5 empty] | 222.1056 | 75.885 | 59.1411 | 277.2221
|  |  | 222.1824 | 0 | 0 | [5 empty] | 222.1824 | 76.075 | 58.711 | 277.7461
|  |  | 222.2592 | 0 | 0 | [5 empty] | 222.2592 | 75.985 | 58.4294 | 278.6091
|  |  | 222.336 | 0 | 0 | [5 empty] | 222.336 | 75.705 | 57.8509 | 280.0578
|  |  | 222.4128 | 0 | 0 | [5 empty] | 222.4128 | 75.33 | 57.1546 | 287.5786
|  |  | 222.4896 | 0 | 0 | [5 empty] | 222.4896 | 74.92 | 56.9344 | 306.134
|  |  | 222.5664 | 0 | 0 | [5 empty] | 222.5664 | 74.715 | 56.7296 | 325.5833
|  |  | 222.6432 | 0 | 0 | [5 empty] | 222.6432 | 74.67 | 57.4822 | 339.1763
|  |  | 222.72 | 0 | 0 | [5 empty] | 222.72 | 75.04 | 58.3373 | 346.882
|  |  | 222.7968 | 0 | 0 | [5 empty] | 222.7968 | 75.73 | 59.3101 | 350.5808
|  |  | 222.8736 | 0 | 0 | [5 empty] | 222.8736 | 76.495 | 60.8102 | 352.8925
|  |  | 222.9504 | 0 | 0 | [5 empty] | 222.9504 | 77.36 | 61.7574 | 356.0673
|  |  | 223.0272 | 0 | 0 | [5 empty] | 223.0272 | 78.235 | 61.7933 | 361.338
|  |  | 223.104 | 0 | 0 | [5 empty] | 223.104 | 78.525 | 62.8992 | 368.3348
|  |  | 223.1808 | 0 | 0 | [5 empty] | 223.1808 | 78.715 | 63.0733 | 375.3317
|  |  | 223.2576 | 0 | 0 | [5 empty] | 223.2576 | 78.625 | 63.3293 | 380.3867
|  |  | 223.3344 | 0 | 0 | [5 empty] | 223.3344 | 78.38 | 63.3344 | 381.8661
|  |  | 223.4112 | 0 | 0 | [5 empty] | 223.4112 | 78.19 | 63.0886 | 379.1229
|  |  | 223.488 | 0 | 0 | [5 empty] | 223.488 | 78.095 | 62.9555 | 372.8967
|  |  | 223.5648 | 0 | 0 | [5 empty] | 223.5648 | 78.045 | 62.7712 | 364.6977
|  |  | 223.6416 | 0 | 0 | [5 empty] | 223.6416 | 77.925 | 63.1654 | 356.4988
|  |  | 223.7184 | 0 | 0 | [5 empty] | 223.7184 | 77.68 | 63.9437 | 349.6869
|  |  | 223.7952 | 0 | 0 | [5 empty] | 223.7952 | 77.645 | 62.9299 | 344.7861
|  |  | 223.872 | 0 | 0 | [5 empty] | 223.872 | 77.555 | 61.0048 | 341.4264
|  |  | 223.9488 | 0 | 0 | [5 empty] | 223.9488 | 77.19 | 59.7043 | 338.0358
|  |  | 224.0256 | 0 | 0 | [5 empty] | 224.0256 | 76.74 | 58.6394 | 333.3199
|  |  | 224.1024 | 0 | 0 | [5 empty] | 224.1024 | 76.34 | 58.4499 | 327.217
|  |  | 224.1792 | 0 | 0 | [5 empty] | 224.1792 | 76.45 | 58.9978 | 320.251
|  |  | 224.256 | 0 | 0 | [5 empty] | 224.256 | 76.98 | 60.8051 | 312.4219
|  |  | 224.3328 | 0 | 0 | [5 empty] | 224.3328 | 77.73 | 62.4794 | 305.2402
|  |  | 224.4096 | 0 | 0 | [5 empty] | 224.4096 | 78.72 | 63.0682 | 299.353
|  |  | 224.4864 | 0 | 0 | [5 empty] | 224.4864 | 79.33 | 64.4915 | 294.7912
|  |  | 224.5632 | 0 | 0 | [5 empty] | 224.5632 | 79.715 | 65.9558 | 291.3698
|  |  | 224.64 | 0 | 0 | [5 empty] | 224.64 | 80.12 | 65.8688 | 288.8732
|  |  | 224.7168 | 0 | 0 | [5 empty] | 224.7168 | 80.395 | 64.8858 | 287.1471
|  |  | 224.7936 | 0 | 0 | [5 empty] | 224.7936 | 80.325 | 64.64 | 285.9758
|  |  | 224.8704 | 0 | 0 | [5 empty] | 224.8704 | 80.185 | 64.5632 | 285.1436
|  |  | 224.9472 | 0 | 0 | [5 empty] | 224.9472 | 80.265 | 64.0256 | 284.4963
|  |  | 225.024 | 0 | 0 | [5 empty] | 225.024 | 80.255 | 64.4198 | 283.849
|  |  | 225.1008 | 0 | 0 | [5 empty] | 225.1008 | 80.22 | 65.1264 | 283.1401
|  |  | 225.1776 | 0 | 0 | [5 empty] | 225.1776 | 80.28 | 65.2851 | 282.9551
|  |  | 225.2544 | 0 | 0 | [5 empty] | 225.2544 | 80.315 | 65.8074 | 285.0819
|  |  | 225.3312 | 0 | 0 | [5 empty] | 225.3312 | 80.73 | 66.5242 | 290.7534
|  |  | 225.408 | 0 | 0 | [5 empty] | 225.408 | 81.325 | 67.2614 | 300.6167
|  |  | 225.4848 | 0 | 0 | [5 empty] | 225.4848 | 81.56 | 68.608 | 314.672
|  |  | 225.5616 | 0 | 0 | [5 empty] | 225.5616 | 81.66 | 69.1968 | 331.1623
|  |  | 225.6384 | 0 | 0 | [5 empty] | 225.6384 | 81.95 | 70.0109 | 347.1286
|  |  | 225.7152 | 0 | 0 | [5 empty] | 225.7152 | 82.365 | 70.359 | 359.5503
|  |  | 225.792 | 0 | 0 | [5 empty] | 225.792 | 82.565 | 70.4614 | 366.3622
|  |  | 225.8688 | 0 | 0 | [5 empty] | 225.8688 | 82.645 | 70.7226 | 367.441
|  |  | 225.9456 | 0 | 0 | [5 empty] | 225.9456 | 82.88 | 70.6048 | 364.6053
|  |  | 226.0224 | 0 | 0 | [5 empty] | 226.0224 | 83.105 | 70.9632 | 360.1976
|  |  | 226.0992 | 0 | 0 | [5 empty] | 226.0992 | 83.305 | 71.6032 | 356.2214
|  |  | 226.176 | 0 | 0 | [5 empty] | 226.176 | 83.52 | 71.2499 | 354.0022
|  |  | 226.2528 | 0 | 0 | [5 empty] | 226.2528 | 83.525 | 70.7686 | 353.4473
|  |  | 226.3296 | 0 | 0 | [5 empty] | 226.3296 | 83.34 | 69.9443 | 353.2316
|  |  | 226.4064 | 0 | 0 | [5 empty] | 226.4064 | 82.91 | 68.9254 | 351.8445
|  |  | 226.4832 | 0 | 0 | [5 empty] | 226.4832 | 82.21 | 68.2803 | 347.8067
|  |  | 226.56 | 0 | 0 | [5 empty] | 226.56 | 81.625 | 66.2221 | 340.7483
|  |  | 226.6368 | 0 | 0 | [5 empty] | 226.6368 | 81.015 | 64.5069 | 331.039
|  |  | 226.7136 | 0 | 0 | [5 empty] | 226.7136 | 80.575 | 63.5238 | 319.7886
|  |  | 226.7904 | 0 | 0 | [5 empty] | 226.7904 | 80.455 | 64.1178 | 307.7677
|  |  | 226.8672 | 0 | 0 | [5 empty] | 226.8672 | 80.73 | 65.0086 | 297.4728
|  |  | 226.944 | 0 | 0 | [5 empty] | 226.944 | 81.04 | 65.6026 | 296.3632
|  |  | 227.0208 | 0 | 0 | [5 empty] | 227.0208 | 81.215 | 65.1981 | 295.1302
|  |  | 227.0976 | 0 | 0 | [5 empty] | 227.0976 | 81.045 | 64.727 | 287.5478
|  |  | 227.1744 | 0 | 0 | [5 empty] | 227.1744 | 80.48 | 64.8243 | 279.9653
|  |  | 227.2512 | 0 | 0 | [5 empty] | 227.2512 | 79.955 | 64.5274 | 273.2459
|  |  | 227.328 | 0 | 0 | [5 empty] | 227.328 | 79.695 | 63.7082 | 267.6053
|  |  | 227.4048 | 0 | 0 | [5 empty] | 227.4048 | 79.48 | 63.4573 | 263.1976
|  |  | 227.4816 | 0 | 0 | [5 empty] | 227.4816 | 79.33 | 63.7082 | 259.8995
|  |  | 227.5584 | 0 | 0 | [5 empty] | 227.5584 | 79.455 | 63.1859 | 257.557
|  |  | 227.6352 | 0 | 0 | [5 empty] | 227.6352 | 79.365 | 63.2883 | 256.0467
|  |  | 227.712 | 0 | 0 | [5 empty] | 227.712 | 79.25 | 62.9146 | 255.3069
|  |  | 227.7888 | 0 | 0 | [5 empty] | 227.7888 | 78.925 | 62.5715 | 255.2144
|  |  | 227.8656 | 0 | 0 | [5 empty] | 227.8656 | 78.515 | 61.5526 | 255.461
|  |  | 227.9424 | 0 | 0 | [5 empty] | 227.9424 | 78 | 60.3238 | 257.9885
|  |  | 228.0192 | 0 | 0 | [5 empty] | 228.0192 | 77.51 | 59.0694 | 270.1944
|  |  | 228.096 | 0 | 0 | [5 empty] | 228.096 | 77.255 | 57.2262 | 288.7807
|  |  | 228.1728 | 0 | 0 | [5 empty] | 228.1728 | 76.91 | 56.8525 | 303.7298
|  |  | 228.2496 | 0 | 0 | [5 empty] | 228.2496 | 76.88 | 56.8678 | 313.1
|  |  | 228.3264 | 0 | 0 | [5 empty] | 228.3264 | 76.97 | 58.1939 | 318.5865
|  |  | 228.4032 | 0 | 0 | [5 empty] | 228.4032 | 77.14 | 60.6259 | 322.4394
|  |  | 228.48 | 0 | 0 | [5 empty] | 228.48 | 77.94 | 61.1072 | 326.8163
|  |  | 228.5568 | 0 | 0 | [5 empty] | 228.5568 | 78.585 | 61.7574 | 332.7959
|  |  | 228.6336 | 0 | 0 | [5 empty] | 228.6336 | 78.845 | 62.5664 | 340.3167
|  |  | 228.7104 | 0 | 0 | [5 empty] | 228.7104 | 79.02 | 62.4794 | 348.1458
|  |  | 228.7872 | 0 | 0 | [5 empty] | 228.7872 | 78.97 | 62.4794 | 354.5262
|  |  | 228.864 | 0 | 0 | [5 empty] | 228.864 | 78.72 | 62.6278 | 357.9475
|  |  | 228.9408 | 0 | 0 | [5 empty] | 228.9408 | 78.645 | 61.5936 | 357.6393
|  |  | 229.0176 | 0 | 0 | [5 empty] | 229.0176 | 78.705 | 60.2573 | 353.8172
|  |  | 229.0944 | 0 | 0 | [5 empty] | 229.0944 | 78.395 | 61.0714 | 347.6218
|  |  | 229.1712 | 0 | 0 | [5 empty] | 229.1712 | 78.13 | 61.7677 | 340.6866
|  |  | 229.248 | 0 | 0 | [5 empty] | 229.248 | 78.195 | 60.4723 | 334.4295
|  |  | 229.3248 | 0 | 0 | [5 empty] | 229.3248 | 77.9 | 60.0883 | 329.5595
|  |  | 229.4016 | 0 | 0 | [5 empty] | 229.4016 | 77.175 | 60.3699 | 325.9532
|  |  | 229.4784 | 0 | 0 | [5 empty] | 229.4784 | 76.5 | 59.5405 | 322.5935
|  |  | 229.5552 | 0 | 0 | [5 empty] | 229.5552 | 75.84 | 58.6598 | 318.0934
|  |  | 229.632 | 0 | 0 | [5 empty] | 229.632 | 75.425 | 56.9549 | 311.8363
|  |  | 229.7088 | 0 | 0 | [5 empty] | 229.7088 | 75.135 | 56.5299 | 304.6545
|  |  | 229.7856 | 0 | 0 | [5 empty] | 229.7856 | 75.535 | 56.1357 | 296.6406
|  |  | 229.8624 | 0 | 0 | [5 empty] | 229.8624 | 76.065 | 56.4378 | 288.8732
|  |  | 229.9392 | 0 | 0 | [5 empty] | 229.9392 | 76.48 | 57.6563 | 282.5544
|  |  | 230.016 | 0 | 0 | [5 empty] | 230.016 | 76.85 | 58.9158 | 277.931
|  |  | 230.0928 | 0 | 0 | [5 empty] | 230.0928 | 77.195 | 59.3613 | 274.8487
|  |  | 230.1696 | 0 | 0 | [5 empty] | 230.1696 | 77.265 | 59.5507 | 273.0301
|  |  | 230.2464 | 0 | 0 | [5 empty] | 230.2464 | 77.13 | 59.6941 | 272.0438
|  |  | 230.3232 | 0 | 0 | [5 empty] | 230.3232 | 76.955 | 59.4381 | 271.5815
|  |  | 230.4 | 0 | 0 | [5 empty] | 230.4 | 76.66 | 59.1514 | 271.4273
|  |  | 230.4768 | 0 | 0 | [5 empty] | 230.4768 | 76.365 | 59.095 | 271.5198
|  |  | 230.5536 | 0 | 0 | [5 empty] | 230.5536 | 76.055 | 59.4944 | 271.828
|  |  | 230.6304 | 0 | 0 | [5 empty] | 230.6304 | 76.07 | 58.9414 | 272.3828
|  |  | 230.7072 | 0 | 0 | [5 empty] | 230.7072 | 76.165 | 58.0813 | 273.3384
|  |  | 230.784 | 0 | 0 | [5 empty] | 230.784 | 76.32 | 57.3696 | 276.1124
|  |  | 230.8608 | 0 | 0 | [5 empty] | 230.8608 | 76.53 | 58.3987 | 282.0921
|  |  | 230.9376 | 0 | 0 | [5 empty] | 230.9376 | 76.82 | 60.503 | 291.678
|  |  | 231.0144 | 0 | 0 | [5 empty] | 231.0144 | 77.33 | 60.9946 | 305.1169
|  |  | 231.0912 | 0 | 0 | [5 empty] | 231.0912 | 77.77 | 60.5901 | 321.3606
|  |  | 231.168 | 0 | 0 | [5 empty] | 231.168 | 77.9 | 61.7216 | 337.8509
|  |  | 231.2448 | 0 | 0 | [5 empty] | 231.2448 | 77.94 | 62.5664 | 351.6904
|  |  | 231.3216 | 0 | 0 | [5 empty] | 231.3216 | 78.085 | 61.9418 | 360.7832
|  |  | 231.3984 | 0 | 0 | [5 empty] | 231.3984 | 78.065 | 62.167 | 364.3895
|  |  | 231.4752 | 0 | 0 | [5 empty] | 231.4752 | 78.215 | 62.7046 | 363.4648
|  |  | 231.552 | 0 | 0 | [5 empty] | 231.552 | 78.42 | 63.6058 | 359.9818
|  |  | 231.6288 | 0 | 0 | [5 empty] | 231.6288 | 78.725 | 64.6349 | 356.0981
|  |  | 231.7056 | 0 | 0 | [5 empty] | 231.7056 | 78.96 | 65.7101 | 353.5706
|  |  | 231.7824 | 0 | 0 | [5 empty] | 231.7824 | 79.04 | 66.2579 | 353.0775
|  |  | 231.8592 | 0 | 0 | [5 empty] | 231.8592 | 79.03 | 65.664 | 353.9405
|  |  | 231.936 | 0 | 0 | [5 empty] | 231.936 | 78.81 | 64.5376 | 354.557
|  |  | 232.0128 | 0 | 0 | [5 empty] | 232.0128 | 78.3 | 63.4675 | 353.2932
|  |  | 232.0896 | 0 | 0 | [5 empty] | 232.0896 | 77.755 | 62.2387 | 349.0088
|  |  | 232.1664 | 0 | 0 | [5 empty] | 232.1664 | 77.475 | 60.0678 | 341.7654
|  |  | 232.2432 | 0 | 0 | [5 empty] | 232.2432 | 77.19 | 58.9414 | 332.3952
|  |  | 232.32 | 0 | 0 | [5 empty] | 232.32 | 77.03 | 59.9859 | 321.8229
|  |  | 232.3968 | 0 | 0 | [5 empty] | 232.3968 | 77.39 | 61.3734 | 312.576
|  |  | 232.4736 | 0 | 0 | [5 empty] | 232.4736 | 77.835 | 62.7763 | 313.2233
|  |  | 232.5504 | 0 | 0 | [5 empty] | 232.5504 | 78.12 | 64.128 | 314.3638
|  |  | 232.6272 | 39.105 | 0.74752 | [5 empty] | 232.6272 | 78.27 | 64.3994 | 308.6615
|  |  | 232.704 | 41.835 | 0.86528 | [5 empty] | 232.704 | 78.25 | 64.0563 | 302.6819
|  |  | 232.7808 | 43.715 | 1.0854 | [5 empty] | 232.7808 | 78.215 | 63.2218 | 297.6269
|  |  | 232.8576 | 45.7 | 1.2851 | [5 empty] | 232.8576 | 78.175 | 62.9248 | 293.7432
|  |  | 232.9344 | 47.765 | 1.4336 | [5 empty] | 232.9344 | 78.295 | 62.8326 | 291.1232
|  |  | 233.0112 | 50.555 | 1.4336 | [5 empty] | 233.0112 | 78.58 | 63.2013 | 289.5204
|  |  | 233.088 | 52.38 | 1.5462 | [5 empty] | 233.088 | 78.89 | 64.3482 | 288.5958
|  |  | 233.1648 | 53.1 | 1.833 | [5 empty] | 233.1648 | 79.395 | 64.7782 | 288.0717
|  |  | 233.2416 | 54.53 | 1.9712 | [5 empty] | 233.2416 | 79.76 | 65.4797 | 288.0101
|  |  | 233.3184 | 56.25 | 2.0378 | [5 empty] | 233.3184 | 79.935 | 65.7408 | 288.2875
|  |  | 233.3952 | 57.395 | 2.1862 | [5 empty] | 233.3952 | 79.865 | 65.8739 | 288.5341
|  |  | 233.472 | 58.5 | 2.2528 | [5 empty] | 233.472 | 79.635 | 65.7408 | 288.904
|  |  | 233.5488 | 59.54 | 2.263 | [5 empty] | 233.5488 | 79.365 | 65.1674 | 294.3288
|  |  | 233.6256 | 60.025 | 2.3501 | [5 empty] | 233.6256 | 79.125 | 64.7168 | 310.6342
|  |  | 233.7024 | 60.07 | 2.4934 | [5 empty] | 233.7024 | 79.26 | 63.488 | 329.0047
|  |  | 233.7792 | 60.695 | 2.4832 | [5 empty] | 233.7792 | 79.43 | 63.6928 | 342.5052
|  |  | 233.856 | 61.3 | 2.4525 | [5 empty] | 233.856 | 79.76 | 65.0138 | 350.9507
|  |  | 233.9328 | 61.375 | 2.5242 | [5 empty] | 233.9328 | 80.36 | 66.6266 | 355.8824
|  |  | 234.0096 | 61.495 | 2.5702 | [5 empty] | 234.0096 | 81.195 | 67.712 | 359.1496
|  |  | 234.0864 | 61.935 | 2.5549 | [5 empty] | 234.0864 | 82.085 | 68.3725 | 362.571
|  |  | 234.1632 | 61.94 | 2.6419 | [5 empty] | 234.1632 | 82.58 | 70.016 | 367.1019
|  |  | 234.24 | 61.97 | 2.7238 | [5 empty] | 234.24 | 82.755 | 71.9206 | 372.6501
|  |  | 234.3168 | 62.21 | 2.7699 | [5 empty] | 234.3168 | 82.705 | 73.0317 | 377.9825
|  |  | 234.3936 | 61.23 | 3.1693 | [5 empty] | 234.3936 | 82.635 | 72.5914 | 381.6196
|  |  | 234.4704 | 62.83 | 2.7546 | [5 empty] | 234.4704 | 82.5 | 71.5418 | 382.3285
|  |  | 234.5472 | 61.58 | 3.1078 | [5 empty] | 234.5472 | 82.305 | 70.7738 | 379.3387
|  |  | 234.624 | 62.625 | 2.8262 | [5 empty] | 234.624 | 82.225 | 69.6371 | 373.0508
|  |  | 234.7008 | 61.795 | 3.0925 | [5 empty] | 234.7008 | 82.105 | 68.9306 | 364.6669
|  |  | 234.7776 | 63.03 | 2.7597 | [5 empty] | 234.7776 | 81.785 | 69.5654 | 355.7899
|  |  | 234.8544 | 61.85 | 3.0874 | [5 empty] | 234.8544 | 81.59 | 68.864 | 347.7143
|  |  | 234.9312 | 62.995 | 2.7597 | [5 empty] | 234.9312 | 81.29 | 67.7222 | 341.0257
|  |  | 235.008 | 61.635 | 3.113 | [5 empty] | 235.008 | 80.63 | 67.1795 | 335.5392
|  |  | 235.0848 | 62.72 | 2.7802 | [5 empty] | 235.0848 | 79.79 | 66.1965 | 330.176
|  |  | 235.1616 | 61.145 | 3.1693 | [5 empty] | 235.1616 | 79.23 | 63.5648 | 323.7956
|  |  | 235.2384 | 62.05 | 2.8211 | [5 empty] | 235.2384 | 78.585 | 62.6483 | 316.1515
|  |  | 235.3152 | 60.855 | 3.1078 | [5 empty] | 235.3152 | 78.49 | 61.5322 | 307.9834
|  |  | 235.392 | 61.43 | 2.816 | [5 empty] | 235.392 | 78.78 | 61.7882 | 299.0447
|  |  | 235.4688 | 61.46 | 2.7136 | [5 empty] | 235.4688 | 78.99 | 63.744 | 290.4143
|  |  | 235.5456 | 61.42 | 2.5907 | [5 empty] | 235.5456 | 79.415 | 64.9882 | 283.0168
|  |  | 235.6224 | 60.94 | 2.5293 | [5 empty] | 235.6224 | 79.93 | 65.0189 | 276.9755
|  |  | 235.6992 | 59.97 | 2.5242 | [5 empty] | 235.6992 | 80.03 | 65.2954 | 272.1054
|  |  | 235.776 | 58.425 | 2.4883 | [5 empty] | 235.776 | 79.97 | 64.9165 | 268.2834
|  |  | 235.8528 | 56.51 | 2.4627 | [5 empty] | 235.8528 | 79.715 | 64.0102 | 265.386
|  |  | 235.9296 | 55.115 | 2.3296 | [5 empty] | 235.9296 | 79.39 | 62.5357 | 263.2901
|  |  | 236.0064 | 53.855 | 2.1555 | [5 empty] | 236.0064 | 78.8 | 61.9418 | 261.8722
|  |  | 236.0832 | 52.88 | 1.9558 | [5 empty] | 236.0832 | 78.295 | 61.5168 | 260.9475
|  |  | 236.16 | 51.6 | 1.8176 | [5 empty] | 236.16 | 77.835 | 61.5424 | 260.3619
|  |  | 236.2368 | 50.035 | 1.7357 | [5 empty] | 236.2368 | 77.585 | 61.1379 | 259.992
|  |  | 236.3136 | 48.22 | 1.7101 | [5 empty] | 236.3136 | 77.46 | 60.2266 | 260.4235
|  |  | 236.3904 | 46.98 | 1.536 | [5 empty] | 236.3904 | 77.285 | 59.9245 | 263.1976
|  |  | 236.4672 | 45.955 | 1.3363 | [5 empty] | 236.4672 | 77.355 | 60.4723 | 269.5163
|  |  | 236.544 | 45.425 | 1.0496 | [5 empty] | 236.544 | 77.62 | 60.9741 | 279.8728
|  |  | 236.6208 | 44.385 | 0.86016 | [5 empty] | 236.6208 | 77.705 | 61.2045 | 294.2672
|  |  | 236.6976 | 42.475 | 0.768 | [5 empty] | 236.6976 | 77.56 | 61.143 | 310.8499
|  |  | 236.7744 | 40.48 | 0.67072 | [5 empty] | 236.7744 | 77.61 | 60.8256 | 326.8779
|  |  | 236.8512 | 0 | 0 | [5 empty] | 236.8512 | 78.14 | 58.9978 | 339.3921
|  |  | 236.928 | 0 | 0 | [5 empty] | 236.928 | 78.145 | 59.0746 | 346.6971
|  |  | 237.0048 | 0 | 0 | [5 empty] | 237.0048 | 78.105 | 59.863 | 349.0396
|  |  | 237.0816 | 0 | 0 | [5 empty] | 237.0816 | 78.065 | 61.4554 | 347.9917
[11 empty] | 237.1584 | 78.185 | 62.464 | 345.5875
[11 empty] | 237.2352 | 78.56 | 62.1312 | 343.7689
[11 empty] | 237.312 | 78.475 | 63.2627 | 343.4607
[11 empty] | 237.3888 | 78.375 | 63.447 | 344.6011
[11 empty] | 237.4656 | 78.105 | 63.0221 | 345.9574
[11 empty] | 237.5424 | 77.47 | 63.104 | 345.8649
[11 empty] | 237.6192 | 77.015 | 61.4758 | 342.9984
[11 empty] | 237.696 | 76.45 | 60.3392 | 336.9262
[11 empty] | 237.7728 | 75.98 | 59.049 | 328.2958
[11 empty] | 237.8496 | 75.67 | 58.1376 | 318.3708
[11 empty] | 237.9264 | 75.795 | 57.7741 | 308.0759
[11 empty] | 238.0032 | 76.155 | 58.4397 | 302.651
[11 empty] | 238.08 | 76.3 | 59.9654 | 304.5312
[11 empty] | 238.1568 | 76.37 | 60.416 | 300.6167
[11 empty] | 238.2336 | 76.35 | 59.5456 | 293.9898
[11 empty] | 238.3104 | 76.205 | 58.0198 | 288.2259
[11 empty] | 238.3872 | 75.92 | 56.9805 | 283.6641
[11 empty] | 238.464 | 75.475 | 57.4874 | 280.4276
[11 empty] | 238.5408 | 75.225 | 58.537 | 278.6091
[11 empty] | 238.6176 | 75.385 | 58.752 | 277.7461
[11 empty] | 238.6944 | 75.76 | 58.5574 | 277.4378
[11 empty] | 238.7712 | 76.205 | 57.9021 | 277.4995
[11 empty] | 238.848 | 76.315 | 58.0813 | 277.9002
[11 empty] | 238.9248 | 76.125 | 58.2861 | 278.4858
[11 empty] | 239.0016 | 75.785 | 57.897 | 279.5338
[11 empty] | 239.0784 | 75.59 | 56.0589 | 285.9142
[11 empty] | 239.1552 | 75.215 | 55.3523 | 303.6682
[11 empty] | 239.232 | 75.005 | 55.4854 | 323.6415
[11 empty] | 239.3088 | 75.05 | 56.064 | 338.005
[11 empty] | 239.3856 | 75.29 | 57.3594 | 346.3581
[11 empty] | 239.4624 | 75.895 | 58.3322 | 350.3342
[11 empty] | 239.5392 | 76.63 | 59.817 | 352.5226
[11 empty] | 239.616 | 77.425 | 61.2966 | 355.4508
[11 empty] | 239.6928 | 78.135 | 62.1926 | 360.3209
[11 empty] | 239.7696 | 78.39 | 63.8976 | 366.9478
[11 empty] | 239.8464 | 78.205 | 66.2886 | 373.883
[11 empty] | 239.9232 | 78.43 | 65.3107 | 379.1845
[11 empty] | 240 | 78.44 | 64.1485 | 381.2189
[11 empty] | 240.0768 | 78.155 | 64.3738 | 379.1229
[11 empty] | 240.1536 | 78.065 | 63.6979 | 373.2357
[11 empty] | 240.2304 | 78.065 | 63.1091 | 365.0984
[11 empty] | 240.3072 | 78.11 | 62.5357 | 356.6221
[11 empty] | 240.384 | 77.845 | 63.447 | 349.317
[11 empty] | 240.4608 | 77.795 | 62.592 | 344.0771
[11 empty] | 240.5376 | 77.685 | 61.1533 | 340.5017
[11 empty] | 240.6144 | 77.14 | 60.8768 | 337.3269
[11 empty] | 240.6912 | 76.685 | 59.8221 | 332.9809
[11 empty] | 240.768 | 76.555 | 57.9891 | 327.0937
[11 empty] | 240.8448 | 76.63 | 58.281 | 320.4051
[11 empty] | 240.9216 | 77.285 | 58.9722 | 312.576
[11 empty] | 240.9984 | 78 | 60.8307 | 305.1785
[11 empty] | 241.0752 | 78.54 | 63.5085 | 299.0447
[11 empty] | 241.152 | 79.3 | 64.2867 | 294.2364
[11 empty] | 241.2288 | 80.185 | 63.3293 | 290.5992
[11 empty] | 241.3056 | 80.335 | 64.3891 | 287.8868
[11 empty] | 241.3824 | 80.23 | 65.3978 | 286.0066
[11 empty] | 241.4592 | 80.28 | 64.9267 | 284.712
[11 empty] | 241.536 | 80.08 | 65.3005 | 283.849
[11 empty] | 241.6128 | 80.04 | 65.2595 | 283.1401
[11 empty] | 241.6896 | 80.29 | 64.2048 | 282.4003
[11 empty] | 241.7664 | 80.4 | 64.2099 | 281.5373
[11 empty] | 241.8432 | 80.475 | 64.2202 | 281.1674
[11 empty] | 241.92 | 80.655 | 63.9539 | 283.0168
[11 empty] | 241.9968 | 80.925 | 65.1008 | 288.5958
[11 empty] | 242.0736 | 81.285 | 67.1027 | 298.3975
[11 empty] | 242.1504 | 81.695 | 67.5994 | 312.2986
[11 empty] | 242.2272 | 81.895 | 67.6147 | 328.604
[11 empty] | 242.304 | 82.105 | 68.3878 | 344.3546
[11 empty] | 242.3808 | 82.31 | 69.4067 | 356.6221
[11 empty] | 242.4576 | 82.52 | 69.6832 | 363.4956
[11 empty] | 242.5344 | 82.63 | 69.8982 | 364.8518
[11 empty] | 242.6112 | 82.89 | 69.7088 | 362.2627
[11 empty] | 242.688 | 83.04 | 70.1491 | 358.0092
[11 empty] | 242.7648 | 83.27 | 70.4717 | 354.0946
[11 empty] | 242.8416 | 83.355 | 70.9274 | 351.6288
[11 empty] | 242.9184 | 83.245 | 70.8557 | 350.7041
[11 empty] | 242.9952 | 82.925 | 70.4256 | 350.2726
[11 empty] | 243.072 | 82.53 | 69.2787 | 348.9472
[11 empty] | 243.1488 | 82.07 | 67.5994 | 345.3101
[11 empty] | 243.2256 | 81.29 | 66.9952 | 338.6215
[11 empty] | 243.3024 | 80.595 | 65.92 | 329.3746
[11 empty] | 243.3792 | 80.33 | 63.7594 | 318.4016
[11 empty] | 243.456 | 80.65 | 62.1261 | 306.3498
[11 empty] | 243.5328 | 81.005 | 62.5459 | 295.2227
[11 empty] | 243.6096 | 81.11 | 64.0051 | 293.2192
[11 empty] | 243.6864 | 81.07 | 64.8909 | 293.4658
[11 empty] | 243.7632 | 80.805 | 65.0906 | 286.4998
[11 empty] | 243.84 | 80.32 | 65.1418 | 278.794
[11 empty] | 243.9168 | 80 | 63.913 | 271.7972
[11 empty] | 243.9936 | 79.68 | 63.2678 | 265.8176
[11 empty] | 244.0704 | 79.515 | 62.6022 | 261.1633
[11 empty] | 244.1472 | 79.405 | 62.4742 | 257.8344
[11 empty] | 244.224 | 79.31 | 62.9555 | 255.646
[11 empty] | 244.3008 | 79.245 | 63.5238 | 254.3822
[11 empty] | 244.3776 | 79.115 | 63.4726 | 253.7966
[11 empty] | 244.4544 | 79.085 | 61.9469 | 253.6733
[11 empty] | 244.5312 | 78.7 | 60.6515 | 253.7658
[11 empty] | 244.608 | 78.175 | 59.4125 | 256.2932
[11 empty] | 244.6848 | 77.48 | 58.7981 | 268.8382
[11 empty] | 244.7616 | 76.92 | 58.3526 | 287.3936
[11 empty] | 244.8384 | 76.66 | 57.8099 | 302.0346
[11 empty] | 244.9152 | 76.505 | 58.7315 | 311.1582
[11 empty] | 244.992 | 76.805 | 59.4381 | 316.4906
[11 empty] | 245.0688 | 77.525 | 59.5405 | 320.1893
[11 empty] | 245.1456 | 78.23 | 60.4928 | 324.4737
[11 empty] | 245.2224 | 78.6 | 62.2285 | 330.5458
[11 empty] | 245.2992 | 79.06 | 62.2592 | 338.2208
[11 empty] | 245.376 | 79.24 | 62.0134 | 346.2039
[11 empty] | 245.4528 | 79.11 | 61.9878 | 352.5843
[11 empty] | 245.5296 | 78.81 | 62.249 | 355.9748
[11 empty] | 245.6064 | 78.675 | 61.481 | 355.5433
[11 empty] | 245.6832 | 78.485 | 61.2403 | 351.6288
[11 empty] | 245.76 | 78.345 | 61.2915 | 345.4642
[11 empty] | 245.8368 | 78.25 | 61.3888 | 338.7448
[11 empty] | 245.9136 | 78.13 | 60.9946 | 332.8576
[11 empty] | 245.9904 | 77.83 | 60.7898 | 328.3882
[11 empty] | 246.0672 | 77.405 | 60.4416 | 324.9977
[11 empty] | 246.144 | 76.84 | 59.4637 | 321.5455
[11 empty] | 246.2208 | 76.465 | 56.8832 | 316.768
[11 empty] | 246.2976 | 76.205 | 54.3078 | 310.1718
[11 empty] | 246.3744 | 75.805 | 54.272 | 302.7743
[11 empty] | 246.4512 | 75.93 | 54.9734 | 294.7912
[11 empty] | 246.528 | 76.1 | 57.0061 | 287.1471
[11 empty] | 246.6048 | 76.33 | 59.0541 | 281.0441
[11 empty] | 246.6816 | 77.01 | 58.7469 | 276.5131
[11 empty] | 246.7584 | 77.425 | 59.0387 | 273.3076
[11 empty] | 246.8352 | 77.4 | 59.9808 | 271.1808
[11 empty] | 246.912 | 77.205 | 60.2368 | 269.8862
[11 empty] | 246.9888 | 77.175 | 58.9363 | 269.3005
[11 empty] | 247.0656 | 77.075 | 57.8253 | 269.2389
[11 empty] | 247.1424 | 76.535 | 58.8032 | 269.578
[11 empty] | 247.2192 | 76.22 | 59.3562 | 270.102
[11 empty] | 247.296 | 76.415 | 58.4448 | 270.7184
[11 empty] | 247.3728 | 76.38 | 58.4858 | 271.5198
[11 empty] | 247.4496 | 76.43 | 58.2042 | 273.9857
[11 empty] | 247.5264 | 76.905 | 57.641 | 279.6263
[11 empty] | 247.6032 | 77.275 | 59.0746 | 289.0889
[11 empty] | 247.68 | 77.225 | 61.8803 | 302.5894
[11 empty] | 247.7568 | 77.415 | 62.6176 | 318.9256
[11 empty] | 247.8336 | 77.975 | 61.7728 | 335.5392
[11 empty] | 247.9104 | 78.315 | 61.3222 | 349.3787
[11 empty] | 247.9872 | 78.485 | 60.7283 | 358.2557
[11 empty] | 248.064 | 78.54 | 60.8922 | 361.7387
[11 empty] | 248.1408 | 78.365 | 62.7405 | 360.814
[11 empty] | 248.2176 | 78.455 | 64.169 | 357.6085
[11 empty] | 248.2944 | 78.795 | 64.9421 | 354.2179
[11 empty] | 248.3712 | 79.1 | 65.4438 | 352.0603
[11 empty] | 248.448 | 79.365 | 64.937 | 351.598
[11 empty] | 248.5248 | 79.3 | 64.4403 | 352.1836
[11 empty] | 248.6016 | 79.235 | 62.5715 | 352.3994
[11 empty] | 248.6784 | 78.89 | 61.2198 | 350.8274
[11 empty] | 248.7552 | 78.08 | 61.6397 | 346.5738
[11 empty] | 248.832 | 77.435 | 61.0458 | 339.7311
[11 empty] | 248.9088 | 77.095 | 59.8374 | 330.9466
[11 empty] | 248.9856 | 77.195 | 59.607 | 320.8366
[11 empty] | 249.0624 | 77.49 | 61.2864 | 311.2198
[11 empty] | 249.1392 | 78.025 | 62.3718 | 310.3876
[11 empty] | 249.216 | 78.455 | 62.9248 | 311.9596
[11 empty] | 249.2928 | 78.71 | 62.6432 | 306.5656
[11 empty] | 249.3696 | 78.82 | 61.5987 | 300.6476
[11 empty] | 249.4464 | 78.53 | 61.8394 | 295.5926
[11 empty] | 249.5232 | 78.385 | 62.208 | 291.7089
[11 empty] | 249.6 | 78.395 | 62.7712 | 289.0581
[11 empty] | 249.6768 | 78.575 | 63.5187 | 287.4245
[11 empty] | 249.7536 | 78.955 | 64.2304 | 286.4998
[11 empty] | 249.8304 | 79.365 | 65.1264 | 286.0683
[11 empty] | 249.9072 | 79.97 | 64.4557 | 286.0683
[11 empty] | 249.984 | 80.07 | 64.9677 | 286.4381
[11 empty] | 250.0608 | 80.065 | 65.152 | 286.8388
[11 empty] | 250.1376 | 79.975 | 64.9165 | 287.2395
[11 empty] | 250.2144 | 79.585 | 65.0445 | 292.387
[11 empty] | 250.2912 | 79.245 | 64.8755 | 308.4149
[11 empty] | 250.368 | 79.38 | 63.6109 | 326.9704
[11 empty] | 250.4448 | 79.55 | 63.9334 | 340.6558
[11 empty] | 250.5216 | 80 | 64.4096 | 349.1629
[11 empty] | 250.5984 | 80.565 | 66.1402 | 354.1255
[11 empty] | 250.6752 | 81.345 | 68.2138 | 357.4543
[11 empty] | 250.752 | 82.115 | 70.2771 | 361.0606
[11 empty] | 250.8288 | 82.79 | 71.6902 | 365.8074
[11 empty] | 250.9056 | 83.4 | 71.127 | 371.448
[11 empty] | 250.9824 | 83.555 | 70.8506 | 376.7804
[11 empty] | 251.0592 | 83.18 | 71.7926 | 380.2942
[11 empty] | 251.136 | 82.86 | 71.5725 | 380.7257
[11 empty] | 251.2128 | 82.65 | 70.8762 | 377.6434
[11 empty] | 251.2896 | 82.41 | 70.441 | 371.5096
[11 empty] | 251.3664 | 82.3 | 69.8624 | 363.434
[11 empty] | 251.4432 | 82.045 | 69.8163 | 355.0193
[11 empty] | 251.52 | 81.85 | 68.8896 | 347.3752
[11 empty] | 251.5968 | 81.66 | 67.2922 | 340.8407
[11 empty] | 251.6736 | 81.245 | 65.664 | 335.1077
[11 empty] | 251.7504 | 80.55 | 64.1997 | 329.1896
[11 empty] | 251.8272 | 79.735 | 62.9197 | 322.3777
[11 empty] | 251.904 | 78.975 | 62.3718 | 314.3638
[11 empty] | 251.9808 | 78.625 | 62.2182 | 306.0724
[11 empty] | 252.0576 | 78.79 | 62.8429 | 297.4728
[11 empty] | 252.1344 | 79.125 | 64.169 | 289.1506
[11 empty] | 252.2112 | 79.695 | 64.4762 | 281.9996
[11 empty] | 252.288 | 80.12 | 64.9011 | 276.02
[11 empty] | 252.3648 | 80.29 | 65.2698 | 271.0575
[11 empty] | 252.4416 | 80.14 | 65.4131 | 266.9888
[11 empty] | 252.5184 | 79.915 | 64.4454 | 263.8449
[11 empty] | 252.5952 | 79.53 | 63.145 | 261.6873
[11 empty] | 252.672 | 78.92 | 62.72 | 260.3311
[11 empty] | 252.7488 | 78.405 | 62.2592 | 259.6221
[11 empty] | 252.8256 | 78.09 | 61.3478 | 259.3139
[11 empty] | 252.9024 | 77.78 | 60.9126 | 259.1598
[11 empty] | 252.9792 | 77.63 | 59.8528 | 259.3755
[11 empty] | 253.056 | 77.415 | 60.0525 | 261.7181
[11 empty] | 253.1328 | 77.44 | 60.713 | 267.5745
[11 empty] | 253.2096 | 77.555 | 61.6141 | 277.6536
[11 empty] | 253.2864 | 77.67 | 61.696 | 291.9554
[11 empty] | 253.3632 | 77.745 | 60.7539 | 308.754
[11 empty] | 253.44 | 77.91 | 59.8886 | 325.0594
[11 empty] | 253.5168 | 77.87 | 60.7334 | 337.8201
[11 empty] | 253.5936 | 77.78 | 61.3376 | 345.3409
[11 empty] | 253.6704 | 77.825 | 61.4042 | 347.7451
[11 empty] | 253.7472 | 77.95 | 61.7114 | 346.6663
[11 empty] | 253.824 | 78.255 | 61.8906 | 344.3546
[11 empty] | 253.9008 | 78.465 | 62.7814 | 342.536
[11 empty] | 253.9776 | 78.585 | 63.1296 | 342.3203
[11 empty] | 254.0544 | 78.625 | 62.4589 | 343.399
[11 empty] | 254.1312 | 78.46 | 61.5526 | 344.6628
[11 empty] | 254.208 | 77.905 | 61.2864 | 344.4778
[11 empty] | 254.2848 | 77.325 | 60.1702 | 341.6421
[11 empty] | 254.3616 | 76.895 | 58.1018 | 335.8474
[11 empty] | 254.4384 | 76.29 | 57.2877 | 327.6177
[11 empty] | 254.5152 | 75.785 | 57.2314 | 318.0625
[11 empty] | 254.592 | 75.855 | 57.1494 | 308.0759
[11 empty] | 254.6688 | 76.16 | 58.0352 | 302.2503
[11 empty] | 254.7456 | 76.555 | 58.455 | 303.699
[11 empty] | 254.8224 | 76.63 | 58.839 | 299.9386
[11 empty] | 254.8992 | 76.265 | 59.5661 | 293.3117
[11 empty] | 254.976 | 75.82 | 59.433 | 287.7327
[11 empty] | 255.0528 | 75.51 | 58.5626 | 283.4483
[11 empty] | 255.1296 | 75.195 | 58.4448 | 280.5201
[11 empty] | 255.2064 | 75.155 | 58.4294 | 278.7016
[11 empty] | 255.2832 | 75.75 | 56.7962 | 277.6536
[11 empty] | 255.36 | 76.3 | 55.9923 | 277.0063
[11 empty] | 255.4368 | 76.165 | 57.8662 | 276.8522
[11 empty] | 255.5136 | 76.14 | 58.5523 | 277.2837
[11 empty] | 255.5904 | 76.285 | 57.3952 | 278.1159
[11 empty] | 255.6672 | 76.17 | 56.4941 | 279.4105
[11 empty] | 255.744 | 75.725 | 56.4275 | 285.2361
[11 empty] | 255.8208 | 75.25 | 56.279 | 302.2812
[11 empty] | 255.8976 | 74.895 | 56.4992 | 322.3777
[11 empty] | 255.9744 | 74.865 | 56.8883 | 337.1111
[11 empty] | 256.0512 | 75.12 | 57.8099 | 345.7416
[11 empty] | 256.128 | 75.9 | 57.9686 | 349.9027
[11 empty] | 256.2048 | 76.76 | 59.0643 | 352.1219
[11 empty] | 256.2816 | 77.44 | 61.1533 | 354.9268
[11 empty] | 256.3584 | 78.005 | 62.8838 | 359.7352
[11 empty] | 256.4352 | 78.395 | 63.9795 | 366.4547
[11 empty] | 256.512 | 78.63 | 64.1382 | 373.5748
[11 empty] | 256.5888 | 78.645 | 63.9283 | 379.0612
[11 empty] | 256.6656 | 78.34 | 64.5734 | 381.2189
[11 empty] | 256.7424 | 78.225 | 64.0102 | 379.2154
[11 empty] | 256.8192 | 78.105 | 63.4982 | 373.4207
[11 empty] | 256.896 | 77.745 | 64.809 | 365.4067
[11 empty] | 256.9728 | 77.685 | 64.9626 | 356.992
[11 empty] | 257.0496 | 77.76 | 64.128 | 349.9335
[11 empty] | 257.1264 | 77.74 | 63.2013 | 344.9094
[11 empty] | 257.2032 | 77.54 | 61.7933 | 341.488
[11 empty] | 257.28 | 76.965 | 61.097 | 338.3441
[11 empty] | 257.3568 | 76.47 | 60.3341 | 333.9672
[11 empty] | 257.4336 | 75.91 | 60.759 | 328.08
[11 empty] | 257.5104 | 76.075 | 60.5952 | 321.2681
[11 empty] | 257.5872 | 76.905 | 60.5082 | 313.4083
[11 empty] | 257.664 | 77.825 | 60.9741 | 305.9491
[11 empty] | 257.7408 | 78.64 | 61.9981 | 299.8153
[11 empty] | 257.8176 | 79.11 | 64.3635 | 295.0378
[11 empty] | 257.8944 | 79.585 | 65.5718 | 291.4315
[11 empty] | 257.9712 | 80 | 65.7664 | 288.8115
[11 empty] | 258.048 | 80.145 | 65.9558 | 286.9313
[11 empty] | 258.1248 | 80.03 | 66.1555 | 285.7292
[11 empty] | 258.2016 | 79.93 | 66.089 | 284.9278
[11 empty] | 258.2784 | 80.085 | 65.1366 | 284.2497
[11 empty] | 258.3552 | 80.355 | 64.0307 | 283.6024
[11 empty] | 258.432 | 80.26 | 64.9984 | 282.8935
[11 empty] | 258.5088 | 80.415 | 64.8294 | 282.5853
[11 empty] | 258.5856 | 80.46 | 65.1981 | 284.3113
[11 empty] | 258.6624 | 80.75 | 66.1197 | 289.3972
[11 empty] | 258.7392 | 81.385 | 66.7443 | 298.5824
[11 empty] | 258.816 | 81.745 | 67.9731 | 311.9904
[11 empty] | 258.8928 | 82.125 | 67.6864 | 328.1725
[11 empty] | 258.9696 | 82.325 | 68.4237 | 344.4162
[11 empty] | 259.0464 | 82.475 | 69.929 | 357.516
[11 empty] | 259.1232 | 82.705 | 70.1235 | 365.1909
[11 empty] | 259.2 | 82.785 | 70.5638 | 367.0711
[11 empty] | 259.2768 | 83.065 | 70.2771 | 364.5128
[11 empty] | 259.3536 | 83.315 | 70.4512 | 359.9202
[11 empty] | 259.4304 | 83.445 | 71.2499 | 355.6666
[11 empty] | 259.5072 | 83.64 | 71.3062 | 353.0466
[11 empty] | 259.584 | 83.555 | 71.4598 | 352.2452
[11 empty] | 259.6608 | 83.315 | 70.9837 | 352.2144
[11 empty] | 259.7376 | 82.86 | 69.7907 | 351.2281
[11 empty] | 259.8144 | 82.37 | 68.0346 | 347.6834
[11 empty] | 259.8912 | 81.815 | 66.263 | 340.964
[11 empty] | 259.968 | 81.12 | 65.2083 | 331.4089
[11 empty] | 260.0448 | 80.71 | 64 | 320.1585
[11 empty] | 260.1216 | 80.635 | 63.9232 | 308.1375
[11 empty] | 260.1984 | 80.765 | 65.2595 | 297.3803
[11 empty] | 260.2752 | 80.855 | 66.6317 | 295.6542
[11 empty] | 260.352 | 81.2 | 65.7408 | 295.2843
[11 empty] | 260.4288 | 81.185 | 64.9728 | 287.7019
[11 empty] | 260.5056 | 80.57 | 65.2032 | 279.7187
[11 empty] | 260.5824 | 79.915 | 65.239 | 272.7219
[11 empty] | 260.6592 | 79.595 | 64.3891 | 266.9272
[11 empty] | 260.736 | 79.56 | 63.3139 | 262.4887
[11 empty] | 260.8128 | 79.42 | 63.5802 | 259.1906
[11 empty] | 260.8896 | 79.415 | 63.7491 | 256.8172
[11 empty] | 260.9664 | 79.385 | 63.6365 | 255.1836
[11 empty] | 261.0432 | 79.185 | 63.7235 | 254.2589
[11 empty] | 261.12 | 79.115 | 62.4998 | 254.0123
[11 empty] | 261.1968 | 78.71 | 61.6038 | 254.1973
[11 empty] | 261.2736 | 78.085 | 60.6618 | 256.0775
[11 empty] | 261.3504 | 77.54 | 59.2896 | 266.8039
[11 empty] | 261.4272 | 77.18 | 57.9072 | 285.421
[11 empty] | 261.504 | 76.95 | 57.2672 | 301.2024
[11 empty] | 261.5808 | 76.95 | 57.2006 | 311.2506
[11 empty] | 261.6576 | 77.055 | 58.0454 | 317.0454
[11 empty] | 261.7344 | 77.33 | 59.5354 | 320.9291
[11 empty] | 261.8112 | 78.205 | 59.6019 | 325.0594
[11 empty] | 261.888 | 78.78 | 60.7283 | 330.7924
[11 empty] | 261.9648 | 79.05 | 62.1363 | 338.1899
[11 empty] | 262.0416 | 79.205 | 62.6022 | 346.1115
[11 empty] | 262.1184 | 79.035 | 63.0477 | 352.8
[11 empty] | 262.1952 | 78.975 | 62.2541 | 356.7146
[11 empty] | 262.272 | 78.775 | 61.6243 | 356.9303
[11 empty] | 262.3488 | 78.57 | 61.0048 | 353.6015
[11 empty] | 262.4256 | 78.495 | 60.5082 | 347.7143
[11 empty] | 262.5024 | 78.595 | 59.6378 | 340.8099
[11 empty] | 262.5792 | 78.53 | 59.3254 | 334.3987
[11 empty] | 262.656 | 78.11 | 59.8426 | 329.3129
[11 empty] | 262.7328 | 77.695 | 59.2896 | 325.5525
[11 empty] | 262.8096 | 77.18 | 57.9533 | 322.1928
[11 empty] | 262.8864 | 76.585 | 56.361 | 317.9701
[11 empty] | 262.9632 | 75.775 | 56.1152 | 312.0521
[11 empty] | 263.04 | 74.965 | 57.4413 | 305.1169
[11 empty] | 263.1168 | 75.165 | 57.303 | 297.2878
[11 empty] | 263.1936 | 75.855 | 56.9037 | 289.3972
[11 empty] | 263.2704 | 76.46 | 57.4618 | 282.7086
[11 empty] | 263.3472 | 76.88 | 58.3731 | 277.5919
[11 empty] | 263.424 | 77.09 | 59.3766 | 274.1398
[11 empty] | 263.5008 | 77.275 | 59.4074 | 272.0438
[11 empty] | 263.5776 | 77.33 | 58.8083 | 270.9958
[11 empty] | 263.6544 | 77.2 | 58.153 | 270.5335
[11 empty] | 263.7312 | 76.74 | 58.6394 | 270.4718
[11 empty] | 263.808 | 76.31 | 59.177 | 270.5335
[11 empty] | 263.8848 | 76.015 | 59.6429 | 270.6876
[11 empty] | 263.9616 | 76.02 | 58.9875 | 271.0575
[11 empty] | 264.0384 | 76.185 | 57.8765 | 271.9205
[11 empty] | 264.1152 | 76.095 | 58.2042 | 274.3247
[11 empty] | 264.192 | 76.195 | 59.1718 | 279.9345
[11 empty] | 264.2688 | 76.7 | 60.1344 | 289.1506
[11 empty] | 264.3456 | 77.285 | 60.6362 | 302.1579
[11 empty] | 264.4224 | 77.615 | 60.887 | 318.0625
[11 empty] | 264.4992 | 77.845 | 61.2506 | 334.522
[11 empty] | 264.576 | 77.99 | 61.8342 | 348.7622
[11 empty] | 264.6528 | 77.995 | 62.2234 | 358.5331
[11 empty] | 264.7296 | 78.02 | 62.3002 | 363.0025
[11 empty] | 264.8064 | 78.32 | 61.7728 | 362.8484
[11 empty] | 264.8832 | 78.67 | 62.2182 | 359.7661
[11 empty] | 264.96 | 78.885 | 63.6314 | 355.9132
[11 empty] | 265.0368 | 79.06 | 64.9011 | 353.0466
[11 empty] | 265.1136 | 79.165 | 65.152 | 352.0295
[11 empty] | 265.1904 | 79.02 | 64.8294 | 352.5843
[11 empty] | 265.2672 | 78.885 | 63.4573 | 353.324
[11 empty] | 265.344 | 78.57 | 62.2592 | 352.5226
[11 empty] | 265.4208 | 78.09 | 61.1994 | 348.978
[11 empty] | 265.4976 | 77.35 | 61.1738 | 342.3819
[11 empty] | 265.5744 | 76.935 | 60.544 | 333.3507
[11 empty] | 265.6512 | 77.005 | 60.4006 | 322.7784
[11 empty] | 265.728 | 77.56 | 60.7539 | 312.5144
[11 empty] | 265.8048 | 78.09 | 62.1005 | 311.2815
[11 empty] | 265.8816 | 78.465 | 63.2934 | 313.9014
[11 empty] | 265.9584 | 78.59 | 63.9078 | 308.8156
[11 empty] | 266.0352 | 78.505 | 64.1587 | 302.651
[11 empty] | 266.112 | 78.515 | 63.1194 | 297.3495
[11 empty] | 266.1888 | 78.62 | 61.9571 | 293.1267
[11 empty] | 266.2656 | 78.645 | 62.1773 | 290.1985
[11 empty] | 266.3424 | 78.735 | 63.3088 | 288.4108
[11 empty] | 266.4192 | 79.14 | 63.7184 | 287.3628
[11 empty] | 266.496 | 79.57 | 64.5325 | 286.808
[11 empty] | 266.5728 | 80.01 | 64.9421 | 286.6847
[11 empty] | 266.6496 | 80.31 | 64.9267 | 286.9313
[11 empty] | 266.7264 | 80.21 | 65.4848 | 287.1779
[11 empty] | 266.8032 | 80.08 | 64.7526 | 287.4245
[11 empty] | 266.88 | 79.68 | 64.6144 | 291.678
[11 empty] | 266.9568 | 79.47 | 63.9027 | 306.843
[11 empty] | 267.0336 | 79.47 | 63.0067 | 325.6142
[11 empty] | 267.1104 | 79.37 | 64.1331 | 339.7928
[11 empty] | 267.1872 | 79.505 | 66.2989 | 348.7314
[11 empty] | 267.264 | 80.29 | 66.9235 | 354.0638
[11 empty] | 267.3408 | 81.19 | 67.9885 | 357.4543
[11 empty] | 267.4176 | 82.01 | 69.6115 | 360.814
[11 empty] | 267.4944 | 82.63 | 71.0093 | 365.1601
[11 empty] | 267.5712 | 82.99 | 72.023 | 370.4616
[11 empty] | 267.648 | 83.095 | 71.9923 | 375.7015
[11 empty] | 267.7248 | 82.97 | 71.6544 | 379.5236
[11 empty] | 267.8016 | 82.905 | 70.2822 | 380.6641
[11 empty] | 267.8784 | 82.695 | 69.6474 | 378.3831
[11 empty] | 267.9552 | 82.345 | 69.6422 | 372.8042
[11 empty] | 268.032 | 82.025 | 69.9443 | 364.8518
[11 empty] | 268.1088 | 81.995 | 69.1558 | 355.9748
[11 empty] | 268.1856 | 82.09 | 67.2205 | 347.5602
[11 empty] | 268.2624 | 81.575 | 67.0515 | 340.4709
[11 empty] | 268.3392 | 80.81 | 66.944 | 334.6453
[11 empty] | 268.416 | 80.16 | 65.2698 | 329.3129
[11 empty] | 268.4928 | 79.445 | 63.5085 | 323.4258
[11 empty] | 268.5696 | 78.705 | 62.9402 | 316.3056
[11 empty] | 268.6464 | 78.56 | 61.7677 | 308.5999
[11 empty] | 268.7232 | 78.75 | 62.1414 | 300.0927
[11 empty] | 268.8 | 79.15 | 63.6928 | 291.3082
[11 empty] | 268.8768 | 79.545 | 65.2544 | 283.5099
[11 empty] | 268.9536 | 80.12 | 64.8858 | 277.0063
[11 empty] | 269.0304 | 80.38 | 64.8141 | 271.7047
[11 empty] | 269.1072 | 80.41 | 64.4198 | 267.5128
[11 empty] | 269.184 | 80.245 | 63.6826 | 264.3997
[11 empty] | 269.2608 | 79.635 | 63.616 | 262.2421
[11 empty] | 269.3376 | 79.025 | 62.9146 | 260.7934
[11 empty] | 269.4144 | 78.725 | 61.4093 | 259.7454
[11 empty] | 269.4912 | 78.355 | 60.841 | 258.9749
[11 empty] | 269.568 | 78.05 | 60.1907 | 258.3584
[11 empty] | 269.6448 | 77.44 | 61.3939 | 258.3276
[11 empty] | 269.7216 | 77.265 | 61.1482 | 260.6085
[11 empty] | 269.7984 | 77.465 | 60.672 | 266.5573
[11 empty] | 269.8752 | 77.685 | 61.1738 | 276.5748
[11 empty] | 269.952 | 77.835 | 61.1789 | 290.6917
[11 empty] | 270.0288 | 77.845 | 60.5286 | 307.3361
[11 empty] | 270.1056 | 77.9 | 59.6429 | 323.5182
[11 empty] | 270.1824 | 77.92 | 59.479 | 336.3098
[11 empty] | 270.2592 | 77.865 | 60.0525 | 344.2004
[11 empty] | 270.336 | 78.055 | 59.7862 | 347.1595
[11 empty] | 270.4128 | 78.065 | 60.9331 | 346.5122
[11 empty] | 270.4896 | 78.165 | 62.2797 | 344.2313
[11 empty] | 270.5664 | 78.65 | 61.7933 | 342.3819
[11 empty] | 270.6432 | 78.84 | 61.8854 | 341.8887
[11 empty] | 270.72 | 78.725 | 62.0339 | 342.7518
[11 empty] | 270.7968 | 78.46 | 61.4912 | 343.9539
[11 empty] | 270.8736 | 77.985 | 60.9075 | 344.0463
[11 empty] | 270.9504 | 77.355 | 60.4928 | 341.5805
[11 empty] | 271.0272 | 76.875 | 58.9005 | 336.0632
[11 empty] | 271.104 | 76.1 | 58.6496 | 327.9259
[11 empty] | 271.1808 | 75.74 | 57.687 | 318.1858
[11 empty] | 271.2576 | 75.615 | 58.2195 | 307.8601
[11 empty] | 271.3344 | 75.84 | 59.2486 | 300.9558
[11 empty] | 271.4112 | 76.355 | 59.0336 | 302.4969
[11 empty] | 271.488 | 76.635 | 58.8493 | 299.9694
[11 empty] | 271.5648 | 76.66 | 58.2963 | 293.5583
[11 empty] | 271.6416 | 76.285 | 57.943 | 287.8252
[11 empty] | 271.7184 | 75.68 | 58.4294 | 283.1709
[11 empty] | 271.7952 | 75.405 | 58.0608 | 279.6263
[11 empty] | 271.872 | 75.355 | 57.7178 | 277.407
[11 empty] | 271.9488 | 75.575 | 57.3901 | 276.2974
[11 empty] | 272.0256 | 75.725 | 58.0659 | 275.8967
[11 empty] | 272.1024 | 75.905 | 58.6189 | 276.0508
[11 empty] | 272.1792 | 76.2 | 58.4499 | 276.5748
[11 empty] | 272.256 | 75.93 | 59.5712 | 277.2529
[11 empty] | 272.3328 | 75.695 | 59.0387 | 278.0543
[11 empty] | 272.4096 | 75.435 | 57.7485 | 283.0168
[11 empty] | 272.4864 | 75.185 | 56.6528 | 299.353
[11 empty] | 272.5632 | 75.055 | 55.7773 | 319.6037
[11 empty] | 272.64 | 74.99 | 56.4019 | 334.7994
[11 empty] | 272.7168 | 75.295 | 57.1136 | 343.8614
[11 empty] | 272.7936 | 75.755 | 58.5574 | 348.4232
[11 empty] | 272.8704 | 76.47 | 60.2419 | 350.8582
[11 empty] | 272.9472 | 77.605 | 60.0576 | 353.7248
[11 empty] | 273.024 | 78.26 | 61.2096 | 358.4099
[11 empty] | 273.1008 | 78.505 | 63.017 | 364.9443
[11 empty] | 273.1776 | 78.85 | 62.889 | 372.0028
[11 empty] | 273.2544 | 78.865 | 62.9043 | 377.5818
[11 empty] | 273.3312 | 78.52 | 63.6006 | 380.0168
[11 empty] | 273.408 | 78.245 | 63.8413 | 378.3831
[11 empty] | 273.4848 | 78.215 | 63.1808 | 372.9275
[11 empty] | 273.5616 | 78.12 | 63.2064 | 365.1293
[11 empty] | 273.6384 | 78.11 | 63.017 | 356.7454
[11 empty] | 273.7152 | 77.96 | 63.3805 | 349.4403
[11 empty] | 273.792 | 77.88 | 62.5869 | 343.9539
[11 empty] | 273.8688 | 77.555 | 62.0698 | 340.1935
[11 empty] | 273.9456 | 77.105 | 61.3018 | 336.957
[11 empty] | 274.0224 | 76.935 | 59.0234 | 332.6726
[11 empty] | 274.0992 | 76.72 | 57.5795 | 326.8163
[11 empty] | 274.176 | 76.875 | 56.919 | 320.2202
[11 empty] | 274.2528 | 77.27 | 58.8134 | 312.4528
[11 empty] | 274.3296 | 77.995 | 60.5594 | 304.8703
[11 empty] | 274.4064 | 78.96 | 61.3222 | 298.5516
[11 empty] | 274.4832 | 79.735 | 62.3053 | 293.5891
[11 empty] | 274.56 | 80.08 | 64.0973 | 289.7979
[11 empty] | 274.6368 | 80.28 | 65.321 | 286.9621
[11 empty] | 274.7136 | 80.35 | 65.4797 | 284.9586
[11 empty] | 274.7904 | 80.32 | 65.2493 | 283.6332
[11 empty] | 274.8672 | 80.31 | 64.9984 | 282.7702
[11 empty] | 274.944 | 80.23 | 65.0701 | 282.1229
[11 empty] | 275.0208 | 80.15 | 65.3414 | 281.5373
[11 empty] | 275.0976 | 80.23 | 65.449 | 280.8283
[11 empty] | 275.1744 | 80.365 | 65.3363 | 280.3352
[11 empty] | 275.2512 | 80.455 | 65.5718 | 281.7222
[11 empty] | 275.328 | 80.955 | 65.6077 | 286.6847
[11 empty] | 275.4048 | 81.565 | 66.4678 | 295.7467
[11 empty] | 275.4816 | 81.9 | 67.5277 | 309.1547
[11 empty] | 275.5584 | 81.925 | 68.4749 | 325.3676
[11 empty] | 275.6352 | 81.975 | 69.6218 | 341.5805
[11 empty] | 275.712 | 82.4 | 69.9187 | 354.7111
[11 empty] | 275.7888 | 82.84 | 69.8368 | 362.5093
[11 empty] | 275.8656 | 83.235 | 69.0586 | 364.5744
[11 empty] | 275.9424 | 83.28 | 69.929 | 362.3552
[11 empty] | 276.0192 | 83.395 | 70.7072 | 358.1324
[11 empty] | 276.096 | 83.585 | 71.2243 | 354.0946
[11 empty] | 276.1728 | 83.71 | 71.5264 | 351.598
[11 empty] | 276.2496 | 83.725 | 71.1834 | 350.5192
[11 empty] | 276.3264 | 83.435 | 70.7277 | 350.026
[11 empty] | 276.4032 | 82.94 | 70.016 | 348.5773
[11 empty] | 276.48 | 82.295 | 69.0944 | 344.971
[11 empty] | 276.5568 | 81.765 | 67.113 | 338.6831
[11 empty] | 276.6336 | 81.14 | 65.6794 | 329.8677
[11 empty] | 276.7104 | 80.59 | 65.0394 | 319.3571
[11 empty] | 276.7872 | 80.54 | 64.6246 | 307.7368
[11 empty] | 276.864 | 80.995 | 64.2355 | 296.209
[11 empty] | 276.9408 | 81.24 | 64.9574 | 291.7089
[11 empty] | 277.0176 | 81.27 | 65.6538 | 292.4795
[11 empty] | 277.0944 | 81.15 | 65.6026 | 286.4998
[11 empty] | 277.1712 | 80.9 | 64.768 | 278.9482
[11 empty] | 277.248 | 80.555 | 63.7184 | 272.0438
[11 empty] | 277.3248 | 80.225 | 62.8326 | 265.9717
[11 empty] | 277.4016 | 79.84 | 63.2371 | 261.0092
[11 empty] | 277.4784 | 79.715 | 63.0938 | 257.2796
[11 empty] | 277.5552 | 79.58 | 63.5802 | 254.6904
[11 empty] | 277.632 | 79.71 | 62.848 | 253.1493
[11 empty] | 277.7088 | 79.615 | 62.4998 | 252.4712
[11 empty] | 277.7856 | 79.19 | 63.2013 | 252.4712
[11 empty] | 277.8624 | 78.79 | 62.505 | 252.7178
[11 empty] | 277.9392 | 78.52 | 60.1293 | 254.2897
[11 empty] | 278.016 | 77.925 | 58.9926 | 264.3072
[11 empty] | 278.0928 | 77.08 | 59.3306 | 282.5544
[11 empty] | 278.1696 | 76.76 | 58.5318 | 298.4283
[11 empty] | 278.2464 | 77.12 | 56.8371 | 308.6924
[11 empty] | 278.3232 | 77.375 | 57.2314 | 314.7028
[11 empty] | 278.4 | 77.445 | 59.8938 | 318.7098
[11 empty] | 278.4768 | 77.955 | 61.5424 | 322.8093
[11 empty] | 278.5536 | 78.59 | 62.3514 | 328.5115
[11 empty] | 278.6304 | 79.02 | 63.0067 | 335.8474
[11 empty] | 278.7072 | 79.365 | 62.464 | 343.7997
[11 empty] | 278.784 | 79.495 | 61.2915 | 350.6116
[11 empty] | 278.8608 | 79.265 | 61.2147 | 354.5878
[11 empty] | 278.9376 | 78.745 | 62.0749 | 354.896
[11 empty] | 279.0144 | 78.455 | 62.0083 | 351.5671
[11 empty] | 279.0912 | 78.31 | 62.0083 | 345.6491
[11 empty] | 279.168 | 78.075 | 62.4998 | 338.8064
[11 empty] | 279.2448 | 78.05 | 61.8701 | 332.5802
[11 empty] | 279.3216 | 78.175 | 59.8272 | 327.8026
[11 empty] | 279.3984 | 77.865 | 58.7366 | 324.258
[11 empty] | 279.4752 | 77.03 | 58.8339 | 321.0215
[11 empty] | 279.552 | 76.135 | 58.5165 | 316.7063
[11 empty] | 279.6288 | 75.795 | 56.1971 | 310.4801
[11 empty] | 279.7056 | 75.69 | 54.5434 | 303.0826
[11 empty] | 279.7824 | 75.885 | 54.4563 | 295.0686
[11 empty] | 279.8592 | 76.015 | 56.4122 | 287.2087
[11 empty] | 279.936 | 76.29 | 58.24 | 280.8283
[11 empty] | 280.0128 | 76.965 | 58.24 | 276.1432
[11 empty] | 280.0896 | 77.455 | 58.2605 | 272.9068
[11 empty] | 280.1664 | 77.535 | 59.0643 | 270.7492
[11 empty] | 280.2432 | 77.415 | 59.3715 | 269.393
[11 empty] | 280.32 | 77.205 | 58.88 | 268.5608
[11 empty] | 280.3968 | 76.92 | 57.9891 | 268.2217
[11 empty] | 280.4736 | 76.65 | 57.6307 | 268.3759
[11 empty] | 280.5504 | 76.495 | 57.5488 | 268.869
[11 empty] | 280.6272 | 76.425 | 57.5693 | 269.6088
[11 empty] | 280.704 | 76.22 | 58.327 | 270.5643
[11 empty] | 280.7808 | 76.23 | 58.4243 | 272.7527
[11 empty] | 280.8576 | 76.38 | 59.1514 | 277.8694
[11 empty] | 280.9344 | 76.865 | 59.8477 | 286.5923
[11 empty] | 281.0112 | 77.35 | 60.3802 | 299.4146
[11 empty] | 281.088 | 77.62 | 60.8051 | 315.4426
[11 empty] | 281.1648 | 77.79 | 61.824 | 332.3028
[11 empty] | 281.2416 | 77.855 | 63.1552 | 346.9437
[11 empty] | 281.3184 | 78.065 | 62.4845 | 356.8995
[11 empty] | 281.3952 | 78.33 | 61.5219 | 361.2148
[11 empty] | 281.472 | 78.385 | 62.1261 | 360.7524
[11 empty] | 281.5488 | 78.54 | 63.104 | 357.4852
[11 empty] | 281.6256 | 78.885 | 63.831 | 353.7864
[11 empty] | 281.7024 | 79.28 | 64.2458 | 351.3206
[11 empty] | 281.7792 | 79.435 | 64.727 | 350.8274
[11 empty] | 281.856 | 79.365 | 64.8192 | 351.6288
[11 empty] | 281.9328 | 79.1 | 64.1946 | 352.2761
[11 empty] | 282.0096 | 78.685 | 63.1245 | 351.1356
[11 empty] | 282.0864 | 78.1 | 62.3462 | 347.1903
[11 empty] | 282.1632 | 77.54 | 61.4861 | 340.3784
[11 empty] | 282.24 | 77.05 | 61.0816 | 331.5322
[11 empty] | 282.3168 | 77.045 | 61.0406 | 321.4531
[11 empty] | 282.3936 | 77.505 | 61.783 | 311.6514
[11 empty] | 282.4704 | 78.09 | 62.5818 | 309.8328
[11 empty] | 282.5472 | 78.665 | 62.4691 | 311.9288
[11 empty] | 282.624 | 78.905 | 62.5357 | 306.8738
[11 empty] | 282.7008 | 79.03 | 61.8086 | 300.74
[11 empty] | 282.7776 | 78.695 | 62.4794 | 295.5001
[11 empty] | 282.8544 | 78.345 | 63.2627 | 291.4315
[11 empty] | 282.9312 | 78.295 | 63.7645 | 288.5649
[11 empty] | 283.008 | 78.455 | 64.3328 | 286.8388
[11 empty] | 283.0848 | 78.745 | 65.1725 | 285.8525
[11 empty] | 283.1616 | 79.445 | 64.6912 | 285.3593
[11 empty] | 283.2384 | 79.955 | 64.6605 | 285.2977
[11 empty] | 283.3152 | 80.08 | 65.3722 | 285.6059
[11 empty] | 283.392 | 80.105 | 65.5002 | 286.0374
[11 empty] | 283.4688 | 79.95 | 65.4131 | 286.3765
[11 empty] | 283.5456 | 79.635 | 65.0854 | 290.0136
[11 empty] | 283.6224 | 79.41 | 64.4096 | 304.3463
[11 empty] | 283.6992 | 79.425 | 63.5648 | 323.2716
[11 empty] | 283.776 | 79.77 | 62.9197 | 338.005
[11 empty] | 283.8528 | 80.125 | 63.6621 | 347.3444
[11 empty] | 283.9296 | 80.425 | 66.0429 | 352.7692
[11 empty] | 284.0064 | 81.06 | 68.3366 | 356.0981
[11 empty] | 284.0832 | 81.84 | 70.3744 | 359.3962
[11 empty] | 284.16 | 82.53 | 71.7414 | 363.8347
[11 empty] | 284.2368 | 83.13 | 71.2499 | 369.4445
[11 empty] | 284.3136 | 83.42 | 70.656 | 375.2084
[11 empty] | 284.3904 | 83.255 | 70.8659 | 379.3695
[11 empty] | 284.4672 | 83 | 70.5331 | 380.5408
[11 empty] | 284.544 | 82.57 | 70.5997 | 377.9825
[11 empty] | 284.6208 | 82.4 | 69.9034 | 372.0336
[11 empty] | 284.6976 | 82.09 | 70.3232 | 363.8039
[11 empty] | 284.7744 | 81.945 | 70.057 | 355.081
[11 empty] | 284.8512 | 81.94 | 68.6131 | 347.2519
[11 empty] | 284.928 | 81.51 | 68.1779 | 340.7791
[11 empty] | 285.0048 | 80.83 | 67.6915 | 335.3851
[11 empty] | 285.0816 | 80.265 | 65.4848 | 329.991
[11 empty] | 285.1584 | 79.48 | 63.9539 | 323.4874
[11 empty] | 285.2352 | 78.685 | 63.1245 | 315.5351
[11 empty] | 285.312 | 78.485 | 62.1363 | 307.0587
[11 empty] | 285.3888 | 78.655 | 62.4486 | 298.1201
[11 empty] | 285.4656 | 78.97 | 63.6621 | 289.4896
[11 empty] | 285.5424 | 79.635 | 63.6723 | 282.1846
[11 empty] | 285.6192 | 80.095 | 64.1331 | 276.2049
[11 empty] | 285.696 | 80.21 | 65.0138 | 271.3657
[11 empty] | 285.7728 | 80.085 | 65.4643 | 267.3895
[11 empty] | 285.8496 | 79.785 | 65.321 | 264.1839
[11 empty] | 285.9264 | 79.46 | 64.0768 | 261.7489
[11 empty] | 286.0032 | 79 | 62.9811 | 260.1769
[11 empty] | 286.08 | 78.625 | 61.6141 | 259.2523
[11 empty] | 286.1568 | 78.155 | 61.271 | 258.8207
[11 empty] | 286.2336 | 77.905 | 60.3136 | 258.7283
[11 empty] | 286.3104 | 77.8 | 59.3562 | 258.9749
[11 empty] | 286.3872 | 77.225 | 60.7232 | 261.0092
[11 empty] | 286.464 | 77.085 | 61.8291 | 266.2799
[11 empty] | 286.5408 | 77.61 | 61.1277 | 275.4651
[11 empty] | 286.6176 | 77.76 | 61.6909 | 288.9348
[11 empty] | 286.6944 | 77.705 | 61.7421 | 305.4559
[11 empty] | 286.7712 | 77.66 | 61.7216 | 322.0695
[11 empty] | 286.848 | 77.775 | 61.5526 | 335.7549
[11 empty] | 286.9248 | 77.555 | 62.5459 | 344.3546
[11 empty] | 287.0016 | 77.63 | 62.4486 | 347.7451
[11 empty] | 287.0784 | 77.995 | 61.7421 | 347.3136
[11 empty] | 287.1552 | 78.165 | 62.7354 | 345.0327
[11 empty] | 287.232 | 78.655 | 62.249 | 342.8442
[11 empty] | 287.3088 | 78.845 | 62.464 | 342.2278
[11 empty] | 287.3856 | 78.635 | 63.5238 | 343.1525
[11 empty] | 287.4624 | 78.315 | 63.5443 | 344.447
[11 empty] | 287.5392 | 77.855 | 62.7558 | 344.5703
[11 empty] | 287.616 | 77.45 | 61.1328 | 342.1353
[11 empty] | 287.6928 | 76.955 | 59.6429 | 336.6796
[11 empty] | 287.7696 | 76.24 | 59.1411 | 328.604
[11 empty] | 287.8464 | 75.875 | 58.025 | 318.9872
[11 empty] | 287.9232 | 75.88 | 57.8048 | 308.8465
[11 empty] | 288 | 76.01 | 59.1462 | 301.8188
[11 empty] | 288.0768 | 76.365 | 59.7094 | 303.1134
[11 empty] | 288.1536 | 76.56 | 59.9245 | 300.5551
[11 empty] | 288.2304 | 76.535 | 59.3664 | 293.774
[11 empty] | 288.3072 | 76.23 | 58.6394 | 287.8252
[11 empty] | 288.384 | 75.875 | 57.7997 | 283.2634
[11 empty] | 288.4608 | 75.625 | 57.3235 | 280.1194
[11 empty] | 288.5376 | 75.52 | 57.3952 | 278.2701
[11 empty] | 288.6144 | 75.985 | 56.0435 | 277.2837
[11 empty] | 288.6912 | 76.18 | 56.7245 | 276.7289
[11 empty] | 288.768 | 76.27 | 57.7075 | 276.4515
[11 empty] | 288.8448 | 76.3 | 58.1325 | 276.5748
[11 empty] | 288.9216 | 76 | 59.136 | 277.1296
[11 empty] | 288.9984 | 75.605 | 59.5661 | 278.1468
[11 empty] | 289.0752 | 75.35 | 58.583 | 282.6469
[11 empty] | 289.152 | 75.09 | 57.6512 | 298.1509
[11 empty] | 289.2288 | 74.825 | 57.431 | 318.679
[11 empty] | 289.3056 | 74.82 | 57.4822 | 334.4604
[11 empty] | 289.3824 | 75.17 | 57.6256 | 344.0463
[11 empty] | 289.4592 | 75.765 | 58.1734 | 348.7931
[11 empty] | 289.536 | 76.705 | 58.3629 | 351.1048
[11 empty] | 289.6128 | 77.525 | 59.735 | 353.6631
[11 empty] | 289.6896 | 78 | 62.0954 | 358.1016
[11 empty] | 289.7664 | 78.52 | 62.8685 | 364.6669
[11 empty] | 289.8432 | 78.545 | 64.6963 | 371.972
[11 empty] | 289.92 | 78.45 | 65.2237 | 377.9208
[11 empty] | 289.9968 | 78.46 | 64.1843 | 380.7874
[11 empty] | 290.0736 | 78.345 | 63.6109 | 379.4619
[11 empty] | 290.1504 | 78.215 | 63.3856 | 374.1912
[11 empty] | 290.2272 | 78.16 | 63.2781 | 366.3005
[11 empty] | 290.304 | 78.385 | 62.1466 | 357.6393
[11 empty] | 290.3808 | 78.27 | 62.4742 | 350.1801
[11 empty] | 290.4576 | 78.005 | 62.6944 | 344.6936
[11 empty] | 290.5344 | 77.495 | 63.0733 | 341.0257
[11 empty] | 290.6112 | 77.115 | 62.0083 | 338.005
[11 empty] | 290.688 | 76.895 | 59.6531 | 333.998
[11 empty] | 290.7648 | 76.685 | 57.5744 | 328.3574
[11 empty] | 290.8416 | 76.715 | 56.9498 | 321.6996
[11 empty] | 290.9184 | 77.035 | 59.0899 | 313.9323
[11 empty] | 290.9952 | 77.655 | 61.4298 | 306.2265
[11 empty] | 291.072 | 78.49 | 62.9965 | 299.7228
[11 empty] | 291.1488 | 79.195 | 64.3891 | 294.5754
[11 empty] | 291.2256 | 79.655 | 65.7357 | 290.6301
[11 empty] | 291.3024 | 80.145 | 65.5206 | 287.7635
[11 empty] | 291.3792 | 80.42 | 65.0445 | 285.76
[11 empty] | 291.456 | 80.21 | 65.7562 | 284.4038
[11 empty] | 291.5328 | 80.1 | 65.3568 | 283.4791
[11 empty] | 291.6096 | 80.17 | 64.8141 | 282.7394
[11 empty] | 291.6864 | 80.295 | 64.4762 | 282.0305
[11 empty] | 291.7632 | 80.355 | 64.7424 | 281.1982
[11 empty] | 291.84 | 80.305 | 65.536 | 280.6742
[11 empty] | 291.9168 | 80.46 | 65.2698 | 281.8763
[11 empty] | 291.9936 | 80.94 | 64.9062 | 286.3765
[11 empty] | 292.0704 | 81.305 | 66.7443 | 294.822
[11 empty] | 292.1472 | 81.455 | 69.1661 | 307.5211
[11 empty] | 292.224 | 81.74 | 69.1046 | 323.2716
[11 empty] | 292.3008 | 82.2 | 68.2854 | 339.4845
[11 empty] | 292.3776 | 82.565 | 68.6131 | 353.1699
[11 empty] | 292.4544 | 82.75 | 69.033 | 361.862
[11 empty] | 292.5312 | 82.77 | 69.6474 | 364.821
[11 empty] | 292.608 | 82.655 | 71.0963 | 363.1258
[11 empty] | 292.6848 | 82.95 | 70.9325 | 358.9647
[11 empty] | 292.7616 | 83.26 | 70.6406 | 354.4953
[11 empty] | 292.8384 | 83.385 | 71.0605 | 351.5055
[11 empty] | 292.9152 | 83.215 | 71.9411 | 350.3959
[11 empty] | 292.992 | 83.055 | 71.4086 | 350.3342
[11 empty] | 293.0688 | 82.845 | 69.9699 | 349.8411
[11 empty] | 293.1456 | 82.31 | 68.6336 | 347.0978
[11 empty] | 293.2224 | 81.6 | 67.5686 | 341.149
[11 empty] | 293.2992 | 80.835 | 66.2426 | 332.1795
[11 empty] | 293.376 | 80.245 | 65.1418 | 321.1448
[11 empty] | 293.4528 | 80.14 | 64.4301 | 308.9698
[11 empty] | 293.5296 | 80.415 | 64.5837 | 297.257
[11 empty] | 293.6064 | 80.62 | 65.6998 | 293.8048
[11 empty] | 293.6832 | 80.94 | 65.5974 | 295.346
[11 empty] | 293.76 | 81.145 | 64.2304 | 288.8423
[11 empty] | 293.8368 | 80.84 | 63.4317 | 280.6434
[11 empty] | 293.9136 | 80.225 | 63.4829 | 273.3076
[11 empty] | 293.9904 | 79.645 | 63.7133 | 267.0813
[11 empty] | 294.0672 | 79.42 | 63.0989 | 262.3037
[11 empty] | 294.144 | 79.42 | 62.4128 | 258.8824
[11 empty] | 294.2208 | 79.49 | 62.1926 | 256.4782
[11 empty] | 294.2976 | 79.415 | 62.6381 | 254.8137
[11 empty] | 294.3744 | 79.27 | 62.8429 | 253.7966
[11 empty] | 294.4512 | 79.095 | 62.0595 | 253.3651
[11 empty] | 294.528 | 78.565 | 61.8803 | 253.2726
[11 empty] | 294.6048 | 78.075 | 60.7386 | 254.5363
[11 empty] | 294.6816 | 77.685 | 58.5728 | 263.9682
[11 empty] | 294.7584 | 77.105 | 57.6461 | 282.277
[11 empty] | 294.8352 | 76.815 | 57.047 | 298.7057
[11 empty] | 294.912 | 76.9 | 56.663 | 309.4013
[11 empty] | 294.9888 | 77.1 | 57.4464 | 315.5351
[11 empty] | 295.0656 | 77.5 | 58.5421 | 319.4804
[11 empty] | 295.1424 | 78.08 | 59.5046 | 323.4566
[11 empty] | 295.2192 | 78.56 | 60.6566 | 329.0047
[11 empty] | 295.296 | 78.87 | 61.9469 | 336.3714
[11 empty] | 295.3728 | 79.045 | 62.6534 | 344.447
[11 empty] | 295.4496 | 78.97 | 62.8173 | 351.413
[11 empty] | 295.5264 | 78.84 | 62.3206 | 355.6358
[11 empty] | 295.6032 | 78.78 | 61.3478 | 356.2831
[11 empty] | 295.68 | 78.525 | 61.1738 | 353.2932
[11 empty] | 295.7568 | 78.49 | 60.5184 | 347.6526
[11 empty] | 295.8336 | 78.42 | 60.1139 | 340.9024
[11 empty] | 295.9104 | 78.15 | 60.6566 | 334.522
[11 empty] | 295.9872 | 77.87 | 60.9024 | 329.3438
[11 empty] | 296.064 | 77.285 | 61.1226 | 325.4601
[11 empty] | 296.1408 | 76.87 | 59.1258 | 322.0695
[11 empty] | 296.2176 | 76.425 | 56.6733 | 317.8776
[11 empty] | 296.2944 | 75.9 | 54.9786 | 312.0829
[11 empty] | 296.3712 | 75.525 | 54.231 | 305.2094
[11 empty] | 296.448 | 75.595 | 54.7123 | 297.6577
[11 empty] | 296.5248 | 76.125 | 55.296 | 289.8903
[11 empty] | 296.6016 | 76.56 | 56.448 | 283.2325
[11 empty] | 296.6784 | 76.875 | 58.1734 | 277.9618
[11 empty] | 296.7552 | 77.145 | 59.2794 | 274.2014
[11 empty] | 296.832 | 77.215 | 59.9245 | 271.7972
[11 empty] | 296.9088 | 77.045 | 60.5901 | 270.5335
[11 empty] | 296.9856 | 77.01 | 59.7658 | 270.0711
[11 empty] | 297.0624 | 76.855 | 58.9722 | 270.0403
[11 empty] | 297.1392 | 76.57 | 59.0746 | 270.2253
[11 empty] | 297.216 | 76.45 | 58.624 | 270.441
[11 empty] | 297.2928 | 76.56 | 57.687 | 270.6568
[11 empty] | 297.3696 | 76.605 | 57.047 | 271.0883
[11 empty] | 297.4464 | 76.36 | 57.8048 | 272.9377
[11 empty] | 297.5232 | 76.46 | 58.496 | 277.9618
[11 empty] | 297.6 | 76.83 | 59.8374 | 286.7155
[11 empty] | 297.6768 | 77.32 | 60.7693 | 299.4146
[11 empty] | 297.7536 | 77.69 | 60.8205 | 315.1652
[11 empty] | 297.8304 | 78.03 | 60.5696 | 331.6863
[11 empty] | 297.9072 | 78.26 | 60.7539 | 346.1423
[11 empty] | 297.984 | 78.405 | 60.3597 | 356.2831
[11 empty] | 298.0608 | 78.465 | 60.3546 | 361.1839
[11 empty] | 298.1376 | 78.255 | 62.0186 | 361.4922
[11 empty] | 298.2144 | 78.395 | 63.447 | 358.903
[11 empty] | 298.2912 | 78.94 | 63.575 | 355.3892
[11 empty] | 298.368 | 79.43 | 63.4266 | 352.5535
[11 empty] | 298.4448 | 79.585 | 63.7798 | 351.2589
[11 empty] | 298.5216 | 79.385 | 64.3635 | 351.3514
[11 empty] | 298.5984 | 79.21 | 63.5802 | 351.6904
[11 empty] | 298.6752 | 79.015 | 61.5782 | 350.889
[11 empty] | 298.752 | 78.44 | 60.4774 | 347.6526
[11 empty] | 298.8288 | 77.635 | 60.6157 | 341.6113
[11 empty] | 298.9056 | 77.295 | 59.6429 | 333.1966
[11 empty] | 298.9824 | 77.26 | 59.2896 | 323.1175
[11 empty] | 299.0592 | 77.51 | 60.4211 | 312.4219
[11 empty] | 299.136 | 78.01 | 61.6499 | 308.3225
[11 empty] | 299.2128 | 78.395 | 62.7712 | 311.4972
[11 empty] | 299.2896 | 78.77 | 62.336 | 307.6752
[11 empty] | 299.3664 | 78.665 | 62.7302 | 301.4798
[11 empty] | 299.4432 | 78.395 | 63.1398 | 296.0549
[11 empty] | 299.52 | 78.255 | 62.7405 | 291.678
[11 empty] | 299.5968 | 78.335 | 62.2848 | 288.5341
[11 empty] | 299.6736 | 78.49 | 62.8787 | 286.5923
[11 empty] | 299.7504 | 78.86 | 63.5699 | 285.421
[11 empty] | 299.8272 | 79.24 | 64.6605 | 284.7429
[11 empty] | 299.904 | 79.49 | 65.8586 | 284.5579
[11 empty] | 299.9808 | 79.765 | 65.8586 | 284.7737
[11 empty] | 300.0576 | 79.92 | 65.408 | 285.0511
[11 empty] | 300.1344 | 79.73 | 65.4387 | 285.2361
[11 empty] | 300.2112 | 79.48 | 64.9984 | 288.1334
[11 empty] | 300.288 | 79.155 | 64.4966 | 301.3873
[11 empty] | 300.3648 | 79.14 | 63.447 | 320.3434
[11 empty] | 300.4416 | 79.28 | 63.2627 | 335.5392
[11 empty] | 300.5184 | 79.56 | 64.0922 | 345.3101
[11 empty] | 300.5952 | 80.14 | 65.3056 | 351.1664
[11 empty] | 300.672 | 81.075 | 66.2835 | 354.9268
[11 empty] | 300.7488 | 81.865 | 68.4186 | 358.4407
[11 empty] | 300.8256 | 82.425 | 70.2515 | 362.7559
[11 empty] | 300.9024 | 82.895 | 70.825 | 367.9958
[11 empty] | 300.9792 | 83.175 | 70.3078 | 373.3282
[11 empty] | 301.056 | 83.105 | 69.7907 | 377.3352
[11 empty] | 301.1328 | 82.855 | 69.0534 | 378.753
[11 empty] | 301.2096 | 82.515 | 68.3162 | 376.9345
[11 empty] | 301.2864 | 82.165 | 68.2906 | 371.8795
[11 empty] | 301.3632 | 81.925 | 68.7514 | 364.5128
[11 empty] | 301.44 | 81.855 | 68.2547 | 356.0673
[11 empty] | 301.5168 | 81.59 | 68.137 | 347.8067
[11 empty] | 301.5936 | 81.095 | 68.265 | 340.5017
[11 empty] | 301.6704 | 80.82 | 66.1146 | 334.2754
[11 empty] | 301.7472 | 80.3 | 63.9386 | 328.5424
[11 empty] | 301.824 | 79.71 | 61.7421 | 322.4394
[11 empty] | 301.9008 | 79.05 | 60.6464 | 315.3501
[11 empty] | 301.9776 | 78.59 | 60.3136 | 307.7368
[11 empty] | 302.0544 | 78.485 | 61.4246 | 299.7537
[11 empty] | 302.1312 | 79.02 | 62.0646 | 291.3082
[11 empty] | 302.208 | 79.54 | 63.0272 | 283.6024
[11 empty] | 302.2848 | 79.88 | 64.1229 | 277.0371
[11 empty] | 302.3616 | 80.145 | 64.5018 | 271.6123
[11 empty] | 302.4384 | 80.205 | 64.3277 | 267.2046
[11 empty] | 302.5152 | 79.97 | 63.8874 | 263.9065
[11 empty] | 302.592 | 79.345 | 64.0102 | 261.6564
[11 empty] | 302.6688 | 79.035 | 62.3258 | 260.2694
[11 empty] | 302.7456 | 78.675 | 61.097 | 259.4372
[11 empty] | 302.8224 | 78.265 | 60.6208 | 258.8516
[11 empty] | 302.8992 | 77.945 | 60.1651 | 258.2659
[11 empty] | 302.976 | 77.6 | 60.1139 | 258.0193
[11 empty] | 303.0528 | 77.13 | 60.969 | 259.5913
[11 empty] | 303.1296 | 77.43 | 59.7606 | 264.7079
[11 empty] | 303.2064 | 77.685 | 60.2061 | 274.0473
[11 empty] | 303.2832 | 77.66 | 61.5475 | 287.7327
[11 empty] | 303.36 | 77.685 | 60.8512 | 304.2847
[11 empty] | 303.4368 | 77.72 | 60.375 | 320.7441
[11 empty] | 303.5136 | 77.625 | 61.4707 | 334.1521
[11 empty] | 303.5904 | 77.645 | 61.8035 | 342.6285
[11 empty] | 303.6672 | 77.765 | 61.481 | 346.0498
[11 empty] | 303.744 | 77.975 | 61.3939 | 345.8032
[11 empty] | 303.8208 | 78.445 | 60.8922 | 343.8306
[11 empty] | 303.8976 | 78.595 | 62.039 | 342.0428
[11 empty] | 303.9744 | 78.59 | 63.1603 | 341.4264
[11 empty] | 304.0512 | 78.7 | 62.6176 | 342.2278
[11 empty] | 304.128 | 78.615 | 61.737 | 343.399
[11 empty] | 304.2048 | 78.37 | 60.5389 | 343.4915
[11 empty] | 304.2816 | 77.635 | 60.2061 | 341.2722
[11 empty] | 304.3584 | 76.865 | 59.8477 | 336.2173
[11 empty] | 304.4352 | 76.24 | 59.0694 | 328.5424
[11 empty] | 304.512 | 75.975 | 57.4669 | 319.1722
[11 empty] | 304.5888 | 76.01 | 57.0112 | 308.9389
[11 empty] | 304.6656 | 76.105 | 58.4448 | 300.74
[11 empty] | 304.7424 | 76.445 | 58.8595 | 301.0174
[11 empty] | 304.8192 | 76.7 | 59.049 | 299.4763
[11 empty] | 304.896 | 76.585 | 59.2077 | 293.1576
[11 empty] | 304.9728 | 76.135 | 59.3562 | 287.4245
[11 empty] | 305.0496 | 75.995 | 57.687 | 282.801
[11 empty] | 305.1264 | 75.675 | 57.3235 | 279.318
[11 empty] | 305.2032 | 75.58 | 57.1955 | 276.9446
[11 empty] | 305.28 | 75.635 | 57.4925 | 275.5268
[11 empty] | 305.3568 | 75.535 | 59.5149 | 274.8795
[11 empty] | 305.4336 | 75.93 | 59.351 | 274.8487
[11 empty] | 305.5104 | 76.345 | 58.4346 | 275.3419
[11 empty] | 305.5872 | 76.425 | 58.0915 | 276.1741
[11 empty] | 305.664 | 76.145 | 58.2758 | 277.0988
[11 empty] | 305.7408 | 75.785 | 57.6768 | 281.0441
[11 empty] | 305.8176 | 75.445 | 56.7654 | 295.7775
[11 empty] | 305.8944 | 75.165 | 56.4992 | 316.1207
[11 empty] | 305.9712 | 75.085 | 56.8218 | 332.0254
[11 empty] | 306.048 | 75.385 | 57.0982 | 341.7654
[11 empty] | 306.1248 | 75.865 | 58.1837 | 346.6971
[11 empty] | 306.2016 | 76.64 | 59.3459 | 349.0705
[11 empty] | 306.2784 | 77.365 | 61.184 | 351.7213
[11 empty] | 306.3552 | 78.155 | 62.208 | 356.2831
[11 empty] | 306.432 | 78.79 | 62.505 | 362.7559
[11 empty] | 306.5088 | 78.78 | 64.0819 | 369.876
[11 empty] | 306.5856 | 78.83 | 63.7082 | 375.7632
[11 empty] | 306.6624 | 78.84 | 62.5203 | 378.5373
[11 empty] | 306.7392 | 78.495 | 63.0477 | 377.3352
[11 empty] | 306.816 | 78.32 | 63.4368 | 372.3418
[11 empty] | 306.8928 | 78.345 | 63.017 | 364.6977
[11 empty] | 306.9696 | 78.295 | 63.104 | 356.4372
[11 empty] | 307.0464 | 78.025 | 64.1997 | 349.1013
[11 empty] | 307.1232 | 77.685 | 64.983 | 343.5223
[11 empty] | 307.2 | 77.53 | 63.744 | 339.6695
[11 empty] | 307.2768 | 77.03 | 63.1398 | 336.3406
[11 empty] | 307.3536 | 76.615 | 61.8035 | 332.1178
[11 empty] | 307.4304 | 76.365 | 60.2522 | 326.3231
[11 empty] | 307.5072 | 76.43 | 59.4534 | 319.5729
[11 empty] | 307.584 | 76.805 | 60.7898 | 311.7747
[11 empty] | 307.6608 | 77.565 | 61.8803 | 304.0689
[11 empty] | 307.7376 | 78.515 | 62.633 | 297.5652
[11 empty] | 307.8144 | 79.36 | 63.6109 | 292.4178
[11 empty] | 307.8912 | 79.945 | 64.5837 | 288.5341
[11 empty] | 307.968 | 80.27 | 65.0547 | 285.6367
[11 empty] | 308.0448 | 80.41 | 65.1878 | 283.5716
[11 empty] | 308.1216 | 80.595 | 64.4915 | 282.2462
[11 empty] | 308.1984 | 80.55 | 64.3584 | 281.3832
[11 empty] | 308.2752 | 80.28 | 64.983 | 280.7667
[11 empty] | 308.352 | 80.435 | 63.9949 | 280.2119
[11 empty] | 308.4288 | 80.35 | 64.5683 | 279.6571
[11 empty] | 308.5056 | 80.165 | 66.0326 | 279.2256
[11 empty] | 308.5824 | 80.32 | 66.0787 | 280.3352
[11 empty] | 308.6592 | 80.695 | 66.4166 | 284.712
[11 empty] | 308.736 | 81.225 | 67.287 | 293.0034
[11 empty] | 308.8128 | 81.84 | 67.3997 | 305.7333
[11 empty] | 308.8896 | 82.135 | 67.7427 | 321.7305
[11 empty] | 308.9664 | 82.21 | 69.12 | 338.3133
[11 empty] | 309.0432 | 82.505 | 69.6883 | 352.3069
[11 empty] | 309.12 | 82.93 | 69.0381 | 361.2456
[11 empty] | 309.1968 | 83.12 | 68.9971 | 364.3895
[11 empty] | 309.2736 | 83.135 | 70.0877 | 362.8175
[11 empty] | 309.3504 | 83.455 | 70.4358 | 358.6873
[11 empty] | 309.4272 | 83.68 | 71.04 | 354.4953
[11 empty] | 309.504 | 83.8 | 71.3523 | 351.7213
[11 empty] | 309.5808 | 83.735 | 71.5878 | 350.6425
[11 empty] | 309.6576 | 83.345 | 72.0179 | 350.3342
[11 empty] | 309.7344 | 83.275 | 69.5859 | 349.317
[11 empty] | 309.8112 | 82.835 | 68.0909 | 346.1731
[11 empty] | 309.888 | 82.035 | 67.4816 | 340.0393
[11 empty] | 309.9648 | 81.445 | 65.9149 | 331.1315
[11 empty] | 310.0416 | 80.925 | 64.7526 | 320.5284
[11 empty] | 310.1184 | 80.595 | 65.2493 | 309.0622
[11 empty] | 310.1952 | 80.745 | 66.0173 | 297.442
[11 empty] | 310.272 | 81.135 | 65.92 | 291.6472
[11 empty] | 310.3488 | 81.345 | 65.7459 | 292.1404
[11 empty] | 310.4256 | 81.215 | 65.8074 | 286.3765
[11 empty] | 310.5024 | 80.875 | 65.1725 | 278.7016
[11 empty] | 310.5792 | 80.48 | 63.9181 | 271.828
[11 empty] | 310.656 | 80.34 | 62.0902 | 265.91
[11 empty] | 310.7328 | 80.135 | 61.8854 | 261.04
[11 empty] | 310.8096 | 79.915 | 62.5818 | 257.3412
[11 empty] | 310.8864 | 79.8 | 62.9453 | 254.6288
[11 empty] | 310.9632 | 79.805 | 62.7046 | 252.8411
[11 empty] | 311.04 | 79.72 | 62.6278 | 251.9164
[11 empty] | 311.1168 | 79.545 | 61.9418 | 251.7623
[11 empty] | 311.1936 | 79.255 | 60.8512 | 252.0397
[11 empty] | 311.2704 | 78.78 | 59.8528 | 253.1801
[11 empty] | 311.3472 | 78.205 | 58.6496 | 261.2866
[11 empty] | 311.424 | 77.51 | 58.112 | 278.979
[11 empty] | 311.5008 | 76.895 | 58.4192 | 295.7159
[11 empty] | 311.5776 | 76.92 | 57.897 | 306.8121
[11 empty] | 311.6544 | 77.1 | 58.3219 | 313.3158
[11 empty] | 311.7312 | 77.47 | 59.3408 | 317.5077
[11 empty] | 311.808 | 78.02 | 60.6771 | 321.5147
[11 empty] | 311.8848 | 78.595 | 61.7574 | 326.9396
[11 empty] | 311.9616 | 78.81 | 63.3446 | 334.1213
[11 empty] | 312.0384 | 79.215 | 62.7712 | 342.1353
[11 empty] | 312.1152 | 79.38 | 62.2182 | 349.3479
[11 empty] | 312.192 | 79.105 | 62.633 | 353.9713
[11 empty] | 312.2688 | 79.01 | 61.7574 | 355.0501
[11 empty] | 312.3456 | 78.755 | 61.6858 | 352.3377
[11 empty] | 312.4224 | 78.685 | 60.9587 | 346.7588
[11 empty] | 312.4992 | 78.565 | 60.969 | 339.8852
[11 empty] | 312.576 | 78.425 | 61.0304 | 333.3199
[11 empty] | 312.6528 | 78.425 | 59.7811 | 328.08
[11 empty] | 312.7296 | 78.015 | 59.4022 | 324.2888
[11 empty] | 312.8064 | 77.435 | 58.5728 | 321.0215
[11 empty] | 312.8832 | 76.785 | 57.5949 | 316.9837
[11 empty] | 312.96 | 76.17 | 56.4941 | 311.1273
[11 empty] | 313.0368 | 75.71 | 55.8182 | 303.9456
[11 empty] | 313.1136 | 75.71 | 56.0896 | 295.9624
[11 empty] | 313.1904 | 76.195 | 56.4378 | 287.7635
[11 empty] | 313.2672 | 76.6 | 57.5795 | 280.8283
[11 empty] | 313.344 | 76.97 | 58.6445 | 275.6809
[11 empty] | 313.4208 | 77.38 | 58.8237 | 272.1363
[11 empty] | 313.4976 | 77.485 | 59.2486 | 269.8862
[11 empty] | 313.5744 | 77.26 | 60.0883 | 268.5916
[11 empty] | 313.6512 | 76.86 | 61.1686 | 267.8827
[11 empty] | 313.728 | 76.85 | 60.0115 | 267.5436
[11 empty] | 313.8048 | 76.87 | 58.3731 | 267.5745
[11 empty] | 313.8816 | 76.775 | 57.6819 | 267.9135
[11 empty] | 313.9584 | 76.66 | 57.4157 | 268.53
[11 empty] | 314.0352 | 76.535 | 57.5181 | 269.4238
[11 empty] | 314.112 | 76.44 | 57.7434 | 271.3965
[11 empty] | 314.1888 | 76.715 | 57.6614 | 276.1124
[11 empty] | 314.2656 | 77.265 | 58.0864 | 284.2497
[11 empty] | 314.3424 | 77.585 | 59.4688 | 296.2399
[11 empty] | 314.4192 | 77.655 | 60.9126 | 311.6514
[11 empty] | 314.496 | 77.955 | 60.9075 | 328.3882
[11 empty] | 314.5728 | 78.215 | 60.9638 | 343.584
[11 empty] | 314.6496 | 78.355 | 60.7181 | 354.557
[11 empty] | 314.7264 | 78.33 | 61.2301 | 360.0435
[11 empty] | 314.8032 | 78.49 | 61.4349 | 360.475
[11 empty] | 314.88 | 78.51 | 63.0118 | 357.6085
[11 empty] | 314.9568 | 78.63 | 64.681 | 353.7248
[11 empty] | 315.0336 | 79.11 | 64.6758 | 350.8274
[11 empty] | 315.1104 | 79.5 | 64.2355 | 349.9027
[11 empty] | 315.1872 | 79.535 | 64.0256 | 350.6116
[11 empty] | 315.264 | 79.22 | 63.8208 | 351.598
[11 empty] | 315.3408 | 78.7 | 63.3395 | 351.1664
[11 empty] | 315.4176 | 78.285 | 61.7984 | 347.93
[11 empty] | 315.4944 | 78.1 | 59.1667 | 341.5188
[11 empty] | 315.5712 | 77.655 | 58.4294 | 332.7343
[11 empty] | 315.648 | 76.955 | 61.0509 | 322.6243
[11 empty] | 315.7248 | 77.16 | 62.4896 | 312.3295
[11 empty] | 315.8016 | 77.925 | 62.6381 | 309.0622
[11 empty] | 315.8784 | 78.36 | 63.8874 | 312.0521
[11 empty] | 315.9552 | 78.71 | 63.6877 | 307.7677
[11 empty] | 316.032 | 78.68 | 63.6774 | 301.4489
[11 empty] | 316.1088 | 78.375 | 64.0819 | 295.9624
[11 empty] | 316.1856 | 78.505 | 62.592 | 291.6164
[11 empty] | 316.2624 | 78.635 | 61.952 | 288.5649
[11 empty] | 316.3392 | 78.67 | 63.145 | 286.6847
[11 empty] | 316.416 | 78.945 | 64.256 | 285.5751
[11 empty] | 316.4928 | 79.445 | 64.6349 | 284.9586
[11 empty] | 316.5696 | 79.8 | 65.2902 | 284.7737
[11 empty] | 316.6464 | 80.075 | 65.408 | 284.9895
[11 empty] | 316.7232 | 79.935 | 66.2374 | 285.3593
[11 empty] | 316.8 | 79.915 | 65.5718 | 285.6059
[11 empty] | 316.8768 | 79.695 | 64.9626 | 288.1642
[11 empty] | 316.9536 | 79.46 | 64.3277 | 300.74
[11 empty] | 317.0304 | 79.46 | 63.2627 | 319.727
[11 empty] | 317.1072 | 79.595 | 63.104 | 335.2618
[11 empty] | 317.184 | 79.81 | 64.4403 | 345.4025
[11 empty] | 317.2608 | 80.37 | 65.7818 | 351.4438
[11 empty] | 317.3376 | 81.08 | 67.5277 | 355.081
[11 empty] | 317.4144 | 81.855 | 69.3094 | 358.2866
[11 empty] | 317.4912 | 82.42 | 71.3472 | 362.386
[11 empty] | 317.568 | 82.89 | 72.1408 | 367.7492
[11 empty] | 317.6448 | 83.165 | 71.8029 | 373.5131
[11 empty] | 317.7216 | 83.13 | 71.6237 | 378.1982
[11 empty] | 317.7984 | 83.145 | 70.2925 | 380.2325
[11 empty] | 317.8752 | 82.925 | 69.5501 | 378.6914
[11 empty] | 317.952 | 82.43 | 70.0262 | 373.4823
[11 empty] | 318.0288 | 82.125 | 70.1594 | 365.5916
[11 empty] | 318.1056 | 81.96 | 70.1491 | 356.6529
[11 empty] | 318.1824 | 81.895 | 69.3658 | 348.2383
[11 empty] | 318.2592 | 81.625 | 68.7616 | 341.3339
[11 empty] | 318.336 | 81.155 | 67.625 | 335.8166
[11 empty] | 318.4128 | 80.425 | 66.2579 | 330.7308
[11 empty] | 318.4896 | 79.605 | 64.6605 | 324.8744
[11 empty] | 318.5664 | 79.01 | 62.8378 | 317.5385
[11 empty] | 318.6432 | 78.66 | 62.0339 | 309.3396
[11 empty] | 318.72 | 78.59 | 63.1603 | 300.4626
[11 empty] | 318.7968 | 78.89 | 64.6656 | 291.3698
[11 empty] | 318.8736 | 79.41 | 65.2902 | 283.4483
[11 empty] | 318.9504 | 80.105 | 64.6656 | 276.9446
[11 empty] | 319.0272 | 80.4 | 65.1008 | 271.7972
[11 empty] | 319.104 | 80.295 | 65.664 | 267.6669
[11 empty] | 319.1808 | 80.01 | 65.4541 | 264.4922
[11 empty] | 319.2576 | 79.745 | 63.9846 | 262.088
[11 empty] | 319.3344 | 79.31 | 62.7251 | 260.3619
[11 empty] | 319.4112 | 78.71 | 62.4435 | 259.2214
[11 empty] | 319.488 | 78.2 | 62.2182 | 258.5125
[11 empty] | 319.5648 | 77.89 | 61.7318 | 258.2043
[11 empty] | 319.6416 | 77.65 | 61.0867 | 258.2659
[11 empty] | 319.7184 | 77.205 | 61.6858 | 259.8995
[11 empty] | 319.7952 | 77.34 | 60.887 | 264.7696
[11 empty] | 319.872 | 77.75 | 60.6976 | 273.4925
[11 empty] | 319.9488 | 77.94 | 61.2813 | 286.4381
[11 empty] | 320.0256 | 77.875 | 61.2762 | 302.4969
[11 empty] | 320.1024 | 77.825 | 61.1226 | 319.203
[11 empty] | 320.1792 | 77.875 | 61.5117 | 333.3816
[11 empty] | 320.256 | 77.965 | 61.481 | 342.9059
[11 empty] | 320.3328 | 78.09 | 60.9894 | 347.2211
[11 empty] | 320.4096 | 78.255 | 60.9024 | 347.1903
[11 empty] | 320.4864 | 78.4 | 61.609 | 345.0327
[11 empty] | 320.5632 | 78.445 | 62.7917 | 342.8134
[11 empty] | 320.64 | 78.485 | 64.0461 | 341.8887
[11 empty] | 320.7168 | 78.63 | 63.7747 | 342.536
[11 empty] | 320.7936 | 78.53 | 63.1142 | 344.0463
[11 empty] | 320.8704 | 78.02 | 63.1194 | 344.632
[11 empty] | 320.9472 | 77.43 | 62.4128 | 342.8751
[11 empty] | 321.024 | 76.665 | 61.9674 | 337.8817
[11 empty] | 321.1008 | 76.175 | 60.4826 | 329.991
[11 empty] | 321.1776 | 75.945 | 58.4397 | 320.3434
[11 empty] | 321.2544 | 75.845 | 58.0301 | 310.0486
[11 empty] | 321.3312 | 76.19 | 58.3424 | 301.7572
[11 empty] | 321.408 | 76.6 | 58.6752 | 302.1887
[11 empty] | 321.4848 | 76.455 | 60.6976 | 301.264
[11 empty] | 321.5616 | 76.365 | 60.8205 | 294.6679
[11 empty] | 321.6384 | 76.355 | 59.0592 | 288.4108
[11 empty] | 321.7152 | 76.175 | 57.4106 | 283.3558
[11 empty] | 321.792 | 75.725 | 57.4208 | 279.7495
[11 empty] | 321.8688 | 75.625 | 57.2211 | 277.5303
[11 empty] | 321.9456 | 75.67 | 57.5642 | 276.3898
[11 empty] | 322.0224 | 75.88 | 58.1734 | 275.8658
[11 empty] | 322.0992 | 76.05 | 59.0234 | 275.6809
[11 empty] | 322.176 | 76.25 | 59.305 | 275.8042
[11 empty] | 322.2528 | 76.505 | 58.24 | 276.2357
[11 empty] | 322.3296 | 76.43 | 57.2416 | 276.9446
[11 empty] | 322.4064 | 75.885 | 57.303 | 280.2427
[11 empty] | 322.4832 | 75.265 | 57.6614 | 293.8665
[11 empty] | 322.56 | 74.99 | 57.1904 | 314.4254
[11 empty] | 322.6368 | 75.035 | 56.6835 | 331.1315
[11 empty] | 322.7136 | 75.22 | 57.4822 | 341.6113
[11 empty] | 322.7904 | 75.92 | 57.687 | 347.0978
[11 empty] | 322.8672 | 76.64 | 59.0234 | 349.7177
[11 empty] | 322.944 | 77.275 | 61.0714 | 352.1528
[11 empty] | 323.0208 | 77.86 | 62.6893 | 356.2214
[11 empty] | 323.0976 | 78.505 | 62.9555 | 362.3552
[11 empty] | 323.1744 | 78.8 | 63.5341 | 369.5061
[11 empty] | 323.2512 | 78.935 | 63.2934 | 375.7632
[11 empty] | 323.328 | 78.775 | 63.36 | 379.3078
[11 empty] | 323.4048 | 78.61 | 62.8941 | 378.8147
[11 empty] | 323.4816 | 78.415 | 62.7917 | 374.3453
[11 empty] | 323.5584 | 78.33 | 62.592 | 366.917
[11 empty] | 323.6352 | 78.3 | 62.4179 | 358.4099
[11 empty] | 323.712 | 78.1 | 63.0989 | 350.5808
[11 empty] | 323.7888 | 77.8 | 63.6467 | 344.6936
[11 empty] | 323.8656 | 77.72 | 62.3258 | 340.6558
[11 empty] | 323.9424 | 77.54 | 60.5747 | 337.6043
[11 empty] | 324.0192 | 77.045 | 59.8016 | 333.9055
[11 empty] | 324.096 | 76.25 | 60.5952 | 328.6348
[11 empty] | 324.1728 | 76.14 | 60.2368 | 322.3161
[11 empty] | 324.2496 | 76.845 | 59.9245 | 314.7953
[11 empty] | 324.3264 | 77.6 | 61.2198 | 306.9971
[11 empty] | 324.4032 | 78.475 | 62.4486 | 300.2777
[11 empty] | 324.48 | 79.335 | 63.2013 | 294.8836
[11 empty] | 324.5568 | 79.865 | 64.4659 | 290.7842
[11 empty] | 324.6336 | 80.195 | 64.9984 | 287.7327
[11 empty] | 324.7104 | 80.285 | 65.7357 | 285.5443
[11 empty] | 324.7872 | 80.22 | 66.0634 | 284.0956
[11 empty] | 324.864 | 80.365 | 64.855 | 283.1709
[11 empty] | 324.9408 | 80.355 | 64.3328 | 282.5236
[11 empty] | 325.0176 | 80.025 | 65.8227 | 281.8763
[11 empty] | 325.0944 | 80.16 | 65.5053 | 281.075
[11 empty] | 325.1712 | 80.445 | 64.7066 | 280.3968
[11 empty] | 325.248 | 80.53 | 64.7322 | 281.1366
[11 empty] | 325.3248 | 80.885 | 64.855 | 285.0819
[11 empty] | 325.4016 | 81.355 | 66.0531 | 292.9726
[11 empty] | 325.4784 | 81.745 | 67.241 | 305.1477
[11 empty] | 325.5552 | 82.015 | 67.7581 | 320.7441
[11 empty] | 325.632 | 81.98 | 69.3197 | 337.2653
[11 empty] | 325.7088 | 82.26 | 70.0314 | 351.598
[11 empty] | 325.7856 | 82.68 | 69.6627 | 361.1839
[11 empty] | 325.8624 | 82.87 | 69.8522 | 365.006
[11 empty] | 325.9392 | 82.86 | 70.784 | 363.9888
[11 empty] | 326.016 | 83.165 | 70.743 | 360.2284
[11 empty] | 326.0928 | 83.365 | 71.255 | 355.8824
[11 empty] | 326.1696 | 83.39 | 72.3405 | 352.5843
[11 empty] | 326.2464 | 83.55 | 71.6698 | 351.074
[11 empty] | 326.3232 | 83.605 | 70.4717 | 350.6733
[11 empty] | 326.4 | 83.205 | 70.0262 | 350.026
[11 empty] | 326.4768 | 82.61 | 69.1968 | 347.6526
[11 empty] | 326.5536 | 81.83 | 68.4442 | 342.3819
[11 empty] | 326.6304 | 81.135 | 67.1795 | 334.1213
[11 empty] | 326.7072 | 80.585 | 65.5514 | 323.5799
[11 empty] | 326.784 | 80.405 | 64.8858 | 311.6822
[11 empty] | 326.8608 | 80.66 | 65.1469 | 299.353
[11 empty] | 326.9376 | 80.86 | 66.3757 | 292.6644
[11 empty] | 327.0144 | 80.98 | 67.0464 | 294.7603
[11 empty] | 327.0912 | 81.19 | 65.9046 | 290.3218
[11 empty] | 327.168 | 80.92 | 65.0906 | 282.277
[11 empty] | 327.2448 | 80.445 | 64.0102 | 274.7254
[11 empty] | 327.3216 | 80.25 | 62.1107 | 268.1293
[11 empty] | 327.3984 | 79.95 | 62.2643 | 262.7969
[11 empty] | 327.4752 | 79.75 | 62.679 | 258.9132
[11 empty] | 327.552 | 79.635 | 63.0118 | 256.2316
[11 empty] | 327.6288 | 79.72 | 62.761 | 254.4747
[11 empty] | 327.7056 | 79.675 | 62.7712 | 253.4267
[11 empty] | 327.7824 | 79.385 | 62.8326 | 252.9644
[11 empty] | 327.8592 | 78.99 | 62.3462 | 252.8102
[11 empty] | 327.936 | 78.355 | 61.9059 | 253.4267
[11 empty] | 328.0128 | 77.78 | 60.6771 | 260.6701
[11 empty] | 328.0896 | 77.22 | 59.6173 | 277.9002
[11 empty] | 328.1664 | 76.8 | 58.7981 | 294.9761
[11 empty] | 328.2432 | 76.65 | 58.6752 | 306.5656
[11 empty] | 328.32 | 76.74 | 59.6531 | 313.3158
[11 empty] | 328.3968 | 77.3 | 59.8733 | 317.5694
[11 empty] | 328.4736 | 78.125 | 60.073 | 321.4839
[11 empty] | 328.5504 | 78.665 | 61.3222 | 326.7546
[11 empty] | 328.6272 | 79.085 | 62.1414 | 333.7823
[11 empty] | 328.704 | 79.33 | 62.3258 | 341.7654
[11 empty] | 328.7808 | 79.37 | 62.0646 | 349.0705
[11 empty] | 328.8576 | 79.27 | 61.4963 | 353.9097
[11 empty] | 328.9344 | 79.01 | 61.1686 | 355.2043
[11 empty] | 329.0112 | 78.735 | 60.9638 | 352.8309
[11 empty] | 329.088 | 78.645 | 60.2266 | 347.5602
[11 empty] | 329.1648 | 78.49 | 60.3546 | 340.964
[11 empty] | 329.2416 | 78.1 | 61.655 | 334.522
[11 empty] | 329.3184 | 77.925 | 61.3683 | 329.1896
[11 empty] | 329.3952 | 77.615 | 60.6054 | 325.0902
[11 empty] | 329.472 | 77.14 | 59.2845 | 321.5455
[11 empty] | 329.5488 | 76.525 | 58.0096 | 317.3536
[11 empty] | 329.6256 | 75.805 | 57.3338 | 311.5589
[11 empty] | 329.7024 | 75.495 | 55.7875 | 304.5004
[11 empty] | 329.7792 | 75.655 | 54.9274 | 296.918
[11 empty] | 329.856 | 76.09 | 55.4957 | 288.9656
[11 empty] | 329.9328 | 76.65 | 56.2381 | 282.0613
[11 empty] | 330.0096 | 77.115 | 57.0931 | 276.6981
[11 empty] | 330.0864 | 77.33 | 58.5011 | 272.8452
[11 empty] | 330.1632 | 77.475 | 58.9517 | 270.1944
[11 empty] | 330.24 | 77.405 | 58.9722 | 268.6533
[11 empty] | 330.3168 | 77.065 | 59.3818 | 267.8827
[11 empty] | 330.3936 | 76.765 | 59.264 | 267.7594
[11 empty] | 330.4704 | 76.55 | 58.9005 | 268.0368
[11 empty] | 330.5472 | 76.395 | 58.496 | 268.4375
[11 empty] | 330.624 | 76.215 | 58.5677 | 268.8998
[11 empty] | 330.7008 | 76.05 | 59.2179 | 269.4547
[11 empty] | 330.7776 | 76.085 | 59.0336 | 270.965
[11 empty] | 330.8544 | 76.415 | 58.496 | 275.311
[11 empty] | 330.9312 | 76.745 | 59.8221 | 283.2634
[11 empty] | 331.008 | 77.32 | 60.375 | 295.2227
[11 empty] | 331.0848 | 77.575 | 60.8973 | 310.665
[11 empty] | 331.1616 | 77.58 | 62.0288 | 327.4327
[11 empty] | 331.2384 | 77.745 | 62.9862 | 342.5668
[11 empty] | 331.3152 | 78.035 | 62.2797 | 353.5398
[11 empty] | 331.392 | 78.285 | 61.1994 | 359.2729
[11 empty] | 331.4688 | 78.48 | 61.143 | 360.1359
[11 empty] | 331.5456 | 78.645 | 62.0134 | 357.7934
[11 empty] | 331.6224 | 78.695 | 64.2253 | 354.372
[11 empty] | 331.6992 | 79.065 | 64.8602 | 351.6596
[11 empty] | 331.776 | 79.31 | 64.9882 | 350.4883
[11 empty] | 331.8528 | 79.31 | 64.9728 | 350.7041
[11 empty] | 331.9296 | 79.195 | 63.8106 | 351.2281
[11 empty] | 332.0064 | 78.86 | 62.5408 | 350.6116
[11 empty] | 332.0832 | 78.365 | 61.2198 | 347.591
[11 empty] | 332.16 | 77.75 | 60.1395 | 341.8579
[11 empty] | 332.2368 | 77.235 | 59.1565 | 333.7823
[11 empty] | 332.3136 | 77.06 | 59.0336 | 324.1039
[11 empty] | 332.3904 | 77.475 | 59.5661 | 313.624
[11 empty] | 332.4672 | 78 | 61.0202 | 308.0143
[11 empty] | 332.544 | 78.455 | 61.9059 | 310.8191
[11 empty] | 332.6208 | 78.86 | 61.8854 | 307.7985
[11 empty] | 332.6976 | 78.77 | 62.3565 | 301.5722
[11 empty] | 332.7744 | 78.435 | 63.063 | 296.0549
[11 empty] | 332.8512 | 78.395 | 62.2592 | 291.5547
[11 empty] | 332.928 | 78.495 | 61.8291 | 288.2567
[11 empty] | 333.0048 | 78.62 | 62.7046 | 286.1916
[11 empty] | 333.0816 | 78.84 | 63.8925 | 284.9895
[11 empty] | 333.1584 | 79.425 | 63.5802 | 284.3422
[11 empty] | 333.2352 | 79.81 | 64.1229 | 284.1264
[11 empty] | 333.312 | 79.98 | 65.1571 | 284.3422
[11 empty] | 333.3888 | 80.165 | 65.0138 | 284.712
[11 empty] | 333.4656 | 80.095 | 64.8243 | 284.9278
[11 empty] | 333.5424 | 79.945 | 64.087 | 286.9621
[11 empty] | 333.6192 | 79.585 | 63.8925 | 298.6132
[11 empty] | 333.696 | 79.35 | 63.5904 | 317.6002
[11 empty] | 333.7728 | 79.35 | 63.9181 | 333.5049
[11 empty] | 333.8496 | 79.82 | 63.8413 | 343.8922
[11 empty] | 333.9264 | 80.34 | 64.9626 | 350.0568
[11 empty] | 334.0032 | 80.975 | 67.0669 | 353.8481
[11 empty] | 334.08 | 81.71 | 69.2531 | 357.2694
[11 empty] | 334.1568 | 82.46 | 70.5331 | 361.523
[11 empty] | 334.2336 | 83.195 | 70.3027 | 366.8553
[11 empty] | 334.3104 | 83.46 | 70.1747 | 372.3418
[11 empty] | 334.3872 | 83.285 | 70.487 | 376.5954
[11 empty] | 334.464 | 82.875 | 71.1117 | 378.2907
[11 empty] | 334.5408 | 82.645 | 70.3693 | 376.6879
[11 empty] | 334.6176 | 82.44 | 69.7702 | 371.8178
[11 empty] | 334.6944 | 82.23 | 69.6474 | 364.5436
[11 empty] | 334.7712 | 82.1 | 69.2531 | 356.2522
[11 empty] | 334.848 | 81.905 | 68.7002 | 348.2074
[11 empty] | 334.9248 | 81.615 | 67.7427 | 341.149
[11 empty] | 335.0016 | 80.96 | 67.5174 | 335.0152
[11 empty] | 335.0784 | 80.415 | 65.6538 | 329.1588
[11 empty] | 335.1552 | 79.89 | 63.2013 | 322.8093
[11 empty] | 335.232 | 79.445 | 60.9178 | 315.4426
[11 empty] | 335.3088 | 78.71 | 61.4451 | 307.3978
[11 empty] | 335.3856 | 78.635 | 61.8291 | 299.1989
[11 empty] | 335.4624 | 79.03 | 62.3923 | 290.7534
[11 empty] | 335.5392 | 79.465 | 63.6621 | 283.1093
[11 empty] | 335.616 | 80.02 | 64.1382 | 276.5748
[11 empty] | 335.6928 | 80.27 | 64.7834 | 271.1191
[11 empty] | 335.7696 | 80.185 | 65.5411 | 266.619
[11 empty] | 335.8464 | 80.105 | 64.5171 | 263.0435
[11 empty] | 335.9232 | 79.755 | 63.616 | 260.516
[11 empty] | 336 | 79.165 | 63.2115 | 258.9749
[11 empty] | 336.0768 | 78.7 | 62.1158 | 258.1118
[11 empty] | 336.1536 | 78.53 | 60.3802 | 257.6803
[11 empty] | 336.2304 | 78.265 | 59.5917 | 257.372
[11 empty] | 336.3072 | 77.89 | 59.8835 | 257.2179
[11 empty] | 336.384 | 77.63 | 60.0064 | 258.42
[11 empty] | 336.4608 | 77.61 | 60.1549 | 262.7661
[11 empty] | 336.5376 | 77.815 | 60.7898 | 271.1191
[11 empty] | 336.6144 | 78.14 | 60.5235 | 283.8798
[11 empty] | 336.6912 | 78.05 | 60.631 | 299.9386
[11 empty] | 336.768 | 77.875 | 60.7488 | 316.6755
[11 empty] | 336.8448 | 77.825 | 61.5526 | 331.0082
[11 empty] | 336.9216 | 77.735 | 62.5152 | 340.5941
[11 empty] | 336.9984 | 77.88 | 61.993 | 344.7552
[11 empty] | 337.0752 | 78.24 | 61.0406 | 344.8477
[11 empty] | 337.152 | 78.675 | 60.5184 | 342.8442
[11 empty] | 337.2288 | 78.725 | 61.8701 | 340.8099
[11 empty] | 337.3056 | 78.715 | 63.1347 | 339.8852
[11 empty] | 337.3824 | 78.96 | 62.2694 | 340.5325
[11 empty] | 337.4592 | 78.69 | 62.336 | 341.7654
[11 empty] | 337.536 | 78.415 | 61.1789 | 342.197
[11 empty] | 337.6128 | 78.085 | 59.3664 | 340.5017
[11 empty] | 337.6896 | 77.43 | 58.4448 | 335.8166
[11 empty] | 337.7664 | 76.685 | 58.2912 | 328.4807
[11 empty] | 337.8432 | 75.98 | 58.5626 | 319.3879
[11 empty] | 337.92 | 75.69 | 59.177 | 309.3705
[11 empty] | 337.9968 | 75.92 | 59.7914 | 300.4626
[11 empty] | 338.0736 | 76.505 | 59.5046 | 299.5687
[11 empty] | 338.1504 | 76.95 | 59.0797 | 298.9214
[11 empty] | 338.2272 | 76.865 | 59.2538 | 292.6952
[11 empty] | 338.304 | 76.575 | 58.7366 | 286.6847
[11 empty] | 338.3808 | 76.135 | 58.153 | 281.8455
[11 empty] | 338.4576 | 75.95 | 57.0829 | 278.2701
[11 empty] | 338.5344 | 75.775 | 56.9549 | 275.8967
[11 empty] | 338.6112 | 75.7 | 57.6 | 274.3864
[11 empty] | 338.688 | 75.87 | 58.0966 | 273.585
[11 empty] | 338.7648 | 76.035 | 58.9261 | 273.3384
[11 empty] | 338.8416 | 76.155 | 59.4278 | 273.6774
[11 empty] | 338.9184 | 76.38 | 58.5574 | 274.4788
[11 empty] | 338.9952 | 76.445 | 57.1904 | 275.496
[11 empty] | 339.072 | 76.13 | 56.8422 | 278.5166
[11 empty] | 339.1488 | 75.435 | 57.8765 | 291.2157
[11 empty] | 339.2256 | 75.015 | 58.1632 | 311.5281
[11 empty] | 339.3024 | 75.245 | 56.7552 | 328.5115
[11 empty] | 339.3792 | 75.705 | 55.8387 | 339.2996
[11 empty] | 339.456 | 76.155 | 56.7347 | 345.0018
[11 empty] | 339.5328 | 76.475 | 59.561 | 347.7451
[11 empty] | 339.6096 | 77.175 | 61.4861 | 350.2418
[11 empty] | 339.6864 | 78.005 | 62.4845 | 354.3104
[11 empty] | 339.7632 | 78.4 | 64.0256 | 360.4442
[11 empty] | 339.84 | 78.75 | 64.1229 | 367.5951
[11 empty] | 339.9168 | 78.93 | 63.6416 | 373.8214
[11 empty] | 339.9936 | 78.93 | 62.8122 | 377.366
[11 empty] | 340.0704 | 78.72 | 62.5459 | 377.0269
[11 empty] | 340.1472 | 78.51 | 62.6893 | 372.7117
[11 empty] | 340.224 | 78.255 | 63.1962 | 365.4991
[11 empty] | 340.3008 | 78.175 | 63.447 | 357.1461
[11 empty] | 340.3776 | 78.39 | 62.5254 | 349.5328
[11 empty] | 340.4544 | 78.33 | 62.2592 | 343.6765
[11 empty] | 340.5312 | 78.01 | 61.9008 | 339.6386
[11 empty] | 340.608 | 77.605 | 60.7283 | 336.4639
[11 empty] | 340.6848 | 77.14 | 59.6634 | 332.611
[11 empty] | 340.7616 | 76.65 | 58.9773 | 327.0629
[11 empty] | 340.8384 | 76.36 | 59.1053 | 320.4667
[11 empty] | 340.9152 | 76.755 | 59.8835 | 312.8535
[11 empty] | 340.992 | 77.7 | 60.2624 | 304.9011
[11 empty] | 341.0688 | 78.48 | 61.7114 | 298.0584
[11 empty] | 341.1456 | 79.1 | 63.6467 | 292.6028
[11 empty] | 341.2224 | 79.805 | 64.0256 | 288.38
[11 empty] | 341.2992 | 80.255 | 64.4198 | 285.2361
[11 empty] | 341.376 | 80.42 | 64.7782 | 282.986
[11 empty] | 341.4528 | 80.395 | 64.7424 | 281.4756
[11 empty] | 341.5296 | 80.305 | 64.2611 | 280.5509
[11 empty] | 341.6064 | 80.195 | 63.9642 | 279.9037
[11 empty] | 341.6832 | 80.13 | 64.0666 | 279.3488
[11 empty] | 341.76 | 80.205 | 64.3072 | 278.7632
[11 empty] | 341.8368 | 80.28 | 64.7117 | 278.2701
[11 empty] | 341.9136 | 80.39 | 64.6707 | 278.9173
[11 empty] | 341.9904 | 80.73 | 64.5683 | 282.5236
[11 empty] | 342.0672 | 81.13 | 65.9098 | 290.0444
[11 empty] | 342.144 | 81.63 | 66.8109 | 301.9729
[11 empty] | 342.2208 | 82.03 | 67.0208 | 317.4153
[11 empty] | 342.2976 | 82.12 | 68.0141 | 334.1213
[11 empty] | 342.3744 | 82.4 | 68.649 | 348.9472
[11 empty] | 342.4512 | 82.675 | 68.8742 | 359.1496
[11 empty] | 342.528 | 82.88 | 68.7974 | 363.434
[11 empty] | 342.6048 | 83.07 | 69.1712 | 362.6326
[11 empty] | 342.6816 | 83.08 | 70.6099 | 358.8105
[11 empty] | 342.7584 | 83.21 | 71.5315 | 354.4337
[11 empty] | 342.8352 | 83.535 | 71.3779 | 351.3206
[11 empty] | 342.912 | 83.635 | 71.1526 | 349.9335
[11 empty] | 342.9888 | 83.465 | 70.8557 | 349.7177
[11 empty] | 343.0656 | 83.15 | 70.0826 | 349.2246
[11 empty] | 343.1424 | 82.675 | 68.9818 | 346.9129
[11 empty] | 343.2192 | 81.9 | 68.4851 | 341.4572
[11 empty] | 343.296 | 81.42 | 66.2989 | 333.0425
[11 empty] | 343.3728 | 80.95 | 64.4198 | 322.5935
[11 empty] | 343.4496 | 80.845 | 63.2064 | 310.9424
[11 empty] | 343.5264 | 81.005 | 63.8362 | 299.0139
[11 empty] | 343.6032 | 81.11 | 65.7101 | 292.3561
[11 empty] | 343.68 | 81.365 | 65.8176 | 293.4658
[11 empty] | 343.7568 | 81.415 | 65.3517 | 288.2875
[11 empty] | 343.8336 | 81.23 | 64.425 | 280.1502
[11 empty] | 343.9104 | 80.845 | 63.5341 | 272.8452
[11 empty] | 343.9872 | 80.515 | 62.464 | 266.5881
[11 empty] | 344.064 | 80.125 | 62.3514 | 261.5331
[11 empty] | 344.1408 | 79.575 | 64.0461 | 257.7111
[11 empty] | 344.2176 | 79.605 | 63.9078 | 254.937
[11 empty] | 344.2944 | 79.815 | 63.0938 | 252.9952
[11 empty] | 344.3712 | 79.6 | 63.6365 | 251.7931
[11 empty] | 344.448 | 79.4 | 63.1757 | 251.3307
[11 empty] | 344.5248 | 79.09 | 62.1466 | 251.3924
[11 empty] | 344.6016 | 78.625 | 60.8102 | 252.163
[11 empty] | 344.6784 | 78.06 | 59.7146 | 258.8516
[11 empty] | 344.7552 | 77.595 | 58.2758 | 275.835
[11 empty] | 344.832 | 77.395 | 56.9395 | 293.3425
[11 empty] | 344.9088 | 77.38 | 56.489 | 305.4251
[11 empty] | 344.9856 | 77.205 | 58.1632 | 312.4836
[11 empty] | 345.0624 | 77.45 | 59.4483 | 316.8296
[11 empty] | 345.1392 | 78.065 | 60.2419 | 320.6208
[11 empty] | 345.216 | 78.77 | 60.544 | 325.7375
[11 empty] | 345.2928 | 79.125 | 61.737 | 332.5802
[11 empty] | 345.3696 | 79.305 | 62.3872 | 340.5941
[11 empty] | 345.4464 | 79.405 | 62.0954 | 347.9608
[11 empty] | 345.5232 | 79.31 | 61.5168 | 353.0466
[11 empty] | 345.6 | 78.985 | 61.4554 | 354.6494
[11 empty] | 345.6768 | 78.6 | 61.9622 | 352.461
[11 empty] | 345.7536 | 78.565 | 61.4554 | 347.2519
[11 empty] | 345.8304 | 78.65 | 60.585 | 340.5017
[11 empty] | 345.9072 | 78.395 | 61.1123 | 333.8439
[11 empty] | 345.984 | 78.31 | 60.5901 | 328.3882
[11 empty] | 346.0608 | 77.99 | 59.6838 | 324.3813
[11 empty] | 346.1376 | 77.245 | 59.7043 | 321.1757
[11 empty] | 346.2144 | 76.48 | 59.0643 | 317.3844
[11 empty] | 346.2912 | 75.935 | 57.4822 | 311.9288
[11 empty] | 346.368 | 75.435 | 56.7347 | 305.0244
[11 empty] | 346.4448 | 75.39 | 56.9088 | 297.3495
[11 empty] | 346.5216 | 76.015 | 56.4122 | 289.1506
[11 empty] | 346.5984 | 76.625 | 56.6835 | 281.9688
[11 empty] | 346.6752 | 76.84 | 58.711 | 276.3898
[11 empty] | 346.752 | 76.97 | 60.5235 | 272.4445
[11 empty] | 346.8288 | 77.345 | 60.0371 | 269.8862
[11 empty] | 346.9056 | 77.485 | 59.5302 | 268.4375
[11 empty] | 346.9824 | 77.46 | 58.6957 | 267.6978
[11 empty] | 347.0592 | 77.11 | 58.8288 | 267.4203
[11 empty] | 347.136 | 76.795 | 58.6803 | 267.4203
[11 empty] | 347.2128 | 76.64 | 57.9789 | 267.6053
[11 empty] | 347.2896 | 76.445 | 58.0864 | 268.006
[11 empty] | 347.3664 | 76.47 | 57.4925 | 268.7149
[11 empty] | 347.4432 | 76.555 | 56.6374 | 270.3794
[11 empty] | 347.52 | 76.595 | 57.2467 | 274.6946
[11 empty] | 347.5968 | 76.8 | 59.1667 | 282.462
[11 empty] | 347.6736 | 77.22 | 60.5286 | 293.9898
[11 empty] | 347.7504 | 77.65 | 60.4826 | 309.0314
[11 empty] | 347.8272 | 77.87 | 61.0714 | 325.6758
[11 empty] | 347.904 | 78.035 | 62.0032 | 341.0873
[11 empty] | 347.9808 | 78.135 | 62.1261 | 352.7076
[11 empty] | 348.0576 | 78.145 | 62.1158 | 359.088
[11 empty] | 348.1344 | 78.25 | 62.2592 | 360.3825
[11 empty] | 348.2112 | 78.5 | 62.5613 | 358.1324
[11 empty] | 348.288 | 79.07 | 62.1056 | 354.4337
[11 empty] | 348.3648 | 79.285 | 63.5034 | 351.2281
[11 empty] | 348.4416 | 79.41 | 64.5069 | 349.8102
[11 empty] | 348.5184 | 79.185 | 65.6794 | 350.1185
[11 empty] | 348.5952 | 79.07 | 64.7014 | 351.0123
[11 empty] | 348.672 | 78.945 | 62.7763 | 350.8274
[11 empty] | 348.7488 | 78.345 | 62.2592 | 348.2074
[11 empty] | 348.8256 | 77.645 | 61.9162 | 342.5668
[11 empty] | 348.9024 | 77.29 | 60.3546 | 334.3987
[11 empty] | 348.9792 | 77.19 | 59.5661 | 324.5662
[11 empty] | 349.056 | 77.4 | 60.375 | 313.9631
[11 empty] | 349.1328 | 78.09 | 60.759 | 308.1067
[11 empty] | 349.2096 | 78.53 | 61.8803 | 311.2506
[11 empty] | 349.2864 | 78.84 | 62.4026 | 308.6307
[11 empty] | 349.3632 | 78.755 | 63.0426 | 302.312
[11 empty] | 349.44 | 78.41 | 63.8822 | 296.6406
[11 empty] | 349.5168 | 78.4 | 62.9606 | 292.0171
[11 empty] | 349.5936 | 78.48 | 62.2234 | 288.6574
[11 empty] | 349.6704 | 78.57 | 62.6842 | 286.5306
[11 empty] | 349.7472 | 78.885 | 63.3754 | 285.2669
[11 empty] | 349.824 | 79.15 | 64.9728 | 284.5887
[11 empty] | 349.9008 | 79.74 | 64.8192 | 284.3113
[11 empty] | 349.9776 | 80.015 | 65.3107 | 284.4346
[11 empty] | 350.0544 | 79.865 | 66.7648 | 284.7737
[11 empty] | 350.1312 | 79.96 | 65.5258 | 284.9586
[11 empty] | 350.208 | 79.895 | 64.2918 | 286.4998
[11 empty] | 350.2848 | 79.435 | 64.4608 | 296.9488
[11 empty] | 350.3616 | 79.26 | 64.087 | 315.6584
[11 empty] | 350.4384 | 79.51 | 63.2269 | 332.1795
[11 empty] | 350.5152 | 79.77 | 63.7542 | 343.3066
[11 empty] | 350.592 | 80.14 | 65.4848 | 349.9335
[11 empty] | 350.6688 | 80.935 | 66.6061 | 353.9097
[11 empty] | 350.7456 | 81.75 | 67.9885 | 357.1461
[11 empty] | 350.8224 | 82.47 | 69.376 | 361.0914
[11 empty] | 350.8992 | 82.915 | 70.825 | 366.2081
[11 empty] | 350.976 | 83.085 | 71.5878 | 371.8178
[11 empty] | 351.0528 | 83.055 | 71.6339 | 376.5954
[11 empty] | 351.1296 | 82.825 | 71.5008 | 379.0612
[11 empty] | 351.2064 | 82.435 | 71.4291 | 378.2599
[11 empty] | 351.2832 | 82.3 | 70.4922 | 373.9755
[11 empty] | 351.36 | 82.04 | 70.3846 | 366.8245
[11 empty] | 351.4368 | 81.905 | 69.9443 | 358.1941
[11 empty] | 351.5136 | 81.76 | 69.376 | 349.5945
[11 empty] | 351.5904 | 81.565 | 68.4544 | 342.1661
[11 empty] | 351.6672 | 81.13 | 67.4918 | 336.1248
[11 empty] | 351.744 | 80.675 | 65.3005 | 330.8541
[11 empty] | 351.8208 | 79.98 | 63.6314 | 325.2751
[11 empty] | 351.8976 | 79.185 | 62.6176 | 318.4324
[11 empty] | 351.9744 | 78.66 | 62.0954 | 310.6342
[11 empty] | 352.0512 | 78.66 | 62.0749 | 302.312
[11 empty] | 352.128 | 78.965 | 63.3907 | 293.3117
[11 empty] | 352.2048 | 79.37 | 64.9933 | 285.0511
[11 empty] | 352.2816 | 79.885 | 65.623 | 278.0851
[11 empty] | 352.3584 | 80.24 | 65.8842 | 272.4137
[11 empty] | 352.4352 | 80.31 | 65.5718 | 267.9443
[11 empty] | 352.512 | 80.135 | 65.1571 | 264.4922
[11 empty] | 352.5888 | 79.81 | 64.1997 | 261.903
[11 empty] | 352.6656 | 79.24 | 63.657 | 260.0845
[11 empty] | 352.7424 | 78.94 | 61.6806 | 258.8824
[11 empty] | 352.8192 | 78.465 | 60.969 | 258.081
[11 empty] | 352.896 | 78.07 | 60.6003 | 257.5878
[11 empty] | 352.9728 | 77.775 | 60.247 | 257.4029
[11 empty] | 353.0496 | 77.5 | 60.1805 | 258.5742
[11 empty] | 353.1264 | 77.42 | 60.7181 | 262.951
[11 empty] | 353.2032 | 77.725 | 60.8768 | 271.1808
[11 empty] | 353.28 | 78 | 61.2506 | 283.6949
[11 empty] | 353.3568 | 78.05 | 60.887 | 299.5687
[11 empty] | 353.4336 | 77.975 | 60.5082 | 316.3056
[11 empty] | 353.5104 | 78.145 | 60.1139 | 330.8849
[11 empty] | 353.5872 | 78.26 | 59.9091 | 340.9948
[11 empty] | 353.664 | 78.22 | 60.375 | 345.7724
[11 empty] | 353.7408 | 78.19 | 61.3069 | 346.2656
[11 empty] | 353.8176 | 78.475 | 61.5936 | 344.3546
[11 empty] | 353.8944 | 78.795 | 61.8547 | 342.2586
[11 empty] | 353.9712 | 78.84 | 62.7405 | 341.2106
[11 empty] | 354.048 | 78.63 | 63.9386 | 341.5805
[11 empty] | 354.1248 | 78.435 | 64.2867 | 342.8442
[11 empty] | 354.2016 | 78.24 | 62.9862 | 343.7381
[11 empty] | 354.2784 | 77.9 | 61.1277 | 342.5052
[11 empty] | 354.3552 | 77.495 | 59.0848 | 338.0975
[11 empty] | 354.432 | 76.895 | 58.0096 | 330.8541
[11 empty] | 354.5088 | 76.085 | 58.6496 | 321.6072
[11 empty] | 354.5856 | 76.125 | 57.3952 | 311.4048
[11 empty] | 354.6624 | 76.295 | 58.025 | 302.0962
[11 empty] | 354.7392 | 76.515 | 59.3664 | 300.9866
[11 empty] | 354.816 | 76.925 | 59.2077 | 301.3257
[11 empty] | 354.8928 | 76.975 | 58.8237 | 295.2227
[11 empty] | 354.9696 | 76.685 | 58.5216 | 288.8732
[11 empty] | 355.0464 | 76.1 | 58.7674 | 283.6332
[11 empty] | 355.1232 | 75.79 | 58.281 | 279.7187
[11 empty] | 355.2 | 75.775 | 57.6717 | 277.1604
[11 empty] | 355.2768 | 75.63 | 58.5523 | 275.7426
[11 empty] | 355.3536 | 75.965 | 58.071 | 275.0953
[11 empty] | 355.4304 | 76.39 | 57.7536 | 274.9412
[11 empty] | 355.5072 | 76.45 | 58.5523 | 275.1877
[11 empty] | 355.584 | 76.35 | 59.1923 | 275.7426
[11 empty] | 355.6608 | 76.32 | 58.2502 | 276.5131
[11 empty] | 355.7376 | 75.96 | 57.6973 | 278.8249
[11 empty] | 355.8144 | 75.29 | 58.3066 | 290.1677
[11 empty] | 355.8912 | 75.27 | 56.6886 | 310.2643
[11 empty] | 355.968 | 75.605 | 54.8557 | 327.9259
[11 empty] | 356.0448 | 75.51 | 56.1664 | 339.4537
[11 empty] | 356.1216 | 75.76 | 57.7126 | 345.7416
[11 empty] | 356.1984 | 76.635 | 57.8714 | 348.8239
[11 empty] | 356.2752 | 77.51 | 59.0899 | 351.2897
[11 empty] | 356.352 | 77.99 | 61.4554 | 355.1118
[11 empty] | 356.4288 | 78.31 | 63.3702 | 360.9681
[11 empty] | 356.5056 | 78.785 | 63.2166 | 367.9958
[11 empty] | 356.5824 | 78.94 | 63.2218 | 374.4686
[11 empty] | 356.6592 | 78.855 | 62.9555 | 378.4756
[11 empty] | 356.736 | 78.74 | 62.3206 | 378.6914
[11 empty] | 356.8128 | 78.6 | 61.9827 | 374.9001
[11 empty] | 356.8896 | 78.395 | 62.464 | 367.965
[11 empty] | 356.9664 | 78.22 | 63.0477 | 359.6119
[11 empty] | 357.0432 | 78.19 | 62.9658 | 351.7213
[11 empty] | 357.12 | 77.92 | 63.5648 | 345.5875
[11 empty] | 357.1968 | 77.81 | 62.3411 | 341.2414
[11 empty] | 357.2736 | 77.52 | 61.1379 | 337.9742
[11 empty] | 357.3504 | 77 | 60.4877 | 334.3062
[11 empty] | 357.4272 | 76.6 | 59.2333 | 329.0355
[11 empty] | 357.504 | 76.41 | 58.7418 | 322.5627
[11 empty] | 357.5808 | 76.575 | 60.0883 | 315.196
[11 empty] | 357.6576 | 77.225 | 61.5475 | 307.2437
[11 empty] | 357.7344 | 78.105 | 62.6637 | 300.216
[11 empty] | 357.8112 | 78.895 | 64.1229 | 294.5446
[11 empty] | 357.888 | 79.63 | 64.6298 | 290.1677
[11 empty] | 357.9648 | 80.21 | 64.4915 | 286.8697
[11 empty] | 358.0416 | 80.44 | 64.471 | 284.4963
[11 empty] | 358.1184 | 80.38 | 64.512 | 282.8935
[11 empty] | 358.1952 | 80.31 | 63.7594 | 281.8455
[11 empty] | 358.272 | 80.33 | 62.7712 | 281.1366
[11 empty] | 358.3488 | 80.25 | 63.1398 | 280.5201
[11 empty] | 358.4256 | 80.185 | 64.1792 | 279.842
[11 empty] | 358.5024 | 80.44 | 63.9027 | 279.1639
[11 empty] | 358.5792 | 80.56 | 63.7696 | 279.4413
[11 empty] | 358.656 | 80.565 | 64.937 | 282.6161
[11 empty] | 358.7328 | 80.93 | 66.391 | 289.5821
[11 empty] | 358.8096 | 81.49 | 67.3434 | 300.8941
[11 empty] | 358.8864 | 81.765 | 68.0141 | 315.9049
[11 empty] | 358.9632 | 81.92 | 68.3366 | 332.4877
[11 empty] | 359.04 | 82.255 | 68.8435 | 347.4369
[11 empty] | 359.1168 | 82.505 | 69.2992 | 358.0708
[11 empty] | 359.1936 | 82.695 | 69.417 | 362.91
[11 empty] | 359.2704 | 82.9 | 69.3146 | 362.6326
[11 empty] | 359.3472 | 83.065 | 69.6474 | 359.1804
[11 empty] | 359.424 | 83.41 | 69.5654 | 354.6803
[11 empty] | 359.5008 | 83.63 | 69.6832 | 351.1664
[11 empty] | 359.5776 | 83.645 | 69.8726 | 349.317
[11 empty] | 359.6544 | 83.405 | 69.9955 | 348.9164
[11 empty] | 359.7312 | 82.995 | 69.6269 | 348.5465
[11 empty] | 359.808 | 82.4 | 68.8998 | 346.4814
[11 empty] | 359.8848 | 81.795 | 67.4611 | 341.6729
[11 empty] | 359.9616 | 81.275 | 65.28 | 334.0288
[11 empty] | 360.0384 | 80.66 | 63.6365 | 324.0422
[11 empty] | 360.1152 | 80.3 | 63.2371 | 312.4836
[11 empty] | 360.192 | 80.375 | 64.2816 | 300.1852
[11 empty] | 360.2688 | 80.66 | 65.321 | 291.5547
[11 empty] | 360.3456 | 80.615 | 66.8262 | 292.7877
[11 empty] | 360.4224 | 80.74 | 66.2938 | 289.8287
[11 empty] | 360.4992 | 80.73 | 64.8858 | 282.0921
[11 empty] | 360.576 | 80.485 | 63.5443 | 274.6946
[11 empty] | 360.6528 | 80.4 | 61.527 | 268.1601
[11 empty] | 360.7296 | 80.155 | 60.7846 | 262.7044
[11 empty] | 360.8064 | 79.775 | 61.5629 | 258.605
[11 empty] | 360.8832 | 79.83 | 61.527 | 255.7076
[11 empty] | 360.96 | 79.965 | 60.969 | 253.7966
[11 empty] | 361.0368 | 79.935 | 60.8512 | 252.687
[11 empty] | 361.1136 | 79.79 | 60.3341 | 252.2554
[11 empty] | 361.1904 | 79.335 | 60.0525 | 252.2863
[11 empty] | 361.2672 | 78.86 | 59.0694 | 252.8411
[11 empty] | 361.344 | 78.315 | 57.8202 | 259.0365
[11 empty] | 361.4208 | 77.575 | 57.5949 | 275.6193
[11 empty] | 361.4976 | 77.08 | 57.2826 | 293.0959
[11 empty] | 361.5744 | 77.095 | 56.3814 | 305.2094
[11 empty] | 361.6512 | 77.23 | 56.7194 | 312.3295
[11 empty] | 361.728 | 77.49 | 58.0352 | 316.6447
[11 empty] | 361.8048 | 78.045 | 59.4074 | 320.4667
[11 empty] | 361.8816 | 78.435 | 61.5322 | 325.5833
[11 empty] | 361.9584 | 78.8 | 62.8787 | 332.4569
[11 empty] | 362.0352 | 79.17 | 62.7661 | 340.44
[11 empty] | 362.112 | 79.25 | 62.4282 | 347.8684
[11 empty] | 362.1888 | 79.135 | 62.0646 | 352.985
[11 empty] | 362.2656 | 78.85 | 62.0339 | 354.6803
[11 empty] | 362.3424 | 78.535 | 62.1414 | 352.7076
[11 empty] | 362.4192 | 78.575 | 61.0099 | 347.7451
[11 empty] | 362.496 | 78.53 | 60.7232 | 341.2722
[11 empty] | 362.5728 | 78.375 | 61.0509 | 334.7994
[11 empty] | 362.6496 | 78.315 | 60.5798 | 329.4362
[11 empty] | 362.7264 | 78.1 | 59.5251 | 325.3676
[11 empty] | 362.8032 | 77.79 | 57.7894 | 322.0079
[11 empty] | 362.88 | 77.05 | 56.8218 | 318.0934
[11 empty] | 362.9568 | 76.38 | 55.552 | 312.6069
[11 empty] | 363.0336 | 75.92 | 54.4819 | 305.6409
[11 empty] | 363.1104 | 75.95 | 53.6371 | 297.9968
[11 empty] | 363.1872 | 76.245 | 54.0058 | 289.8595
[11 empty] | 363.264 | 76.54 | 55.5878 | 282.6469
[11 empty] | 363.3408 | 76.985 | 56.9702 | 277.0063
[11 empty] | 363.4176 | 77.225 | 58.7213 | 272.9068
[11 empty] | 363.4944 | 77.41 | 59.3613 | 270.1328
[11 empty] | 363.5712 | 77.39 | 59.5712 | 268.53
[11 empty] | 363.648 | 77.175 | 59.5866 | 267.6978
[11 empty] | 363.7248 | 77.205 | 58.0301 | 267.4203
[11 empty] | 363.8016 | 77.08 | 57.0778 | 267.5436
[11 empty] | 363.8784 | 76.93 | 56.4173 | 267.8827
[11 empty] | 363.9552 | 76.635 | 57.1085 | 268.3759
[11 empty] | 364.032 | 76.5 | 57.641 | 269.054
[11 empty] | 364.1088 | 76.585 | 57.088 | 270.4102
[11 empty] | 364.1856 | 76.415 | 58.3782 | 274.2631
[11 empty] | 364.2624 | 76.71 | 59.3664 | 281.5065
[11 empty] | 364.3392 | 77.425 | 59.433 | 292.6644
[11 empty] | 364.416 | 77.73 | 60.1805 | 307.4286
[11 empty] | 364.4928 | 77.86 | 60.9024 | 323.9806
[11 empty] | 364.5696 | 77.99 | 61.8342 | 339.5462
[11 empty] | 364.6464 | 78.225 | 61.5219 | 351.4747
[11 empty] | 364.7232 | 78.47 | 60.6566 | 358.2249
[11 empty] | 364.8 | 78.41 | 61.568 | 359.8277
[11 empty] | 364.8768 | 78.41 | 63.1245 | 357.7934
[11 empty] | 364.9536 | 78.67 | 64.0051 | 354.2487
[11 empty] | 365.0304 | 79.06 | 64.4096 | 351.074
[11 empty] | 365.1072 | 79.32 | 64.8858 | 349.5328
[11 empty] | 365.184 | 79.415 | 64.6042 | 349.6869
[11 empty] | 365.2608 | 79.2 | 64.384 | 350.4267
[11 empty] | 365.3376 | 78.885 | 63.2576 | 350.2109
[11 empty] | 365.4144 | 78.27 | 62.848 | 347.7143
[11 empty] | 365.4912 | 77.835 | 61.3069 | 342.3203
[11 empty] | 365.568 | 77.585 | 59.6173 | 334.4295
[11 empty] | 365.6448 | 77.43 | 58.9466 | 324.8128
[11 empty] | 365.7216 | 77.525 | 60.0013 | 314.2713
[11 empty] | 365.7984 | 78.06 | 60.887 | 307.1512
[11 empty] | 365.8752 | 78.285 | 62.9094 | 309.5246
[11 empty] | 365.952 | 78.385 | 64.4864 | 307.7677
[11 empty] | 366.0288 | 78.63 | 63.4522 | 301.4489
[11 empty] | 366.1056 | 78.705 | 62.0595 | 295.6542
[11 empty] | 366.1824 | 78.5 | 62.0339 | 290.8766
[11 empty] | 366.2592 | 78.545 | 61.4912 | 287.332
[11 empty] | 366.336 | 78.685 | 61.5219 | 285.0203
[11 empty] | 366.4128 | 78.825 | 62.7917 | 283.6332
[11 empty] | 366.4896 | 79.115 | 64.0358 | 282.8318
[11 empty] | 366.5664 | 79.51 | 64.8397 | 282.4928
[11 empty] | 366.6432 | 79.9 | 64.8141 | 282.5853
[11 empty] | 366.72 | 79.965 | 64.9933 | 282.9551
[11 empty] | 366.7968 | 79.715 | 65.6742 | 283.2017
[11 empty] | 366.8736 | 79.545 | 65.0445 | 284.4655
[11 empty] | 366.9504 | 79.35 | 64.0358 | 294.0206
[11 empty] | 367.0272 | 79.125 | 63.7798 | 312.5452
[11 empty] | 367.104 | 79.11 | 63.6621 | 329.4362
[11 empty] | 367.1808 | 79.56 | 63.1808 | 340.8716
[11 empty] | 367.2576 | 80.26 | 63.3958 | 347.7143
[11 empty] | 367.3344 | 80.86 | 65.6794 | 351.8137
[11 empty] | 367.4112 | 81.715 | 67.2768 | 355.1118
[11 empty] | 367.488 | 82.325 | 69.3709 | 359.1188
[11 empty] | 367.5648 | 82.775 | 70.8864 | 364.2662
[11 empty] | 367.6416 | 83.265 | 70.3437 | 369.9685
[11 empty] | 367.7184 | 83.455 | 69.1149 | 374.8385
[11 empty] | 367.7952 | 83.095 | 69.5808 | 377.366
[11 empty] | 367.872 | 82.62 | 70.3283 | 376.503
[11 empty] | 367.9488 | 82.185 | 70.7686 | 372.1569
[11 empty] | 368.0256 | 82.1 | 69.7344 | 365.1293
[11 empty] | 368.1024 | 82.065 | 68.8691 | 356.6837
[11 empty] | 368.1792 | 81.78 | 69.0893 | 348.3307
[11 empty] | 368.256 | 81.4 | 68.905 | 341.0565
[11 empty] | 368.3328 | 81.065 | 67.2205 | 334.9844
[11 empty] | 368.4096 | 80.59 | 65.5002 | 329.5287
[11 empty] | 368.4864 | 79.97 | 63.4266 | 323.5799
[11 empty] | 368.5632 | 79.265 | 61.7728 | 316.3981
[11 empty] | 368.64 | 78.565 | 61.7779 | 308.3841
[11 empty] | 368.7168 | 78.4 | 62.3258 | 300.0619
[11 empty] | 368.7936 | 78.98 | 62.4179 | 291.2773
[11 empty] | 368.8704 | 79.54 | 63.2166 | 283.325
[11 empty] | 368.9472 | 79.94 | 64.4301 | 276.6364
[11 empty] | 369.024 | 80.25 | 64.9472 | 271.1191
[11 empty] | 369.1008 | 80.52 | 64.3942 | 266.6806
[11 empty] | 369.1776 | 80.52 | 63.5648 | 263.1051
[11 empty] | 369.2544 | 79.9 | 63.8515 | 260.4852
[11 empty] | 369.3312 | 79.37 | 63.063 | 258.6666
[11 empty] | 369.408 | 78.975 | 61.7984 | 257.557
[11 empty] | 369.4848 | 78.41 | 61.568 | 256.9097
[11 empty] | 369.5616 | 78.1 | 60.8614 | 256.5398
[11 empty] | 369.6384 | 77.9 | 60.2368 | 256.4165
[11 empty] | 369.7152 | 77.77 | 59.6275 | 257.4953
[11 empty] | 369.792 | 77.88 | 59.1667 | 261.564
[11 empty] | 369.8688 | 78.03 | 60.1702 | 269.4855
[11 empty] | 369.9456 | 78.345 | 60.1805 | 281.7222
[11 empty] | 370.0224 | 78.29 | 60.5235 | 297.442
[11 empty] | 370.0992 | 78.12 | 60.7437 | 314.3329
[11 empty] | 370.176 | 78.035 | 61.44 | 329.0972
[11 empty] | 370.2528 | 78.015 | 61.9469 | 339.2688
[11 empty] | 370.3296 | 78.09 | 61.7011 | 344.1696
[11 empty] | 370.4064 | 78.145 | 61.9264 | 344.7552
[11 empty] | 370.4832 | 78.36 | 62.3462 | 342.9059
[11 empty] | 370.56 | 78.55 | 63.3702 | 340.7174
[11 empty] | 370.6368 | 78.86 | 63.2371 | 339.577
[11 empty] | 370.7136 | 78.975 | 63.0016 | 339.916
[11 empty] | 370.7904 | 78.81 | 62.9043 | 341.3031
[11 empty] | 370.8672 | 78.51 | 62.121 | 342.1661
[11 empty] | 370.944 | 78.14 | 60.6054 | 340.8716
[11 empty] | 371.0208 | 77.615 | 59.392 | 336.6488
[11 empty] | 371.0976 | 76.945 | 58.6906 | 329.4979
[11 empty] | 371.1744 | 76.265 | 58.3987 | 320.3434
[11 empty] | 371.2512 | 76.075 | 58.0813 | 310.326
[11 empty] | 371.328 | 76.185 | 59.1974 | 300.9558
[11 empty] | 371.4048 | 76.635 | 59.4995 | 299.168
[11 empty] | 371.4816 | 76.81 | 60.2624 | 299.5071
[11 empty] | 371.5584 | 76.72 | 60.5082 | 293.5583
[11 empty] | 371.6352 | 76.76 | 58.7418 | 287.2704
[11 empty] | 371.712 | 76.155 | 59.1411 | 282.0921
[11 empty] | 371.7888 | 75.7 | 59.2896 | 278.2392
[11 empty] | 371.8656 | 75.75 | 58.2093 | 275.7117
[11 empty] | 371.9424 | 75.815 | 58.025 | 274.2631
[11 empty] | 372.0192 | 76.23 | 57.2774 | 273.5541
[11 empty] | 372.096 | 76.5 | 57.7331 | 273.3076
[11 empty] | 372.1728 | 76.65 | 58.327 | 273.5541
[11 empty] | 372.2496 | 76.66 | 58.4141 | 274.1706
[11 empty] | 372.3264 | 76.29 | 58.9773 | 275.0336
[11 empty] | 372.4032 | 75.93 | 58.9107 | 277.1296
[11 empty] | 372.48 | 75.64 | 57.7178 | 287.6402
[11 empty] | 372.5568 | 75.23 | 57.3696 | 307.4594
[11 empty] | 372.6336 | 75.145 | 56.919 | 325.4601
[11 empty] | 372.7104 | 75.225 | 57.5334 | 337.3577
[11 empty] | 372.7872 | 75.61 | 58.88 | 343.8614
[11 empty] | 372.864 | 76.61 | 58.711 | 346.9745
[11 empty] | 372.9408 | 77.48 | 59.6275 | 349.3479
[11 empty] | 373.0176 | 78.3 | 60.4211 | 353.0158
[11 empty] | 373.0944 | 78.685 | 62.2029 | 358.7181
[11 empty] | 373.1712 | 78.715 | 64.1946 | 365.8074
[11 empty] | 373.248 | 78.645 | 65.111 | 372.4343
[11 empty] | 373.3248 | 78.55 | 64.9318 | 376.5954
[11 empty] | 373.4016 | 78.555 | 63.7286 | 377.0269
[11 empty] | 373.4784 | 78.48 | 63.0938 | 373.4515
[11 empty] | 373.5552 | 78.42 | 62.6432 | 366.6396
[11 empty] | 373.632 | 78.375 | 62.7302 | 358.1941
[11 empty] | 373.7088 | 78.375 | 62.5664 | 350.1801
[11 empty] | 373.7856 | 78.21 | 62.6995 | 343.923
[11 empty] | 373.8624 | 77.96 | 62.2438 | 339.5462
[11 empty] | 373.9392 | 77.74 | 60.6925 | 336.3098
[11 empty] | 374.016 | 77.3 | 59.3459 | 332.7035
[11 empty] | 374.0928 | 76.795 | 58.7981 | 327.556
[11 empty] | 374.1696 | 76.43 | 59.1002 | 321.0832
[11 empty] | 374.2464 | 76.655 | 59.8528 | 313.6548
[11 empty] | 374.3232 | 77.375 | 60.7283 | 305.61
[11 empty] | 374.4 | 78.165 | 61.7574 | 298.4283
[11 empty] | 374.4768 | 78.83 | 63.2883 | 292.6336
[11 empty] | 374.5536 | 79.565 | 63.9437 | 288.1334
[11 empty] | 374.6304 | 79.93 | 65.0752 | 284.7737
[11 empty] | 374.7072 | 80.29 | 64.6912 | 282.3079
[11 empty] | 374.784 | 80.36 | 64.3891 | 280.6434
[11 empty] | 374.8608 | 80.185 | 64.6707 | 279.5646
[11 empty] | 374.9376 | 80 | 64.9062 | 278.8249
[11 empty] | 375.0144 | 79.995 | 64.5734 | 278.2084
[11 empty] | 375.0912 | 80.13 | 64.0819 | 277.5611
[11 empty] | 375.168 | 80.11 | 64.3738 | 276.9138
[11 empty] | 375.2448 | 80.035 | 65.2032 | 277.0988
[11 empty] | 375.3216 | 80.255 | 65.4694 | 280.0578
[11 empty] | 375.3984 | 80.84 | 65.5974 | 286.7464
[11 empty] | 375.4752 | 81.175 | 67.241 | 297.7194
[11 empty] | 375.552 | 81.665 | 66.816 | 312.5452
[11 empty] | 375.6288 | 82.005 | 66.5088 | 329.1896
[11 empty] | 375.7056 | 82.14 | 68.0243 | 344.4778
[11 empty] | 375.7824 | 82.3 | 69.3094 | 355.6049
[11 empty] | 375.8592 | 82.43 | 69.6064 | 361.0298
[11 empty] | 375.936 | 82.685 | 69.3197 | 361.1531
[11 empty] | 376.0128 | 82.84 | 69.5808 | 357.8242
[11 empty] | 376.0896 | 82.89 | 70.743 | 353.509
[11 empty] | 376.1664 | 83.22 | 70.5843 | 350.1801
[11 empty] | 376.2432 | 83.4 | 70.0979 | 348.454
[11 empty] | 376.32 | 83.39 | 69.5808 | 348.0841
[11 empty] | 376.3968 | 83.145 | 68.7053 | 347.8684
[11 empty] | 376.4736 | 82.625 | 67.7376 | 346.019
[11 empty] | 376.5504 | 81.825 | 67.2614 | 341.3339
[11 empty] | 376.6272 | 81.155 | 65.7869 | 333.7206
[11 empty] | 376.704 | 80.555 | 64.7629 | 323.7032
[11 empty] | 376.7808 | 80.425 | 63.7082 | 312.2062
[11 empty] | 376.8576 | 80.635 | 63.7286 | 300.0311
[11 empty] | 376.9344 | 80.74 | 65.0394 | 291.4623
[11 empty] | 377.0112 | 81.005 | 65.1725 | 292.5411
[11 empty] | 377.088 | 81.235 | 64.2202 | 289.4896
[11 empty] | 377.1648 | 81.03 | 63.657 | 281.5989
[11 empty] | 377.2416 | 80.71 | 62.4077 | 274.1706
[11 empty] | 377.3184 | 80.025 | 62.8736 | 267.6978
[11 empty] | 377.3952 | 79.765 | 62.4845 | 262.3654
[11 empty] | 377.472 | 79.555 | 62.6483 | 258.3584
[11 empty] | 377.5488 | 79.36 | 63.3651 | 255.5227
[11 empty] | 377.6256 | 79.435 | 63.145 | 253.6425
[11 empty] | 377.7024 | 79.535 | 62.4896 | 252.5637
[11 empty] | 377.7792 | 79.435 | 61.9622 | 252.1013
[11 empty] | 377.856 | 79.135 | 61.1891 | 252.1013
[11 empty] | 377.9328 | 78.64 | 60.585 | 252.5945
[11 empty] | 378.0096 | 78.14 | 59.2077 | 257.9885
[11 empty] | 378.0864 | 77.505 | 58.1837 | 273.924
[11 empty] | 378.1632 | 77.165 | 57.1136 | 291.7397
[11 empty] | 378.24 | 77.035 | 56.9651 | 304.408
[11 empty] | 378.3168 | 76.975 | 57.9174 | 311.9904
[11 empty] | 378.3936 | 77.41 | 58.2195 | 316.5522
[11 empty] | 378.4704 | 78.145 | 58.4294 | 320.3126
[11 empty] | 378.5472 | 78.695 | 59.7914 | 325.2135
[11 empty] | 378.624 | 79.125 | 61.056 | 331.9021
[11 empty] | 378.7008 | 79.355 | 61.824 | 339.8544
[11 empty] | 378.7776 | 79.305 | 62.2541 | 347.4677
[11 empty] | 378.8544 | 79.09 | 62.4384 | 352.985
[11 empty] | 378.9312 | 78.82 | 62.377 | 355.081
[11 empty] | 379.008 | 78.665 | 61.8445 | 353.4165
[11 empty] | 379.0848 | 78.43 | 62.0595 | 348.5773
[11 empty] | 379.1616 | 78.38 | 61.737 | 341.9812
[11 empty] | 379.2384 | 78.275 | 61.655 | 335.2926
[11 empty] | 379.3152 | 78.045 | 61.6346 | 329.6828
[11 empty] | 379.392 | 77.78 | 60.9331 | 325.5217
[11 empty] | 379.4688 | 77.445 | 59.1923 | 322.2236
[11 empty] | 379.5456 | 76.84 | 57.9328 | 318.494
[11 empty] | 379.6224 | 76.225 | 56.8115 | 313.1925
[11 empty] | 379.6992 | 75.535 | 56.6938 | 306.319
[11 empty] | 379.776 | 75.36 | 56.5606 | 298.6749
[11 empty] | 379.8528 | 75.87 | 56.1101 | 290.476
[11 empty] | 379.9296 | 76.465 | 56.4838 | 283.1401
[11 empty] | 380.0064 | 76.825 | 57.8355 | 277.3762
[11 empty] | 380.0832 | 77.175 | 58.8288 | 273.2767
[11 empty] | 380.16 | 77.43 | 59.4074 | 270.5335
[11 empty] | 380.2368 | 77.59 | 59.1514 | 268.869
[11 empty] | 380.3136 | 77.485 | 58.7469 | 268.0368
[11 empty] | 380.3904 | 77.19 | 58.6547 | 267.7594
[11 empty] | 380.4672 | 76.7 | 59.1616 | 267.821
[11 empty] | 380.544 | 76.62 | 58.2502 | 268.1601
[11 empty] | 380.6208 | 76.585 | 57.856 | 268.6224
[11 empty] | 380.6976 | 76.475 | 58.1478 | 269.2389
[11 empty] | 380.7744 | 76.41 | 58.1018 | 270.5335
[11 empty] | 380.8512 | 76.54 | 58.1837 | 274.0781
[11 empty] | 380.928 | 76.99 | 58.5523 | 280.9825
[11 empty] | 381.0048 | 77.43 | 59.8938 | 291.7089
[11 empty] | 381.0816 | 77.58 | 61.399 | 306.1957
[11 empty] | 381.1584 | 77.865 | 61.2454 | 322.8093
[11 empty] | 381.2352 | 78.1 | 61.6141 | 338.714
[11 empty] | 381.312 | 78.115 | 62.3309 | 351.074
[11 empty] | 381.3888 | 78.115 | 62.7046 | 358.2866
[11 empty] | 381.4656 | 78.195 | 62.9862 | 360.2592
[11 empty] | 381.5424 | 78.315 | 63.6723 | 358.379
[11 empty] | 381.6192 | 78.715 | 63.6774 | 354.8036
[11 empty] | 381.696 | 79.285 | 63.2115 | 351.5671
[11 empty] | 381.7728 | 79.58 | 63.2781 | 349.9335
[11 empty] | 381.8496 | 79.57 | 63.7184 | 350.026
[11 empty] | 381.9264 | 79.425 | 63.2986 | 350.8274
[11 empty] | 382.0032 | 78.985 | 63.1501 | 350.8274
[11 empty] | 382.08 | 78.255 | 63.4419 | 348.5465
[11 empty] | 382.1568 | 77.715 | 62.505 | 343.3682
[11 empty] | 382.2336 | 77.2 | 61.5424 | 335.6008
[11 empty] | 382.3104 | 76.925 | 60.9792 | 326.0765
[11 empty] | 382.3872 | 77.065 | 61.3171 | 315.5967
[11 empty] | 382.464 | 77.73 | 61.4502 | 308.0759
[11 empty] | 382.5408 | 78.255 | 62.2234 | 310.326
[11 empty] | 382.6176 | 78.66 | 62.6586 | 309.1547
[11 empty] | 382.6944 | 78.915 | 61.8906 | 302.9284
[11 empty] | 382.7712 | 78.615 | 62.5203 | 297.1646
[11 empty] | 382.848 | 78.33 | 63.0221 | 292.387
[11 empty] | 382.9248 | 78.23 | 63.104 | 288.7807
[11 empty] | 383.0016 | 78.505 | 62.551 | 286.469
[11 empty] | 383.0784 | 78.98 | 62.4384 | 285.0819
[11 empty] | 383.1552 | 79.38 | 63.2883 | 284.3113
[11 empty] | 383.232 | 79.58 | 64.9523 | 284.0339
[11 empty] | 383.3088 | 79.885 | 65.4541 | 284.1572
[11 empty] | 383.3856 | 80.12 | 65.1571 | 284.4963
[11 empty] | 383.4624 | 79.73 | 66.5037 | 284.7429
[11 empty] | 383.5392 | 79.405 | 66.6419 | 285.6984
[11 empty] | 383.616 | 79.205 | 66.1094 | 294.0514
[11 empty] | 383.6928 | 79.13 | 64.9882 | 312.1754
[11 empty] | 383.7696 | 79.16 | 64.297 | 329.4979
[11 empty] | 383.8464 | 79.335 | 64.809 | 341.4264
[11 empty] | 383.9232 | 80.115 | 64.5939 | 348.6389
[11 empty] | 384 | 80.94 | 65.5104 | 352.8925
[11 empty] | 384.0768 | 81.58 | 68.1574 | 356.1289
[11 empty] | 384.1536 | 82.23 | 70.4358 | 359.951
[11 empty] | 384.2304 | 82.755 | 72.0282 | 364.8827
[11 empty] | 384.3072 | 83.205 | 71.6339 | 370.4925
[11 empty] | 384.384 | 83.26 | 71.2909 | 375.4858
[11 empty] | 384.4608 | 83.07 | 70.7328 | 378.3215
[11 empty] | 384.5376 | 82.86 | 69.8675 | 377.9208
[11 empty] | 384.6144 | 82.485 | 69.9904 | 374.0371
[11 empty] | 384.6912 | 82.305 | 69.5757 | 367.2252
[11 empty] | 384.768 | 82.12 | 69.5603 | 358.8722
[11 empty] | 384.8448 | 81.975 | 69.1558 | 350.3959
[11 empty] | 384.9216 | 81.725 | 68.5312 | 342.8442
[11 empty] | 384.9984 | 81.27 | 67.8298 | 336.5563
[11 empty] | 385.0752 | 80.865 | 65.7818 | 331.0699
[11 empty] | 385.152 | 80.05 | 64.7475 | 325.3984
[11 empty] | 385.2288 | 79.455 | 62.3462 | 318.5557
[11 empty] | 385.3056 | 78.82 | 61.7984 | 310.6958
[11 empty] | 385.3824 | 78.585 | 62.3616 | 302.4969
[11 empty] | 385.4592 | 79.05 | 62.6227 | 293.7124
[11 empty] | 385.536 | 79.535 | 63.4573 | 285.5443
[11 empty] | 385.6128 | 79.98 | 64.5683 | 278.6399
[11 empty] | 385.6896 | 80.33 | 65.1059 | 272.9685
[11 empty] | 385.7664 | 80.495 | 64.9114 | 268.345
[11 empty] | 385.8432 | 80.45 | 64.1331 | 264.6771
[11 empty] | 385.92 | 80.05 | 63.6877 | 261.9339
[11 empty] | 385.9968 | 79.49 | 63.2525 | 260.0845
[11 empty] | 386.0736 | 78.935 | 62.761 | 258.944
[11 empty] | 386.1504 | 78.675 | 61.2608 | 258.2043
[11 empty] | 386.2272 | 78.315 | 60.9075 | 257.7727
[11 empty] | 386.304 | 78.095 | 60.0883 | 257.5878
[11 empty] | 386.3808 | 77.775 | 60.0576 | 258.2659
[11 empty] | 386.4576 | 77.51 | 60.7488 | 261.7797
[11 empty] | 386.5344 | 77.655 | 61.3888 | 269.054
[11 empty] | 386.6112 | 77.855 | 62.0749 | 280.6742
[11 empty] | 386.688 | 77.88 | 62.0339 | 296.0241
[11 empty] | 386.7648 | 77.905 | 61.0765 | 312.9151
[11 empty] | 386.8416 | 78.25 | 59.9603 | 328.1725
[11 empty] | 386.9184 | 78.255 | 60.2573 | 339.1147
[11 empty] | 386.9952 | 78.22 | 60.6413 | 344.7861
[11 empty] | 387.072 | 78.29 | 60.8717 | 345.8649
[11 empty] | 387.1488 | 78.18 | 62.3155 | 344.1696
[11 empty] | 387.2256 | 78.38 | 63.0067 | 341.9195
[11 empty] | 387.3024 | 78.725 | 62.7917 | 340.5941
[11 empty] | 387.3792 | 78.875 | 62.6637 | 340.9024
[11 empty] | 387.456 | 78.745 | 62.8122 | 342.2894
[11 empty] | 387.5328 | 78.43 | 62.4077 | 343.2758
[11 empty] | 387.6096 | 77.99 | 61.3274 | 342.3203
[11 empty] | 387.6864 | 77.375 | 60.4416 | 338.4674
[11 empty] | 387.7632 | 76.635 | 60.0064 | 331.5938
[11 empty] | 387.84 | 76.195 | 58.6445 | 322.6552
[11 empty] | 387.9168 | 76.145 | 57.3338 | 312.6993
[11 empty] | 387.9936 | 76.2 | 58.0198 | 302.9284
[11 empty] | 388.0704 | 76.255 | 59.8067 | 299.9078
[11 empty] | 388.1472 | 76.675 | 59.6173 | 301.0174
[11 empty] | 388.224 | 76.755 | 59.6736 | 295.5617
[11 empty] | 388.3008 | 76.59 | 59.0387 | 289.0889
[11 empty] | 388.3776 | 76.22 | 58.327 | 283.7565
[11 empty] | 388.4544 | 75.67 | 58.5882 | 279.7495
[11 empty] | 388.5312 | 75.355 | 58.6854 | 277.0988
[11 empty] | 388.608 | 75.885 | 56.5914 | 275.5884
[11 empty] | 388.6848 | 76.25 | 56.1562 | 274.8179
[11 empty] | 388.7616 | 76.33 | 57.4618 | 274.5405
[11 empty] | 388.8384 | 76.445 | 58.4141 | 274.6946
[11 empty] | 388.9152 | 76.615 | 57.9891 | 275.2802
[11 empty] | 388.992 | 76.435 | 57.9123 | 276.1124
[11 empty] | 389.0688 | 76.125 | 57.2262 | 277.8077
[11 empty] | 389.1456 | 75.71 | 56.2893 | 286.9005
[11 empty] | 389.2224 | 75.43 | 55.2806 | 306.2265
[11 empty] | 389.2992 | 75.075 | 56.1101 | 324.9361
[11 empty] | 389.376 | 75.035 | 57.6 | 337.5735
[11 empty] | 389.4528 | 75.595 | 57.815 | 344.5703
[11 empty] | 389.5296 | 76.45 | 57.9635 | 347.93
[11 empty] | 389.6064 | 77.135 | 59.8579 | 350.1801
[11 empty] | 389.6832 | 77.67 | 62.5818 | 353.5706
[11 empty] | 389.76 | 78.19 | 64.169 | 359.0571
[11 empty] | 389.8368 | 78.705 | 64.2458 | 366.054
[11 empty] | 389.9136 | 78.98 | 63.7952 | 372.8967
[11 empty] | 389.9904 | 78.975 | 63.5904 | 377.4585
[11 empty] | 390.0672 | 78.55 | 64.5018 | 378.3831
[11 empty] | 390.144 | 78.2 | 64.9779 | 375.2392
[11 empty] | 390.2208 | 78.17 | 64.3328 | 368.7047
[11 empty] | 390.2976 | 78.285 | 63.488 | 360.3209
[11 empty] | 390.3744 | 78.255 | 63.2371 | 352.0911
[11 empty] | 390.4512 | 78.185 | 63.2115 | 345.5258
[11 empty] | 390.528 | 78.06 | 62.464 | 340.9024
[11 empty] | 390.6048 | 77.755 | 61.4042 | 337.6043
[11 empty] | 390.6816 | 77.18 | 60.8768 | 334.1213
[11 empty] | 390.7584 | 76.715 | 59.6224 | 329.1588
[11 empty] | 390.8352 | 76.505 | 58.496 | 322.8401
[11 empty] | 390.912 | 76.785 | 58.4192 | 315.5659
[11 empty] | 390.9888 | 77.625 | 58.8339 | 307.5211
[11 empty] | 391.0656 | 78.395 | 60.3341 | 300.2469
[11 empty] | 391.1424 | 79 | 62.4486 | 294.3596
[11 empty] | 391.2192 | 79.44 | 64.5376 | 289.7362
[11 empty] | 391.296 | 80.005 | 64.6349 | 286.3148
[11 empty] | 391.3728 | 80.325 | 64.7117 | 283.849
[11 empty] | 391.4496 | 80.245 | 65.3261 | 282.1846
[11 empty] | 391.5264 | 80.24 | 64.8909 | 281.1366
[11 empty] | 391.6032 | 80.3 | 63.7747 | 280.3968
[11 empty] | 391.68 | 80.095 | 64.3533 | 279.7804
[11 empty] | 391.7568 | 80.18 | 64.2355 | 279.2256
[11 empty] | 391.8336 | 80.5 | 63.5443 | 278.5783
[11 empty] | 391.9104 | 80.815 | 62.7917 | 278.6399
[11 empty] | 391.9872 | 80.975 | 63.3139 | 281.2907
[11 empty] | 392.064 | 81.31 | 64.471 | 287.6402
[11 empty] | 392.1408 | 81.33 | 67.6762 | 298.2433
[11 empty] | 392.2176 | 81.425 | 69.3862 | 312.8535
[11 empty] | 392.2944 | 81.73 | 69.0995 | 329.4054
[11 empty] | 392.3712 | 82.17 | 69.0995 | 344.8785
[11 empty] | 392.448 | 82.68 | 68.6746 | 356.4372
[11 empty] | 392.5248 | 82.96 | 68.4186 | 362.3552
[11 empty] | 392.6016 | 83.01 | 69.0995 | 362.9408
[11 empty] | 392.6784 | 83.055 | 69.9238 | 359.7969
[11 empty] | 392.7552 | 83.485 | 69.6371 | 355.3584
[11 empty] | 392.832 | 83.685 | 70.2003 | 351.598
[11 empty] | 392.9088 | 83.735 | 70.5946 | 349.5636
[11 empty] | 392.9856 | 83.63 | 70.5536 | 349.0088
[11 empty] | 393.0624 | 83.295 | 70.2106 | 348.6698
[11 empty] | 393.1392 | 82.85 | 69.2224 | 346.9745
[11 empty] | 393.216 | 82.24 | 68.1677 | 342.5977
[11 empty] | 393.2928 | 81.59 | 66.4678 | 335.1385
[11 empty] | 393.3696 | 80.97 | 65.1008 | 325.2443
[11 empty] | 393.4464 | 80.585 | 64.5632 | 313.7473
[11 empty] | 393.5232 | 80.535 | 65.5258 | 301.3873
[11 empty] | 393.6 | 80.89 | 65.961 | 291.8938
[11 empty] | 393.6768 | 81.155 | 66.2835 | 292.387
[11 empty] | 393.7536 | 81.185 | 66.1914 | 290.1369
[11 empty] | 393.8304 | 81.045 | 65.1725 | 282.277
[11 empty] | 393.9072 | 80.435 | 65.2749 | 274.6638
[11 empty] | 393.984 | 80.135 | 63.9232 | 268.006
[11 empty] | 394.0608 | 80.065 | 62.2643 | 262.5503
[11 empty] | 394.1376 | 79.93 | 61.4758 | 258.42
[11 empty] | 394.2144 | 79.73 | 61.8854 | 255.4302
[11 empty] | 394.2912 | 79.705 | 62.5664 | 253.3959
[11 empty] | 394.368 | 79.93 | 61.8291 | 252.1321
[11 empty] | 394.4448 | 79.745 | 61.6294 | 251.5465
[11 empty] | 394.5216 | 79.375 | 61.225 | 251.4232
[11 empty] | 394.5984 | 78.81 | 60.8563 | 251.7623
[11 empty] | 394.6752 | 78.3 | 59.392 | 256.1083
[11 empty] | 394.752 | 77.695 | 58.3424 | 271.1191
[11 empty] | 394.8288 | 77.45 | 56.6477 | 289.3047
[11 empty] | 394.9056 | 77.185 | 56.6733 | 302.6202
[11 empty] | 394.9824 | 77.255 | 57.0522 | 310.5725
[11 empty] | 395.0592 | 77.535 | 58.1274 | 315.2885
[11 empty] | 395.136 | 77.985 | 59.4944 | 319.0489
[11 empty] | 395.2128 | 78.515 | 60.7642 | 323.7648
[11 empty] | 395.2896 | 78.815 | 62.5664 | 330.2376
[11 empty] | 395.3664 | 79.205 | 62.6637 | 338.0975
[11 empty] | 395.4432 | 79.525 | 61.6858 | 345.7416
[11 empty] | 395.52 | 79.405 | 61.5782 | 351.413
[11 empty] | 395.5968 | 79.05 | 61.8803 | 353.7864
[11 empty] | 395.6736 | 79.015 | 60.6566 | 352.461
[11 empty] | 395.7504 | 78.76 | 60.6771 | 347.9608
[11 empty] | 395.8272 | 78.695 | 60.3494 | 341.6113
[11 empty] | 395.904 | 78.685 | 59.817 | 334.9844
[11 empty] | 395.9808 | 78.375 | 60.3136 | 329.2205
[11 empty] | 396.0576 | 77.875 | 60.7898 | 324.8128
[11 empty] | 396.1344 | 77.425 | 60.0218 | 321.2373
[11 empty] | 396.2112 | 76.885 | 58.5421 | 317.4153
[11 empty] | 396.288 | 76.36 | 56.6682 | 312.2678
[11 empty] | 396.3648 | 75.835 | 55.4803 | 305.5176
[11 empty] | 396.4416 | 75.67 | 55.0758 | 298.1509
[11 empty] | 396.5184 | 75.88 | 56.064 | 290.1369
[11 empty] | 396.5952 | 76.19 | 57.7434 | 282.7394
[11 empty] | 396.672 | 76.665 | 58.7264 | 276.7597
[11 empty] | 396.7488 | 77.12 | 59.3459 | 272.3828
[11 empty] | 396.8256 | 77.565 | 58.9056 | 269.393
[11 empty] | 396.9024 | 77.55 | 59.177 | 267.6053
[11 empty] | 396.9792 | 77.315 | 59.2896 | 266.7422
[11 empty] | 397.056 | 77.095 | 58.8083 | 266.4957
[11 empty] | 397.1328 | 76.875 | 58.4243 | 266.619
[11 empty] | 397.2096 | 76.655 | 58.3219 | 266.8964
[11 empty] | 397.2864 | 76.555 | 57.7331 | 267.2971
[11 empty] | 397.3632 | 76.425 | 57.8048 | 267.7902
[11 empty] | 397.44 | 76.205 | 58.6445 | 268.9307
[11 empty] | 397.5168 | 76.205 | 59.1002 | 272.352
[11 empty] | 397.5936 | 76.985 | 57.6512 | 279.1947
[11 empty] | 397.6704 | 77.495 | 58.624 | 289.8287
[11 empty] | 397.7472 | 77.75 | 60.032 | 304.2538
[11 empty] | 397.824 | 78.065 | 60.2829 | 320.775
[11 empty] | 397.9008 | 78.365 | 60.6259 | 336.6796
[11 empty] | 397.9776 | 78.3 | 61.6755 | 349.2862
[11 empty] | 398.0544 | 78.14 | 62.4384 | 356.8379
[11 empty] | 398.1312 | 78.11 | 62.9146 | 359.2112
[11 empty] | 398.208 | 78.41 | 62.9606 | 357.6701
[11 empty] | 398.2848 | 79.12 | 62.121 | 354.1871
[11 empty] | 398.3616 | 79.53 | 62.6534 | 350.7966
[11 empty] | 398.4384 | 79.65 | 63.872 | 348.8239
[11 empty] | 398.5152 | 79.595 | 64.5786 | 348.5465
[11 empty] | 398.592 | 79.38 | 64.5325 | 349.2246
[11 empty] | 398.6688 | 79.08 | 63.6621 | 349.2862
[11 empty] | 398.7456 | 78.85 | 61.3222 | 347.3136
[11 empty] | 398.8224 | 78.245 | 60.6515 | 342.4744
[11 empty] | 398.8992 | 77.48 | 60.7898 | 334.9536
[11 empty] | 398.976 | 77.305 | 59.7402 | 325.4909
[11 empty] | 399.0528 | 77.53 | 59.8938 | 314.8261
[11 empty] | 399.1296 | 78 | 61.0816 | 306.4731
[11 empty] | 399.2064 | 78.425 | 62.3565 | 308.23
[11 empty] | 399.2832 | 78.855 | 62.6739 | 307.8601
[11 empty] | 399.36 | 79.05 | 62.3718 | 301.6955
[11 empty] | 399.4368 | 78.785 | 62.9811 | 295.7467
[11 empty] | 399.5136 | 78.465 | 63.232 | 290.7842
[11 empty] | 399.5904 | 78.515 | 62.1824 | 286.993
[11 empty] | 399.6672 | 78.695 | 62.039 | 284.4963
[11 empty] | 399.744 | 78.915 | 63.2166 | 282.986
[11 empty] | 399.8208 | 79.245 | 64.4147 | 282.1229
[11 empty] | 399.8976 | 79.74 | 64.7629 | 281.7531
[11 empty] | 399.9744 | 80.1 | 65.0086 | 281.8147
[11 empty] | 400.0512 | 80.26 | 65.1622 | 282.1537
[11 empty] | 400.128 | 80.11 | 65.664 | 282.4003
[11 empty] | 400.2048 | 79.915 | 65.2954 | 283.325
[11 empty] | 400.2816 | 79.635 | 64.727 | 291.4623
[11 empty] | 400.3584 | 79.48 | 63.8208 | 309.4937
[11 empty] | 400.4352 | 79.545 | 63.0374 | 326.9704
[11 empty] | 400.512 | 79.82 | 63.2115 | 339.1147
[11 empty] | 400.5888 | 80.285 | 64.2765 | 346.5122
[11 empty] | 400.6656 | 80.84 | 66.6419 | 350.9815
[11 empty] | 400.7424 | 81.65 | 68.8691 | 354.4029
[11 empty] | 400.8192 | 82.45 | 70.4666 | 358.379
[11 empty] | 400.896 | 82.915 | 71.8234 | 363.3723
[11 empty] | 400.9728 | 83.235 | 72.1818 | 368.9821
[11 empty] | 401.0496 | 83.365 | 71.8848 | 373.9755
[11 empty] | 401.1264 | 83 | 72.5606 | 376.842
[11 empty] | 401.2032 | 82.815 | 71.465 | 376.5646
[11 empty] | 401.28 | 82.755 | 69.9597 | 372.7734
[11 empty] | 401.3568 | 82.61 | 69.1814 | 366.0848
[11 empty] | 401.4336 | 82.26 | 69.8112 | 357.7626
[11 empty] | 401.5104 | 82.015 | 69.6525 | 349.2862
[11 empty] | 401.5872 | 81.93 | 68.2496 | 341.6729
[11 empty] | 401.664 | 81.515 | 67.5277 | 335.2618
[11 empty] | 401.7408 | 80.83 | 66.7699 | 329.652
[11 empty] | 401.8176 | 80.015 | 65.5872 | 323.9189
[11 empty] | 401.8944 | 79.545 | 62.9453 | 317.1995
[11 empty] | 401.9712 | 78.915 | 62.1875 | 309.4013
[11 empty] | 402.048 | 78.745 | 62.2746 | 301.2948
[11 empty] | 402.1248 | 78.945 | 63.6621 | 292.5719
[11 empty] | 402.2016 | 79.385 | 64.8038 | 284.2805
[11 empty] | 402.2784 | 80.135 | 64.4506 | 277.1604
[11 empty] | 402.3552 | 80.525 | 64.9421 | 271.2732
[11 empty] | 402.432 | 80.55 | 65.4387 | 266.4957
[11 empty] | 402.5088 | 80.52 | 64.4301 | 262.7969
[11 empty] | 402.5856 | 80.03 | 64.2714 | 260.0228
[11 empty] | 402.6624 | 79.515 | 63.7235 | 258.2043
[11 empty] | 402.7392 | 78.935 | 63.2371 | 257.0638
[11 empty] | 402.816 | 78.44 | 62.8224 | 256.3549
[11 empty] | 402.8928 | 78.225 | 61.609 | 255.8617
[11 empty] | 402.9696 | 78.035 | 60.9741 | 255.5535
[11 empty] | 403.0464 | 77.87 | 60.416 | 256.2624
[11 empty] | 403.1232 | 77.67 | 60.7488 | 259.8379
[11 empty] | 403.2 | 77.97 | 60.7181 | 267.1429
[11 empty] | 403.2768 | 78.33 | 60.759 | 278.7324
[11 empty] | 403.3536 | 78.315 | 61.1174 | 294.0206
[11 empty] | 403.4304 | 78.135 | 60.9178 | 310.8499
[11 empty] | 403.5072 | 78.235 | 60.4262 | 325.984
[11 empty] | 403.584 | 78.175 | 61.0355 | 336.8954
[11 empty] | 403.6608 | 78 | 62.0493 | 342.7209
[11 empty] | 403.7376 | 78.135 | 62.167 | 344.0463
[11 empty] | 403.8144 | 78.6 | 61.5014 | 342.5977
[11 empty] | 403.8912 | 78.87 | 61.8854 | 340.4709
[11 empty] | 403.968 | 78.845 | 63.3754 | 339.1763
[11 empty] | 404.0448 | 78.955 | 63.4214 | 339.2688
[11 empty] | 404.1216 | 79.08 | 62.1773 | 340.5017
[11 empty] | 404.1984 | 78.82 | 61.143 | 341.4264
[11 empty] | 404.2752 | 77.975 | 61.6038 | 340.5633
[11 empty] | 404.352 | 77.15 | 61.5475 | 336.8646
[11 empty] | 404.4288 | 76.565 | 60.4928 | 330.2376
[11 empty] | 404.5056 | 76.255 | 58.6342 | 321.4839
[11 empty] | 404.5824 | 76.11 | 57.6614 | 311.5281
[11 empty] | 404.6592 | 76.49 | 56.9958 | 301.6031
[11 empty] | 404.736 | 76.965 | 57.1494 | 297.8427
[11 empty] | 404.8128 | 77.09 | 58.2963 | 299.1064
[11 empty] | 404.8896 | 76.93 | 59.2128 | 294.0822
[11 empty] | 404.9664 | 76.445 | 59.7094 | 287.7327
[11 empty] | 405.0432 | 76.185 | 58.6701 | 282.462
[11 empty] | 405.12 | 75.92 | 58.0301 | 278.3625
[11 empty] | 405.1968 | 75.68 | 57.9072 | 275.5576
[11 empty] | 405.2736 | 75.94 | 56.6682 | 273.8624
[11 empty] | 405.3504 | 76.125 | 57.0317 | 272.9993
[11 empty] | 405.4272 | 76.065 | 58.7827 | 272.6911
[11 empty] | 405.504 | 76.085 | 60.247 | 272.9685
[11 empty] | 405.5808 | 76.285 | 59.7965 | 273.7083
[11 empty] | 405.6576 | 76.24 | 59.3306 | 274.5713
[11 empty] | 405.7344 | 75.96 | 59.136 | 276.0816
[11 empty] | 405.8112 | 75.67 | 58.1888 | 284.4963
[11 empty] | 405.888 | 75.41 | 57.344 | 303.4832
[11 empty] | 405.9648 | 75.355 | 56.6221 | 322.4086
[11 empty] | 406.0416 | 75.43 | 56.9549 | 335.5084
[11 empty] | 406.1184 | 75.845 | 57.6614 | 342.9367
[11 empty] | 406.1952 | 76.52 | 58.6547 | 346.543
[11 empty] | 406.272 | 77.16 | 60.6157 | 348.8547
[11 empty] | 406.3488 | 77.785 | 62.5562 | 352.2144
[11 empty] | 406.4256 | 78.465 | 63.4112 | 357.5776
[11 empty] | 406.5024 | 78.845 | 63.7747 | 364.482
[11 empty] | 406.5792 | 78.925 | 64.1894 | 371.2939
[11 empty] | 406.656 | 78.715 | 65.0394 | 376.0098
[11 empty] | 406.7328 | 78.645 | 64.512 | 377.1811
[11 empty] | 406.8096 | 78.58 | 63.8464 | 374.1912
[11 empty] | 406.8864 | 78.44 | 63.6211 | 367.7492
[11 empty] | 406.9632 | 78.34 | 63.6877 | 359.5195
[11 empty] | 407.04 | 78.43 | 63.0835 | 351.3514
[11 empty] | 407.1168 | 78.31 | 62.9811 | 344.6936
[11 empty] | 407.1936 | 78.015 | 62.8326 | 339.9469
[11 empty] | 407.2704 | 77.645 | 62.0851 | 336.5563
[11 empty] | 407.3472 | 77.255 | 60.759 | 333.135
[11 empty] | 407.424 | 76.745 | 60.0013 | 328.3266
[11 empty] | 407.5008 | 76.56 | 58.8032 | 322.1003
[11 empty] | 407.5776 | 76.92 | 58.3987 | 315.011
[11 empty] | 407.6544 | 77.86 | 58.6291 | 306.9971
[11 empty] | 407.7312 | 78.335 | 61.2762 | 299.7228
[11 empty] | 407.808 | 79.08 | 62.5306 | 293.774
[11 empty] | 407.8848 | 79.81 | 63.232 | 289.2122
[11 empty] | 407.9616 | 80.16 | 64.4147 | 285.76
[11 empty] | 408.0384 | 80.46 | 64.4557 | 283.2634
[11 empty] | 408.1152 | 80.58 | 64.0358 | 281.5373
[11 empty] | 408.192 | 80.39 | 64.4915 | 280.3968
[11 empty] | 408.2688 | 80.05 | 65.5514 | 279.7187
[11 empty] | 408.3456 | 80.02 | 65.2902 | 279.1947
[11 empty] | 408.4224 | 80.4 | 63.744 | 278.6091
[11 empty] | 408.4992 | 80.695 | 62.6842 | 277.9618
[11 empty] | 408.576 | 80.59 | 63.6467 | 277.9618
[11 empty] | 408.6528 | 80.585 | 64.896 | 280.4893
[11 empty] | 408.7296 | 80.94 | 66.1299 | 286.6231
[11 empty] | 408.8064 | 81.45 | 67.4304 | 297.0413
[11 empty] | 408.8832 | 81.89 | 67.6966 | 311.5281
[11 empty] | 408.96 | 82.27 | 67.287 | 328.1108
[11 empty] | 409.0368 | 82.55 | 67.8042 | 343.8306
[11 empty] | 409.1136 | 82.8 | 68.4134 | 355.6974
[11 empty] | 409.1904 | 82.995 | 68.5824 | 361.9545
[11 empty] | 409.2672 | 83.03 | 69.3146 | 362.7559
[11 empty] | 409.344 | 83 | 70.5946 | 359.7661
[11 empty] | 409.4208 | 83.275 | 70.953 | 355.4508
[11 empty] | 409.4976 | 83.565 | 71.209 | 351.8445
[11 empty] | 409.5744 | 83.59 | 71.7875 | 349.8102
[11 empty] | 409.6512 | 83.475 | 71.6954 | 349.2246
[11 empty] | 409.728 | 83.42 | 70.0826 | 348.8855
[11 empty] | 409.8048 | 82.97 | 69.0381 | 347.1903
[11 empty] | 409.8816 | 82.315 | 68.1626 | 342.8751
[11 empty] | 409.9584 | 81.5 | 67.5994 | 335.6008
[11 empty] | 410.0352 | 80.915 | 66.2118 | 325.9224
[11 empty] | 410.112 | 80.55 | 65.4746 | 314.7336
[11 empty] | 410.1888 | 80.55 | 65.792 | 302.6819
[11 empty] | 410.2656 | 80.78 | 66.4166 | 292.4795
[11 empty] | 410.3424 | 80.91 | 67.3536 | 291.7089
[11 empty] | 410.4192 | 81.205 | 66.1965 | 290.2602
[11 empty] | 410.496 | 81.03 | 65.6947 | 282.7702
[11 empty] | 410.5728 | 80.61 | 65.1776 | 275.3419
[11 empty] | 410.6496 | 80.215 | 64.5478 | 268.7457
[11 empty] | 410.7264 | 79.875 | 64.0154 | 263.1976
[11 empty] | 410.8032 | 79.79 | 63.4061 | 258.8516
[11 empty] | 410.88 | 79.715 | 63.3958 | 255.6768
[11 empty] | 410.9568 | 79.63 | 63.7082 | 253.5192
[11 empty] | 411.0336 | 79.675 | 63.2986 | 252.2246
[11 empty] | 411.1104 | 79.59 | 62.7251 | 251.7314
[11 empty] | 411.1872 | 79.225 | 62.6688 | 251.7623
[11 empty] | 411.264 | 78.83 | 61.9008 | 252.1013
[11 empty] | 411.3408 | 78.415 | 60.032 | 255.6151
[11 empty] | 411.4176 | 77.7 | 59.4022 | 269.4547
[11 empty] | 411.4944 | 77.265 | 58.4602 | 287.7635
[11 empty] | 411.5712 | 77.135 | 57.641 | 301.6647
[11 empty] | 411.648 | 77.305 | 57.3491 | 310.1718
[11 empty] | 411.7248 | 77.795 | 57.0419 | 315.2576
[11 empty] | 411.8016 | 78.185 | 58.5165 | 319.1722
[11 empty] | 411.8784 | 78.62 | 60.544 | 323.8264
[11 empty] | 411.9552 | 79.06 | 61.8086 | 330.176
[11 empty] | 412.032 | 79.55 | 61.4707 | 337.8817
[11 empty] | 412.1088 | 79.615 | 61.7472 | 345.5875
[11 empty] | 412.1856 | 79.475 | 61.8547 | 351.598
[11 empty] | 412.2624 | 79.385 | 60.9024 | 354.4645
[11 empty] | 412.3392 | 79.125 | 60.8 | 353.6015
[11 empty] | 412.416 | 78.955 | 60.5696 | 349.3479
[11 empty] | 412.4928 | 78.695 | 61.1738 | 343.06
[11 empty] | 412.5696 | 78.31 | 62.2592 | 336.2789
[11 empty] | 412.6464 | 78.285 | 61.1789 | 330.3301
[11 empty] | 412.7232 | 78.22 | 59.5661 | 325.8299
[11 empty] | 412.8 | 77.825 | 58.5011 | 322.4394
[11 empty] | 412.8768 | 77.205 | 57.5386 | 318.9256
[11 empty] | 412.9536 | 76.57 | 56.4838 | 314.0247
[11 empty] | 413.0304 | 75.98 | 55.5674 | 307.3053
[11 empty] | 413.1072 | 75.83 | 54.743 | 299.7228
[11 empty] | 413.184 | 75.925 | 55.7568 | 291.4315
[11 empty] | 413.2608 | 76.085 | 57.7792 | 283.6949
[11 empty] | 413.3376 | 76.7 | 57.9123 | 277.5611
[11 empty] | 413.4144 | 77.23 | 58.1427 | 273.1842
[11 empty] | 413.4912 | 77.525 | 58.5165 | 270.3794
[11 empty] | 413.568 | 77.435 | 59.6019 | 268.6533
[11 empty] | 413.6448 | 77.18 | 60.2061 | 267.6361
[11 empty] | 413.7216 | 77.16 | 59.0694 | 267.1738
[11 empty] | 413.7984 | 76.9 | 58.5574 | 267.0813
[11 empty] | 413.8752 | 76.465 | 59.049 | 267.3279
[11 empty] | 413.952 | 76.22 | 59.1923 | 267.8519
[11 empty] | 414.0288 | 76.505 | 57.7229 | 268.5916
[11 empty] | 414.1056 | 76.45 | 57.6819 | 269.7629
[11 empty] | 414.1824 | 76.38 | 58.2042 | 272.7835
[11 empty] | 414.2592 | 76.69 | 58.9517 | 279.0406
[11 empty] | 414.336 | 77.145 | 60.0576 | 288.9656
[11 empty] | 414.4128 | 77.39 | 61.5526 | 302.836
[11 empty] | 414.4896 | 77.635 | 61.9776 | 319.2955
[11 empty] | 414.5664 | 78.05 | 61.8701 | 335.7241
[11 empty] | 414.6432 | 78 | 62.9811 | 349.2554
[11 empty] | 414.72 | 78.045 | 62.6227 | 357.6701
[11 empty] | 414.7968 | 78.32 | 61.8342 | 360.6291
[11 empty] | 414.8736 | 78.53 | 62.3258 | 359.1804
[11 empty] | 414.9504 | 78.8 | 63.3549 | 355.5125
[11 empty] | 415.0272 | 79.28 | 63.6672 | 351.9062
[11 empty] | 415.104 | 79.51 | 64.6093 | 349.8411
[11 empty] | 415.1808 | 79.65 | 64.681 | 349.6869
[11 empty] | 415.2576 | 79.5 | 64.4813 | 350.6116
[11 empty] | 415.3344 | 79.195 | 63.5187 | 350.9507
[11 empty] | 415.4112 | 78.775 | 62.2029 | 349.1321
[11 empty] | 415.488 | 78.14 | 61.6294 | 344.2929
[11 empty] | 415.5648 | 77.465 | 61.3734 | 336.6796
[11 empty] | 415.6416 | 77.015 | 61.4554 | 327.217
[11 empty] | 415.7184 | 77.24 | 61.4093 | 316.6755
[11 empty] | 415.7952 | 77.83 | 61.8342 | 308.0759
[11 empty] | 415.872 | 78.16 | 63.7338 | 309.4013
[11 empty] | 415.9488 | 78.56 | 64.2765 | 309.4629
[11 empty] | 416.0256 | 78.81 | 63.8208 | 303.36
[11 empty] | 416.1024 | 78.87 | 62.8582 | 297.4111
[11 empty] | 416.1792 | 78.715 | 62.4128 | 292.4178
[11 empty] | 416.256 | 78.735 | 61.7677 | 288.5958
[11 empty] | 416.3328 | 78.73 | 62.3411 | 286.1299
[11 empty] | 416.4096 | 78.855 | 63.7645 | 284.6504
[11 empty] | 416.4864 | 79.12 | 65.0598 | 283.7874
[11 empty] | 416.5632 | 79.59 | 65.3466 | 283.3867
[11 empty] | 416.64 | 80.1 | 64.9523 | 283.3867
[11 empty] | 416.7168 | 80.325 | 64.7834 | 283.7257
[11 empty] | 416.7936 | 80.29 | 64.7475 | 284.0031
[11 empty] | 416.8704 | 80.07 | 64.4762 | 284.6504
[11 empty] | 416.9472 | 79.815 | 64.0819 | 291.4315
[11 empty] | 417.024 | 79.625 | 63.5392 | 308.7232
[11 empty] | 417.1008 | 79.705 | 62.8838 | 326.7546
[11 empty] | 417.1776 | 79.78 | 63.8054 | 339.577
[11 empty] | 417.2544 | 80.01 | 65.4541 | 347.4985
[11 empty] | 417.3312 | 80.72 | 66.4781 | 352.1219
[11 empty] | 417.408 | 81.6 | 68.0141 | 355.42
[11 empty] | 417.4848 | 82.26 | 70.1901 | 359.088
[11 empty] | 417.5616 | 82.92 | 71.1424 | 363.9272
[11 empty] | 417.6384 | 83.335 | 71.3523 | 369.6294
[11 empty] | 417.7152 | 83.395 | 71.2038 | 374.9618
[11 empty] | 417.792 | 83.27 | 70.7123 | 378.3215
[11 empty] | 417.8688 | 83.02 | 70.1798 | 378.4756
[11 empty] | 417.9456 | 82.465 | 71.0861 | 374.931
[11 empty] | 418.0224 | 82.17 | 70.7738 | 368.2732
[11 empty] | 418.0992 | 82.275 | 69.161 | 359.7969
[11 empty] | 418.176 | 81.985 | 69.6627 | 351.1356
[11 empty] | 418.2528 | 81.675 | 69.6269 | 343.4607
[11 empty] | 418.3296 | 81.37 | 68.3315 | 337.2036
[11 empty] | 418.4064 | 80.935 | 66.345 | 331.9329
[11 empty] | 418.4832 | 80.27 | 64.512 | 326.4772
[11 empty] | 418.56 | 79.555 | 62.8941 | 319.8195
[11 empty] | 418.6368 | 79 | 61.5885 | 311.8671
[11 empty] | 418.7136 | 78.69 | 62.1363 | 303.4832
[11 empty] | 418.7904 | 78.915 | 63.1757 | 294.4829
[11 empty] | 418.8672 | 79.475 | 63.6416 | 286.0991
[11 empty] | 418.944 | 79.955 | 64.4557 | 279.0098
[11 empty] | 419.0208 | 80.29 | 65.4285 | 273.2459
[11 empty] | 419.0976 | 80.525 | 65.3158 | 268.6533
[11 empty] | 419.1744 | 80.525 | 64.553 | 265.0161
[11 empty] | 419.2512 | 80.27 | 63.7184 | 262.1496
[11 empty] | 419.328 | 79.715 | 63.4317 | 260.0228
[11 empty] | 419.4048 | 79.245 | 62.5203 | 258.605
[11 empty] | 419.4816 | 78.715 | 61.993 | 257.6803
[11 empty] | 419.5584 | 78.365 | 61.3018 | 257.1871
[11 empty] | 419.6352 | 78.28 | 60.0218 | 256.9713
[11 empty] | 419.712 | 77.975 | 59.9603 | 257.5262
[11 empty] | 419.7888 | 77.675 | 60.713 | 260.5776
[11 empty] | 419.8656 | 77.61 | 61.993 | 267.2046
[11 empty] | 419.9424 | 78.015 | 61.6755 | 277.9002
[11 empty] | 420.0192 | 78.18 | 60.8154 | 292.6028
[11 empty] | 420.096 | 77.945 | 60.503 | 309.278
[11 empty] | 420.1728 | 77.995 | 60.16 | 324.9361
[11 empty] | 420.2496 | 78.145 | 60.201 | 336.8646
[11 empty] | 420.3264 | 78.2 | 60.5184 | 343.6148
[11 empty] | 420.4032 | 78.385 | 60.3597 | 345.6183
[11 empty] | 420.48 | 78.665 | 60.2931 | 344.3237
[11 empty] | 420.5568 | 78.86 | 61.015 | 341.8887
[11 empty] | 420.6336 | 79 | 61.9213 | 340.3167
[11 empty] | 420.7104 | 79.085 | 62.5306 | 340.2859
[11 empty] | 420.7872 | 78.945 | 62.72 | 341.6421
[11 empty] | 420.864 | 78.62 | 62.5818 | 342.8751
[11 empty] | 420.9408 | 78.19 | 61.6806 | 342.4127
[11 empty] | 421.0176 | 77.78 | 60.1395 | 339.1147
[11 empty] | 421.0944 | 77.225 | 58.7674 | 332.6418
[11 empty] | 421.1712 | 76.605 | 57.9738 | 323.8573
[11 empty] | 421.248 | 76.225 | 57.6 | 313.9323
[11 empty] | 421.3248 | 75.97 | 59.1206 | 303.9456
[11 empty] | 421.4016 | 76.235 | 60.1344 | 299.0756
[11 empty] | 421.4784 | 76.795 | 59.5763 | 300.74
[11 empty] | 421.5552 | 76.97 | 59.5763 | 296.3015
[11 empty] | 421.632 | 76.665 | 59.904 | 289.4896
[11 empty] | 421.7088 | 76.365 | 58.7366 | 283.7257
[11 empty] | 421.7856 | 75.875 | 58.368 | 279.318
[11 empty] | 421.8624 | 75.655 | 57.7229 | 276.3898
[11 empty] | 421.9392 | 75.7 | 57.3235 | 274.787
[11 empty] | 422.016 | 75.865 | 57.769 | 273.9857
[11 empty] | 422.0928 | 76.14 | 58.5011 | 273.6466
[11 empty] | 422.1696 | 76.365 | 58.9261 | 273.6158
[11 empty] | 422.2464 | 76.425 | 59.2691 | 273.924
[11 empty] | 422.3232 | 76.525 | 58.1018 | 274.5405
[11 empty] | 422.4 | 76.355 | 57.1392 | 275.8042
[11 empty] | 422.4768 | 75.84 | 57.2314 | 282.9243
[11 empty] | 422.5536 | 75.45 | 56.8013 | 301.2024
[11 empty] | 422.6304 | 75.195 | 56.8166 | 320.8983
[11 empty] | 422.7072 | 75.44 | 56.3814 | 334.8611
[11 empty] | 422.784 | 75.94 | 56.4224 | 342.8442
[11 empty] | 422.8608 | 76.445 | 57.897 | 346.6663
[11 empty] | 422.9376 | 76.93 | 60.5542 | 348.9164
[11 empty] | 423.0144 | 77.645 | 62.1312 | 351.9678
[11 empty] | 423.0912 | 78.44 | 62.3923 | 357.0844
[11 empty] | 423.168 | 78.95 | 62.551 | 363.958
[11 empty] | 423.2448 | 78.965 | 63.8259 | 371.0164
[11 empty] | 423.3216 | 78.94 | 63.9078 | 376.2872
[11 empty] | 423.3984 | 78.68 | 64.041 | 378.0749
[11 empty] | 423.4752 | 78.4 | 64.0666 | 375.7015
[11 empty] | 423.552 | 78.45 | 63.2064 | 369.5678
[11 empty] | 423.6288 | 78.615 | 62.3616 | 361.338
[11 empty] | 423.7056 | 78.53 | 62.6074 | 352.9542
[11 empty] | 423.7824 | 78.455 | 62.5766 | 345.8957
[11 empty] | 423.8592 | 78.225 | 62.4794 | 340.8716
[11 empty] | 423.936 | 77.715 | 62.5715 | 337.481
[11 empty] | 424.0128 | 77.43 | 60.7898 | 334.2754
[11 empty] | 424.0896 | 77.25 | 58.5882 | 329.7444
[11 empty] | 424.1664 | 76.92 | 57.943 | 323.7032
[11 empty] | 424.2432 | 76.88 | 58.8749 | 316.7372
[11 empty] | 424.32 | 77.27 | 61.1891 | 308.754
[11 empty] | 424.3968 | 78.1 | 62.4077 | 301.264
[11 empty] | 424.4736 | 78.82 | 64.1843 | 295.0994
[11 empty] | 424.5504 | 79.565 | 65.0957 | 290.2602
[11 empty] | 424.6272 | 80.19 | 65.1981 | 286.5614
[11 empty] | 424.704 | 80.48 | 65.3056 | 283.8798
[11 empty] | 424.7808 | 80.335 | 66.1453 | 282.0613
[11 empty] | 424.8576 | 80.04 | 66.7955 | 280.8592
[11 empty] | 424.9344 | 80.005 | 65.8227 | 280.0578
[11 empty] | 425.0112 | 80.085 | 64.8038 | 279.4105
[11 empty] | 425.088 | 80.175 | 64.2509 | 278.7324
[11 empty] | 425.1648 | 80.325 | 64.3174 | 277.9926
[11 empty] | 425.2416 | 80.44 | 64.7782 | 277.8385
[11 empty] | 425.3184 | 80.68 | 64.8243 | 280.0886
[11 empty] | 425.3952 | 81.16 | 65.0803 | 285.945
[11 empty] | 425.472 | 81.61 | 66.3757 | 295.9624
[11 empty] | 425.5488 | 82.045 | 66.9286 | 310.0794
[11 empty] | 425.6256 | 82.165 | 67.497 | 326.5388
[11 empty] | 425.7024 | 82.435 | 68.2086 | 342.3511
[11 empty] | 425.7792 | 82.595 | 69.6371 | 354.6803
[11 empty] | 425.856 | 82.685 | 70.7226 | 361.523
[11 empty] | 425.9328 | 82.895 | 70.5638 | 362.8792
[11 empty] | 426.0096 | 83.29 | 69.6166 | 360.2592
[11 empty] | 426.0864 | 83.585 | 69.4733 | 355.9132
[11 empty] | 426.1632 | 83.725 | 70.2874 | 351.9678
[11 empty] | 426.24 | 83.83 | 70.5946 | 349.6561
[11 empty] | 426.3168 | 83.92 | 70.0314 | 348.8855
[11 empty] | 426.3936 | 83.67 | 69.6269 | 348.5773
[11 empty] | 426.4704 | 83.14 | 69.2326 | 347.1595
[11 empty] | 426.5472 | 82.415 | 68.6797 | 343.3374
[11 empty] | 426.624 | 81.825 | 67.0515 | 336.5255
[11 empty] | 426.7008 | 81.34 | 64.8755 | 326.9088
[11 empty] | 426.7776 | 80.93 | 63.9078 | 315.4426
[11 empty] | 426.8544 | 80.755 | 64.7066 | 302.9593
[11 empty] | 426.9312 | 80.91 | 65.7971 | 292.2945
[11 empty] | 427.008 | 81.205 | 66.263 | 291.863
[11 empty] | 427.0848 | 81.36 | 66.432 | 291.4006
[11 empty] | 427.1616 | 81.45 | 64.9318 | 283.7874
[11 empty] | 427.2384 | 81.245 | 63.3805 | 275.8658
[11 empty] | 427.3152 | 80.85 | 62.5818 | 268.8074
[11 empty] | 427.392 | 80.645 | 61.655 | 262.9818
[11 empty] | 427.4688 | 80.445 | 61.4605 | 258.5742
[11 empty] | 427.5456 | 80.27 | 61.9213 | 255.5227
[11 empty] | 427.6224 | 80.135 | 62.7405 | 253.4267
[11 empty] | 427.6992 | 80.095 | 63.0016 | 252.0397
[11 empty] | 427.776 | 79.945 | 62.9453 | 251.2383
[11 empty] | 427.8528 | 79.59 | 62.633 | 250.93
[11 empty] | 427.9296 | 79.27 | 61.0509 | 250.8992
[11 empty] | 428.0064 | 78.775 | 59.5098 | 253.8582
[11 empty] | 428.0832 | 78.2 | 58.1888 | 267.0813
[11 empty] | 428.16 | 77.57 | 57.9482 | 285.5443
[11 empty] | 428.2368 | 77.175 | 58.112 | 299.8153
[11 empty] | 428.3136 | 77.13 | 58.5933 | 308.6615
[11 empty] | 428.3904 | 77.55 | 58.6138 | 313.809
[11 empty] | 428.4672 | 78.29 | 58.583 | 317.5077
[11 empty] | 428.544 | 78.8 | 59.9296 | 321.9771
[11 empty] | 428.6208 | 79.075 | 61.7421 | 328.2033
[11 empty] | 428.6976 | 79.375 | 62.6432 | 335.9399
[11 empty] | 428.7744 | 79.77 | 61.5782 | 343.7997
[11 empty] | 428.8512 | 79.695 | 61.3837 | 349.9027
[11 empty] | 428.928 | 79.28 | 61.8189 | 352.9233
[11 empty] | 429.0048 | 79.095 | 61.0099 | 352.1836
[11 empty] | 429.0816 | 78.915 | 60.6464 | 348.0841
[11 empty] | 429.1584 | 78.73 | 60.6822 | 341.8887
[11 empty] | 429.2352 | 78.55 | 60.9997 | 335.231
[11 empty] | 429.312 | 78.345 | 61.2096 | 329.3438
[11 empty] | 429.3888 | 78.185 | 60.5338 | 324.7511
[11 empty] | 429.4656 | 77.915 | 58.8032 | 321.1448
[11 empty] | 429.5424 | 77.425 | 57.129 | 317.4461
[11 empty] | 429.6192 | 76.94 | 55.3114 | 312.4836
[11 empty] | 429.696 | 76.41 | 54.1184 | 305.8875
[11 empty] | 429.7728 | 75.87 | 54.7123 | 298.5824
[11 empty] | 429.8496 | 75.745 | 56.4378 | 290.6917
[11 empty] | 429.9264 | 76.165 | 57.5283 | 283.1401
[11 empty] | 430.0032 | 76.74 | 58.0403 | 276.9138
[11 empty] | 430.08 | 77.31 | 58.3424 | 272.1671
[11 empty] | 430.1568 | 77.655 | 58.5626 | 268.8074
[11 empty] | 430.2336 | 77.425 | 59.9552 | 266.6498
[11 empty] | 430.3104 | 77.17 | 60.5594 | 265.5093
[11 empty] | 430.3872 | 77.115 | 59.4483 | 265.1703
[11 empty] | 430.464 | 76.965 | 58.3526 | 265.3244
[11 empty] | 430.5408 | 76.975 | 56.9651 | 265.6943
[11 empty] | 430.6176 | 76.82 | 57.2058 | 266.095
[11 empty] | 430.6944 | 76.56 | 58.3168 | 266.5265
[11 empty] | 430.7712 | 76.38 | 58.8902 | 267.3279
[11 empty] | 430.848 | 76.515 | 58.3373 | 270.102
[11 empty] | 430.9248 | 76.805 | 59.0234 | 276.3282
[11 empty] | 431.0016 | 77.175 | 60.5696 | 286.4381
[11 empty] | 431.0784 | 77.77 | 60.5286 | 300.5551
[11 empty] | 431.1552 | 77.92 | 61.3222 | 317.1378
[11 empty] | 431.232 | 78.285 | 61.4451 | 333.474
[11 empty] | 431.3088 | 78.49 | 61.4554 | 346.6971
[11 empty] | 431.3856 | 78.375 | 61.9725 | 354.9577
[11 empty] | 431.4624 | 78.44 | 62.0083 | 357.9783
[11 empty] | 431.5392 | 78.53 | 62.6278 | 356.9612
[11 empty] | 431.616 | 78.77 | 63.7542 | 353.8481
[11 empty] | 431.6928 | 79.345 | 63.7798 | 350.6116
[11 empty] | 431.7696 | 79.525 | 64.9626 | 348.6081
[11 empty] | 431.8464 | 79.66 | 65.2186 | 348.2074
[11 empty] | 431.9232 | 79.695 | 64.343 | 348.7006
[11 empty] | 432 | 79.345 | 63.8157 | 348.8239
[11 empty] | 432.0768 | 78.745 | 63.5904 | 347.0978
[11 empty] | 432.1536 | 78.205 | 62.4691 | 342.7209
[11 empty] | 432.2304 | 77.675 | 61.2762 | 335.7549
[11 empty] | 432.3072 | 77.41 | 60.3955 | 326.7854
[11 empty] | 432.384 | 77.53 | 60.5235 | 316.3365
[11 empty] | 432.4608 | 77.93 | 61.7267 | 306.8121
[11 empty] | 432.5376 | 78.395 | 62.9965 | 306.843
[11 empty] | 432.6144 | 78.805 | 63.5085 | 307.9834
[11 empty] | 432.6912 | 78.87 | 64.128 | 302.3428
[11 empty] | 432.768 | 78.875 | 63.5546 | 296.3323
[11 empty] | 432.8448 | 78.845 | 62.4077 | 291.2157
[11 empty] | 432.9216 | 78.96 | 61.225 | 287.3012
[11 empty] | 432.9984 | 79.045 | 61.44 | 284.6196
[11 empty] | 433.0752 | 79.115 | 62.7712 | 283.0168
[11 empty] | 433.152 | 79.375 | 64.2304 | 282.1229
[11 empty] | 433.2288 | 79.8 | 64.9728 | 281.6914
[11 empty] | 433.3056 | 80.175 | 65.3773 | 281.6914
[11 empty] | 433.3824 | 80.28 | 65.879 | 281.9996
[11 empty] | 433.4592 | 80.28 | 65.6282 | 282.3079
[11 empty] | 433.536 | 79.995 | 65.4746 | 282.8627
[11 empty] | 433.6128 | 79.85 | 64.2867 | 289.3972
[11 empty] | 433.6896 | 79.78 | 63.1757 | 306.4731
[11 empty] | 433.7664 | 79.53 | 63.8464 | 324.4429
[11 empty] | 433.8432 | 79.635 | 64.7629 | 337.2961
[11 empty] | 433.92 | 80.25 | 64.7834 | 345.2176
[11 empty] | 433.9968 | 80.94 | 66.0634 | 349.9335
[11 empty] | 434.0736 | 81.645 | 68.5261 | 353.4165
[11 empty] | 434.1504 | 82.35 | 70.3846 | 357.3002
[11 empty] | 434.2272 | 83.035 | 71.1782 | 362.2935
[11 empty] | 434.304 | 83.435 | 71.3011 | 367.965
[11 empty] | 434.3808 | 83.47 | 71.1373 | 373.1124
[11 empty] | 434.4576 | 83.355 | 70.5997 | 376.2563
[11 empty] | 434.5344 | 83.135 | 70.0006 | 376.4105
[11 empty] | 434.6112 | 82.765 | 70.1952 | 373.1432
[11 empty] | 434.688 | 82.655 | 69.3504 | 366.9478
[11 empty] | 434.7648 | 82.435 | 69.3965 | 358.9955
[11 empty] | 434.8416 | 82.25 | 69.2634 | 350.7041
[11 empty] | 434.9184 | 81.93 | 69.1866 | 343.1216
[11 empty] | 434.9952 | 81.745 | 67.2717 | 336.5255
[11 empty] | 435.072 | 81.325 | 65.5104 | 330.6692
[11 empty] | 435.1488 | 80.695 | 63.5443 | 324.7511
[11 empty] | 435.2256 | 79.915 | 62.1722 | 318.0009
[11 empty] | 435.3024 | 79.215 | 61.3222 | 310.1718
[11 empty] | 435.3792 | 78.925 | 61.3683 | 302.1887
[11 empty] | 435.456 | 79.11 | 62.7251 | 293.7432
[11 empty] | 435.5328 | 79.595 | 63.7645 | 285.5135
[11 empty] | 435.6096 | 80.21 | 64.2202 | 278.3625
[11 empty] | 435.6864 | 80.32 | 66.0224 | 272.3212
[11 empty] | 435.7632 | 80.32 | 66.8467 | 267.2971
[11 empty] | 435.84 | 80.51 | 65.321 | 263.3209
[11 empty] | 435.9168 | 80.42 | 63.5802 | 260.3311
[11 empty] | 435.9936 | 80.02 | 62.5459 | 258.3276
[11 empty] | 436.0704 | 79.39 | 62.5357 | 257.1871
[11 empty] | 436.1472 | 78.81 | 62.3667 | 256.509
[11 empty] | 436.224 | 78.55 | 61.1994 | 256.0775
[11 empty] | 436.3008 | 78.025 | 61.6397 | 255.7384
[11 empty] | 436.3776 | 77.87 | 60.8614 | 255.9542
[11 empty] | 436.4544 | 77.67 | 61.1584 | 258.605
[11 empty] | 436.5312 | 77.915 | 61.1123 | 265.0161
[11 empty] | 436.608 | 78.28 | 61.1584 | 275.7426
[11 empty] | 436.6848 | 78.32 | 61.6653 | 290.5068
[11 empty] | 436.7616 | 78.34 | 60.7539 | 307.3053
[11 empty] | 436.8384 | 78.34 | 60.8154 | 322.9942
[11 empty] | 436.9152 | 78.27 | 61.7114 | 334.9227
[11 empty] | 436.992 | 78.17 | 62.2746 | 341.7346
[11 empty] | 437.0688 | 78.16 | 62.8173 | 343.7689
[11 empty] | 437.1456 | 78.335 | 62.9402 | 342.7209
[11 empty] | 437.2224 | 78.6 | 63.1757 | 340.625
[11 empty] | 437.2992 | 79.055 | 62.6278 | 339.2688
[11 empty] | 437.376 | 79.32 | 62.2336 | 339.2996
[11 empty] | 437.4528 | 79.31 | 61.8803 | 340.3784
[11 empty] | 437.5296 | 78.86 | 62.3616 | 341.3031
[11 empty] | 437.6064 | 78.305 | 61.8701 | 340.7174
[11 empty] | 437.6832 | 77.83 | 60.4723 | 337.5427
[11 empty] | 437.76 | 77.22 | 59.5302 | 331.5014
[11 empty] | 437.8368 | 76.785 | 57.8714 | 323.1175
[11 empty] | 437.9136 | 76.595 | 56.361 | 313.5007
[11 empty] | 437.9904 | 76.45 | 57.3542 | 303.4216
[11 empty] | 438.0672 | 76.71 | 58.1427 | 297.5344
[11 empty] | 438.144 | 77.035 | 58.5011 | 298.7057
[11 empty] | 438.2208 | 77.015 | 58.9466 | 294.7295
[11 empty] | 438.2976 | 76.715 | 59.049 | 288.2875
[11 empty] | 438.3744 | 76.425 | 58.2093 | 282.8935
[11 empty] | 438.4512 | 76.11 | 57.3952 | 278.7016
[11 empty] | 438.528 | 75.92 | 56.9037 | 275.6193
[11 empty] | 438.6048 | 75.71 | 57.8304 | 273.6774
[11 empty] | 438.6816 | 75.88 | 58.2246 | 272.537
[11 empty] | 438.7584 | 76.27 | 57.7792 | 272.0746
[11 empty] | 438.8352 | 76.285 | 58.6086 | 272.1979
[11 empty] | 438.912 | 76.41 | 58.6445 | 272.9377
[11 empty] | 438.9888 | 76.435 | 58.2195 | 273.9857
[11 empty] | 439.0656 | 76.175 | 57.8662 | 275.3419
[11 empty] | 439.1424 | 75.93 | 56.7603 | 281.6914
[11 empty] | 439.2192 | 75.535 | 56.3405 | 299.1989
[11 empty] | 439.296 | 75.185 | 56.7808 | 319.0489
[11 empty] | 439.3728 | 75.17 | 57.1341 | 333.4432
[11 empty] | 439.4496 | 75.55 | 57.7229 | 341.8271
[11 empty] | 439.5264 | 76.32 | 58.1786 | 345.9265
[11 empty] | 439.6032 | 77.135 | 59.4534 | 348.2999
[11 empty] | 439.68 | 77.775 | 61.5731 | 351.3822
[11 empty] | 439.7568 | 78.36 | 62.9197 | 356.4372
[11 empty] | 439.8336 | 78.645 | 64.3482 | 363.2491
[11 empty] | 439.9104 | 78.84 | 64.7424 | 370.3075
[11 empty] | 439.9872 | 78.86 | 64.1843 | 375.6091
[11 empty] | 440.064 | 78.63 | 64.1178 | 377.5201
[11 empty] | 440.1408 | 78.485 | 63.703 | 375.3008
[11 empty] | 440.2176 | 78.49 | 62.9197 | 369.4445
[11 empty] | 440.2944 | 78.415 | 63.0989 | 361.4613
[11 empty] | 440.3712 | 78.505 | 62.8582 | 353.2316
[11 empty] | 440.448 | 78.49 | 62.3462 | 346.2964
[11 empty] | 440.5248 | 78.225 | 62.1261 | 341.2722
[11 empty] | 440.6016 | 77.925 | 61.3427 | 337.6968
[11 empty] | 440.6784 | 77.425 | 60.6822 | 334.3371
[11 empty] | 440.7552 | 76.925 | 59.9808 | 329.8061
[11 empty] | 440.832 | 76.785 | 58.3731 | 323.734
[11 empty] | 440.9088 | 76.755 | 59.1923 | 316.768
[11 empty] | 440.9856 | 77.45 | 59.7094 | 308.8156
[11 empty] | 441.0624 | 78.36 | 60.3136 | 301.3873
[11 empty] | 441.1392 | 79.02 | 62.3002 | 295.2535
[11 empty] | 441.216 | 79.685 | 63.7184 | 290.476
[11 empty] | 441.2928 | 80.105 | 64.937 | 286.8697
[11 empty] | 441.3696 | 80.495 | 64.681 | 284.2497
[11 empty] | 441.4464 | 80.685 | 64.2458 | 282.462
[11 empty] | 441.5232 | 80.635 | 64.2406 | 281.2291
[11 empty] | 441.6 | 80.46 | 64.3123 | 280.3968
[11 empty] | 441.6768 | 80.3 | 64.7475 | 279.8112
[11 empty] | 441.7536 | 80.315 | 64.855 | 279.318
[11 empty] | 441.8304 | 80.405 | 64.9574 | 278.7324
[11 empty] | 441.9072 | 80.49 | 65.0394 | 278.5166
[11 empty] | 441.984 | 80.49 | 66.0122 | 280.3044
[11 empty] | 442.0608 | 80.99 | 66.4013 | 285.6059
[11 empty] | 442.1376 | 81.615 | 67.0157 | 295.0994
[11 empty] | 442.2144 | 81.965 | 67.8093 | 308.7848
[11 empty] | 442.2912 | 82.085 | 68.439 | 325.1826
[11 empty] | 442.368 | 82.33 | 69.1814 | 341.3647
[11 empty] | 442.4448 | 82.615 | 69.9034 | 354.2487
[11 empty] | 442.5216 | 82.795 | 70.3642 | 361.6771
[11 empty] | 442.5984 | 83.06 | 69.9699 | 363.3415
[11 empty] | 442.6752 | 83.255 | 70.0928 | 360.7216
[11 empty] | 442.752 | 83.495 | 70.4358 | 356.2831
[11 empty] | 442.8288 | 83.715 | 70.7226 | 352.3377
[11 empty] | 442.9056 | 83.845 | 70.953 | 349.9952
[11 empty] | 442.9824 | 83.94 | 70.1798 | 349.317
[11 empty] | 443.0592 | 83.58 | 70.2106 | 349.2554
[11 empty] | 443.136 | 82.885 | 70.4102 | 348.115
[11 empty] | 443.2128 | 82.38 | 68.9152 | 344.2313
[11 empty] | 443.2896 | 81.825 | 67.3894 | 337.2653
[11 empty] | 443.3664 | 81.1 | 66.4678 | 327.8026
[11 empty] | 443.4432 | 80.63 | 65.577 | 316.7372
[11 empty] | 443.52 | 80.78 | 64.5888 | 304.9011
[11 empty] | 443.5968 | 80.99 | 65.2493 | 294.298
[11 empty] | 443.6736 | 81.375 | 65.1827 | 292.4486
[11 empty] | 443.7504 | 81.615 | 64.7424 | 291.6164
[11 empty] | 443.8272 | 81.42 | 64.681 | 284.2189
[11 empty] | 443.904 | 80.89 | 65.024 | 276.6672
[11 empty] | 443.9808 | 80.365 | 64.9574 | 270.0403
[11 empty] | 444.0576 | 80.145 | 63.6314 | 264.4613
[11 empty] | 444.1344 | 79.96 | 63.0067 | 260.0228
[11 empty] | 444.2112 | 79.84 | 63.3139 | 256.6015
[11 empty] | 444.288 | 79.845 | 63.2422 | 254.1356
[11 empty] | 444.3648 | 79.745 | 63.6109 | 252.502
[11 empty] | 444.4416 | 79.59 | 63.5187 | 251.7006
[11 empty] | 444.5184 | 79.435 | 62.3821 | 251.6082
[11 empty] | 444.5952 | 79.09 | 61.3069 | 251.8856
[11 empty] | 444.672 | 78.545 | 60.416 | 254.2281
[11 empty] | 444.7488 | 77.96 | 59.4944 | 265.8792
[11 empty] | 444.8256 | 77.57 | 58.1018 | 284.373
[11 empty] | 444.9024 | 77.125 | 58.4704 | 299.4146
[11 empty] | 444.9792 | 77.265 | 57.9072 | 308.8773
[11 empty] | 445.056 | 77.745 | 57.344 | 314.3638
[11 empty] | 445.1328 | 77.98 | 59.2282 | 318.2166
[11 empty] | 445.2096 | 78.59 | 60.1702 | 322.501
[11 empty] | 445.2864 | 79.315 | 60.3648 | 328.4191
[11 empty] | 445.3632 | 79.71 | 60.9946 | 335.9091
[11 empty] | 445.44 | 79.675 | 62.1005 | 343.7997
[11 empty] | 445.5168 | 79.485 | 62.3411 | 350.2726
[11 empty] | 445.5936 | 79.32 | 61.4093 | 353.8789
[11 empty] | 445.6704 | 79.01 | 61.2045 | 353.7556
[11 empty] | 445.7472 | 79.045 | 60.1242 | 350.1493
[11 empty] | 445.824 | 78.95 | 59.776 | 344.108
[11 empty] | 445.9008 | 78.635 | 60.8256 | 337.1728
[11 empty] | 445.9776 | 78.205 | 62.1978 | 330.8233
[11 empty] | 446.0544 | 77.75 | 62.7763 | 325.9224
[11 empty] | 446.1312 | 77.44 | 61.3069 | 322.3161
[11 empty] | 446.208 | 76.985 | 59.4586 | 319.018
[11 empty] | 446.2848 | 76.55 | 57.1034 | 314.5795
[11 empty] | 446.3616 | 76.05 | 55.3421 | 308.2917
[11 empty] | 446.4384 | 75.555 | 55.6134 | 300.9558
[11 empty] | 446.5152 | 75.66 | 56.2637 | 292.7569
[11 empty] | 446.592 | 76.175 | 56.5709 | 284.6504
[11 empty] | 446.6688 | 76.74 | 57.2109 | 277.9926
[11 empty] | 446.7456 | 77.175 | 57.9789 | 273.2151
[11 empty] | 446.8224 | 77.595 | 57.8611 | 270.102
[11 empty] | 446.8992 | 77.54 | 59.095 | 268.2526
[11 empty] | 446.976 | 77.325 | 59.8579 | 267.2354
[11 empty] | 447.0528 | 77.25 | 59.0541 | 266.6806
[11 empty] | 447.1296 | 76.97 | 58.7571 | 266.434
[11 empty] | 447.2064 | 76.655 | 58.5318 | 266.4032
[11 empty] | 447.2832 | 76.405 | 58.5062 | 266.6806
[11 empty] | 447.36 | 76.21 | 58.9773 | 267.2971
[11 empty] | 447.4368 | 76.215 | 58.5933 | 268.3759
[11 empty] | 447.5136 | 76.3 | 58.2502 | 271.2116
[11 empty] | 447.5904 | 76.675 | 58.1376 | 277.1604
[11 empty] | 447.6672 | 77.34 | 58.5062 | 286.7155
[11 empty] | 447.744 | 77.765 | 59.6224 | 300.0619
[11 empty] | 447.8208 | 77.93 | 60.4774 | 316.2132
[11 empty] | 447.8976 | 78.235 | 60.6771 | 332.7343
[11 empty] | 447.9744 | 78.385 | 61.3069 | 346.8204
[11 empty] | 448.0512 | 78.28 | 61.9213 | 356.1906
[11 empty] | 448.128 | 78.165 | 62.7558 | 360.1051
[11 empty] | 448.2048 | 78.32 | 62.9862 | 359.3962
[11 empty] | 448.2816 | 78.6 | 63.6979 | 355.9748
[11 empty] | 448.3584 | 79.15 | 63.7901 | 352.1219
[11 empty] | 448.4352 | 79.24 | 65.4336 | 349.5636
[11 empty] | 448.512 | 79.395 | 65.4848 | 349.0705
[11 empty] | 448.5888 | 79.49 | 64.5837 | 349.9644
[11 empty] | 448.6656 | 79.3 | 63.5546 | 350.7041
[11 empty] | 448.7424 | 78.77 | 62.935 | 349.5945
[11 empty] | 448.8192 | 78.115 | 62.2899 | 345.495
[11 empty] | 448.896 | 77.615 | 61.1277 | 338.3749
[11 empty] | 448.9728 | 77.335 | 59.863 | 329.0047
[11 empty] | 449.0496 | 77.61 | 59.008 | 318.3399
[11 empty] | 449.1264 | 78.065 | 59.945 | 308.754
[11 empty] | 449.2032 | 78.61 | 61.1635 | 308.8773
[11 empty] | 449.28 | 79.125 | 61.6602 | 310.4492
[11 empty] | 449.3568 | 79.1 | 62.5562 | 304.7778
[11 empty] | 449.4336 | 78.84 | 63.1142 | 298.6132
[11 empty] | 449.5104 | 78.705 | 62.5408 | 293.3425
[11 empty] | 449.5872 | 78.53 | 62.6278 | 289.2739
[11 empty] | 449.664 | 78.71 | 62.4077 | 286.5306
[11 empty] | 449.7408 | 79.07 | 62.4179 | 284.8354
[11 empty] | 449.8176 | 79.25 | 64.0512 | 283.8798
[11 empty] | 449.8944 | 79.65 | 64.7936 | 283.3867
[11 empty] | 449.9712 | 79.99 | 65.2902 | 283.325
[11 empty] | 450.048 | 80.25 | 65.2339 | 283.5716
[11 empty] | 450.1248 | 80.285 | 65.0906 | 283.849
[11 empty] | 450.2016 | 80.1 | 64.809 | 284.2805
[11 empty] | 450.2784 | 79.725 | 64.6144 | 289.4896
[11 empty] | 450.3552 | 79.555 | 63.7645 | 305.5792
[11 empty] | 450.432 | 79.405 | 63.8157 | 324.1655
[11 empty] | 450.5088 | 79.84 | 62.8122 | 337.9434
[11 empty] | 450.5856 | 80.31 | 63.5955 | 346.543
[11 empty] | 450.6624 | 80.69 | 66.3142 | 351.598
[11 empty] | 450.7392 | 81.47 | 68.2189 | 354.9268
[11 empty] | 450.816 | 82.37 | 69.4221 | 358.3174
[11 empty] | 450.8928 | 82.965 | 70.6355 | 362.7867
[11 empty] | 450.9696 | 83.195 | 71.424 | 368.2424
[11 empty] | 451.0464 | 83.34 | 70.6662 | 373.6672
[11 empty] | 451.1232 | 83.36 | 69.8675 | 377.5201
[11 empty] | 451.2 | 82.785 | 71.3626 | 378.4448
[11 empty] | 451.2768 | 82.3 | 71.9923 | 375.7015
[11 empty] | 451.3536 | 82.29 | 70.615 | 369.5061
[11 empty] | 451.4304 | 82.245 | 69.9187 | 361.0914
[11 empty] | 451.5072 | 82.065 | 69.9597 | 352.0295
[11 empty] | 451.584 | 81.84 | 69.6422 | 343.8922
[11 empty] | 451.6608 | 81.59 | 68.352 | 337.2653
[11 empty] | 451.7376 | 81.26 | 66.2221 | 331.9329
[11 empty] | 451.8144 | 80.66 | 64.1638 | 326.7854
[11 empty] | 451.8912 | 79.82 | 63.0426 | 320.7441
[11 empty] | 451.968 | 79.045 | 62.4589 | 313.2542
[11 empty] | 452.0448 | 78.84 | 61.9264 | 305.1477
[11 empty] | 452.1216 | 79.25 | 61.8138 | 296.1474
[11 empty] | 452.1984 | 79.76 | 62.6842 | 287.2704
[11 empty] | 452.2752 | 80.175 | 64.1434 | 279.6571
[11 empty] | 452.352 | 80.465 | 65.111 | 273.4925
[11 empty] | 452.4288 | 80.51 | 65.9405 | 268.6224
[11 empty] | 452.5056 | 80.38 | 66.1402 | 264.8004
[11 empty] | 452.5824 | 80.225 | 64.7834 | 261.9339
[11 empty] | 452.6592 | 79.905 | 63.2986 | 259.8071
[11 empty] | 452.736 | 79.425 | 62.1926 | 258.2968
[11 empty] | 452.8128 | 78.88 | 61.6806 | 257.1871
[11 empty] | 452.8896 | 78.49 | 61.1482 | 256.4474
[11 empty] | 452.9664 | 78.29 | 60.503 | 256.0775
[11 empty] | 453.0432 | 78.07 | 60.0781 | 256.3857
[11 empty] | 453.12 | 78 | 59.735 | 259.0981
[11 empty] | 453.1968 | 77.95 | 60.8051 | 265.386
[11 empty] | 453.2736 | 78.17 | 61.6294 | 275.6501
[11 empty] | 453.3504 | 78.465 | 60.9485 | 289.8287
[11 empty] | 453.4272 | 78.225 | 61.0253 | 306.2573
[11 empty] | 453.504 | 78.11 | 61.1789 | 322.1312
[11 empty] | 453.5808 | 78.27 | 60.9638 | 334.707
[11 empty] | 453.6576 | 78.415 | 60.6874 | 342.2278
[11 empty] | 453.7344 | 78.43 | 61.0406 | 344.7552
[11 empty] | 453.8112 | 78.305 | 62.6893 | 343.8614
[11 empty] | 453.888 | 78.62 | 62.9248 | 341.6113
[11 empty] | 453.9648 | 79.145 | 61.8598 | 339.7928
[11 empty] | 454.0416 | 79.31 | 61.7062 | 339.3612
[11 empty] | 454.1184 | 79.12 | 62.3514 | 340.3167
[11 empty] | 454.1952 | 78.79 | 62.3872 | 341.5805
[11 empty] | 454.272 | 78.31 | 61.8035 | 341.3955
[11 empty] | 454.3488 | 77.705 | 60.8819 | 338.714
[11 empty] | 454.4256 | 77.215 | 59.433 | 332.8268
[11 empty] | 454.5024 | 76.435 | 59.2538 | 324.3504
[11 empty] | 454.5792 | 76.06 | 58.1786 | 314.5487
[11 empty] | 454.656 | 76.14 | 58.0454 | 304.2847
[11 empty] | 454.7328 | 76.39 | 59.0029 | 297.9043
[11 empty] | 454.8096 | 76.78 | 59.2538 | 299.6304
[11 empty] | 454.8864 | 76.965 | 59.305 | 296.5789
[11 empty] | 454.9632 | 76.785 | 59.4534 | 289.7979
[11 empty] | 455.04 | 76.57 | 58.5165 | 283.7565
[11 empty] | 455.1168 | 76.205 | 57.7075 | 278.9482
[11 empty] | 455.1936 | 75.94 | 57.0573 | 275.5576
[11 empty] | 455.2704 | 75.945 | 56.6067 | 273.6158
[11 empty] | 455.3472 | 75.81 | 57.9379 | 272.7835
[11 empty] | 455.424 | 75.855 | 59.3101 | 272.537
[11 empty] | 455.5008 | 76.445 | 58.2758 | 272.6294
[11 empty] | 455.5776 | 76.685 | 58.0864 | 272.9993
[11 empty] | 455.6544 | 76.43 | 58.6752 | 273.4925
[11 empty] | 455.7312 | 76.44 | 56.6835 | 274.3555
[11 empty] | 455.808 | 75.98 | 56.1766 | 280.0886
[11 empty] | 455.8848 | 75.43 | 56.0128 | 297.1954
[11 empty] | 455.9616 | 75.165 | 55.5674 | 317.4769
[11 empty] | 456.0384 | 75.42 | 55.0195 | 332.4261
[11 empty] | 456.1152 | 75.595 | 56.5862 | 341.2722
[11 empty] | 456.192 | 76.115 | 58.2246 | 345.6183
[11 empty] | 456.2688 | 77.02 | 59.0234 | 348.0225
[11 empty] | 456.3456 | 77.595 | 61.3478 | 350.9199
[11 empty] | 456.4224 | 78.065 | 63.2166 | 355.6666
[11 empty] | 456.4992 | 78.56 | 63.5904 | 362.2627
[11 empty] | 456.576 | 78.995 | 63.0118 | 369.352
[11 empty] | 456.6528 | 79.085 | 63.0784 | 374.9001
[11 empty] | 456.7296 | 79.06 | 62.7098 | 377.2735
[11 empty] | 456.8064 | 78.85 | 62.5715 | 375.6399
[11 empty] | 456.8832 | 78.765 | 61.9674 | 370.1842
[11 empty] | 456.96 | 78.645 | 61.8752 | 362.2935
[11 empty] | 457.0368 | 78.51 | 62.08 | 353.7248
[11 empty] | 457.1136 | 78.52 | 61.7933 | 346.3272
[11 empty] | 457.1904 | 78.37 | 61.5834 | 340.964
[11 empty] | 457.2672 | 77.905 | 61.7779 | 337.2653
[11 empty] | 457.344 | 77.56 | 60.6208 | 334.0597
[11 empty] | 457.4208 | 77.365 | 58.5011 | 329.8369
[11 empty] | 457.4976 | 76.955 | 57.984 | 323.9806
[11 empty] | 457.5744 | 76.73 | 59.0029 | 317.292
[11 empty] | 457.6512 | 77.425 | 59.52 | 309.4629
[11 empty] | 457.728 | 78.36 | 60.2317 | 301.8188
[11 empty] | 457.8048 | 79.16 | 61.5526 | 295.4384
[11 empty] | 457.8816 | 79.83 | 62.8634 | 290.4143
[11 empty] | 457.9584 | 80.315 | 63.7389 | 286.5923
[11 empty] | 458.0352 | 80.51 | 64.6451 | 283.7565
[11 empty] | 458.112 | 80.525 | 65.0854 | 281.7222
[11 empty] | 458.1888 | 80.51 | 64.855 | 280.3968
[11 empty] | 458.2656 | 80.43 | 64.599 | 279.5338
[11 empty] | 458.3424 | 80.355 | 64.727 | 278.9173
[11 empty] | 458.4192 | 80.425 | 64.7936 | 278.2701
[11 empty] | 458.496 | 80.585 | 64.599 | 277.4687
[11 empty] | 458.5728 | 80.575 | 65.0394 | 277.0679
[11 empty] | 458.6496 | 80.575 | 65.5514 | 278.794
[11 empty] | 458.7264 | 80.885 | 66.0992 | 284.0956
[11 empty] | 458.8032 | 81.41 | 67.113 | 293.4966
[11 empty] | 458.88 | 81.805 | 68.0294 | 307.1512
[11 empty] | 458.9568 | 82.04 | 68.5568 | 323.4566
[11 empty] | 459.0336 | 82.385 | 68.5978 | 339.5462
[11 empty] | 459.1104 | 82.725 | 69.1354 | 352.461
[11 empty] | 459.1872 | 83.065 | 68.9971 | 360.1051
[11 empty] | 459.264 | 83.32 | 68.5773 | 362.2319
[11 empty] | 459.3408 | 83.4 | 69.1098 | 360.1667
[11 empty] | 459.4176 | 83.52 | 69.9648 | 356.0365
[11 empty] | 459.4944 | 83.725 | 70.5331 | 351.9678
[11 empty] | 459.5712 | 83.86 | 70.6458 | 349.2554
[11 empty] | 459.648 | 83.7 | 71.1373 | 348.115
[11 empty] | 459.7248 | 83.39 | 71.1731 | 347.591
[11 empty] | 459.8016 | 83.11 | 69.7395 | 346.4197
[11 empty] | 459.8784 | 82.59 | 68.1882 | 343.1216
[11 empty] | 459.9552 | 81.995 | 66.7238 | 336.7721
[11 empty] | 460.032 | 81.365 | 65.3517 | 327.8026
[11 empty] | 460.1088 | 80.935 | 64.0666 | 316.9221
[11 empty] | 460.1856 | 81.065 | 63.1142 | 304.747
[11 empty] | 460.2624 | 81.23 | 64.1382 | 292.9418
[11 empty] | 460.3392 | 81.41 | 65.3107 | 289.6746
[11 empty] | 460.416 | 81.585 | 65.3005 | 291.1232
[11 empty] | 460.4928 | 81.465 | 65.2493 | 284.8045
[11 empty] | 460.5696 | 81.22 | 64.4506 | 276.9138
[11 empty] | 460.6464 | 80.805 | 63.6826 | 269.6396
[11 empty] | 460.7232 | 80.47 | 63.0477 | 263.3517
[11 empty] | 460.8 | 80.205 | 62.935 | 258.4509
[11 empty] | 460.8768 | 80 | 63.2166 | 254.9062
[11 empty] | 460.9536 | 80.015 | 63.0733 | 252.5945
[11 empty] | 461.0304 | 80.25 | 61.9674 | 251.2074
[11 empty] | 461.1072 | 80.255 | 61.5526 | 250.5293
[11 empty] | 461.184 | 79.855 | 61.9622 | 250.3136
[11 empty] | 461.2608 | 79.255 | 62.0134 | 250.2828
[11 empty] | 461.3376 | 78.65 | 60.8666 | 252.3787
[11 empty] | 461.4144 | 78.155 | 59.0387 | 263.999
[11 empty] | 461.4912 | 77.48 | 58.9107 | 282.3387
[11 empty] | 461.568 | 77.03 | 59.4176 | 297.3495
[11 empty] | 461.6448 | 77.36 | 58.0813 | 306.843
[11 empty] | 461.7216 | 77.73 | 58.4192 | 312.4219
[11 empty] | 461.7984 | 78.285 | 59.2794 | 316.3365
[11 empty] | 461.8752 | 78.875 | 60.585 | 320.7133
[11 empty] | 461.952 | 79.375 | 61.6858 | 326.693
[11 empty] | 462.0288 | 79.78 | 62.0186 | 334.2446
[11 empty] | 462.1056 | 79.795 | 62.5357 | 342.1353
[11 empty] | 462.1824 | 79.465 | 63.4368 | 348.6081
[11 empty] | 462.2592 | 79.165 | 63.4829 | 352.0911
[11 empty] | 462.336 | 78.945 | 62.935 | 351.8445
[11 empty] | 462.4128 | 78.935 | 61.911 | 348.1458
[11 empty] | 462.4896 | 78.805 | 61.6397 | 342.1661
[11 empty] | 462.5664 | 78.745 | 61.353 | 335.4775
[11 empty] | 462.6432 | 78.565 | 61.353 | 329.467
[11 empty] | 462.72 | 78.235 | 61.1942 | 324.7511
[11 empty] | 462.7968 | 77.9 | 60.2163 | 321.1448
[11 empty] | 462.8736 | 77.305 | 59.3613 | 317.6002
[11 empty] | 462.9504 | 76.585 | 58.5062 | 312.8535
[11 empty] | 463.0272 | 76.07 | 56.9805 | 306.3498
[11 empty] | 463.104 | 75.705 | 56.1152 | 298.9523
[11 empty] | 463.1808 | 75.935 | 55.6237 | 291.0924
[11 empty] | 463.2576 | 76.315 | 56.3456 | 283.4483
[11 empty] | 463.3344 | 76.54 | 58.4397 | 277.1296
[11 empty] | 463.4112 | 77.07 | 58.9363 | 272.352
[11 empty] | 463.488 | 77.615 | 58.7469 | 268.8998
[11 empty] | 463.5648 | 77.625 | 59.6326 | 266.619
[11 empty] | 463.6416 | 77.39 | 60.3648 | 265.2627
[11 empty] | 463.7184 | 77.32 | 59.4893 | 264.6154
[11 empty] | 463.7952 | 77.285 | 58.0403 | 264.6154
[11 empty] | 463.872 | 76.995 | 57.7792 | 265.0161
[11 empty] | 463.9488 | 76.77 | 57.856 | 265.571
[11 empty] | 464.0256 | 76.325 | 59.3613 | 266.095
[11 empty] | 464.1024 | 76.3 | 59.2742 | 266.8039
[11 empty] | 464.1792 | 76.495 | 58.3475 | 269.1156
[11 empty] | 464.256 | 76.775 | 58.5779 | 274.6638
[11 empty] | 464.3328 | 77.325 | 59.1309 | 284.0648
[11 empty] | 464.4096 | 77.695 | 60.3187 | 297.5961
[11 empty] | 464.4864 | 77.91 | 61.015 | 314.0247
[11 empty] | 464.5632 | 78.025 | 62.1619 | 330.7616
[11 empty] | 464.64 | 78.075 | 63.2525 | 344.8169
[11 empty] | 464.7168 | 78.35 | 62.4486 | 353.9713
[11 empty] | 464.7936 | 78.38 | 62.6995 | 357.7009
[11 empty] | 464.8704 | 78.585 | 62.848 | 357.0228
[11 empty] | 464.9472 | 78.875 | 63.2781 | 354.0022
[11 empty] | 465.024 | 79.28 | 63.6672 | 350.6733
[11 empty] | 465.1008 | 79.49 | 64.7936 | 348.5465
[11 empty] | 465.1776 | 79.555 | 65.5616 | 348.0841
[11 empty] | 465.2544 | 79.565 | 64.8858 | 348.6081
[11 empty] | 465.3312 | 79.385 | 63.5034 | 348.8239
[11 empty] | 465.408 | 78.995 | 62.2848 | 347.2519
[11 empty] | 465.4848 | 78.415 | 61.568 | 343.1216
[11 empty] | 465.5616 | 77.81 | 60.9331 | 336.3714
[11 empty] | 465.6384 | 77.55 | 59.7197 | 327.6793
[11 empty] | 465.7152 | 77.655 | 59.5763 | 317.5694
[11 empty] | 465.792 | 77.94 | 61.0253 | 307.706
[11 empty] | 465.8688 | 78.3 | 62.7712 | 306.1957
[11 empty] | 465.9456 | 78.735 | 63.4675 | 308.0759
[11 empty] | 466.0224 | 79.01 | 63.2115 | 302.8668
[11 empty] | 466.0992 | 78.985 | 62.679 | 296.8255
[11 empty] | 466.176 | 78.955 | 61.5731 | 291.678
[11 empty] | 466.2528 | 78.895 | 61.4554 | 287.6711
[11 empty] | 466.3296 | 78.845 | 62.1261 | 284.9278
[11 empty] | 466.4064 | 79.04 | 62.9094 | 283.2325
[11 empty] | 466.4832 | 79.495 | 63.3037 | 282.277
[11 empty] | 466.56 | 79.845 | 64.2816 | 281.7839
[11 empty] | 466.6368 | 80.15 | 64.9062 | 281.8147
[11 empty] | 466.7136 | 80.035 | 66.391 | 282.1537
[11 empty] | 466.7904 | 80.245 | 65.5206 | 282.5236
[11 empty] | 466.8672 | 80.375 | 64.1382 | 282.8935
[11 empty] | 466.944 | 80.18 | 63.5802 | 287.7019
[11 empty] | 467.0208 | 79.82 | 63.5802 | 303.4216
[11 empty] | 467.0976 | 79.62 | 63.4675 | 322.0387
[11 empty] | 467.1744 | 79.785 | 63.5699 | 335.8474
[11 empty] | 467.2512 | 80.255 | 64.169 | 344.4778
[11 empty] | 467.328 | 81.075 | 64.7322 | 349.5328
[11 empty] | 467.4048 | 81.875 | 66.6266 | 352.9542
[11 empty] | 467.4816 | 82.455 | 69.3811 | 356.5913
[11 empty] | 467.5584 | 82.785 | 72.2637 | 361.4305
[11 empty] | 467.6352 | 83.135 | 73.0368 | 367.1944
[11 empty] | 467.712 | 83.52 | 71.7466 | 372.6501
[11 empty] | 467.7888 | 83.495 | 70.9222 | 376.2563
[11 empty] | 467.8656 | 83.165 | 70.6355 | 376.8112
[11 empty] | 467.9424 | 82.91 | 70.185 | 373.7289
[11 empty] | 468.0192 | 82.675 | 69.8778 | 367.5643
[11 empty] | 468.096 | 82.395 | 69.8624 | 359.6428
[11 empty] | 468.1728 | 82.37 | 69.0125 | 351.413
[11 empty] | 468.2496 | 82.085 | 68.7565 | 343.8614
[11 empty] | 468.3264 | 81.5 | 69.0944 | 337.4502
[11 empty] | 468.4032 | 81.095 | 67.6506 | 331.748
[11 empty] | 468.48 | 80.4 | 66.3552 | 325.8607
[11 empty] | 468.5568 | 79.665 | 64.8704 | 319.0489
[11 empty] | 468.6336 | 79.39 | 61.9827 | 311.0041
[11 empty] | 468.7104 | 79.11 | 61.2659 | 302.7127
[11 empty] | 468.7872 | 79.185 | 62.249 | 294.1439
[11 empty] | 468.864 | 79.6 | 63.1654 | 285.8525
[11 empty] | 468.9408 | 80.19 | 63.5136 | 278.7632
[11 empty] | 469.0176 | 80.62 | 64.0973 | 272.8144
[11 empty] | 469.0944 | 80.475 | 66.089 | 267.8519
[11 empty] | 469.1712 | 80.29 | 66.4371 | 263.6908
[11 empty] | 469.248 | 80.24 | 64.7629 | 260.4235
[11 empty] | 469.3248 | 79.675 | 64.6656 | 258.1118
[11 empty] | 469.4016 | 79.24 | 63.616 | 256.6939
[11 empty] | 469.4784 | 79.045 | 61.5373 | 255.985
[11 empty] | 469.5552 | 78.51 | 61.6499 | 255.7076
[11 empty] | 469.632 | 78.13 | 61.6192 | 255.5535
[11 empty] | 469.7088 | 77.835 | 61.353 | 255.7384
[11 empty] | 469.7856 | 77.425 | 62.1261 | 257.9268
[11 empty] | 469.8624 | 77.535 | 62.3053 | 263.5367
[11 empty] | 469.9392 | 78.1 | 61.2813 | 273.3076
[11 empty] | 470.016 | 78.165 | 61.7165 | 287.332
[11 empty] | 470.0928 | 78.215 | 60.8154 | 304.0381
[11 empty] | 470.1696 | 78.26 | 60.3443 | 320.3434
[11 empty] | 470.2464 | 78.145 | 61.2608 | 333.3507
[11 empty] | 470.3232 | 78.225 | 61.0253 | 341.2414
[11 empty] | 470.4 | 78.165 | 61.8189 | 344.108
[11 empty] | 470.4768 | 78.235 | 62.5203 | 343.3066
[11 empty] | 470.5536 | 78.715 | 61.8547 | 341.0873
[11 empty] | 470.6304 | 79.01 | 62.249 | 339.4845
[11 empty] | 470.7072 | 78.985 | 63.3446 | 339.2379
[11 empty] | 470.784 | 78.87 | 63.6723 | 340.2551
[11 empty] | 470.8608 | 78.81 | 62.2285 | 341.488
[11 empty] | 470.9376 | 78.635 | 60.1805 | 341.3031
[11 empty] | 471.0144 | 78.05 | 59.4176 | 338.4982
[11 empty] | 471.0912 | 77.06 | 60.0269 | 332.8268
[11 empty] | 471.168 | 76.64 | 58.1888 | 324.7511
[11 empty] | 471.2448 | 76.255 | 57.2774 | 315.381
[11 empty] | 471.3216 | 75.86 | 58.9466 | 305.5484
[11 empty] | 471.3984 | 76.105 | 59.9398 | 298.5824
[11 empty] | 471.4752 | 76.635 | 59.6378 | 299.3838
[11 empty] | 471.552 | 76.925 | 59.2538 | 296.4556
[11 empty] | 471.6288 | 76.73 | 59.392 | 289.7979
[11 empty] | 471.7056 | 76.555 | 58.199 | 284.0956
[11 empty] | 471.7824 | 76.115 | 57.9123 | 279.7187
[11 empty] | 471.8592 | 75.96 | 56.8422 | 276.6364
[11 empty] | 471.936 | 75.71 | 57.6307 | 274.6946
[11 empty] | 472.0128 | 75.585 | 59.0029 | 273.4925
[11 empty] | 472.0896 | 75.955 | 58.839 | 272.8144
[11 empty] | 472.1664 | 76.425 | 58.2298 | 272.5986
[11 empty] | 472.2432 | 76.445 | 59.0285 | 272.9993
[11 empty] | 472.32 | 76.325 | 59.1514 | 273.8932
[11 empty] | 472.3968 | 76.115 | 58.7571 | 275.1569
[11 empty] | 472.4736 | 75.91 | 57.6768 | 279.9653
[11 empty] | 472.5504 | 75.475 | 57.4618 | 295.7467
[11 empty] | 472.6272 | 75.16 | 57.4208 | 316.1207
[11 empty] | 472.704 | 75.065 | 57.9072 | 331.6863
[11 empty] | 472.7808 | 75.415 | 57.9635 | 341.0257
[11 empty] | 472.8576 | 76.1 | 58.2861 | 345.7108
[11 empty] | 472.9344 | 76.875 | 59.3818 | 348.1458
[11 empty] | 473.0112 | 77.575 | 61.5066 | 350.9199
[11 empty] | 473.088 | 78.26 | 62.8838 | 355.5741
[11 empty] | 473.1648 | 78.715 | 63.8771 | 362.1703
[11 empty] | 473.2416 | 78.88 | 64.5939 | 369.4753
[11 empty] | 473.3184 | 78.885 | 64.6605 | 375.3933
[11 empty] | 473.3952 | 78.745 | 64.4147 | 378.1674
[11 empty] | 473.472 | 78.47 | 64.5069 | 376.7187
[11 empty] | 473.5488 | 78.535 | 63.4419 | 371.3863
[11 empty] | 473.6256 | 78.59 | 62.9197 | 363.4648
[11 empty] | 473.7024 | 78.605 | 62.8685 | 354.9577
[11 empty] | 473.7792 | 78.625 | 62.4384 | 347.591
[11 empty] | 473.856 | 78.3 | 62.8122 | 342.2278
[11 empty] | 473.9328 | 77.97 | 62.1773 | 338.5907
[11 empty] | 474.0096 | 77.455 | 61.5219 | 335.4467
[11 empty] | 474.0864 | 76.7 | 61.9366 | 331.2856
[11 empty] | 474.1632 | 76.555 | 60.1856 | 325.4292
[11 empty] | 474.24 | 76.875 | 58.7418 | 318.6173
[11 empty] | 474.3168 | 77.395 | 59.7555 | 310.6958
[11 empty] | 474.3936 | 78.165 | 60.6362 | 302.9593
[11 empty] | 474.4704 | 78.885 | 62.1312 | 296.4865
[11 empty] | 474.5472 | 79.645 | 63.1501 | 291.3698
[11 empty] | 474.624 | 80.3 | 63.3958 | 287.5169
[11 empty] | 474.7008 | 80.385 | 65.0854 | 284.712
[11 empty] | 474.7776 | 80.445 | 65.8074 | 282.7394
[11 empty] | 474.8544 | 80.42 | 65.6179 | 281.4448
[11 empty] | 474.9312 | 80.375 | 65.193 | 280.5818
[11 empty] | 475.008 | 80.5 | 64.2355 | 279.9345
[11 empty] | 475.0848 | 80.48 | 64.4147 | 279.3488
[11 empty] | 475.1616 | 80.71 | 63.6826 | 278.7016
[11 empty] | 475.2384 | 80.755 | 64.0256 | 278.3317
[11 empty] | 475.3152 | 80.67 | 65.0035 | 279.6879
[11 empty] | 475.392 | 80.755 | 66.6778 | 284.4346
[11 empty] | 475.4688 | 81.255 | 67.6762 | 293.1267
[11 empty] | 475.5456 | 81.89 | 67.5686 | 306.1649
[11 empty] | 475.6224 | 82.155 | 67.8246 | 322.2853
[11 empty] | 475.6992 | 82.185 | 69.3658 | 338.6831
[11 empty] | 475.776 | 82.545 | 69.8931 | 352.3377
[11 empty] | 475.8528 | 82.875 | 69.847 | 360.814
[11 empty] | 475.9296 | 82.875 | 70.697 | 363.4032
[11 empty] | 476.0064 | 83.12 | 70.5434 | 361.4305
[11 empty] | 476.0832 | 83.5 | 70.1645 | 356.9612
[11 empty] | 476.16 | 83.47 | 71.6493 | 352.4918
[11 empty] | 476.2368 | 83.56 | 72.0435 | 349.6253
[11 empty] | 476.3136 | 83.76 | 71.3523 | 348.7931
[11 empty] | 476.3904 | 83.79 | 69.7037 | 348.7931
[11 empty] | 476.4672 | 83.38 | 68.5005 | 347.9917
[11 empty] | 476.544 | 82.72 | 67.4867 | 344.8169
[11 empty] | 476.6208 | 81.94 | 66.8006 | 338.4674
[11 empty] | 476.6976 | 81.22 | 65.7203 | 329.1896
[11 empty] | 476.7744 | 80.78 | 64.4864 | 318.0625
[11 empty] | 476.8512 | 80.725 | 64.5888 | 306.0416
[11 empty] | 476.928 | 80.95 | 65.6435 | 294.8836
[11 empty] | 477.0048 | 81.24 | 66.0224 | 292.2328
[11 empty] | 477.0816 | 81.48 | 65.7869 | 292.6028
[11 empty] | 477.1584 | 81.465 | 64.9472 | 285.4826
[11 empty] | 477.2352 | 81.015 | 64.6963 | 277.4995
[11 empty] | 477.312 | 80.64 | 63.7645 | 270.5027
[11 empty] | 477.3888 | 80.275 | 63.273 | 264.6771
[11 empty] | 477.4656 | 80.155 | 62.7814 | 260.1769
[11 empty] | 477.5424 | 80.19 | 62.4794 | 256.7864
[11 empty] | 477.6192 | 80.125 | 62.7712 | 254.2589
[11 empty] | 477.696 | 80.09 | 62.9453 | 252.4712
[11 empty] | 477.7728 | 79.74 | 63.9232 | 251.4232
[11 empty] | 477.8496 | 79.51 | 63.1859 | 251.0842
[11 empty] | 477.9264 | 79.035 | 62.2438 | 251.1766
[11 empty] | 478.0032 | 78.61 | 60.759 | 252.8102
[11 empty] | 478.08 | 78.145 | 59.305 | 262.8894
[11 empty] | 478.1568 | 77.6 | 58.6854 | 281.1366
[11 empty] | 478.2336 | 77.265 | 58.24 | 296.9796
[11 empty] | 478.3104 | 77.31 | 57.8355 | 307.1204
[11 empty] | 478.3872 | 77.51 | 58.0045 | 313.0384
[11 empty] | 478.464 | 77.87 | 58.9056 | 316.9837
[11 empty] | 478.5408 | 78.36 | 60.6618 | 321.2065
[11 empty] | 478.6176 | 78.885 | 61.9315 | 326.9704
[11 empty] | 478.6944 | 79.3 | 62.423 | 334.3679
[11 empty] | 478.7712 | 79.615 | 62.0493 | 342.2894
[11 empty] | 478.848 | 79.74 | 61.6038 | 349.1321
[11 empty] | 478.9248 | 79.32 | 62.3872 | 353.2007
[11 empty] | 479.0016 | 79.02 | 62.0288 | 353.6015
[11 empty] | 479.0784 | 78.715 | 61.7933 | 350.4883
[11 empty] | 479.1552 | 78.645 | 61.0714 | 344.7861
[11 empty] | 479.232 | 78.695 | 60.5542 | 338.005
[11 empty] | 479.3088 | 78.5 | 60.8358 | 331.5938
[11 empty] | 479.3856 | 78.21 | 60.8563 | 326.4772
[11 empty] | 479.4624 | 78.03 | 59.2691 | 322.6552
[11 empty] | 479.5392 | 77.42 | 58.4141 | 319.4188
[11 empty] | 479.616 | 76.475 | 58.3629 | 315.3501
[11 empty] | 479.6928 | 75.79 | 57.4925 | 309.4937
[11 empty] | 479.7696 | 75.415 | 56.5914 | 302.4661
[11 empty] | 479.8464 | 75.275 | 57.5693 | 294.6679
[11 empty] | 479.9232 | 75.77 | 58.0147 | 286.5306
[11 empty] | 480 | 76.38 | 58.3219 | 279.5646
[11 empty] | 480.0768 | 77.065 | 57.9226 | 274.2322
[11 empty] | 480.1536 | 77.355 | 58.8595 | 270.5643
[11 empty] | 480.2304 | 77.275 | 60.1395 | 268.345
[11 empty] | 480.3072 | 77.305 | 59.8938 | 267.1738
[11 empty] | 480.384 | 77.325 | 58.624 | 266.6806
[11 empty] | 480.4608 | 76.865 | 59.2794 | 266.4957
[11 empty] | 480.5376 | 76.435 | 59.9142 | 266.4648
[11 empty] | 480.6144 | 76.42 | 59.0131 | 266.5881
[11 empty] | 480.6912 | 76.415 | 58.2554 | 266.8964
[11 empty] | 480.768 | 76.33 | 58.0301 | 267.6669
[11 empty] | 480.8448 | 76.405 | 57.6307 | 270.0095
[11 empty] | 480.9216 | 76.585 | 58.3424 | 275.496
[11 empty] | 480.9984 | 77.09 | 59.2845 | 284.4963
[11 empty] | 481.0752 | 77.44 | 60.8461 | 297.3187
[11 empty] | 481.152 | 77.775 | 60.9126 | 313.1617
[11 empty] | 481.2288 | 78.225 | 60.2214 | 329.6828
[11 empty] | 481.3056 | 78.235 | 61.5014 | 344.1696
[11 empty] | 481.3824 | 78.185 | 61.9878 | 354.2487
[11 empty] | 481.4592 | 78.035 | 62.8634 | 359.0571
[11 empty] | 481.536 | 78.225 | 63.1194 | 359.1188
[11 empty] | 481.6128 | 78.405 | 64.3482 | 356.2214
[11 empty] | 481.6896 | 78.915 | 64.4608 | 352.4302
[11 empty] | 481.7664 | 79.53 | 63.4317 | 349.5328
[11 empty] | 481.8432 | 79.71 | 63.0682 | 348.5157
[11 empty] | 481.92 | 79.41 | 63.7235 | 349.0396
[11 empty] | 481.9968 | 79.065 | 63.9795 | 349.8719
[11 empty] | 482.0736 | 78.705 | 63.063 | 349.2554
[11 empty] | 482.1504 | 78.315 | 61.4246 | 345.8649
[11 empty] | 482.2272 | 78.06 | 59.1309 | 339.4229
[11 empty] | 482.304 | 77.605 | 58.7827 | 330.5458
[11 empty] | 482.3808 | 77.575 | 59.008 | 320.0969
[11 empty] | 482.4576 | 77.835 | 60.6925 | 309.617
[11 empty] | 482.5344 | 78.395 | 61.8496 | 307.3361
[11 empty] | 482.6112 | 78.845 | 62.7251 | 310.3568
[11 empty] | 482.688 | 79.055 | 63.1347 | 305.5792
[11 empty] | 482.7648 | 78.85 | 63.6416 | 299.2605
[11 empty] | 482.8416 | 78.79 | 63.0221 | 293.7432
[11 empty] | 482.9184 | 78.825 | 62.2746 | 289.3355
[11 empty] | 482.9952 | 78.715 | 62.7968 | 286.284
[11 empty] | 483.072 | 79.035 | 62.3821 | 284.4038
[11 empty] | 483.1488 | 79.6 | 62.4333 | 283.325
[11 empty] | 483.2256 | 79.855 | 63.9078 | 282.7086
[11 empty] | 483.3024 | 80.04 | 65.239 | 282.5544
[11 empty] | 483.3792 | 80.265 | 65.5053 | 282.7394
[11 empty] | 483.456 | 80.31 | 65.2339 | 282.986
[11 empty] | 483.5328 | 80.11 | 64.9728 | 283.2325
[11 empty] | 483.6096 | 79.805 | 64.5069 | 287.2395
[11 empty] | 483.6864 | 79.535 | 63.7952 | 302.0654
[11 empty] | 483.7632 | 79.17 | 64.6502 | 320.8983
[11 empty] | 483.84 | 79.465 | 64.041 | 335.2926
[11 empty] | 483.9168 | 80.045 | 64.0717 | 344.447
[11 empty] | 483.9936 | 80.77 | 65.065 | 349.9335
[11 empty] | 484.0704 | 81.54 | 66.8979 | 353.4781
[11 empty] | 484.1472 | 82.315 | 68.7616 | 356.9303
[11 empty] | 484.224 | 82.95 | 70.2362 | 361.3072
[11 empty] | 484.3008 | 83.385 | 70.6458 | 366.6704
[11 empty] | 484.3776 | 83.475 | 70.743 | 372.0336
[11 empty] | 484.4544 | 83.24 | 71.2038 | 375.9173
[11 empty] | 484.5312 | 82.93 | 71.3882 | 377.1502
[11 empty] | 484.608 | 82.72 | 70.5485 | 375.0234
[11 empty] | 484.6848 | 82.68 | 69.0739 | 369.5678
[11 empty] | 484.7616 | 82.41 | 69.2019 | 361.6463
[11 empty] | 484.8384 | 82.115 | 69.2224 | 352.7384
[11 empty] | 484.9152 | 82.115 | 67.6454 | 344.3546
[11 empty] | 484.992 | 81.825 | 66.9338 | 337.2036
[11 empty] | 485.0688 | 81.285 | 66.2272 | 331.3164
[11 empty] | 485.1456 | 80.82 | 63.7338 | 326.0149
[11 empty] | 485.2224 | 80.025 | 62.1978 | 320.2202
[11 empty] | 485.2992 | 79 | 62.4691 | 313.1617
[11 empty] | 485.376 | 78.685 | 62.1824 | 305.4559
[11 empty] | 485.4528 | 79.21 | 61.4707 | 297.0104
[11 empty] | 485.5296 | 79.595 | 63.0118 | 288.2567
[11 empty] | 485.6064 | 79.95 | 64.4557 | 280.4276
[11 empty] | 485.6832 | 80.45 | 64.6861 | 273.8007
[11 empty] | 485.76 | 80.675 | 64.9882 | 268.4683
[11 empty] | 485.8368 | 80.58 | 65.2493 | 264.3072
[11 empty] | 485.9136 | 80.275 | 64.9062 | 261.2557
[11 empty] | 485.9904 | 79.735 | 64.5632 | 259.1906
[11 empty] | 486.0672 | 79.23 | 63.7747 | 257.8036
[11 empty] | 486.144 | 78.82 | 62.9914 | 256.8481
[11 empty] | 486.2208 | 78.645 | 61.5731 | 256.17
[11 empty] | 486.2976 | 78.445 | 60.8358 | 255.646
[11 empty] | 486.3744 | 78.31 | 59.9091 | 255.6151
[11 empty] | 486.4512 | 78.085 | 59.5968 | 257.7111
[11 empty] | 486.528 | 78.01 | 60.1754 | 263.3517
[11 empty] | 486.6048 | 78.26 | 60.4467 | 272.9685
[11 empty] | 486.6816 | 78.43 | 60.3648 | 286.6847
[11 empty] | 486.7584 | 78.445 | 59.3459 | 302.9901
[11 empty] | 486.8352 | 78.265 | 59.6941 | 319.0797
[11 empty] | 486.912 | 78.28 | 60.2061 | 332.1487
[11 empty] | 486.9888 | 78.475 | 59.945 | 340.3784
[11 empty] | 487.0656 | 78.59 | 60.0678 | 343.6765
[11 empty] | 487.1424 | 78.49 | 61.3734 | 343.3682
[11 empty] | 487.2192 | 78.515 | 62.7354 | 341.3955
[11 empty] | 487.296 | 78.7 | 63.3395 | 339.4537
[11 empty] | 487.3728 | 78.935 | 62.9402 | 338.7756
[11 empty] | 487.4496 | 78.925 | 62.7302 | 339.3921
[11 empty] | 487.5264 | 78.625 | 62.9043 | 340.44
[11 empty] | 487.6032 | 78.245 | 62.0186 | 340.625
[11 empty] | 487.68 | 77.795 | 60.585 | 338.3749
[11 empty] | 487.7568 | 77.435 | 58.5728 | 333.2274
[11 empty] | 487.8336 | 76.905 | 57.3184 | 325.4292
[11 empty] | 487.9104 | 76.19 | 57.5539 | 315.8741
[11 empty] | 487.9872 | 75.945 | 57.9738 | 305.61
[11 empty] | 488.064 | 76.32 | 58.2246 | 297.7194
[11 empty] | 488.1408 | 76.745 | 58.5523 | 298.3975
[11 empty] | 488.2176 | 76.69 | 60.1293 | 296.8871
[11 empty] | 488.2944 | 76.66 | 59.7555 | 290.476
[11 empty] | 488.3712 | 76.24 | 59.7146 | 284.5579
[11 empty] | 488.448 | 75.99 | 58.5626 | 279.6879
[11 empty] | 488.5248 | 75.945 | 56.9498 | 275.9583
[11 empty] | 488.6016 | 75.79 | 57.1238 | 273.6158
[11 empty] | 488.6784 | 75.96 | 56.9856 | 272.4445
[11 empty] | 488.7552 | 76.19 | 57.1904 | 272.0746
[11 empty] | 488.832 | 76.425 | 57.7587 | 272.1671
[11 empty] | 488.9088 | 76.695 | 57.5795 | 272.6294
[11 empty] | 488.9856 | 76.59 | 57.9686 | 273.3692
[11 empty] | 489.0624 | 76.195 | 58.7162 | 274.1706
[11 empty] | 489.1392 | 75.725 | 58.9363 | 278.1468
[11 empty] | 489.216 | 75.37 | 58.24 | 293.1267
[11 empty] | 489.2928 | 75.48 | 56.32 | 313.6857
[11 empty] | 489.3696 | 75.69 | 55.2909 | 329.7444
[11 empty] | 489.4464 | 75.645 | 57.0778 | 339.6078
[11 empty] | 489.5232 | 76.09 | 58.5421 | 344.6628
[11 empty] | 489.6 | 76.945 | 59.4637 | 347.1903
[11 empty] | 489.6768 | 77.605 | 61.5322 | 349.8411
[11 empty] | 489.7536 | 78.12 | 63.5443 | 354.2179
[11 empty] | 489.8304 | 78.645 | 64.0461 | 360.5058
[11 empty] | 489.9072 | 79.1 | 63.3958 | 367.5026
[11 empty] | 489.984 | 79.205 | 63.0477 | 373.3898
[11 empty] | 490.0608 | 78.92 | 63.529 | 376.3488
[11 empty] | 490.1376 | 78.77 | 63.1962 | 375.3933
[11 empty] | 490.2144 | 78.735 | 62.505 | 370.5541
[11 empty] | 490.2912 | 78.55 | 62.6176 | 363.0333
[11 empty] | 490.368 | 78.565 | 62.2848 | 354.557
[11 empty] | 490.4448 | 78.655 | 61.481 | 346.9745
[11 empty] | 490.5216 | 78.51 | 61.225 | 341.3031
[11 empty] | 490.5984 | 78.035 | 61.6704 | 337.3577
[11 empty] | 490.6752 | 77.42 | 61.7318 | 334.1213
[11 empty] | 490.752 | 77.145 | 59.6787 | 330.176
[11 empty] | 490.8288 | 76.87 | 58.4653 | 324.4737
[11 empty] | 490.9056 | 76.575 | 59.5968 | 317.9084
[11 empty] | 490.9824 | 77.05 | 60.5594 | 310.2643
[11 empty] | 491.0592 | 78.185 | 60.2214 | 302.5277
[11 empty] | 491.136 | 79.065 | 61.2454 | 295.9008
[11 empty] | 491.2128 | 79.59 | 63.4829 | 290.6301
[11 empty] | 491.2896 | 80.095 | 64.5683 | 286.5923
[11 empty] | 491.3664 | 80.48 | 64.896 | 283.5716
[11 empty] | 491.4432 | 80.635 | 65.1213 | 281.414
[11 empty] | 491.52 | 80.425 | 65.9917 | 279.9653
[11 empty] | 491.5968 | 80.285 | 65.8995 | 279.0098
[11 empty] | 491.6736 | 80.335 | 65.0138 | 278.3625
[11 empty] | 491.7504 | 80.445 | 64.5786 | 277.7769
[11 empty] | 491.8272 | 80.37 | 65.4438 | 277.0679
[11 empty] | 491.904 | 80.485 | 65.4131 | 276.4823
[11 empty] | 491.9808 | 80.825 | 64.4966 | 277.6536
[11 empty] | 492.0576 | 81.115 | 65.2237 | 282.277
[11 empty] | 492.1344 | 81.44 | 67.1283 | 290.9691
[11 empty] | 492.2112 | 81.98 | 67.369 | 304.0381
[11 empty] | 492.288 | 82.335 | 67.0515 | 320.066
[11 empty] | 492.3648 | 82.41 | 68.3878 | 336.4639
[11 empty] | 492.4416 | 82.515 | 70.1696 | 350.0568
[11 empty] | 492.5184 | 82.76 | 70.6048 | 358.5023
[11 empty] | 492.5952 | 82.985 | 70.5126 | 361.2456
[11 empty] | 492.672 | 83.2 | 70.6662 | 359.5503
[11 empty] | 492.7488 | 83.58 | 70.7328 | 355.5741
[11 empty] | 492.8256 | 83.81 | 71.1066 | 351.5055
[11 empty] | 492.9024 | 84.05 | 70.8454 | 348.7006
[11 empty] | 492.9792 | 84.13 | 70.2413 | 347.3752
[11 empty] | 493.056 | 83.665 | 70.6662 | 346.8204
[11 empty] | 493.1328 | 83.18 | 70.2208 | 345.6799
[11 empty] | 493.2096 | 82.66 | 68.9203 | 342.5977
[11 empty] | 493.2864 | 82.03 | 67.6915 | 336.6796
[11 empty] | 493.3632 | 81.385 | 66.4576 | 328.1417
[11 empty] | 493.44 | 81.04 | 64.9523 | 317.6927
[11 empty] | 493.5168 | 81.225 | 63.785 | 305.9491
[11 empty] | 493.5936 | 81.415 | 64.3635 | 293.9898
[11 empty] | 493.6704 | 81.57 | 65.3926 | 288.5341
[11 empty] | 493.7472 | 81.75 | 65.536 | 289.8903
[11 empty] | 493.824 | 81.55 | 65.8176 | 284.4963
[11 empty] | 493.9008 | 81.02 | 66.1504 | 276.883
[11 empty] | 493.9776 | 80.82 | 64.2714 | 269.7937
[11 empty] | 494.0544 | 80.51 | 63.3856 | 263.5983
[11 empty] | 494.1312 | 80.29 | 62.9658 | 258.5433
[11 empty] | 494.208 | 80.07 | 63.4368 | 254.8137
[11 empty] | 494.2848 | 80.155 | 63.017 | 252.4712
[11 empty] | 494.3616 | 80.22 | 62.3411 | 251.2074
[11 empty] | 494.4384 | 79.86 | 62.8531 | 250.7759
[11 empty] | 494.5152 | 79.48 | 62.4333 | 250.9609
[11 empty] | 494.592 | 78.965 | 61.993 | 251.3307
[11 empty] | 494.6688 | 78.495 | 60.9536 | 251.454
[11 empty] | 494.7456 | 78.13 | 59.5354 | 251.2691
[11 empty] | 494.8224 | 77.845 | 57.7792 | 250.8684
[11 empty] | 494.8992 | 77.25 | 57.257 | 250.3752
[11 empty] | 494.976 | 76.685 | 57.0675 | 249.8512
[11 empty] | 495.0528 | 76.22 | 56.105 | 249.1731
[11 empty] | 495.1296 | 75.68 | 55.1168 | 248.1252
[11 empty] | 495.2064 | 75.14 | 53.3606 | 246.6765
[11 empty] | 495.2832 | 74.64 | 50.8467 | 244.7963
[11 empty] | 495.36 | 73.675 | 49.152 | 242.577
[11 empty] | 495.4368 | 72.165 | 49.408 | 239.8029
[11 empty] | 495.5136 | 70.505 | 48.9216 | 236.5665
[11 empty] | 495.5904 | 68.87 | 48.3686 | 232.8678
[11 empty] | 495.6672 | 67.725 | 45.2045 | 228.7375
[11 empty] | 495.744 | 66.075 | 44.4774 | 224.0832
[11 empty] | 495.8208 | 64.18 | 44.8666 | 218.9358
[11 empty] | 495.8976 | 63.025 | 39.3626 | 213.326
[11 empty] | 495.9744 | 61.29 | 37.5654 | 207.5004
[11 empty] | 496.0512 | 0 | 0 | 0
[11 empty] | 496.128 | 0 | 0 | 0
[11 empty] | 496.2048 | 0 | 0 | 0
[11 empty] | 496.2816 | 0 | 0 | 0

## Test
XLPRO ™ Version 4.06
Operation Settings Area |  |  | 0 | R/min | 0 | R/min | 0 | R | 4.0244e-05 | R | 0.53052 | R | 11693.7794 | uR/pls
|  |  | 0 | R/min | 0 | R/min | 0 | R | 2.0404 | R | 1.0614 | R | 11703.8777 | uR/pls
Units |  |  | 63.6559 | R/min | 32.9397 | R/min | 0.2799 | R | 4.4641 | R |  |  | 11703.8777 | uR/pls
Radiation: |  | R | 63.6659 | R/min | 32.9397 | R/min | 1.4025 | R | 6.6091 | R |  |  | 11703.8777 | uR/pls
/Time: |  | /min | 63.6659 | R/min | 32.9397 | R/min | 2.5253 | R | 8.6555 | R |  |  | 11703.8777 | uR/pls
Scale: |  |  | 63.676 | R/min | 63.6458 | R/min | 3.6542 | R | 9.5932 | R |  |  | 2108.5157 | uR/pls
Pulse Rate: |  | 15-9-6-5 | 14.4174 | R/min | 63.6458 | R/min | 4.4872 | R | 9.5932 | R
Mode |  |  |  |  | 63.6458 | R/min | 4.4872 | R | 9.5932 | R
Units Display: |  | Rescale |  |  | 63.6458 | R/min | 4.4872 | R | [6 empty] | 2.816 | R | 220 | ms | 60.5 | kVp
Cell Update: |  | Incremental |  |  | 63.6458 | R/min | 4.9752 | R | [6 empty] | 63.03 | R/min | 88.5 | kVp
Auto Zero: |  | Disabled |  |  | 63.6458 | R/min | 6.098 | R | [6 empty] | 63.04 | R/min | 88.4 | kVp
Capture Mode: |  | Auto |  |  | 63.6559 | R/min | 7.2218 | R
Screen Tracking: |  | Yes |  |  | 63.6559 | R/min | 8.3446 | R
kVp |  |  |  |  | 63.6559 | R/min | 9.0197 | R
Config: | kVSDmA | Dose+Time+kV |  |  | 63.6559 | R/min | 9.0197 | R
Scale: |  | 9095 |  |  | 63.6559 | R/min | 9.0197 | R
Start delay (ms): |  | 0
Width (ms): |  | 4914.6
Filtration (mm-Al): |  | 2.5 | 0.020432 | R | 369 | ms | 62.7 | kVp
Dental threshold: |  | 0 | Peak= | 62.7 | kVp | PPV= | 57.3 | kV | Control= | 32834
Device Index |  | 0 | Rel Dose= | 0.55552 | /s | Average= | 61.06 | kV
Monitor Configuration |  |  | Pk Freq= | 59.4059 | Hz | Rate= | 63.0608 | Hz | Avg. mA= | N/A
Serial No: |  | 00008 | Width= | 369.4848 | ms | Pulses= | 24 |  | mAs= | N/A
ROM Ver: |  | 4.01 | ms | kV | Dose/s |  |  | LINES | Hz | kV
Detector: |  | 6 cc | -0.3072 | 0 | 0 |  |  | -0.3072 | 0 | 3.2498e-15
Chamb Kq: |  | 1.02 | -0.2304 | 0 | 0 |  |  | 387.2256 | 2.1583 | 0.064857
Atm Corr: |  | Automatic | -0.1536 | 0 | 0 |  |  | 62.7 | 4.3166 | 0.10326
Temp (°C): |  | 24.8 | -0.0768 | 0 | 0 |  |  | 62.7 | 6.4749 | 0.10244
Press (kPa): |  | 100 | 0 | 36.89 | 0.768 |  |  | 61.06 | 8.6333 | 0.07382
|  |  | 0.0768 | 38.38 | 0.77824 |  |  | 61.06 | 10.7916 | 0.056461
|  |  | 0.1536 | 38.835 | 0.91648 |  |  | 0.55552 | 12.9499 | 0.06314
|  |  | 0.2304 | 38.73 | 1.152 |  |  | 0.55552 | 15.1082 | 0.068268
|  |  | 0.3072 | 39.72 | 1.1674 |  |  | 5056 | 17.2665 | 0.10443
|  |  | 0.384 | 40.22 | 1.2493 |  |  |  | 19.4248 | 0.16292
|  |  | 0.4608 | 40.72 | 1.3261 |  |  | 0 | 21.5832 | 0.1883
|  |  | 0.5376 | 41.08 | 1.3824 |  |  | 0 | 23.7415 | 0.14831
|  |  | 0.6144 | 41.705 | 1.3722 |  |  |  | 25.8998 | 0.061977
|  |  | 0.6912 | 42.64 | 1.2954 |  |  |  | 28.0581 | 0.08161
|  |  | 0.768 | 42.785 | 1.3619 |  |  |  | 30.2164 | 0.12448
|  |  | 0.8448 | 42.97 | 1.4336 |  |  |  | 32.3747 | 0.073198
|  |  | 0.9216 | 43.19 | 1.4848 |  |  |  | 34.533 | 0.06621
|  |  | 0.9984 | 43.13 | 1.5821 |  |  |  | 36.6914 | 0.21233
|  |  | 1.0752 | 42.9 | 1.6794 |  |  |  | 38.8497 | 0.2719
|  |  | 1.152 | 43.035 | 1.6691 |  |  |  | 41.008 | 0.18535
|  |  | 1.2288 | 43.645 | 1.577 |  |  |  | 43.1663 | 0.13107
|  |  | 1.3056 | 44.01 | 1.5514 |  |  |  | 45.3246 | 0.35989
|  |  | 1.3824 | 43.71 | 1.6691 |  |  |  | 47.4829 | 0.46907
|  |  | 1.4592 | 43.75 | 1.6538 |  |  |  | 49.6413 | 0.28252
|  |  | 1.536 | 43.195 | 1.7613 |  |  |  | 51.7996 | 0.25229
|  |  | 1.6128 | 42.405 | 1.9712 |  |  |  | 53.9579 | 1.0263
|  |  | 1.6896 | 42.895 | 1.8278 |  |  |  | 56.1162 | 1.8169
|  |  | 1.7664 | 42.86 | 1.8074 |  |  |  | 58.2745 | 2.3561
|  |  | 1.8432 | 42.595 | 1.8381 |  |  |  | 60.4328 | 2.4481
|  |  | 1.92 | 42.92 | 1.6896 |  |  |  | 62.5912 | 2.0537
|  |  | 1.9968 | 43.275 | 1.5667 |  |  |  | 64.7495 | 1.3103
|  |  | 2.0736 | 42.505 | 1.6947 |  |  |  | 66.9078 | 0.47911
|  |  | 2.1504 | 41.85 | 1.7715 |  |  |  | 69.0661 | 0.24452
|  |  | 2.2272 | 41.28 | 1.7613 |  |  |  | 71.2244 | 0.55704
|  |  | 2.304 | 40.95 | 1.6538 |  |  |  | 73.3827 | 0.53172
|  |  | 2.3808 | 40.535 | 1.5155 |  |  |  | 75.541 | 0.2735
|  |  | 2.4576 | 39.83 | 1.3722 |  |  |  | 77.6994 | 0.03826
|  |  | 2.5344 | 39.115 | 1.1827 |  |  |  | 79.8577 | 0.23148
|  |  | 2.6112 | 38 | 1.1469 |  |  |  | 82.016 | 0.25165
|  |  | 2.688 | 36.755 | 1.1418 |  |  |  | 84.1743 | 0.13264
|  |  | 2.7648 | 36.45 | 0.92672 |  |  |  | 86.3326 | 0.030744
|  |  | 2.8416 | 35.9 | 0.768 |  |  |  | 88.4909 | 0.13238
|  |  | 2.9184 | 0 | 0 |  |  |  | 90.6493 | 0.13723
|  |  | 2.9952 | 0 | 0 |  |  |  | 92.8076 | 0.067836
|  |  | 3.072 | 0 | 0 |  |  |  | 94.9659 | 0.045486
|  |  | 3.1488 | 0 | 0 |  |  |  | 97.1242 | 0.095904
|  |  | 3.2256 | 0 | 0 |  |  |  | 99.2825 | 0.10681
|  |  | 3.3024 | 0 | 0 |  |  |  | 101.4408 | 0.10414
|  |  | 3.3792 | 0 | 0 |  |  |  | 103.5991 | 0.11854
|  |  | 3.456 | 0 | 0 |  |  |  | 105.7575 | 0.13342
|  |  | 3.5328 | 0 | 0 |  |  |  | 107.9158 | 0.10992
|  |  | 3.6096 | 0 | 0 |  |  |  | 110.0741 | 0.022972
|  |  | 3.6864 | 0 | 0 |  |  |  | 112.2324 | 0.23772
|  |  | 3.7632 | 0 | 0 |  |  |  | 114.3907 | 0.57411
|  |  | 3.84 | 0 | 0 |  |  |  | 116.549 | 0.93042
|  |  | 3.9168 | 0 | 0 |  |  |  | 118.7074 | 1.1808
|  |  | 3.9936 | 0 | 0 |  |  |  | 120.8657 | 1.2151
|  |  | 4.0704 | 0 | 0 |  |  |  | 123.024 | 0.99545
|  |  | 4.1472 | 0 | 0 |  |  |  | 125.1823 | 0.58247
|  |  | 4.224 | 0 | 0 |  |  |  | 127.3406 | 0.11443
|  |  | 4.3008 | 0 | 0 |  |  |  | 129.4989 | 0.25248
|  |  | 4.3776 | 0 | 0 |  |  |  | 131.6572 | 0.41247
|  |  | 4.4544 | 0 | 0 |  |  |  | 133.8156 | 0.35413
|  |  | 4.5312 | 0 | 0 |  |  |  | 135.9739 | 0.15824
|  |  | 4.608 | 0 | 0 |  |  |  | 138.1322 | 0.066097
|  |  | 4.6848 | 0 | 0 |  |  |  | 140.2905 | 0.18048
|  |  | 4.7616 | 0 | 0 |  |  |  | 142.4488 | 0.16823
|  |  | 4.8384 | 0 | 0 |  |  |  | 144.6071 | 0.063873
|  |  | 4.9152 | 0 | 0 |  |  |  | 146.7655 | 0.079738
|  |  | 4.992 | 0 | 0 |  |  |  | 148.9238 | 0.14662
|  |  | 5.0688 | 0 | 0 |  |  |  | 151.0821 | 0.12987
|  |  | 5.1456 | 0 | 0 |  |  |  | 153.2404 | 0.061018
|  |  | 5.2224 | 0 | 0 |  |  |  | 155.3987 | 0.070696
|  |  | 5.2992 | 0 | 0 |  |  |  | 157.557 | 0.1105
|  |  | 5.376 | 0 | 0 |  |  |  | 159.7154 | 0.10078
|  |  | 5.4528 | 0 | 0 |  |  |  | 161.8737 | 0.065799
|  |  | 5.5296 | 0 | 0 |  |  |  | 164.032 | 0.053581
|  |  | 5.6064 | 0 | 0 |  |  |  | 166.1903 | 0.049177
|  |  | 5.6832 | 0 | 0 |  |  |  | 168.3486 | 0.020072
|  |  | 5.76 | 0 | 0 |  |  |  | 170.5069 | 0.055188
|  |  | 5.8368 | 0 | 0 |  |  |  | 172.6652 | 0.18169
|  |  | 5.9136 | 0 | 0 |  |  |  | 174.8236 | 0.36904
|  |  | 5.9904 | 0 | 0 |  |  |  | 176.9819 | 0.57399
|  |  | 6.0672 | 0 | 0 |  |  |  | 179.1402 | 0.7178
|  |  | 6.144 | 0 | 0 |  |  |  | 181.2985 | 0.72558
|  |  | 6.2208 | 0 | 0 |  |  |  | 183.4568 | 0.56994
|  |  | 6.2976 | 0 | 0 |  |  |  | 185.6151 | 0.29389
|  |  | 6.3744 | 0 | 0 |  |  |  | 187.7735 | 0.057526
|  |  | 6.4512 | 0 | 0 |  |  |  | 189.9318 | 0.24134
|  |  | 6.528 | 0 | 0 |  |  |  | 192.0901 | 0.305
|  |  | 6.6048 | 0 | 0 |  |  |  | 194.2484 | 0.21759
|  |  | 6.6816 | 0 | 0 |  |  |  | 196.4067 | 0.058972
|  |  | 6.7584 | 0 | 0 |  |  |  | 198.565 | 0.13452
|  |  | 6.8352 | 0 | 0 |  |  |  | 200.7233 | 0.20796
|  |  | 6.912 | 0 | 0 |  |  |  | 202.8817 | 0.17606
|  |  | 6.9888 | 0 | 0 |  |  |  | 205.04 | 0.082263
|  |  | 7.0656 | 0 | 0 |  |  |  | 207.1983 | 0.066042
|  |  | 7.1424 | 0 | 0 |  |  |  | 209.3566 | 0.09928
|  |  | 7.2192 | 0 | 0 |  |  |  | 211.5149 | 0.066767
|  |  | 7.296 | 0 | 0 |  |  |  | 213.6732 | 0.036634
|  |  | 7.3728 | 0 | 0 |  |  |  | 215.8316 | 0.10525
|  |  | 7.4496 | 0 | 0 |  |  |  | 217.9899 | 0.13365
|  |  | 7.5264 | 0 | 0 |  |  |  | 220.1482 | 0.10444
|  |  | 7.6032 | 0 | 0 |  |  |  | 222.3065 | 0.045839
|  |  | 7.68 | 0 | 0 |  |  |  | 224.4648 | 0.02768
|  |  | 7.7568 | 0 | 0 |  |  |  | 226.6231 | 0.039067
|  |  | 7.8336 | 0 | 0 |  |  |  | 228.7814 | 0.027774
|  |  | 7.9104 | 0 | 0 |  |  |  | 230.9398 | 0.040322
|  |  | 7.9872 | 0 | 0 |  |  |  | 233.0981 | 0.10139
|  |  | 8.064 | 0 | 0 |  |  |  | 235.2564 | 0.21126
|  |  | 8.1408 | 0 | 0 |  |  |  | 237.4147 | 0.34493
|  |  | 8.2176 | 0 | 0 |  |  |  | 239.573 | 0.44278
|  |  | 8.2944 | 0 | 0 |  |  |  | 241.7313 | 0.44517
|  |  | 8.3712 | 0 | 0 |  |  |  | 243.8897 | 0.32941
|  |  | 8.448 | 0 | 0 |  |  |  | 246.048 | 0.128
|  |  | 8.5248 | 0 | 0 |  |  |  | 248.2063 | 0.082867
|  |  | 8.6016 | 0 | 0 |  |  |  | 250.3646 | 0.21942
|  |  | 8.6784 | 0 | 0 |  |  |  | 252.5229 | 0.23232
|  |  | 8.7552 | 0 | 0 |  |  |  | 254.6812 | 0.13136
|  |  | 8.832 | 0 | 0 |  |  |  | 256.8396 | 0.023951
|  |  | 8.9088 | 0 | 0 |  |  |  | 258.9979 | 0.14309
|  |  | 8.9856 | 0 | 0 |  |  |  | 261.1562 | 0.17357
|  |  | 9.0624 | 0 | 0 |  |  |  | 263.3145 | 0.10571
|  |  | 9.1392 | 0 | 0 |  |  |  | 265.4728 | 0.017376
|  |  | 9.216 | 0 | 0 |  |  |  | 267.6311 | 0.12573
|  |  | 9.2928 | 0 | 0 |  |  |  | 269.7894 | 0.16695
|  |  | 9.3696 | 0 | 0 |  |  |  | 271.9478 | 0.12867
|  |  | 9.4464 | 0 | 0 |  |  |  | 274.1061 | 0.0448
|  |  | 9.5232 | 0 | 0 |  |  |  | 276.2644 | 0.041768
|  |  | 9.6 | 0 | 0 |  |  |  | 278.4227 | 0.071567
|  |  | 9.6768 | 0 | 0 |  |  |  | 280.581 | 0.055351
|  |  | 9.7536 | 0 | 0 |  |  |  | 282.7393 | 0.062214
|  |  | 9.8304 | 0 | 0 |  |  |  | 284.8977 | 0.099753
|  |  | 9.9072 | 0 | 0 |  |  |  | 287.056 | 0.1058
|  |  | 9.984 | 0 | 0 |  |  |  | 289.2143 | 0.077469
|  |  | 10.0608 | 0 | 0 |  |  |  | 291.3726 | 0.0508
|  |  | 10.1376 | 0 | 0 |  |  |  | 293.5309 | 0.0506
|  |  | 10.2144 | 0 | 0 |  |  |  | 295.6892 | 0.082612
|  |  | 10.2912 | 0 | 0 |  |  |  | 297.8475 | 0.15822
|  |  | 10.368 | 0 | 0 |  |  |  | 300.0059 | 0.23096
|  |  | 10.4448 | 0 | 0 |  |  |  | 302.1642 | 0.24904
|  |  | 10.5216 | 0 | 0 |  |  |  | 304.3225 | 0.19146
|  |  | 10.5984 | 0 | 0 |  |  |  | 306.4808 | 0.098135
|  |  | 10.6752 | 0 | 0 |  |  |  | 308.6391 | 0.12052
|  |  | 10.752 | 0 | 0 |  |  |  | 310.7974 | 0.18144
|  |  | 10.8288 | 0 | 0 |  |  |  | 312.9558 | 0.1676
|  |  | 10.9056 | 0 | 0 |  |  |  | 315.1141 | 0.090752
|  |  | 10.9824 | 0 | 0 |  |  |  | 317.2724 | 0.088735
|  |  | 11.0592 | 0 | 0 |  |  |  | 319.4307 | 0.15669
|  |  | 11.136 | 0 | 0 |  |  |  | 321.589 | 0.16224
|  |  | 11.2128 | 0 | 0 |  |  |  | 323.7473 | 0.10656
|  |  | 11.2896 | 0 | 0 |  |  |  | 325.9056 | 0.092608
|  |  | 11.3664 | 0 | 0 |  |  |  | 328.064 | 0.15149
|  |  | 11.4432 | 0 | 0 |  |  |  | 330.2223 | 0.16767
|  |  | 11.52 | 0 | 0 |  |  |  | 332.3806 | 0.11975
|  |  | 11.5968 | 0 | 0 |  |  |  | 334.5389 | 0.039461
|  |  | 11.6736 | 0 | 0 |  |  |  | 336.6972 | 0.040211
|  |  | 11.7504 | 0 | 0 |  |  |  | 338.8555 | 0.063056
|  |  | 11.8272 | 0 | 0 |  |  |  | 341.0139 | 0.045439
|  |  | 11.904 | 0 | 0 |  |  |  | 343.1722 | 0.053152
|  |  | 11.9808 | 0 | 0 |  |  |  | 345.3305 | 0.087557
|  |  | 12.0576 | 0 | 0 |  |  |  | 347.4888 | 0.096483
|  |  | 12.1344 | 0 | 0 |  |  |  | 349.6471 | 0.084298
|  |  | 12.2112 | 0 | 0 |  |  |  | 351.8054 | 0.075679
|  |  | 12.288 | 0 | 0 |  |  |  | 353.9638 | 0.072354
|  |  | 12.3648 | 0 | 0 |  |  |  | 356.1221 | 0.097297
|  |  | 12.4416 | 0 | 0 |  |  |  | 358.2804 | 0.17132
|  |  | 12.5184 | 0 | 0 |  |  |  | 360.4387 | 0.24143
|  |  | 12.5952 | 0 | 0 |  |  |  | 362.597 | 0.25504
|  |  | 12.672 | 0 | 0 |  |  |  | 364.7553 | 0.19083
|  |  | 12.7488 | 0 | 0 |  |  |  | 366.9136 | 0.069411
|  |  | 12.8256 | 0 | 0 |  |  |  | 369.072 | 0.063034
|  |  | 12.9024 | 0 | 0 |  |  |  | 371.2303 | 0.14102
|  |  | 12.9792 | 0 | 0 |  |  |  | 373.3886 | 0.13932
|  |  | 13.056 | 0 | 0 |  |  |  | 375.5469 | 0.070074
|  |  | 13.1328 | 0 | 0 |  |  |  | 377.7052 | 0.034846
|  |  | 13.2096 | 0 | 0 |  |  |  | 379.8635 | 0.085984
|  |  | 13.2864 | 0 | 0 |  |  |  | 382.0219 | 0.071591
|  |  | 13.3632 | 0 | 0 |  |  |  | 384.1802 | 0.041407
|  |  | 13.44 | 0 | 0 |  |  |  | 386.3385 | 0.13913
|  |  | 13.5168 | 0 | 0 |  |  |  | 388.4968 | 0.21873
|  |  | 13.5936 | 0 | 0 |  |  |  | 390.6551 | 0.22791
|  |  | 13.6704 | 0 | 0 |  |  |  | 392.8134 | 0.16199
|  |  | 13.7472 | 0 | 0 |  |  |  | 394.9717 | 0.057575
|  |  | 13.824 | 0 | 0 |  |  |  | 397.1301 | 0.058375
|  |  | 13.9008 | 0 | 0 |  |  |  | 399.2884 | 0.10537
|  |  | 13.9776 | 0 | 0 |  |  |  | 401.4467 | 0.093043
|  |  | 14.0544 | 0 | 0 |  |  |  | 403.605 | 0.042574
|  |  | 14.1312 | 0 | 0 |  |  |  | 405.7633 | 0.049334
|  |  | 14.208 | 0 | 0 |  |  |  | 407.9216 | 0.093676
|  |  | 14.2848 | 0 | 0 |  |  |  | 410.08 | 0.11392
|  |  | 14.3616 | 0 | 0 |  |  |  | 412.2383 | 0.11309
|  |  | 14.4384 | 0 | 0 |  |  |  | 414.3966 | 0.099819
|  |  | 14.5152 | 0 | 0 |  |  |  | 416.5549 | 0.085936
|  |  | 14.592 | 0 | 0 |  |  |  | 418.7132 | 0.091921
|  |  | 14.6688 | 0 | 0 |  |  |  | 420.8715 | 0.10708
|  |  | 14.7456 | 0 | 0 |  |  |  | 423.0298 | 0.097159
|  |  | 14.8224 | 0 | 0 |  |  |  | 425.1882 | 0.05245
|  |  | 14.8992 | 0 | 0 |  |  |  | 427.3465 | 0.047505
|  |  | 14.976 | 0 | 0 |  |  |  | 429.5048 | 0.10381
|  |  | 15.0528 | 0 | 0 |  |  |  | 431.6631 | 0.12191
|  |  | 15.1296 | 0 | 0 |  |  |  | 433.8214 | 0.082664
|  |  | 15.2064 | 0 | 0 |  |  |  | 435.9797 | 0.005608
|  |  | 15.2832 | 0 | 0 |  |  |  | 438.1381 | 0.085182
|  |  | 15.36 | 0 | 0 |  |  |  | 440.2964 | 0.13676
|  |  | 15.4368 | 0 | 0 |  |  |  | 442.4547 | 0.1358
|  |  | 15.5136 | 0 | 0 |  |  |  | 444.613 | 0.098232
|  |  | 15.5904 | 0 | 0 |  |  |  | 446.7713 | 0.073954
|  |  | 15.6672 | 0 | 0 |  |  |  | 448.9296 | 0.079563
|  |  | 15.744 | 0 | 0 |  |  |  | 451.088 | 0.066055
|  |  | 15.8208 | 0 | 0 |  |  |  | 453.2463 | 0.025403
|  |  | 15.8976 | 0 | 0 |  |  |  | 455.4046 | 0.032561
|  |  | 15.9744 | 0 | 0 |  |  |  | 457.5629 | 0.056252
|  |  | 16.0512 | 0 | 0 |  |  |  | 459.7212 | 0.043151
|  |  | 16.128 | 0 | 0 |  |  |  | 461.8795 | 0.012518
|  |  | 16.2048 | 0 | 0 |  |  |  | 464.0378 | 0.044058
|  |  | 16.2816 | 0 | 0 |  |  |  | 466.1962 | 0.056295
|  |  | 16.3584 | 36.905 | 0.80896 |  |  |  | 468.3545 | 0.029883
|  |  | 16.4352 | 38.48 | 0.97792 |  |  |  | 470.5128 | 0.02232
|  |  | 16.512 | 39.92 | 1.1059 |  |  |  | 472.6711 | 0.065952
|  |  | 16.5888 | 41.075 | 1.2646 |  |  |  | 474.8294 | 0.075986
|  |  | 16.6656 | 42.86 | 1.2902 |  |  |  | 476.9877 | 0.059617
|  |  | 16.7424 | 43.08 | 1.6026 |  |  |  | 479.1461 | 0.086479
|  |  | 16.8192 | 43.845 | 1.7562 |  |  |  | 481.3044 | 0.13823
|  |  | 16.896 | 44.96 | 1.751 |  |  |  | 483.4627 | 0.15345
|  |  | 16.9728 | 45.905 | 1.7357 |  |  |  | 485.621 | 0.11339
|  |  | 17.0496 | 46.78 | 1.7459 |  |  |  | 487.7793 | 0.034362
|  |  | 17.1264 | 47.515 | 1.7152 |  |  |  | 489.9376 | 0.050554
|  |  | 17.2032 | 47.85 | 1.7562 |  |  |  | 492.0959 | 0.097352
|  |  | 17.28 | 47.98 | 1.8688 |  |  |  | 494.2543 | 0.093645
|  |  | 17.3568 | 48.225 | 1.9251 |  |  |  | 496.4126 | 0.062155
|  |  | 17.4336 | 48.01 | 2.0941 |  |  |  | 498.5709 | 0.068745
|  |  | 17.5104 | 48.22 | 2.1402 |  |  |  | 500.7292 | 0.091568
|  |  | 17.5872 | 48.585 | 2.1402 |  |  |  | 502.8875 | 0.079393
|  |  | 17.664 | 48.95 | 2.1504 |  |  |  | 505.0458 | 0.032843
|  |  | 17.7408 | 48.71 | 2.2682 |  |  |  | 507.2042 | 0.02224
|  |  | 17.8176 | 48.595 | 2.3501 |  |  |  | 509.3625 | 0.057269
|  |  | 17.8944 | 49.155 | 2.2426 |  |  |  | 511.5208 | 0.061674
|  |  | 17.9712 | 49.79 | 2.0941 |  |  |  | 513.6791 | 0.055527
|  |  | 18.048 | 49.85 | 2.0941 |  |  |  | 515.8374 | 0.064098
|  |  | 18.1248 | 49.445 | 2.217 |  |  |  | 517.9957 | 0.065298
|  |  | 18.2016 | 49.555 | 2.1862 |  |  |  | 520.154 | 0.041842
|  |  | 18.2784 | 49.405 | 2.2272 |  |  |  | 522.3124 | 0.032403
|  |  | 18.3552 | 49.63 | 2.1504 |  |  |  | 524.4707 | 0.064056
|  |  | 18.432 | 49.465 | 2.1453 |  |  |  | 526.629 | 0.073469
|  |  | 18.5088 | 48.825 | 2.2682 |  |  |  | 528.7873 | 0.051057
|  |  | 18.5856 | 48.765 | 2.2426 |  |  |  | 530.9456 | 0.041441
|  |  | 18.6624 | 48.895 | 2.1555 |  |  |  | 533.1039 | 0.072925
|  |  | 18.7392 | 48.63 | 2.1504 |  |  |  | 535.2623 | 0.083832
|  |  | 18.816 | 48.19 | 2.1709 |  |  |  | 537.4206 | 0.059401
|  |  | 18.8928 | 48.215 | 2.0326 |  |  |  | 539.5789 | 0.020949
|  |  | 18.9696 | 47.98 | 1.9046 |  |  |  | 541.7372 | 0.030412
|  |  | 19.0464 | 47.045 | 1.8842 |  |  |  | 543.8955 | 0.031588
|  |  | 19.1232 | 46.13 | 1.751 |  |  |  | 546.0538 | 0.028556
|  |  | 19.2 | 44.965 | 1.6077 |  |  |  | 548.2122 | 0.081511
|  |  | 19.2768 | 44.055 | 1.4387 |  |  |  | 550.3705 | 0.12816
|  |  | 19.3536 | 42.72 | 1.3773 |  |  |  | 552.5288 | 0.13521
|  |  | 19.4304 | 41.135 | 1.3722 |  |  |  | 554.6871 | 0.096944
|  |  | 19.5072 | 40.725 | 1.1213 |  |  |  | 556.8454 | 0.040284
|  |  | 19.584 | 40.67 | 0.86016 |  |  |  | 559.0037 | 0.048233
|  |  | 19.6608 | 39.905 | 0.69632 |  |  |  | 561.162 | 0.069588
|  |  | 19.7376 | 38.275 | 0.65536 |  |  |  | 563.3204 | 0.062965
|  |  | 19.8144 | 38.225 | 0.47616 |  |  |  | 565.4787 | 0.069895
|  |  | 19.8912 | 0 | 0 |  |  |  | 567.637 | 0.097877
|  |  | 19.968 | 0 | 0 |  |  |  | 569.7953 | 0.10101
|  |  | 20.0448 | 0 | 0 |  |  |  | 571.9536 | 0.063624
|  |  | 20.1216 | 0 | 0 |  |  |  | 574.1119 | 0.009573
|  |  | 20.1984 | 0 | 0 |  |  |  | 576.2703 | 0.054307
|  |  | 20.2752 | 0 | 0 |  |  |  | 578.4286 | 0.072372
|  |  | 20.352 | 0 | 0 |  |  |  | 580.5869 | 0.057485
|  |  | 20.4288 | 0 | 0 |  |  |  | 582.7452 | 0.069178
|  |  | 20.5056 | 0 | 0 |  |  |  | 584.9035 | 0.11012
|  |  | 20.5824 | 0 | 0 |  |  |  | 587.0618 | 0.12079
|  |  | 20.6592 | 0 | 0 |  |  |  | 589.2201 | 0.082258
|  |  | 20.736 | 0 | 0 |  |  |  | 591.3785 | 0.0086385
|  |  | 20.8128 | 0 | 0 |  |  |  | 593.5368 | 0.067754
|  |  | 20.8896 | 0 | 0 |  |  |  | 595.6951 | 0.11214
|  |  | 20.9664 | 0 | 0 |  |  |  | 597.8534 | 0.10884
|  |  | 21.0432 | 0 | 0 |  |  |  | 600.0117 | 0.066374
|  |  | 21.12 | 0 | 0 |  |  |  | 602.17 | 0.012936
|  |  | 21.1968 | 0 | 0 |  |  |  | 604.3284 | 0.033187
|  |  | 21.2736 | 0 | 0 |  |  |  | 606.4867 | 0.046781
|  |  | 21.3504 | 0 | 0 |  |  |  | 608.645 | 0.038226
|  |  | 21.4272 | 0 | 0 |  |  |  | 610.8033 | 0.026071
|  |  | 21.504 | 0 | 0 |  |  |  | 612.9616 | 0.034985
|  |  | 21.5808 | 0 | 0 |  |  |  | 615.1199 | 0.065929
|  |  | 21.6576 | 0 | 0 |  |  |  | 617.2782 | 0.098961
|  |  | 21.7344 | 0 | 0 |  |  |  | 619.4366 | 0.1104
|  |  | 21.8112 | 0 | 0 |  |  |  | 621.5949 | 0.083176
|  |  | 21.888 | 0 | 0 |  |  |  | 623.7532 | 0.019989
|  |  | 21.9648 | 0 | 0 |  |  |  | 625.9115 | 0.059984
|  |  | 22.0416 | 0 | 0 |  |  |  | 628.0698 | 0.11843
|  |  | 22.1184 | 0 | 0 |  |  |  | 630.2281 | 0.13081
|  |  | 22.1952 | 0 | 0 |  |  |  | 632.3865 | 0.091716
|  |  | 22.272 | 0 | 0 |  |  |  | 634.5448 | 0.021918
|  |  | 22.3488 | 0 | 0 |  |  |  | 636.7031 | 0.051051
|  |  | 22.4256 | 0 | 0 |  |  |  | 638.8614 | 0.087374
|  |  | 22.5024 | 0 | 0 |  |  |  | 641.0197 | 0.0793
|  |  | 22.5792 | 0 | 0 |  |  |  | 643.178 | 0.045079
|  |  | 22.656 | 0 | 0 |  |  |  | 645.3364 | 0.049403
|  |  | 22.7328 | 0 | 0 |  |  |  | 647.4947 | 0.075322
|  |  | 22.8096 | 0 | 0 |  |  |  | 649.653 | 0.072428
|  |  | 22.8864 | 0 | 0 |  |  |  | 651.8113 | 0.039187
|  |  | 22.9632 | 0 | 0 |  |  |  | 653.9696 | 0.012442
|  |  | 23.04 | 0 | 0 |  |  |  | 656.1279 | 0.057378
|  |  | 23.1168 | 0 | 0 |  |  |  | 658.2862 | 0.092633
|  |  | 23.1936 | 0 | 0 |  |  |  | 660.4446 | 0.11311
|  |  | 23.2704 | 0 | 0 |  |  |  | 662.6029 | 0.11387
|  |  | 23.3472 | 0 | 0 |  |  |  | 664.7612 | 0.090001
|  |  | 23.424 | 0 | 0 |  |  |  | 666.9195 | 0.047265
|  |  | 23.5008 | 0 | 0 |  |  |  | 669.0778 | 0.033365
|  |  | 23.5776 | 0 | 0 |  |  |  | 671.2361 | 0.0651
|  |  | 23.6544 | 0 | 0 |  |  |  | 673.3945 | 0.074682
|  |  | 23.7312 | 0 | 0 |  |  |  | 675.5528 | 0.05703
|  |  | 23.808 | 0 | 0 |  |  |  | 677.7111 | 0.046436
|  |  | 23.8848 | 0 | 0 |  |  |  | 679.8694 | 0.072403
|  |  | 23.9616 | 0 | 0 |  |  |  | 682.0277 | 0.092919
|  |  | 24.0384 | 0 | 0 |  |  |  | 684.186 | 0.089733
|  |  | 24.1152 | 0 | 0 |  |  |  | 686.3443 | 0.069022
|  |  | 24.192 | 0 | 0 |  |  |  | 688.5027 | 0.048677
|  |  | 24.2688 | 0 | 0 |  |  |  | 690.661 | 0.042335
|  |  | 24.3456 | 0 | 0 |  |  |  | 692.8193 | 0.043303
|  |  | 24.4224 | 0 | 0 |  |  |  | 694.9776 | 0.044285
|  |  | 24.4992 | 0 | 0 |  |  |  | 697.1359 | 0.045253
|  |  | 24.576 | 0 | 0 |  |  |  | 699.2942 | 0.045685
|  |  | 24.6528 | 0 | 0 |  |  |  | 701.4526 | 0.043468
|  |  | 24.7296 | 0 | 0 |  |  |  | 703.6109 | 0.033435
|  |  | 24.8064 | 0 | 0 |  |  |  | 705.7692 | 0.010686
|  |  | 24.8832 | 0 | 0 |  |  |  | 707.9275 | 0.024781
|  |  | 24.96 | 0 | 0 |  |  |  | 710.0858 | 0.056944
|  |  | 25.0368 | 0 | 0 |  |  |  | 712.2441 | 0.069846
|  |  | 25.1136 | 0 | 0 |  |  |  | 714.4024 | 0.052938
|  |  | 25.1904 | 0 | 0 |  |  |  | 716.5608 | 0.027844
|  |  | 25.2672 | 0 | 0 |  |  |  | 718.7191 | 0.067128
|  |  | 25.344 | 0 | 0 |  |  |  | 720.8774 | 0.10565
|  |  | 25.4208 | 0 | 0 |  |  |  | 723.0357 | 0.10974
|  |  | 25.4976 | 0 | 0 |  |  |  | 725.194 | 0.084113
|  |  | 25.5744 | 0 | 0 |  |  |  | 727.3523 | 0.079076
|  |  | 25.6512 | 0 | 0 |  |  |  | 729.5107 | 0.1146
|  |  | 25.728 | 0 | 0 |  |  |  | 731.669 | 0.13368
|  |  | 25.8048 | 0 | 0 |  |  |  | 733.8273 | 0.11248
|  |  | 25.8816 | 0 | 0 |  |  |  | 735.9856 | 0.060277
|  |  | 25.9584 | 0 | 0 |  |  |  | 738.1439 | 0.0045162
|  |  | 26.0352 | 0 | 0 |  |  |  | 740.3022 | 0.027009
|  |  | 26.112 | 0 | 0 |  |  |  | 742.4606 | 0.020512
|  |  | 26.1888 | 0 | 0 |  |  |  | 744.6189 | 0.015635
|  |  | 26.2656 | 0 | 0 |  |  |  | 746.7772 | 0.057332
|  |  | 26.3424 | 0 | 0 |  |  |  | 748.9355 | 0.079309
|  |  | 26.4192 | 0 | 0 |  |  |  | 751.0938 | 0.070428
|  |  | 26.496 | 0 | 0 |  |  |  | 753.2521 | 0.043052
|  |  | 26.5728 | 0 | 0 |  |  |  | 755.4104 | 0.034189
|  |  | 26.6496 | 0 | 0 |  |  |  | 757.5688 | 0.037494
|  |  | 26.7264 | 0 | 0 |  |  |  | 759.7271 | 0.017626
|  |  | 26.8032 | 0 | 0 |  |  |  | 761.8854 | 0.030151
|  |  | 26.88 | 0 | 0 |  |  |  | 764.0437 | 0.073239
|  |  | 26.9568 | 0 | 0 |  |  |  | 766.202 | 0.094441
|  |  | 27.0336 | 0 | 0 |  |  |  | 768.3603 | 0.084551
|  |  | 27.1104 | 0 | 0 |  |  |  | 770.5187 | 0.059082
|  |  | 27.1872 | 0 | 0 |  |  |  | 772.677 | 0.053626
|  |  | 27.264 | 0 | 0 |  |  |  | 774.8353 | 0.060641
|  |  | 27.3408 | 0 | 0 |  |  |  | 776.9936 | 0.048651
|  |  | 27.4176 | 0 | 0 |  |  |  | 779.1519 | 0.018648
|  |  | 27.4944 | 0 | 0 |  |  |  | 781.3102 | 0.010812
|  |  | 27.5712 | 0 | 0 |  |  |  | 783.4685 | 0.018369
|  |  | 27.648 | 0 | 0 |  |  |  | 785.6269 | 0.0016706
|  |  | 27.7248 | 0 | 0 |  |  |  | 787.7852 | 0.037766
|  |  | 27.8016 | 0 | 0 |  |  |  | 789.9435 | 0.067047
|  |  | 27.8784 | 0 | 0 |  |  |  | 792.1018 | 0.069216
|  |  | 27.9552 | 0 | 0 |  |  |  | 794.2601 | 0.041302
|  |  | 28.032 | 0 | 0 |  |  |  | 796.4184 | 0.028088
|  |  | 28.1088 | 0 | 0 |  |  |  | 798.5768 | 0.061918
|  |  | 28.1856 | 0 | 0 |  |  |  | 800.7351 | 0.072144
|  |  | 28.2624 | 0 | 0 |  |  |  | 802.8934 | 0.049569
|  |  | 28.3392 | 0 | 0 |  |  |  | 805.0517 | 0.034165
|  |  | 28.416 | 0 | 0 |  |  |  | 807.21 | 0.063809
|  |  | 28.4928 | 0 | 0 |  |  |  | 809.3683 | 0.073441
|  |  | 28.5696 | 0 | 0 |  |  |  | 811.5266 | 0.044749
|  |  | 28.6464 | 0 | 0 |  |  |  | 813.685 | 0.0097539
|  |  | 28.7232 | 0 | 0 |  |  |  | 815.8433 | 0.061725
|  |  | 28.8 | 0 | 0 |  |  |  | 818.0016 | 0.083801
|  |  | 28.8768 | 0 | 0 |  |  |  | 820.1599 | 0.064966
|  |  | 28.9536 | 0 | 0 |  |  |  | 822.3182 | 0.0186
|  |  | 29.0304 | 0 | 0 |  |  |  | 824.4765 | 0.039516
|  |  | 29.1072 | 0 | 0 |  |  |  | 826.6349 | 0.067669
|  |  | 29.184 | 0 | 0 |  |  |  | 828.7932 | 0.060353
|  |  | 29.2608 | 0 | 0 |  |  |  | 830.9515 | 0.028231
|  |  | 29.3376 | 0 | 0 |  |  |  | 833.1098 | 0.031335
|  |  | 29.4144 | 0 | 0 |  |  |  | 835.2681 | 0.056203
|  |  | 29.4912 | 0 | 0 |  |  |  | 837.4264 | 0.060789
|  |  | 29.568 | 0 | 0 |  |  |  | 839.5848 | 0.054675
|  |  | 29.6448 | 0 | 0 |  |  |  | 841.7431 | 0.051295
|  |  | 29.7216 | 0 | 0 |  |  |  | 843.9014 | 0.040143
|  |  | 29.7984 | 0 | 0 |  |  |  | 846.0597 | 0.017441
|  |  | 29.8752 | 0 | 0 |  |  |  | 848.218 | 0.044652
|  |  | 29.952 | 0 | 0 |  |  |  | 850.3763 | 0.078546
|  |  | 30.0288 | 0 | 0 |  |  |  | 852.5346 | 0.08491
|  |  | 30.1056 | 0 | 0 |  |  |  | 854.693 | 0.06007
|  |  | 30.1824 | 0 | 0 |  |  |  | 856.8513 | 0.030471
|  |  | 30.2592 | 0 | 0 |  |  |  | 859.0096 | 0.044779
|  |  | 30.336 | 0 | 0 |  |  |  | 861.1679 | 0.051477
|  |  | 30.4128 | 0 | 0 |  |  |  | 863.3262 | 0.027273
|  |  | 30.4896 | 0 | 0 |  |  |  | 865.4845 | 0.019479
|  |  | 30.5664 | 0 | 0 |  |  |  | 867.6429 | 0.05407
|  |  | 30.6432 | 0 | 0 |  |  |  | 869.8012 | 0.060581
|  |  | 30.72 | 0 | 0 |  |  |  | 871.9595 | 0.041065
|  |  | 30.7968 | 0 | 0 |  |  |  | 874.1178 | 0.051622
|  |  | 30.8736 | 0 | 0 |  |  |  | 876.2761 | 0.085107
|  |  | 30.9504 | 0 | 0 |  |  |  | 878.4344 | 0.087727
|  |  | 31.0272 | 0 | 0 |  |  |  | 880.5927 | 0.046599
|  |  | 31.104 | 0 | 0 |  |  |  | 882.7511 | 0.023958
|  |  | 31.1808 | 0 | 0 |  |  |  | 884.9094 | 0.092708
|  |  | 31.2576 | 0 | 0 |  |  |  | 887.0677 | 0.12998
|  |  | 31.3344 | 0 | 0 |  |  |  | 889.226 | 0.12359
|  |  | 31.4112 | 0 | 0 |  |  |  | 891.3843 | 0.088644
|  |  | 31.488 | 0 | 0 |  |  |  | 893.5426 | 0.06624
|  |  | 31.5648 | 0 | 0 |  |  |  | 895.701 | 0.070801
|  |  | 31.6416 | 0 | 0 |  |  |  | 897.8593 | 0.065365
|  |  | 31.7184 | 0 | 0 |  |  |  | 900.0176 | 0.055538
|  |  | 31.7952 | 0 | 0 |  |  |  | 902.1759 | 0.075461
|  |  | 31.872 | 0 | 0 |  |  |  | 904.3342 | 0.097673
|  |  | 31.9488 | 0 | 0 |  |  |  | 906.4925 | 0.092315
|  |  | 32.0256 | 0 | 0 |  |  |  | 908.6508 | 0.058384
|  |  | 32.1024 | 0 | 0 |  |  |  | 910.8092 | 0.01559
|  |  | 32.1792 | 0 | 0 |  |  |  | 912.9675 | 0.026526
|  |  | 32.256 | 0 | 0 |  |  |  | 915.1258 | 0.039735
|  |  | 32.3328 | 0 | 0 |  |  |  | 917.2841 | 0.041547
|  |  | 32.4096 | 0 | 0 |  |  |  | 919.4424 | 0.051016
|  |  | 32.4864 | 0 | 0 |  |  |  | 921.6007 | 0.058476
|  |  | 32.5632 | 0 | 0 |  |  |  | 923.7591 | 0.045886
|  |  | 32.64 | 0 | 0 |  |  |  | 925.9174 | 0.018333
|  |  | 32.7168 | 0 | 0 |  |  |  | 928.0757 | 0.047087
|  |  | 32.7936 | 35.825 | 0.61952 |  |  |  | 930.234 | 0.089267
|  |  | 32.8704 | 38.795 | 0.61952 |  |  |  | 932.3923 | 0.11364
|  |  | 32.9472 | 40.755 | 0.71168 |  |  |  | 934.5506 | 0.11402
|  |  | 33.024 | 41.255 | 1.0138 |  |  |  | 936.709 | 0.093485
|  |  | 33.1008 | 42.735 | 1.2186 |  |  |  | 938.8673 | 0.061161
|  |  | 33.1776 | 44.55 | 1.3414 |  |  |  | 941.0256 | 0.03413
|  |  | 33.2544 | 46.015 | 1.4746 |  |  |  | 943.1839 | 0.040481
|  |  | 33.3312 | 46.675 | 1.7562 |  |  |  | 945.3422 | 0.061604
|  |  | 33.408 | 47.96 | 1.833 |  |  |  | 947.5005 | 0.081736
|  |  | 33.4848 | 49.535 | 1.7971 |  |  |  | 949.6588 | 0.10388
|  |  | 33.5616 | 49.7 | 2.0224 |  |  |  | 951.8172 | 0.12562
|  |  | 33.6384 | 50.145 | 2.1555 |  |  |  | 953.9755 | 0.13669
|  |  | 33.7152 | 51.18 | 2.0992 |  |  |  | 956.1338 | 0.1296
|  |  | 33.792 | 52.1 | 2.0275 |  |  |  | 958.2921 | 0.10758
|  |  | 33.8688 | 52.03 | 2.1504 |  |  |  | 960.4504 | 0.086535
|  |  | 33.9456 | 51.765 | 2.3142 |  |  |  | 962.6087 | 0.086227
|  |  | 34.0224 | 52.11 | 2.3398 |  |  |  | 964.7671 | 0.10297
|  |  | 34.0992 | 52.6 | 2.3296 |  |  |  | 966.9254 | 0.11716
|  |  | 34.176 | 53.14 | 2.3194 |  |  |  | 969.0837 | 0.11577
|  |  | 34.2528 | 53.595 | 2.304 |  |  |  | 971.242 | 0.094495
|  |  | 34.3296 | 53.84 | 2.3194 |  |  |  | 973.4003 | 0.061258
|  |  | 34.4064 | 53.855 | 2.3859 |  |  |  | 975.5586 | 0.042373
|  |  | 34.4832 | 53.745 | 2.4525 |  |  |  | 977.7169 | 0.048008
|  |  | 34.56 | 53.725 | 2.5139 |  |  |  | 979.8753 | 0.044401
|  |  | 34.6368 | 53.625 | 2.5498 |  |  |  | 982.0336 | 0.041966
|  |  | 34.7136 | 53.585 | 2.5651 |  |  |  | 984.1919 | 0.075598
|  |  | 34.7904 | 53.725 | 2.5139 |  |  |  | 986.3502 | 0.11056
|  |  | 34.8672 | 53.785 | 2.5037 |  |  |  | 988.5085 | 0.11767
|  |  | 34.944 | 53.355 | 2.6317 |  |  |  | 990.6668 | 0.091312
|  |  | 35.0208 | 52.96 | 2.7443 |  |  |  | 992.8252 | 0.050227
|  |  | 35.0976 | 53.04 | 2.6982 |  |  |  | 994.9835 | 0.042365
|  |  | 35.1744 | 53.38 | 2.5549 |  |  |  | 997.1418 | 0.055647
|  |  | 35.2512 | 53.125 | 2.5344 |  |  |  | 999.3001 | 0.047621
|  |  | 35.328 | 52.87 | 2.5293 |  |  |  | 1001.4584 | 0.028673
|  |  | 35.4048 | 52.79 | 2.4678 |  |  |  | 1003.6167 | 0.029412
|  |  | 35.4816 | 52.865 | 2.345 |  |  |  | 1005.7751 | 0.025849
|  |  | 35.5584 | 52.57 | 2.304 |  |  |  | 1007.9334 | 0.005889
|  |  | 35.6352 | 51.315 | 2.4474 |  |  |  | 1010.0917 | 0.051832
|  |  | 35.712 | 50.735 | 2.3347 |  |  |  | 1012.25 | 0.093656
|  |  | 35.7888 | 49.835 | 2.2016 |  |  |  | 1014.4083 | 0.1129
|  |  | 35.8656 | 48.68 | 2.0634 |  |  |  | 1016.5666 | 0.11336
|  |  | 35.9424 | 47.85 | 1.8278 |  |  |  | 1018.7249 | 0.11759
|  |  | 36.0192 | 46.44 | 1.7101 |  |  |  | 1020.8833 | 0.1269
|  |  | 36.096 | 45.765 | 1.4592 |  |  |  | 1023.0416 | 0.11439
|  |  | 36.1728 | 44.845 | 1.2851 |  |  |  | 1025.1999 | 0.067555
|  |  | 36.2496 | 43.63 | 1.1776 |  |  |  | 1027.3582 | 0.025443
|  |  | 36.3264 | 42.415 | 1.065 |  |  |  | 1029.5165 | 0.071
|  |  | 36.4032 | 40.9 | 1.0035 |  |  |  | 1031.6748 | 0.086276
|  |  | 36.48 | 39.945 | 0.8704 |  |  |  | 1033.8332 | 0.047317
|  |  | 36.5568 | 38.34 | 0.79872 |  |  |  | 1035.9915 | 0.036641
|  |  | 36.6336 | 37.875 | 0.57856 |  |  |  | 1038.1498 | 0.12322
|  |  | 36.7104 | 0 | 0 |  |  |  | 1040.3081 | 0.17963
|  |  | 36.7872 | 0 | 0 |  |  |  | 1042.4664 | 0.18275
|  |  | 36.864 | 0 | 0 |  |  |  | 1044.6247 | 0.1365
|  |  | 36.9408 | 0 | 0 |  |  |  | 1046.783 | 0.068976
|  |  | 37.0176 | 0 | 0 |  |  |  | 1048.9414 | 0.037134
|  |  | 37.0944 | 0 | 0 |  |  |  | 1051.0997 | 0.061551
|  |  | 37.1712 | 0 | 0 |  |  |  | 1053.258 | 0.075213
|  |  | 37.248 | 0 | 0 |  |  |  | 1055.4163 | 0.073453
|  |  | 37.3248 | 0 | 0 |  |  |  | 1057.5746 | 0.051268
|  |  | 37.4016 | 0 | 0 |  |  |  | 1059.7329 | 0.00046516
|  |  | 37.4784 | 0 | 0 |  |  |  | 1061.8913 | 0.073873
|  |  | 37.5552 | 0 | 0 |  |  |  | 1064.0496 | 0.14744
|  |  | 37.632 | 0 | 0 |  |  |  | 1066.2079 | 0.19223
|  |  | 37.7088 | 0 | 0 |  |  |  | 1068.3662 | 0.19159
|  |  | 37.7856 | 0 | 0 |  |  |  | 1070.5245 | 0.1506
|  |  | 37.8624 | 0 | 0 |  |  |  | 1072.6828 | 0.093851
|  |  | 37.9392 | 0 | 0 |  |  |  | 1074.8411 | 0.051465
|  |  | 38.016 | 0 | 0 |  |  |  | 1076.9995 | 0.039994
|  |  | 38.0928 | 0 | 0 |  |  |  | 1079.1578 | 0.062936
|  |  | 38.1696 | 0 | 0 |  |  |  | 1081.3161 | 0.10381
|  |  | 38.2464 | 0 | 0 |  |  |  | 1083.4744 | 0.13256
|  |  | 38.3232 | 0 | 0 |  |  |  | 1085.6327 | 0.12898
|  |  | 38.4 | 0 | 0 |  |  |  | 1087.791 | 0.095408
|  |  | 38.4768 | 0 | 0 |  |  |  | 1089.9494 | 0.069444
|  |  | 38.5536 | 0 | 0 |  |  |  | 1092.1077 | 0.090635
|  |  | 38.6304 | 0 | 0 |  |  |  | 1094.266 | 0.1122
|  |  | 38.7072 | 0 | 0 |  |  |  | 1096.4243 | 0.10426
|  |  | 38.784 | 0 | 0 |  |  |  | 1098.5826 | 0.069869
|  |  | 38.8608 | 0 | 0 |  |  |  | 1100.7409 | 0.028642
|  |  | 38.9376 | 0 | 0 |  |  |  | 1102.8993 | 0.02433
|  |  | 39.0144 | 0 | 0
|  |  | 39.0912 | 0 | 0
|  |  | 39.168 | 0 | 0
|  |  | 39.2448 | 0 | 0
|  |  | 39.3216 | 0 | 0
|  |  | 39.3984 | 0 | 0
|  |  | 39.4752 | 0 | 0
|  |  | 39.552 | 0 | 0
|  |  | 39.6288 | 0 | 0
|  |  | 39.7056 | 0 | 0
|  |  | 39.7824 | 0 | 0
|  |  | 39.8592 | 0 | 0
|  |  | 39.936 | 0 | 0
|  |  | 40.0128 | 0 | 0
|  |  | 40.0896 | 0 | 0
|  |  | 40.1664 | 0 | 0
|  |  | 40.2432 | 0 | 0
|  |  | 40.32 | 0 | 0
|  |  | 40.3968 | 0 | 0
|  |  | 40.4736 | 0 | 0
|  |  | 40.5504 | 0 | 0
|  |  | 40.6272 | 0 | 0
|  |  | 40.704 | 0 | 0
|  |  | 40.7808 | 0 | 0
|  |  | 40.8576 | 0 | 0
|  |  | 40.9344 | 0 | 0
|  |  | 41.0112 | 0 | 0
|  |  | 41.088 | 0 | 0
|  |  | 41.1648 | 0 | 0
|  |  | 41.2416 | 0 | 0
|  |  | 41.3184 | 0 | 0
|  |  | 41.3952 | 0 | 0
|  |  | 41.472 | 0 | 0
|  |  | 41.5488 | 0 | 0
|  |  | 41.6256 | 0 | 0
|  |  | 41.7024 | 0 | 0
|  |  | 41.7792 | 0 | 0
|  |  | 41.856 | 0 | 0
|  |  | 41.9328 | 0 | 0
|  |  | 42.0096 | 0 | 0
|  |  | 42.0864 | 0 | 0
|  |  | 42.1632 | 0 | 0
|  |  | 42.24 | 0 | 0
|  |  | 42.3168 | 0 | 0
|  |  | 42.3936 | 0 | 0
|  |  | 42.4704 | 0 | 0
|  |  | 42.5472 | 0 | 0
|  |  | 42.624 | 0 | 0
|  |  | 42.7008 | 0 | 0
|  |  | 42.7776 | 0 | 0
|  |  | 42.8544 | 0 | 0
|  |  | 42.9312 | 0 | 0
|  |  | 43.008 | 0 | 0
|  |  | 43.0848 | 0 | 0
|  |  | 43.1616 | 0 | 0
|  |  | 43.2384 | 0 | 0
|  |  | 43.3152 | 0 | 0
|  |  | 43.392 | 0 | 0
|  |  | 43.4688 | 0 | 0
|  |  | 43.5456 | 0 | 0
|  |  | 43.6224 | 0 | 0
|  |  | 43.6992 | 0 | 0
|  |  | 43.776 | 0 | 0
|  |  | 43.8528 | 0 | 0
|  |  | 43.9296 | 0 | 0
|  |  | 44.0064 | 0 | 0
|  |  | 44.0832 | 0 | 0
|  |  | 44.16 | 0 | 0
|  |  | 44.2368 | 0 | 0
|  |  | 44.3136 | 0 | 0
|  |  | 44.3904 | 0 | 0
|  |  | 44.4672 | 0 | 0
|  |  | 44.544 | 0 | 0
|  |  | 44.6208 | 0 | 0
|  |  | 44.6976 | 0 | 0
|  |  | 44.7744 | 0 | 0
|  |  | 44.8512 | 0 | 0
|  |  | 44.928 | 0 | 0
|  |  | 45.0048 | 0 | 0
|  |  | 45.0816 | 0 | 0
|  |  | 45.1584 | 0 | 0
|  |  | 45.2352 | 0 | 0
|  |  | 45.312 | 0 | 0
|  |  | 45.3888 | 0 | 0
|  |  | 45.4656 | 0 | 0
|  |  | 45.5424 | 0 | 0
|  |  | 45.6192 | 0 | 0
|  |  | 45.696 | 0 | 0
|  |  | 45.7728 | 0 | 0
|  |  | 45.8496 | 0 | 0
|  |  | 45.9264 | 0 | 0
|  |  | 46.0032 | 0 | 0
|  |  | 46.08 | 0 | 0
|  |  | 46.1568 | 0 | 0
|  |  | 46.2336 | 0 | 0
|  |  | 46.3104 | 0 | 0
|  |  | 46.3872 | 0 | 0
|  |  | 46.464 | 0 | 0
|  |  | 46.5408 | 0 | 0
|  |  | 46.6176 | 0 | 0
|  |  | 46.6944 | 0 | 0
|  |  | 46.7712 | 0 | 0
|  |  | 46.848 | 0 | 0
|  |  | 46.9248 | 0 | 0
|  |  | 47.0016 | 0 | 0
|  |  | 47.0784 | 0 | 0
|  |  | 47.1552 | 0 | 0
|  |  | 47.232 | 0 | 0
|  |  | 47.3088 | 0 | 0
|  |  | 47.3856 | 0 | 0
|  |  | 47.4624 | 0 | 0
|  |  | 47.5392 | 0 | 0
|  |  | 47.616 | 0 | 0
|  |  | 47.6928 | 0 | 0
|  |  | 47.7696 | 0 | 0
|  |  | 47.8464 | 0 | 0
|  |  | 47.9232 | 0 | 0
|  |  | 48 | 0 | 0
|  |  | 48.0768 | 0 | 0
|  |  | 48.1536 | 0 | 0
|  |  | 48.2304 | 0 | 0
|  |  | 48.3072 | 0 | 0
|  |  | 48.384 | 0 | 0
|  |  | 48.4608 | 0 | 0
|  |  | 48.5376 | 0 | 0
|  |  | 48.6144 | 0 | 0
|  |  | 48.6912 | 0 | 0
|  |  | 48.768 | 0 | 0
|  |  | 48.8448 | 0 | 0
|  |  | 48.9216 | 0 | 0
|  |  | 48.9984 | 0 | 0
|  |  | 49.0752 | 0 | 0
|  |  | 49.152 | 0 | 0
|  |  | 49.2288 | 0 | 0
|  |  | 49.3056 | 0 | 0
|  |  | 49.3824 | 37.775 | 0.47616
|  |  | 49.4592 | 38.655 | 0.7424
|  |  | 49.536 | 41.05 | 0.86528
|  |  | 49.6128 | 43.4 | 0.95744
|  |  | 49.6896 | 44.77 | 1.1674
|  |  | 49.7664 | 46.605 | 1.2954
|  |  | 49.8432 | 48.185 | 1.4592
|  |  | 49.92 | 48.935 | 1.7306
|  |  | 49.9968 | 50.3 | 1.8586
|  |  | 50.0736 | 51.55 | 1.9712
|  |  | 50.1504 | 52.3 | 2.1248
|  |  | 50.2272 | 53.03 | 2.217
|  |  | 50.304 | 53.85 | 2.2733
|  |  | 50.3808 | 54.485 | 2.2886
|  |  | 50.4576 | 55.19 | 2.2426
|  |  | 50.5344 | 55.175 | 2.3347
|  |  | 50.6112 | 55.365 | 2.3808
|  |  | 50.688 | 55.63 | 2.4269
|  |  | 50.7648 | 56.06 | 2.4371
|  |  | 50.8416 | 56.06 | 2.5344
|  |  | 50.9184 | 56.36 | 2.5549
|  |  | 50.9952 | 56.6 | 2.6061
|  |  | 51.072 | 56.595 | 2.6982
|  |  | 51.1488 | 56.89 | 2.6778
|  |  | 51.2256 | 57.105 | 2.6675
|  |  | 51.3024 | 57.51 | 2.6112
|  |  | 51.3792 | 57.575 | 2.5958
|  |  | 51.456 | 57.34 | 2.6522
|  |  | 51.5328 | 57.81 | 2.56
|  |  | 51.6096 | 57.79 | 2.5446
|  |  | 51.6864 | 57.165 | 2.6573
|  |  | 51.7632 | 57.22 | 2.6112
|  |  | 51.84 | 56.95 | 2.6419
|  |  | 51.9168 | 56.245 | 2.7648
|  |  | 51.9936 | 55.96 | 2.7648
|  |  | 52.0704 | 56.15 | 2.6214
|  |  | 52.1472 | 56.245 | 2.5242
|  |  | 52.224 | 56.445 | 2.3654
|  |  | 52.3008 | 55.805 | 2.3347
|  |  | 52.3776 | 54.865 | 2.2682
|  |  | 52.4544 | 53.435 | 2.2477
|  |  | 52.5312 | 51.875 | 2.2118
|  |  | 52.608 | 50.865 | 2.007
|  |  | 52.6848 | 49.815 | 1.8176
|  |  | 52.7616 | 48.49 | 1.7101
|  |  | 52.8384 | 47.685 | 1.495
|  |  | 52.9152 | 46.645 | 1.3517
|  |  | 52.992 | 45.6 | 1.1827
|  |  | 53.0688 | 43.71 | 1.1418
|  |  | 53.1456 | 42.02 | 1.1008
|  |  | 53.2224 | 40.845 | 0.96768
|  |  | 53.2992 | 39.89 | 0.77312
|  |  | 53.376 | 38.19 | 0.7168
|  |  | 53.4528 | 36.395 | 0.60928
|  |  | 53.5296 | 0 | 0
|  |  | 53.6064 | 0 | 0
|  |  | 53.6832 | 0 | 0
|  |  | 53.76 | 0 | 0
|  |  | 53.8368 | 0 | 0
|  |  | 53.9136 | 0 | 0
|  |  | 53.9904 | 0 | 0
|  |  | 54.0672 | 0 | 0
|  |  | 54.144 | 0 | 0
|  |  | 54.2208 | 0 | 0
|  |  | 54.2976 | 0 | 0
|  |  | 54.3744 | 0 | 0
|  |  | 54.4512 | 0 | 0
|  |  | 54.528 | 0 | 0
|  |  | 54.6048 | 0 | 0
|  |  | 54.6816 | 0 | 0
|  |  | 54.7584 | 0 | 0
|  |  | 54.8352 | 0 | 0
|  |  | 54.912 | 0 | 0
|  |  | 54.9888 | 0 | 0
|  |  | 55.0656 | 0 | 0
|  |  | 55.1424 | 0 | 0
|  |  | 55.2192 | 0 | 0
|  |  | 55.296 | 0 | 0
|  |  | 55.3728 | 0 | 0
|  |  | 55.4496 | 0 | 0
|  |  | 55.5264 | 0 | 0
|  |  | 55.6032 | 0 | 0
|  |  | 55.68 | 0 | 0
|  |  | 55.7568 | 0 | 0
|  |  | 55.8336 | 0 | 0
|  |  | 55.9104 | 0 | 0
|  |  | 55.9872 | 0 | 0
|  |  | 56.064 | 0 | 0
|  |  | 56.1408 | 0 | 0
|  |  | 56.2176 | 0 | 0
|  |  | 56.2944 | 0 | 0
|  |  | 56.3712 | 0 | 0
|  |  | 56.448 | 0 | 0
|  |  | 56.5248 | 0 | 0
|  |  | 56.6016 | 0 | 0
|  |  | 56.6784 | 0 | 0
|  |  | 56.7552 | 0 | 0
|  |  | 56.832 | 0 | 0
|  |  | 56.9088 | 0 | 0
|  |  | 56.9856 | 0 | 0
|  |  | 57.0624 | 0 | 0
|  |  | 57.1392 | 0 | 0
|  |  | 57.216 | 0 | 0
|  |  | 57.2928 | 0 | 0
|  |  | 57.3696 | 0 | 0
|  |  | 57.4464 | 0 | 0
|  |  | 57.5232 | 0 | 0
|  |  | 57.6 | 0 | 0
|  |  | 57.6768 | 0 | 0
|  |  | 57.7536 | 0 | 0
|  |  | 57.8304 | 0 | 0
|  |  | 57.9072 | 0 | 0
|  |  | 57.984 | 0 | 0
|  |  | 58.0608 | 0 | 0
|  |  | 58.1376 | 0 | 0
|  |  | 58.2144 | 0 | 0
|  |  | 58.2912 | 0 | 0
|  |  | 58.368 | 0 | 0
|  |  | 58.4448 | 0 | 0
|  |  | 58.5216 | 0 | 0
|  |  | 58.5984 | 0 | 0
|  |  | 58.6752 | 0 | 0
|  |  | 58.752 | 0 | 0
|  |  | 58.8288 | 0 | 0
|  |  | 58.9056 | 0 | 0
|  |  | 58.9824 | 0 | 0
|  |  | 59.0592 | 0 | 0
|  |  | 59.136 | 0 | 0
|  |  | 59.2128 | 0 | 0
|  |  | 59.2896 | 0 | 0
|  |  | 59.3664 | 0 | 0
|  |  | 59.4432 | 0 | 0
|  |  | 59.52 | 0 | 0
|  |  | 59.5968 | 0 | 0
|  |  | 59.6736 | 0 | 0
|  |  | 59.7504 | 0 | 0
|  |  | 59.8272 | 0 | 0
|  |  | 59.904 | 0 | 0
|  |  | 59.9808 | 0 | 0
|  |  | 60.0576 | 0 | 0
|  |  | 60.1344 | 0 | 0
|  |  | 60.2112 | 0 | 0
|  |  | 60.288 | 0 | 0
|  |  | 60.3648 | 0 | 0
|  |  | 60.4416 | 0 | 0
|  |  | 60.5184 | 0 | 0
|  |  | 60.5952 | 0 | 0
|  |  | 60.672 | 0 | 0
|  |  | 60.7488 | 0 | 0
|  |  | 60.8256 | 0 | 0
|  |  | 60.9024 | 0 | 0
|  |  | 60.9792 | 0 | 0
|  |  | 61.056 | 0 | 0
|  |  | 61.1328 | 0 | 0
|  |  | 61.2096 | 0 | 0
|  |  | 61.2864 | 0 | 0
|  |  | 61.3632 | 0 | 0
|  |  | 61.44 | 0 | 0
|  |  | 61.5168 | 0 | 0
|  |  | 61.5936 | 0 | 0
|  |  | 61.6704 | 0 | 0
|  |  | 61.7472 | 0 | 0
|  |  | 61.824 | 0 | 0
|  |  | 61.9008 | 0 | 0
|  |  | 61.9776 | 0 | 0
|  |  | 62.0544 | 0 | 0
|  |  | 62.1312 | 0 | 0
|  |  | 62.208 | 0 | 0
|  |  | 62.2848 | 0 | 0
|  |  | 62.3616 | 0 | 0
|  |  | 62.4384 | 0 | 0
|  |  | 62.5152 | 0 | 0
|  |  | 62.592 | 0 | 0
|  |  | 62.6688 | 0 | 0
|  |  | 62.7456 | 0 | 0
|  |  | 62.8224 | 0 | 0
|  |  | 62.8992 | 0 | 0
|  |  | 62.976 | 0 | 0
|  |  | 63.0528 | 0 | 0
|  |  | 63.1296 | 0 | 0
|  |  | 63.2064 | 0 | 0
|  |  | 63.2832 | 0 | 0
|  |  | 63.36 | 0 | 0
|  |  | 63.4368 | 0 | 0
|  |  | 63.5136 | 0 | 0
|  |  | 63.5904 | 0 | 0
|  |  | 63.6672 | 0 | 0
|  |  | 63.744 | 0 | 0
|  |  | 63.8208 | 0 | 0
|  |  | 63.8976 | 0 | 0
|  |  | 63.9744 | 0 | 0
|  |  | 64.0512 | 0 | 0
|  |  | 64.128 | 0 | 0
|  |  | 64.2048 | 0 | 0
|  |  | 64.2816 | 0 | 0
|  |  | 64.3584 | 0 | 0
|  |  | 64.4352 | 0 | 0
|  |  | 64.512 | 0 | 0
|  |  | 64.5888 | 0 | 0
|  |  | 64.6656 | 0 | 0
|  |  | 64.7424 | 0 | 0
|  |  | 64.8192 | 0 | 0
|  |  | 64.896 | 0 | 0
|  |  | 64.9728 | 0 | 0
|  |  | 65.0496 | 0 | 0
|  |  | 65.1264 | 0 | 0
|  |  | 65.2032 | 0 | 0
|  |  | 65.28 | 0 | 0
|  |  | 65.3568 | 0 | 0
|  |  | 65.4336 | 0 | 0
|  |  | 65.5104 | 0 | 0
|  |  | 65.5872 | 0 | 0
|  |  | 65.664 | 0 | 0
|  |  | 65.7408 | 0 | 0
|  |  | 65.8176 | 0 | 0
|  |  | 65.8944 | 0 | 0
|  |  | 65.9712 | 0 | 0
|  |  | 66.048 | 38.07 | 0.7168
|  |  | 66.1248 | 40.15 | 0.90624
|  |  | 66.2016 | 42.71 | 0.97792
|  |  | 66.2784 | 44.855 | 1.1469
|  |  | 66.3552 | 46.715 | 1.3261
|  |  | 66.432 | 48.465 | 1.4797
|  |  | 66.5088 | 50.15 | 1.6077
|  |  | 66.5856 | 51.425 | 1.7562
|  |  | 66.6624 | 53.08 | 1.8381
|  |  | 66.7392 | 54.28 | 1.9763
|  |  | 66.816 | 55.275 | 2.0736
|  |  | 66.8928 | 55.735 | 2.2016
|  |  | 66.9696 | 56.085 | 2.2784
|  |  | 67.0464 | 56.565 | 2.304
|  |  | 67.1232 | 57.27 | 2.2938
|  |  | 67.2 | 57.84 | 2.304
|  |  | 67.2768 | 57.905 | 2.4013
|  |  | 67.3536 | 57.865 | 2.519
|  |  | 67.4304 | 58.4 | 2.5139
|  |  | 67.5072 | 58.72 | 2.5242
|  |  | 67.584 | 59.245 | 2.4986
|  |  | 67.6608 | 59.85 | 2.4576
|  |  | 67.7376 | 59.85 | 2.519
|  |  | 67.8144 | 59.555 | 2.6368
|  |  | 67.8912 | 59.65 | 2.6573
|  |  | 67.968 | 60 | 2.6163
|  |  | 68.0448 | 59.92 | 2.6573
|  |  | 68.1216 | 59.495 | 2.7546
|  |  | 68.1984 | 59.635 | 2.7187
|  |  | 68.2752 | 60.095 | 2.6061
|  |  | 68.352 | 60.16 | 2.5805
|  |  | 68.4288 | 59.915 | 2.6061
|  |  | 68.5056 | 59.34 | 2.688
|  |  | 68.5824 | 59.095 | 2.6982
|  |  | 68.6592 | 58.9 | 2.688
|  |  | 68.736 | 58.465 | 2.6931
|  |  | 68.8128 | 58.17 | 2.6573
|  |  | 68.8896 | 57.835 | 2.6214
|  |  | 68.9664 | 57.16 | 2.6214
|  |  | 69.0432 | 56.425 | 2.56
|  |  | 69.12 | 55.36 | 2.473
|  |  | 69.1968 | 54.31 | 2.3091
|  |  | 69.2736 | 52.96 | 2.1658
|  |  | 69.3504 | 51.615 | 2.0173
|  |  | 69.4272 | 50.18 | 1.92
|  |  | 69.504 | 48.86 | 1.8022
|  |  | 69.5808 | 48.33 | 1.5206
|  |  | 69.6576 | 47.455 | 1.3056
|  |  | 69.7344 | 45.7 | 1.2493
|  |  | 69.8112 | 44.135 | 1.1674
|  |  | 69.888 | 43.075 | 0.99328
|  |  | 69.9648 | 41.785 | 0.84992
|  |  | 70.0416 | 41.575 | 0.5888
|  |  | 70.1184 | 40.4 | 0.46592
|  |  | 70.1952 | 0 | 0
|  |  | 70.272 | 0 | 0
|  |  | 70.3488 | 0 | 0
|  |  | 70.4256 | 0 | 0
|  |  | 70.5024 | 0 | 0
|  |  | 70.5792 | 0 | 0
|  |  | 70.656 | 0 | 0
|  |  | 70.7328 | 0 | 0
|  |  | 70.8096 | 0 | 0
|  |  | 70.8864 | 0 | 0
|  |  | 70.9632 | 0 | 0
|  |  | 71.04 | 0 | 0
|  |  | 71.1168 | 0 | 0
|  |  | 71.1936 | 0 | 0
|  |  | 71.2704 | 0 | 0
|  |  | 71.3472 | 0 | 0
|  |  | 71.424 | 0 | 0
|  |  | 71.5008 | 0 | 0
|  |  | 71.5776 | 0 | 0
|  |  | 71.6544 | 0 | 0
|  |  | 71.7312 | 0 | 0
|  |  | 71.808 | 0 | 0
|  |  | 71.8848 | 0 | 0
|  |  | 71.9616 | 0 | 0
|  |  | 72.0384 | 0 | 0
|  |  | 72.1152 | 0 | 0
|  |  | 72.192 | 0 | 0
|  |  | 72.2688 | 0 | 0
|  |  | 72.3456 | 0 | 0
|  |  | 72.4224 | 0 | 0
|  |  | 72.4992 | 0 | 0
|  |  | 72.576 | 0 | 0
|  |  | 72.6528 | 0 | 0
|  |  | 72.7296 | 0 | 0
|  |  | 72.8064 | 0 | 0
|  |  | 72.8832 | 0 | 0
|  |  | 72.96 | 0 | 0
|  |  | 73.0368 | 0 | 0
|  |  | 73.1136 | 0 | 0
|  |  | 73.1904 | 0 | 0
|  |  | 73.2672 | 0 | 0
|  |  | 73.344 | 0 | 0
|  |  | 73.4208 | 0 | 0
|  |  | 73.4976 | 0 | 0
|  |  | 73.5744 | 0 | 0
|  |  | 73.6512 | 0 | 0
|  |  | 73.728 | 0 | 0
|  |  | 73.8048 | 0 | 0
|  |  | 73.8816 | 0 | 0
|  |  | 73.9584 | 0 | 0
|  |  | 74.0352 | 0 | 0
|  |  | 74.112 | 0 | 0
|  |  | 74.1888 | 0 | 0
|  |  | 74.2656 | 0 | 0
|  |  | 74.3424 | 0 | 0
|  |  | 74.4192 | 0 | 0
|  |  | 74.496 | 0 | 0
|  |  | 74.5728 | 0 | 0
|  |  | 74.6496 | 0 | 0
|  |  | 74.7264 | 0 | 0
|  |  | 74.8032 | 0 | 0
|  |  | 74.88 | 0 | 0
|  |  | 74.9568 | 0 | 0
|  |  | 75.0336 | 0 | 0
|  |  | 75.1104 | 0 | 0
|  |  | 75.1872 | 0 | 0
|  |  | 75.264 | 0 | 0
|  |  | 75.3408 | 0 | 0
|  |  | 75.4176 | 0 | 0
|  |  | 75.4944 | 0 | 0
|  |  | 75.5712 | 0 | 0
|  |  | 75.648 | 0 | 0
|  |  | 75.7248 | 0 | 0
|  |  | 75.8016 | 0 | 0
|  |  | 75.8784 | 0 | 0
|  |  | 75.9552 | 0 | 0
|  |  | 76.032 | 0 | 0
|  |  | 76.1088 | 0 | 0
|  |  | 76.1856 | 0 | 0
|  |  | 76.2624 | 0 | 0
|  |  | 76.3392 | 0 | 0
|  |  | 76.416 | 0 | 0
|  |  | 76.4928 | 0 | 0
|  |  | 76.5696 | 0 | 0
|  |  | 76.6464 | 0 | 0
|  |  | 76.7232 | 0 | 0
|  |  | 76.8 | 0 | 0
|  |  | 76.8768 | 0 | 0
|  |  | 76.9536 | 0 | 0
|  |  | 77.0304 | 0 | 0
|  |  | 77.1072 | 0 | 0
|  |  | 77.184 | 0 | 0
|  |  | 77.2608 | 0 | 0
|  |  | 77.3376 | 0 | 0
|  |  | 77.4144 | 0 | 0
|  |  | 77.4912 | 0 | 0
|  |  | 77.568 | 0 | 0
|  |  | 77.6448 | 0 | 0
|  |  | 77.7216 | 0 | 0
|  |  | 77.7984 | 0 | 0
|  |  | 77.8752 | 0 | 0
|  |  | 77.952 | 0 | 0
|  |  | 78.0288 | 0 | 0
|  |  | 78.1056 | 0 | 0
|  |  | 78.1824 | 0 | 0
|  |  | 78.2592 | 0 | 0
|  |  | 78.336 | 0 | 0
|  |  | 78.4128 | 0 | 0
|  |  | 78.4896 | 0 | 0
|  |  | 78.5664 | 0 | 0
|  |  | 78.6432 | 0 | 0
|  |  | 78.72 | 0 | 0
|  |  | 78.7968 | 0 | 0
|  |  | 78.8736 | 0 | 0
|  |  | 78.9504 | 0 | 0
|  |  | 79.0272 | 0 | 0
|  |  | 79.104 | 0 | 0
|  |  | 79.1808 | 0 | 0
|  |  | 79.2576 | 0 | 0
|  |  | 79.3344 | 0 | 0
|  |  | 79.4112 | 0 | 0
|  |  | 79.488 | 0 | 0
|  |  | 79.5648 | 0 | 0
|  |  | 79.6416 | 0 | 0
|  |  | 79.7184 | 0 | 0
|  |  | 79.7952 | 0 | 0
|  |  | 79.872 | 0 | 0
|  |  | 79.9488 | 0 | 0
|  |  | 80.0256 | 0 | 0
|  |  | 80.1024 | 0 | 0
|  |  | 80.1792 | 0 | 0
|  |  | 80.256 | 0 | 0
|  |  | 80.3328 | 0 | 0
|  |  | 80.4096 | 0 | 0
|  |  | 80.4864 | 0 | 0
|  |  | 80.5632 | 0 | 0
|  |  | 80.64 | 0 | 0
|  |  | 80.7168 | 0 | 0
|  |  | 80.7936 | 0 | 0
|  |  | 80.8704 | 0 | 0
|  |  | 80.9472 | 0 | 0
|  |  | 81.024 | 0 | 0
|  |  | 81.1008 | 0 | 0
|  |  | 81.1776 | 0 | 0
|  |  | 81.2544 | 0 | 0
|  |  | 81.3312 | 0 | 0
|  |  | 81.408 | 0 | 0
|  |  | 81.4848 | 0 | 0
|  |  | 81.5616 | 0 | 0
|  |  | 81.6384 | 0 | 0
|  |  | 81.7152 | 0 | 0
|  |  | 81.792 | 0 | 0
|  |  | 81.8688 | 0 | 0
|  |  | 81.9456 | 0 | 0
|  |  | 82.0224 | 0 | 0
|  |  | 82.0992 | 0 | 0
|  |  | 82.176 | 0 | 0
|  |  | 82.2528 | 0 | 0
|  |  | 82.3296 | 0 | 0
|  |  | 82.4064 | 0 | 0
|  |  | 82.4832 | 0 | 0
|  |  | 82.56 | 0 | 0
|  |  | 82.6368 | 0 | 0
|  |  | 82.7136 | 38.565 | 0.88064
|  |  | 82.7904 | 41.32 | 0.99328
|  |  | 82.8672 | 43.93 | 1.0752
|  |  | 82.944 | 46.075 | 1.2134
|  |  | 83.0208 | 48.04 | 1.367
|  |  | 83.0976 | 49.92 | 1.4694
|  |  | 83.1744 | 51.77 | 1.577
|  |  | 83.2512 | 52.79 | 1.792
|  |  | 83.328 | 54.225 | 1.9098
|  |  | 83.4048 | 55.985 | 1.9149
|  |  | 83.4816 | 56.645 | 2.0838
|  |  | 83.5584 | 56.99 | 2.2784
|  |  | 83.6352 | 57.625 | 2.3347
|  |  | 83.712 | 58.315 | 2.3296
|  |  | 83.7888 | 58.795 | 2.3603
|  |  | 83.8656 | 59.23 | 2.3706
|  |  | 83.9424 | 59.69 | 2.3501
|  |  | 84.0192 | 59.815 | 2.4115
|  |  | 84.096 | 60.34 | 2.4166
|  |  | 84.1728 | 60.725 | 2.4525
|  |  | 84.2496 | 60.275 | 2.6368
|  |  | 84.3264 | 60.385 | 2.6931
|  |  | 84.4032 | 60.995 | 2.6317
|  |  | 84.48 | 61.65 | 2.5549
|  |  | 84.5568 | 61.72 | 2.5907
|  |  | 84.6336 | 61.665 | 2.6163
|  |  | 84.7104 | 61.755 | 2.5958
|  |  | 84.7872 | 61.365 | 2.6778
|  |  | 84.864 | 61.01 | 2.7648
|  |  | 84.9408 | 60.765 | 2.8006
|  |  | 85.0176 | 61.155 | 2.7034
|  |  | 85.0944 | 61.355 | 2.6522
|  |  | 85.1712 | 60.69 | 2.7443
|  |  | 85.248 | 60.08 | 2.8211
|  |  | 85.3248 | 59.935 | 2.7955
|  |  | 85.4016 | 60.29 | 2.6317
|  |  | 85.4784 | 60.31 | 2.5549
|  |  | 85.5552 | 59.93 | 2.5395
|  |  | 85.632 | 59.1 | 2.5754
|  |  | 85.7088 | 58.32 | 2.5344
|  |  | 85.7856 | 57.45 | 2.4166
|  |  | 85.8624 | 56.19 | 2.2733
|  |  | 85.9392 | 55.095 | 2.0634
|  |  | 86.016 | 53.97 | 1.8637
|  |  | 86.0928 | 52.505 | 1.7715
|  |  | 86.1696 | 50.6 | 1.7664
|  |  | 86.2464 | 48.84 | 1.7408
|  |  | 86.3232 | 47.155 | 1.6896
|  |  | 86.4 | 46.43 | 1.4336
|  |  | 86.4768 | 45.405 | 1.239
|  |  | 86.5536 | 43.97 | 1.1264
|  |  | 86.6304 | 42.61 | 0.97792
|  |  | 86.7072 | 42.34 | 0.69632
|  |  | 86.784 | 40.635 | 0.59904
|  |  | 86.8608 | 0 | 0
|  |  | 86.9376 | 0 | 0
|  |  | 87.0144 | 0 | 0
|  |  | 87.0912 | 0 | 0
|  |  | 87.168 | 0 | 0
|  |  | 87.2448 | 0 | 0
|  |  | 87.3216 | 0 | 0
|  |  | 87.3984 | 0 | 0
|  |  | 87.4752 | 0 | 0
|  |  | 87.552 | 0 | 0
|  |  | 87.6288 | 0 | 0
|  |  | 87.7056 | 0 | 0
|  |  | 87.7824 | 0 | 0
|  |  | 87.8592 | 0 | 0
|  |  | 87.936 | 0 | 0
|  |  | 88.0128 | 0 | 0
|  |  | 88.0896 | 0 | 0
|  |  | 88.1664 | 0 | 0
|  |  | 88.2432 | 0 | 0
|  |  | 88.32 | 0 | 0
|  |  | 88.3968 | 0 | 0
|  |  | 88.4736 | 0 | 0
|  |  | 88.5504 | 0 | 0
|  |  | 88.6272 | 0 | 0
|  |  | 88.704 | 0 | 0
|  |  | 88.7808 | 0 | 0
|  |  | 88.8576 | 0 | 0
|  |  | 88.9344 | 0 | 0
|  |  | 89.0112 | 0 | 0
|  |  | 89.088 | 0 | 0
|  |  | 89.1648 | 0 | 0
|  |  | 89.2416 | 0 | 0
|  |  | 89.3184 | 0 | 0
|  |  | 89.3952 | 0 | 0
|  |  | 89.472 | 0 | 0
|  |  | 89.5488 | 0 | 0
|  |  | 89.6256 | 0 | 0
|  |  | 89.7024 | 0 | 0
|  |  | 89.7792 | 0 | 0
|  |  | 89.856 | 0 | 0
|  |  | 89.9328 | 0 | 0
|  |  | 90.0096 | 0 | 0
|  |  | 90.0864 | 0 | 0
|  |  | 90.1632 | 0 | 0
|  |  | 90.24 | 0 | 0
|  |  | 90.3168 | 0 | 0
|  |  | 90.3936 | 0 | 0
|  |  | 90.4704 | 0 | 0
|  |  | 90.5472 | 0 | 0
|  |  | 90.624 | 0 | 0
|  |  | 90.7008 | 0 | 0
|  |  | 90.7776 | 0 | 0
|  |  | 90.8544 | 0 | 0
|  |  | 90.9312 | 0 | 0
|  |  | 91.008 | 0 | 0
|  |  | 91.0848 | 0 | 0
|  |  | 91.1616 | 0 | 0
|  |  | 91.2384 | 0 | 0
|  |  | 91.3152 | 0 | 0
|  |  | 91.392 | 0 | 0
|  |  | 91.4688 | 0 | 0
|  |  | 91.5456 | 0 | 0
|  |  | 91.6224 | 0 | 0
|  |  | 91.6992 | 0 | 0
|  |  | 91.776 | 0 | 0
|  |  | 91.8528 | 0 | 0
|  |  | 91.9296 | 0 | 0
|  |  | 92.0064 | 0 | 0
|  |  | 92.0832 | 0 | 0
|  |  | 92.16 | 0 | 0
|  |  | 92.2368 | 0 | 0
|  |  | 92.3136 | 0 | 0
|  |  | 92.3904 | 0 | 0
|  |  | 92.4672 | 0 | 0
|  |  | 92.544 | 0 | 0
|  |  | 92.6208 | 0 | 0
|  |  | 92.6976 | 0 | 0
|  |  | 92.7744 | 0 | 0
|  |  | 92.8512 | 0 | 0
|  |  | 92.928 | 0 | 0
|  |  | 93.0048 | 0 | 0
|  |  | 93.0816 | 0 | 0
|  |  | 93.1584 | 0 | 0
|  |  | 93.2352 | 0 | 0
|  |  | 93.312 | 0 | 0
|  |  | 93.3888 | 0 | 0
|  |  | 93.4656 | 0 | 0
|  |  | 93.5424 | 0 | 0
|  |  | 93.6192 | 0 | 0
|  |  | 93.696 | 0 | 0
|  |  | 93.7728 | 0 | 0
|  |  | 93.8496 | 0 | 0
|  |  | 93.9264 | 0 | 0
|  |  | 94.0032 | 0 | 0
|  |  | 94.08 | 0 | 0
|  |  | 94.1568 | 0 | 0
|  |  | 94.2336 | 0 | 0
|  |  | 94.3104 | 0 | 0
|  |  | 94.3872 | 0 | 0
|  |  | 94.464 | 0 | 0
|  |  | 94.5408 | 0 | 0
|  |  | 94.6176 | 0 | 0
|  |  | 94.6944 | 0 | 0
|  |  | 94.7712 | 0 | 0
|  |  | 94.848 | 0 | 0
|  |  | 94.9248 | 0 | 0
|  |  | 95.0016 | 0 | 0
|  |  | 95.0784 | 0 | 0
|  |  | 95.1552 | 0 | 0
|  |  | 95.232 | 0 | 0
|  |  | 95.3088 | 0 | 0
|  |  | 95.3856 | 0 | 0
|  |  | 95.4624 | 0 | 0
|  |  | 95.5392 | 0 | 0
|  |  | 95.616 | 0 | 0
|  |  | 95.6928 | 0 | 0
|  |  | 95.7696 | 0 | 0
|  |  | 95.8464 | 0 | 0
|  |  | 95.9232 | 0 | 0
|  |  | 96 | 0 | 0
|  |  | 96.0768 | 0 | 0
|  |  | 96.1536 | 0 | 0
|  |  | 96.2304 | 0 | 0
|  |  | 96.3072 | 0 | 0
|  |  | 96.384 | 0 | 0
|  |  | 96.4608 | 0 | 0
|  |  | 96.5376 | 0 | 0
|  |  | 96.6144 | 0 | 0
|  |  | 96.6912 | 0 | 0
|  |  | 96.768 | 0 | 0
|  |  | 96.8448 | 0 | 0
|  |  | 96.9216 | 0 | 0
|  |  | 96.9984 | 0 | 0
|  |  | 97.0752 | 0 | 0
|  |  | 97.152 | 0 | 0
|  |  | 97.2288 | 0 | 0
|  |  | 97.3056 | 0 | 0
|  |  | 97.3824 | 0 | 0
|  |  | 97.4592 | 0 | 0
|  |  | 97.536 | 0 | 0
|  |  | 97.6128 | 0 | 0
|  |  | 97.6896 | 0 | 0
|  |  | 97.7664 | 0 | 0
|  |  | 97.8432 | 0 | 0
|  |  | 97.92 | 0 | 0
|  |  | 97.9968 | 0 | 0
|  |  | 98.0736 | 0 | 0
|  |  | 98.1504 | 0 | 0
|  |  | 98.2272 | 0 | 0
|  |  | 98.304 | 0 | 0
|  |  | 98.3808 | 0 | 0
|  |  | 98.4576 | 0 | 0
|  |  | 98.5344 | 0 | 0
|  |  | 98.6112 | 0 | 0
|  |  | 98.688 | 0 | 0
|  |  | 98.7648 | 0 | 0
|  |  | 98.8416 | 0 | 0
|  |  | 98.9184 | 0 | 0
|  |  | 98.9952 | 0 | 0
|  |  | 99.072 | 0 | 0
|  |  | 99.1488 | 0 | 0
|  |  | 99.2256 | 0 | 0
|  |  | 99.3024 | 0 | 0
|  |  | 99.3792 | 40.25 | 0.7936
|  |  | 99.456 | 42.615 | 0.93184
|  |  | 99.5328 | 45.04 | 1.0547
|  |  | 99.6096 | 47.725 | 1.0906
|  |  | 99.6864 | 49.67 | 1.239
|  |  | 99.7632 | 50.865 | 1.495
|  |  | 99.84 | 52.715 | 1.5974
|  |  | 99.9168 | 54.465 | 1.6794
|  |  | 99.9936 | 55.665 | 1.8381
|  |  | 100.0704 | 57.09 | 1.9098
|  |  | 100.1472 | 58.06 | 2.0122
|  |  | 100.224 | 58.845 | 2.1043
|  |  | 100.3008 | 59.56 | 2.1555
|  |  | 100.3776 | 60.055 | 2.2016
|  |  | 100.4544 | 60.425 | 2.2426
|  |  | 100.5312 | 60.205 | 2.3552
|  |  | 100.608 | 60.14 | 2.3808
|  |  | 100.6848 | 60.815 | 2.3398
|  |  | 100.7616 | 61.175 | 2.4115
|  |  | 100.8384 | 61.46 | 2.4934
|  |  | 100.9152 | 61.69 | 2.5805
|  |  | 100.992 | 62.035 | 2.5805
|  |  | 101.0688 | 60.755 | 2.9389
|  |  | 101.1456 | 62.175 | 2.5702
|  |  | 101.2224 | 61.4 | 2.8365
|  |  | 101.2992 | 62.595 | 2.5754
|  |  | 101.376 | 61.505 | 2.8928
|  |  | 101.4528 | 62.78 | 2.601
|  |  | 101.5296 | 60.68 | 3.1744
|  |  | 101.6064 | 61.965 | 2.7699
|  |  | 101.6832 | 61.945 | 2.7597
|  |  | 101.76 | 61.87 | 2.7494
|  |  | 101.8368 | 61.41 | 2.7955
|  |  | 101.9136 | 61.685 | 2.6931
|  |  | 101.9904 | 61.725 | 2.6419
|  |  | 102.0672 | 61.385 | 2.6419
|  |  | 102.144 | 61.01 | 2.6368
|  |  | 102.2208 | 60.315 | 2.6778
|  |  | 102.2976 | 59.775 | 2.6522
|  |  | 102.3744 | 59.25 | 2.5651
|  |  | 102.4512 | 58.725 | 2.3859
|  |  | 102.528 | 57.695 | 2.217
|  |  | 102.6048 | 55.34 | 2.2579
|  |  | 102.6816 | 53.695 | 2.1555
|  |  | 102.7584 | 52.505 | 1.9968
|  |  | 102.8352 | 51.325 | 1.8534
|  |  | 102.912 | 49.875 | 1.7715
|  |  | 102.9888 | 48.265 | 1.6896
|  |  | 103.0656 | 47.035 | 1.5258
|  |  | 103.1424 | 46.06 | 1.3261
|  |  | 103.2192 | 45.315 | 1.0957
|  |  | 103.296 | 43.615 | 0.98816
|  |  | 103.3728 | 41.64 | 0.90112
|  |  | 103.4496 | 39.645 | 0.8192
|  |  | 103.5264 | 0 | 0
|  |  | 103.6032 | 0 | 0
|  |  | 103.68 | 0 | 0
|  |  | 103.7568 | 0 | 0
|  |  | 103.8336 | 0 | 0
|  |  | 103.9104 | 0 | 0
|  |  | 103.9872 | 0 | 0
|  |  | 104.064 | 0 | 0
|  |  | 104.1408 | 0 | 0
|  |  | 104.2176 | 0 | 0
|  |  | 104.2944 | 0 | 0
|  |  | 104.3712 | 0 | 0
|  |  | 104.448 | 0 | 0
|  |  | 104.5248 | 0 | 0
|  |  | 104.6016 | 0 | 0
|  |  | 104.6784 | 0 | 0
|  |  | 104.7552 | 0 | 0
|  |  | 104.832 | 0 | 0
|  |  | 104.9088 | 0 | 0
|  |  | 104.9856 | 0 | 0
|  |  | 105.0624 | 0 | 0
|  |  | 105.1392 | 0 | 0
|  |  | 105.216 | 0 | 0
|  |  | 105.2928 | 0 | 0
|  |  | 105.3696 | 0 | 0
|  |  | 105.4464 | 0 | 0
|  |  | 105.5232 | 0 | 0
|  |  | 105.6 | 0 | 0
|  |  | 105.6768 | 0 | 0
|  |  | 105.7536 | 0 | 0
|  |  | 105.8304 | 0 | 0
|  |  | 105.9072 | 0 | 0
|  |  | 105.984 | 0 | 0
|  |  | 106.0608 | 0 | 0
|  |  | 106.1376 | 0 | 0
|  |  | 106.2144 | 0 | 0
|  |  | 106.2912 | 0 | 0
|  |  | 106.368 | 0 | 0
|  |  | 106.4448 | 0 | 0
|  |  | 106.5216 | 0 | 0
|  |  | 106.5984 | 0 | 0
|  |  | 106.6752 | 0 | 0
|  |  | 106.752 | 0 | 0
|  |  | 106.8288 | 0 | 0
|  |  | 106.9056 | 0 | 0
|  |  | 106.9824 | 0 | 0
|  |  | 107.0592 | 0 | 0
|  |  | 107.136 | 0 | 0
|  |  | 107.2128 | 0 | 0
|  |  | 107.2896 | 0 | 0
|  |  | 107.3664 | 0 | 0
|  |  | 107.4432 | 0 | 0
|  |  | 107.52 | 0 | 0
|  |  | 107.5968 | 0 | 0
|  |  | 107.6736 | 0 | 0
|  |  | 107.7504 | 0 | 0
|  |  | 107.8272 | 0 | 0
|  |  | 107.904 | 0 | 0
|  |  | 107.9808 | 0 | 0
|  |  | 108.0576 | 0 | 0
|  |  | 108.1344 | 0 | 0
|  |  | 108.2112 | 0 | 0
|  |  | 108.288 | 0 | 0
|  |  | 108.3648 | 0 | 0
|  |  | 108.4416 | 0 | 0
|  |  | 108.5184 | 0 | 0
|  |  | 108.5952 | 0 | 0
|  |  | 108.672 | 0 | 0
|  |  | 108.7488 | 0 | 0
|  |  | 108.8256 | 0 | 0
|  |  | 108.9024 | 0 | 0
|  |  | 108.9792 | 0 | 0
|  |  | 109.056 | 0 | 0
|  |  | 109.1328 | 0 | 0
|  |  | 109.2096 | 0 | 0
|  |  | 109.2864 | 0 | 0
|  |  | 109.3632 | 0 | 0
|  |  | 109.44 | 0 | 0
|  |  | 109.5168 | 0 | 0
|  |  | 109.5936 | 0 | 0
|  |  | 109.6704 | 0 | 0
|  |  | 109.7472 | 0 | 0
|  |  | 109.824 | 0 | 0
|  |  | 109.9008 | 0 | 0
|  |  | 109.9776 | 0 | 0
|  |  | 110.0544 | 0 | 0
|  |  | 110.1312 | 0 | 0
|  |  | 110.208 | 0 | 0
|  |  | 110.2848 | 0 | 0
|  |  | 110.3616 | 0 | 0
|  |  | 110.4384 | 0 | 0
|  |  | 110.5152 | 0 | 0
|  |  | 110.592 | 0 | 0
|  |  | 110.6688 | 0 | 0
|  |  | 110.7456 | 0 | 0
|  |  | 110.8224 | 0 | 0
|  |  | 110.8992 | 0 | 0
|  |  | 110.976 | 0 | 0
|  |  | 111.0528 | 0 | 0
|  |  | 111.1296 | 0 | 0
|  |  | 111.2064 | 0 | 0
|  |  | 111.2832 | 0 | 0
|  |  | 111.36 | 0 | 0
|  |  | 111.4368 | 0 | 0
|  |  | 111.5136 | 0 | 0
|  |  | 111.5904 | 0 | 0
|  |  | 111.6672 | 0 | 0
|  |  | 111.744 | 0 | 0
|  |  | 111.8208 | 0 | 0
|  |  | 111.8976 | 0 | 0
|  |  | 111.9744 | 0 | 0
|  |  | 112.0512 | 0 | 0
|  |  | 112.128 | 0 | 0
|  |  | 112.2048 | 0 | 0
|  |  | 112.2816 | 0 | 0
|  |  | 112.3584 | 0 | 0
|  |  | 112.4352 | 0 | 0
|  |  | 112.512 | 0 | 0
|  |  | 112.5888 | 0 | 0
|  |  | 112.6656 | 0 | 0
|  |  | 112.7424 | 0 | 0
|  |  | 112.8192 | 0 | 0
|  |  | 112.896 | 0 | 0
|  |  | 112.9728 | 0 | 0
|  |  | 113.0496 | 0 | 0
|  |  | 113.1264 | 0 | 0
|  |  | 113.2032 | 0 | 0
|  |  | 113.28 | 0 | 0
|  |  | 113.3568 | 0 | 0
|  |  | 113.4336 | 0 | 0
|  |  | 113.5104 | 0 | 0
|  |  | 113.5872 | 0 | 0
|  |  | 113.664 | 0 | 0
|  |  | 113.7408 | 0 | 0
|  |  | 113.8176 | 0 | 0
|  |  | 113.8944 | 0 | 0
|  |  | 113.9712 | 0 | 0
|  |  | 114.048 | 0 | 0
|  |  | 114.1248 | 0 | 0
|  |  | 114.2016 | 0 | 0
|  |  | 114.2784 | 0 | 0
|  |  | 114.3552 | 0 | 0
|  |  | 114.432 | 0 | 0
|  |  | 114.5088 | 0 | 0
|  |  | 114.5856 | 0 | 0
|  |  | 114.6624 | 0 | 0
|  |  | 114.7392 | 0 | 0
|  |  | 114.816 | 0 | 0
|  |  | 114.8928 | 0 | 0
|  |  | 114.9696 | 0 | 0
|  |  | 115.0464 | 0 | 0
|  |  | 115.1232 | 0 | 0
|  |  | 115.2 | 0 | 0
|  |  | 115.2768 | 0 | 0
|  |  | 115.3536 | 0 | 0
|  |  | 115.4304 | 0 | 0
|  |  | 115.5072 | 0 | 0
|  |  | 115.584 | 0 | 0
|  |  | 115.6608 | 0 | 0
|  |  | 115.7376 | 0 | 0
|  |  | 115.8144 | 0 | 0
|  |  | 115.8912 | 0 | 0
|  |  | 115.968 | 0 | 0
|  |  | 116.0448 | 40.73 | 0.75776
|  |  | 116.1216 | 43.02 | 0.9216
|  |  | 116.1984 | 45.215 | 1.0854
|  |  | 116.2752 | 46.925 | 1.2954
|  |  | 116.352 | 49.185 | 1.367
|  |  | 116.4288 | 50.255 | 1.6435
|  |  | 116.5056 | 52.05 | 1.7766
|  |  | 116.5824 | 54.28 | 1.7869
|  |  | 116.6592 | 55.795 | 1.8944
|  |  | 116.736 | 56.575 | 2.089
|  |  | 116.8128 | 57.515 | 2.2118
|  |  | 116.8896 | 58.465 | 2.263
|  |  | 116.9664 | 59.5 | 2.2579
|  |  | 117.0432 | 60.45 | 2.2426
|  |  | 117.12 | 60.54 | 2.345
|  |  | 117.1968 | 60.47 | 2.4474
|  |  | 117.2736 | 60.84 | 2.4474
|  |  | 117.3504 | 61.07 | 2.4832
|  |  | 117.4272 | 61.06 | 2.5651
|  |  | 117.504 | 61.625 | 2.56
|  |  | 117.5808 | 61.815 | 2.6061
|  |  | 117.6576 | 61.775 | 2.6829
|  |  | 117.7344 | 62.26 | 2.647
|  |  | 117.8112 | 61.7 | 2.8877
|  |  | 117.888 | 62.83 | 2.647
|  |  | 117.9648 | 62.35 | 2.8109
|  |  | 118.0416 | 62.055 | 2.8979
|  |  | 118.1184 | 61.72 | 3.0003
|  |  | 118.1952 | 62.77 | 2.7085
|  |  | 118.272 | 61.635 | 3.0259
|  |  | 118.3488 | 62.7 | 2.6982
|  |  | 118.4256 | 61.505 | 2.9952
|  |  | 118.5024 | 62.155 | 2.7494
|  |  | 118.5792 | 61.44 | 2.9235
|  |  | 118.656 | 62.24 | 2.6624
|  |  | 118.7328 | 60.81 | 2.9901
|  |  | 118.8096 | 61.765 | 2.6112
|  |  | 118.8864 | 61.155 | 2.6317
|  |  | 118.9632 | 60.595 | 2.6061
|  |  | 119.04 | 59.725 | 2.5754
|  |  | 119.1168 | 58.98 | 2.4422
|  |  | 119.1936 | 57.945 | 2.2989
|  |  | 119.2704 | 56.765 | 2.1197
|  |  | 119.3472 | 55.28 | 1.9917
|  |  | 119.424 | 53.605 | 1.92
|  |  | 119.5008 | 52.35 | 1.7818
|  |  | 119.5776 | 51.06 | 1.664
|  |  | 119.6544 | 50.075 | 1.4797
|  |  | 119.7312 | 48.715 | 1.3619
|  |  | 119.808 | 46.96 | 1.2954
|  |  | 119.8848 | 45.54 | 1.1725
|  |  | 119.9616 | 44.135 | 1.0138
|  |  | 120.0384 | 43.485 | 0.768
|  |  | 120.1152 | 41.135 | 0.73728
|  |  | 120.192 | 0 | 0
|  |  | 120.2688 | 0 | 0
|  |  | 120.3456 | 0 | 0
|  |  | 120.4224 | 0 | 0
|  |  | 120.4992 | 0 | 0
|  |  | 120.576 | 0 | 0
|  |  | 120.6528 | 0 | 0
|  |  | 120.7296 | 0 | 0
|  |  | 120.8064 | 0 | 0
|  |  | 120.8832 | 0 | 0
|  |  | 120.96 | 0 | 0
|  |  | 121.0368 | 0 | 0
|  |  | 121.1136 | 0 | 0
|  |  | 121.1904 | 0 | 0
|  |  | 121.2672 | 0 | 0
|  |  | 121.344 | 0 | 0
|  |  | 121.4208 | 0 | 0
|  |  | 121.4976 | 0 | 0
|  |  | 121.5744 | 0 | 0
|  |  | 121.6512 | 0 | 0
|  |  | 121.728 | 0 | 0
|  |  | 121.8048 | 0 | 0
|  |  | 121.8816 | 0 | 0
|  |  | 121.9584 | 0 | 0
|  |  | 122.0352 | 0 | 0
|  |  | 122.112 | 0 | 0
|  |  | 122.1888 | 0 | 0
|  |  | 122.2656 | 0 | 0
|  |  | 122.3424 | 0 | 0
|  |  | 122.4192 | 0 | 0
|  |  | 122.496 | 0 | 0
|  |  | 122.5728 | 0 | 0
|  |  | 122.6496 | 0 | 0
|  |  | 122.7264 | 0 | 0
|  |  | 122.8032 | 0 | 0
|  |  | 122.88 | 0 | 0
|  |  | 122.9568 | 0 | 0
|  |  | 123.0336 | 0 | 0
|  |  | 123.1104 | 0 | 0
|  |  | 123.1872 | 0 | 0
|  |  | 123.264 | 0 | 0
|  |  | 123.3408 | 0 | 0
|  |  | 123.4176 | 0 | 0
|  |  | 123.4944 | 0 | 0
|  |  | 123.5712 | 0 | 0
|  |  | 123.648 | 0 | 0
|  |  | 123.7248 | 0 | 0
|  |  | 123.8016 | 0 | 0
|  |  | 123.8784 | 0 | 0
|  |  | 123.9552 | 0 | 0
|  |  | 124.032 | 0 | 0
|  |  | 124.1088 | 0 | 0
|  |  | 124.1856 | 0 | 0
|  |  | 124.2624 | 0 | 0
|  |  | 124.3392 | 0 | 0
|  |  | 124.416 | 0 | 0
|  |  | 124.4928 | 0 | 0
|  |  | 124.5696 | 0 | 0
|  |  | 124.6464 | 0 | 0
|  |  | 124.7232 | 0 | 0
|  |  | 124.8 | 0 | 0
|  |  | 124.8768 | 0 | 0
|  |  | 124.9536 | 0 | 0
|  |  | 125.0304 | 0 | 0
|  |  | 125.1072 | 0 | 0
|  |  | 125.184 | 0 | 0
|  |  | 125.2608 | 0 | 0
|  |  | 125.3376 | 0 | 0
|  |  | 125.4144 | 0 | 0
|  |  | 125.4912 | 0 | 0
|  |  | 125.568 | 0 | 0
|  |  | 125.6448 | 0 | 0
|  |  | 125.7216 | 0 | 0
|  |  | 125.7984 | 0 | 0
|  |  | 125.8752 | 0 | 0
|  |  | 125.952 | 0 | 0
|  |  | 126.0288 | 0 | 0
|  |  | 126.1056 | 0 | 0
|  |  | 126.1824 | 0 | 0
|  |  | 126.2592 | 0 | 0
|  |  | 126.336 | 0 | 0
|  |  | 126.4128 | 0 | 0
|  |  | 126.4896 | 0 | 0
|  |  | 126.5664 | 0 | 0
|  |  | 126.6432 | 0 | 0
|  |  | 126.72 | 0 | 0
|  |  | 126.7968 | 0 | 0
|  |  | 126.8736 | 0 | 0
|  |  | 126.9504 | 0 | 0
|  |  | 127.0272 | 0 | 0
|  |  | 127.104 | 0 | 0
|  |  | 127.1808 | 0 | 0
|  |  | 127.2576 | 0 | 0
|  |  | 127.3344 | 0 | 0
|  |  | 127.4112 | 0 | 0
|  |  | 127.488 | 0 | 0
|  |  | 127.5648 | 0 | 0
|  |  | 127.6416 | 0 | 0
|  |  | 127.7184 | 0 | 0
|  |  | 127.7952 | 0 | 0
|  |  | 127.872 | 0 | 0
|  |  | 127.9488 | 0 | 0
|  |  | 128.0256 | 0 | 0
|  |  | 128.1024 | 0 | 0
|  |  | 128.1792 | 0 | 0
|  |  | 128.256 | 0 | 0
|  |  | 128.3328 | 0 | 0
|  |  | 128.4096 | 0 | 0
|  |  | 128.4864 | 0 | 0
|  |  | 128.5632 | 0 | 0
|  |  | 128.64 | 0 | 0
|  |  | 128.7168 | 0 | 0
|  |  | 128.7936 | 0 | 0
|  |  | 128.8704 | 0 | 0
|  |  | 128.9472 | 0 | 0
|  |  | 129.024 | 0 | 0
|  |  | 129.1008 | 0 | 0
|  |  | 129.1776 | 0 | 0
|  |  | 129.2544 | 0 | 0
|  |  | 129.3312 | 0 | 0
|  |  | 129.408 | 0 | 0
|  |  | 129.4848 | 0 | 0
|  |  | 129.5616 | 0 | 0
|  |  | 129.6384 | 0 | 0
|  |  | 129.7152 | 0 | 0
|  |  | 129.792 | 0 | 0
|  |  | 129.8688 | 0 | 0
|  |  | 129.9456 | 0 | 0
|  |  | 130.0224 | 0 | 0
|  |  | 130.0992 | 0 | 0
|  |  | 130.176 | 0 | 0
|  |  | 130.2528 | 0 | 0
|  |  | 130.3296 | 0 | 0
|  |  | 130.4064 | 0 | 0
|  |  | 130.4832 | 0 | 0
|  |  | 130.56 | 0 | 0
|  |  | 130.6368 | 0 | 0
|  |  | 130.7136 | 0 | 0
|  |  | 130.7904 | 0 | 0
|  |  | 130.8672 | 0 | 0
|  |  | 130.944 | 0 | 0
|  |  | 131.0208 | 0 | 0
|  |  | 131.0976 | 0 | 0
|  |  | 131.1744 | 0 | 0
|  |  | 131.2512 | 0 | 0
|  |  | 131.328 | 0 | 0
|  |  | 131.4048 | 0 | 0
|  |  | 131.4816 | 0 | 0
|  |  | 131.5584 | 0 | 0
|  |  | 131.6352 | 0 | 0
|  |  | 131.712 | 0 | 0
|  |  | 131.7888 | 0 | 0
|  |  | 131.8656 | 0 | 0
|  |  | 131.9424 | 0 | 0
|  |  | 132.0192 | 0 | 0
|  |  | 132.096 | 0 | 0
|  |  | 132.1728 | 0 | 0
|  |  | 132.2496 | 0 | 0
|  |  | 132.3264 | 0 | 0
|  |  | 132.4032 | 0 | 0
|  |  | 132.48 | 0 | 0
|  |  | 132.5568 | 0 | 0
|  |  | 132.6336 | 0 | 0
|  |  | 132.7104 | 40.59 | 0.82432
|  |  | 132.7872 | 43.305 | 0.91648
|  |  | 132.864 | 45.15 | 1.1162
|  |  | 132.9408 | 47.4 | 1.239
|  |  | 133.0176 | 49.59 | 1.367
|  |  | 133.0944 | 51.615 | 1.4746
|  |  | 133.1712 | 53.09 | 1.623
|  |  | 133.248 | 55.065 | 1.6794
|  |  | 133.3248 | 56.635 | 1.7766
|  |  | 133.4016 | 57.44 | 1.9712
|  |  | 133.4784 | 58.44 | 2.0838
|  |  | 133.5552 | 59.525 | 2.135
|  |  | 133.632 | 60.415 | 2.1709
|  |  | 133.7088 | 60.48 | 2.2989
|  |  | 133.7856 | 60.235 | 2.4627
|  |  | 133.8624 | 60.405 | 2.5293
|  |  | 133.9392 | 60.775 | 2.5293
|  |  | 134.016 | 61.76 | 2.4371
|  |  | 134.0928 | 62.395 | 2.4013
|  |  | 134.1696 | 61.205 | 2.7546
|  |  | 134.2464 | 62.225 | 2.5754
|  |  | 134.3232 | 61.57 | 2.8928
|  |  | 134.4 | 63.5 | 2.5037
|  |  | 134.4768 | 62.47 | 2.7955
|  |  | 134.5536 | 62.405 | 2.8416
|  |  | 134.6304 | 62.32 | 2.8928
|  |  | 134.7072 | 62.375 | 2.8928
|  |  | 134.784 | 62.32 | 2.9338
|  |  | 134.8608 | 62.105 | 2.9952
|  |  | 134.9376 | 61.88 | 3.0618
|  |  | 135.0144 | 63.045 | 2.6982
|  |  | 135.0912 | 61.84 | 2.9952
|  |  | 135.168 | 62.965 | 2.6778
|  |  | 135.2448 | 61.68 | 2.9645
|  |  | 135.3216 | 62.065 | 2.7648
|  |  | 135.3984 | 60.92 | 3.0208
|  |  | 135.4752 | 61.915 | 2.6368
|  |  | 135.552 | 61.485 | 2.6163
|  |  | 135.6288 | 61.315 | 2.5395
|  |  | 135.7056 | 60.94 | 2.4371
|  |  | 135.7824 | 59.8 | 2.3859
|  |  | 135.8592 | 58.165 | 2.3552
|  |  | 135.936 | 56.945 | 2.2016
|  |  | 136.0128 | 55.46 | 2.0736
|  |  | 136.0896 | 54.205 | 1.9046
|  |  | 136.1664 | 52.685 | 1.8176
|  |  | 136.2432 | 51.38 | 1.6947
|  |  | 136.32 | 49.915 | 1.623
|  |  | 136.3968 | 49.075 | 1.3978
|  |  | 136.4736 | 47.135 | 1.3312
|  |  | 136.5504 | 45.07 | 1.321
|  |  | 136.6272 | 44.1 | 1.1213
|  |  | 136.704 | 43.555 | 0.82432
|  |  | 136.7808 | 42.685 | 0.61952
|  |  | 136.8576 | 39.81 | 0.62976
|  |  | 136.9344 | 0 | 0
|  |  | 137.0112 | 0 | 0
|  |  | 137.088 | 0 | 0
|  |  | 137.1648 | 0 | 0
|  |  | 137.2416 | 0 | 0
|  |  | 137.3184 | 0 | 0
|  |  | 137.3952 | 0 | 0
|  |  | 137.472 | 0 | 0
|  |  | 137.5488 | 0 | 0
|  |  | 137.6256 | 0 | 0
|  |  | 137.7024 | 0 | 0
|  |  | 137.7792 | 0 | 0
|  |  | 137.856 | 0 | 0
|  |  | 137.9328 | 0 | 0
|  |  | 138.0096 | 0 | 0
|  |  | 138.0864 | 0 | 0
|  |  | 138.1632 | 0 | 0
|  |  | 138.24 | 0 | 0
|  |  | 138.3168 | 0 | 0
|  |  | 138.3936 | 0 | 0
|  |  | 138.4704 | 0 | 0
|  |  | 138.5472 | 0 | 0
|  |  | 138.624 | 0 | 0
|  |  | 138.7008 | 0 | 0
|  |  | 138.7776 | 0 | 0
|  |  | 138.8544 | 0 | 0
|  |  | 138.9312 | 0 | 0
|  |  | 139.008 | 0 | 0
|  |  | 139.0848 | 0 | 0
|  |  | 139.1616 | 0 | 0
|  |  | 139.2384 | 0 | 0
|  |  | 139.3152 | 0 | 0
|  |  | 139.392 | 0 | 0
|  |  | 139.4688 | 0 | 0
|  |  | 139.5456 | 0 | 0
|  |  | 139.6224 | 0 | 0
|  |  | 139.6992 | 0 | 0
|  |  | 139.776 | 0 | 0
|  |  | 139.8528 | 0 | 0
|  |  | 139.9296 | 0 | 0
|  |  | 140.0064 | 0 | 0
|  |  | 140.0832 | 0 | 0
|  |  | 140.16 | 0 | 0
|  |  | 140.2368 | 0 | 0
|  |  | 140.3136 | 0 | 0
|  |  | 140.3904 | 0 | 0
|  |  | 140.4672 | 0 | 0
|  |  | 140.544 | 0 | 0
|  |  | 140.6208 | 0 | 0
|  |  | 140.6976 | 0 | 0
|  |  | 140.7744 | 0 | 0
|  |  | 140.8512 | 0 | 0
|  |  | 140.928 | 0 | 0
|  |  | 141.0048 | 0 | 0
|  |  | 141.0816 | 0 | 0
|  |  | 141.1584 | 0 | 0
|  |  | 141.2352 | 0 | 0
|  |  | 141.312 | 0 | 0
|  |  | 141.3888 | 0 | 0
|  |  | 141.4656 | 0 | 0
|  |  | 141.5424 | 0 | 0
|  |  | 141.6192 | 0 | 0
|  |  | 141.696 | 0 | 0
|  |  | 141.7728 | 0 | 0
|  |  | 141.8496 | 0 | 0
|  |  | 141.9264 | 0 | 0
|  |  | 142.0032 | 0 | 0
|  |  | 142.08 | 0 | 0
|  |  | 142.1568 | 0 | 0
|  |  | 142.2336 | 0 | 0
|  |  | 142.3104 | 0 | 0
|  |  | 142.3872 | 0 | 0
|  |  | 142.464 | 0 | 0
|  |  | 142.5408 | 0 | 0
|  |  | 142.6176 | 0 | 0
|  |  | 142.6944 | 0 | 0
|  |  | 142.7712 | 0 | 0
|  |  | 142.848 | 0 | 0
|  |  | 142.9248 | 0 | 0
|  |  | 143.0016 | 0 | 0
|  |  | 143.0784 | 0 | 0
|  |  | 143.1552 | 0 | 0
|  |  | 143.232 | 0 | 0
|  |  | 143.3088 | 0 | 0
|  |  | 143.3856 | 0 | 0
|  |  | 143.4624 | 0 | 0
|  |  | 143.5392 | 0 | 0
|  |  | 143.616 | 0 | 0
|  |  | 143.6928 | 0 | 0
|  |  | 143.7696 | 0 | 0
|  |  | 143.8464 | 0 | 0
|  |  | 143.9232 | 0 | 0
|  |  | 144 | 0 | 0
|  |  | 144.0768 | 0 | 0
|  |  | 144.1536 | 0 | 0
|  |  | 144.2304 | 0 | 0
|  |  | 144.3072 | 0 | 0
|  |  | 144.384 | 0 | 0
|  |  | 144.4608 | 0 | 0
|  |  | 144.5376 | 0 | 0
|  |  | 144.6144 | 0 | 0
|  |  | 144.6912 | 0 | 0
|  |  | 144.768 | 0 | 0
|  |  | 144.8448 | 0 | 0
|  |  | 144.9216 | 0 | 0
|  |  | 144.9984 | 0 | 0
|  |  | 145.0752 | 0 | 0
|  |  | 145.152 | 0 | 0
|  |  | 145.2288 | 0 | 0
|  |  | 145.3056 | 0 | 0
|  |  | 145.3824 | 0 | 0
|  |  | 145.4592 | 0 | 0
|  |  | 145.536 | 0 | 0
|  |  | 145.6128 | 0 | 0
|  |  | 145.6896 | 0 | 0
|  |  | 145.7664 | 0 | 0
|  |  | 145.8432 | 0 | 0
|  |  | 145.92 | 0 | 0
|  |  | 145.9968 | 0 | 0
|  |  | 146.0736 | 0 | 0
|  |  | 146.1504 | 0 | 0
|  |  | 146.2272 | 0 | 0
|  |  | 146.304 | 0 | 0
|  |  | 146.3808 | 0 | 0
|  |  | 146.4576 | 0 | 0
|  |  | 146.5344 | 0 | 0
|  |  | 146.6112 | 0 | 0
|  |  | 146.688 | 0 | 0
|  |  | 146.7648 | 0 | 0
|  |  | 146.8416 | 0 | 0
|  |  | 146.9184 | 0 | 0
|  |  | 146.9952 | 0 | 0
|  |  | 147.072 | 0 | 0
|  |  | 147.1488 | 0 | 0
|  |  | 147.2256 | 0 | 0
|  |  | 147.3024 | 0 | 0
|  |  | 147.3792 | 0 | 0
|  |  | 147.456 | 0 | 0
|  |  | 147.5328 | 0 | 0
|  |  | 147.6096 | 0 | 0
|  |  | 147.6864 | 0 | 0
|  |  | 147.7632 | 0 | 0
|  |  | 147.84 | 0 | 0
|  |  | 147.9168 | 0 | 0
|  |  | 147.9936 | 0 | 0
|  |  | 148.0704 | 0 | 0
|  |  | 148.1472 | 0 | 0
|  |  | 148.224 | 0 | 0
|  |  | 148.3008 | 0 | 0
|  |  | 148.3776 | 0 | 0
|  |  | 148.4544 | 0 | 0
|  |  | 148.5312 | 0 | 0
|  |  | 148.608 | 0 | 0
|  |  | 148.6848 | 0 | 0
|  |  | 148.7616 | 0 | 0
|  |  | 148.8384 | 0 | 0
|  |  | 148.9152 | 0 | 0
|  |  | 148.992 | 0 | 0
|  |  | 149.0688 | 0 | 0
|  |  | 149.1456 | 0 | 0
|  |  | 149.2224 | 0 | 0
|  |  | 149.2992 | 0 | 0
|  |  | 149.376 | 40.085 | 0.87552
|  |  | 149.4528 | 43.065 | 0.9472
|  |  | 149.5296 | 45.55 | 1.065
|  |  | 149.6064 | 47.185 | 1.2698
|  |  | 149.6832 | 49.045 | 1.4234
|  |  | 149.76 | 51.385 | 1.4899
|  |  | 149.8368 | 53.175 | 1.6128
|  |  | 149.9136 | 53.94 | 1.8688
|  |  | 149.9904 | 55.735 | 1.9149
|  |  | 150.0672 | 57.23 | 1.9968
|  |  | 150.144 | 58.095 | 2.1453
|  |  | 150.2208 | 59.455 | 2.1555
|  |  | 150.2976 | 59.755 | 2.2835
|  |  | 150.3744 | 60.23 | 2.3501
|  |  | 150.4512 | 60.76 | 2.3706
|  |  | 150.528 | 61.44 | 2.3603
|  |  | 150.6048 | 61.775 | 2.3859
|  |  | 150.6816 | 61.74 | 2.4627
|  |  | 150.7584 | 61.725 | 2.5293
|  |  | 150.8352 | 61.95 | 2.5651
|  |  | 150.912 | 62.645 | 2.5395
|  |  | 150.9888 | 61.505 | 2.9389
|  |  | 151.0656 | 63.085 | 2.5958
|  |  | 151.1424 | 62.01 | 2.9491
|  |  | 151.2192 | 62.815 | 2.7648
|  |  | 151.296 | 62.975 | 2.7546
|  |  | 151.3728 | 62.065 | 3.0157
|  |  | 151.4496 | 62.01 | 3.0413
|  |  | 151.5264 | 61.74 | 3.1283
|  |  | 151.6032 | 63.375 | 2.7085
|  |  | 151.68 | 61.925 | 3.0874
|  |  | 151.7568 | 63.31 | 2.6982
|  |  | 151.8336 | 61.11 | 3.2666
|  |  | 151.9104 | 62.77 | 2.6982
|  |  | 151.9872 | 61.68 | 2.9286
|  |  | 152.064 | 62.16 | 2.7238
|  |  | 152.1408 | 61.24 | 2.8928
|  |  | 152.2176 | 61.845 | 2.6163
|  |  | 152.2944 | 61.715 | 2.5242
|  |  | 152.3712 | 61.015 | 2.4883
|  |  | 152.448 | 59.815 | 2.4576
|  |  | 152.5248 | 57.95 | 2.4832
|  |  | 152.6016 | 56.575 | 2.3501
|  |  | 152.6784 | 55.51 | 2.1453
|  |  | 152.7552 | 54.295 | 1.9763
|  |  | 152.832 | 53.265 | 1.8022
|  |  | 152.9088 | 51.735 | 1.7254
|  |  | 152.9856 | 50.345 | 1.6333
|  |  | 153.0624 | 49.14 | 1.4643
|  |  | 153.1392 | 47.765 | 1.2954
|  |  | 153.216 | 46.72 | 1.1162
|  |  | 153.2928 | 45.595 | 0.9728
|  |  | 153.3696 | 44.36 | 0.79872
|  |  | 153.4464 | 42.415 | 0.72192
|  |  | 153.5232 | 40.735 | 0.59904
|  |  | 153.6 | 0 | 0
|  |  | 153.6768 | 0 | 0
|  |  | 153.7536 | 0 | 0
|  |  | 153.8304 | 0 | 0
|  |  | 153.9072 | 0 | 0
|  |  | 153.984 | 0 | 0
|  |  | 154.0608 | 0 | 0
|  |  | 154.1376 | 0 | 0
|  |  | 154.2144 | 0 | 0
|  |  | 154.2912 | 0 | 0
|  |  | 154.368 | 0 | 0
|  |  | 154.4448 | 0 | 0
|  |  | 154.5216 | 0 | 0
|  |  | 154.5984 | 0 | 0
|  |  | 154.6752 | 0 | 0
|  |  | 154.752 | 0 | 0
|  |  | 154.8288 | 0 | 0
|  |  | 154.9056 | 0 | 0
|  |  | 154.9824 | 0 | 0
|  |  | 155.0592 | 0 | 0
|  |  | 155.136 | 0 | 0
|  |  | 155.2128 | 0 | 0
|  |  | 155.2896 | 0 | 0
|  |  | 155.3664 | 0 | 0
|  |  | 155.4432 | 0 | 0
|  |  | 155.52 | 0 | 0
|  |  | 155.5968 | 0 | 0
|  |  | 155.6736 | 0 | 0
|  |  | 155.7504 | 0 | 0
|  |  | 155.8272 | 0 | 0
|  |  | 155.904 | 0 | 0
|  |  | 155.9808 | 0 | 0
|  |  | 156.0576 | 0 | 0
|  |  | 156.1344 | 0 | 0
|  |  | 156.2112 | 0 | 0
|  |  | 156.288 | 0 | 0
|  |  | 156.3648 | 0 | 0
|  |  | 156.4416 | 0 | 0
|  |  | 156.5184 | 0 | 0
|  |  | 156.5952 | 0 | 0
|  |  | 156.672 | 0 | 0
|  |  | 156.7488 | 0 | 0
|  |  | 156.8256 | 0 | 0
|  |  | 156.9024 | 0 | 0
|  |  | 156.9792 | 0 | 0
|  |  | 157.056 | 0 | 0
|  |  | 157.1328 | 0 | 0
|  |  | 157.2096 | 0 | 0
|  |  | 157.2864 | 0 | 0
|  |  | 157.3632 | 0 | 0
|  |  | 157.44 | 0 | 0
|  |  | 157.5168 | 0 | 0
|  |  | 157.5936 | 0 | 0
|  |  | 157.6704 | 0 | 0
|  |  | 157.7472 | 0 | 0
|  |  | 157.824 | 0 | 0
|  |  | 157.9008 | 0 | 0
|  |  | 157.9776 | 0 | 0
|  |  | 158.0544 | 0 | 0
|  |  | 158.1312 | 0 | 0
|  |  | 158.208 | 0 | 0
|  |  | 158.2848 | 0 | 0
|  |  | 158.3616 | 0 | 0
|  |  | 158.4384 | 0 | 0
|  |  | 158.5152 | 0 | 0
|  |  | 158.592 | 0 | 0
|  |  | 158.6688 | 0 | 0
|  |  | 158.7456 | 0 | 0
|  |  | 158.8224 | 0 | 0
|  |  | 158.8992 | 0 | 0
|  |  | 158.976 | 0 | 0
|  |  | 159.0528 | 0 | 0
|  |  | 159.1296 | 0 | 0
|  |  | 159.2064 | 0 | 0
|  |  | 159.2832 | 0 | 0
|  |  | 159.36 | 0 | 0
|  |  | 159.4368 | 0 | 0
|  |  | 159.5136 | 0 | 0
|  |  | 159.5904 | 0 | 0
|  |  | 159.6672 | 0 | 0
|  |  | 159.744 | 0 | 0
|  |  | 159.8208 | 0 | 0
|  |  | 159.8976 | 0 | 0
|  |  | 159.9744 | 0 | 0
|  |  | 160.0512 | 0 | 0
|  |  | 160.128 | 0 | 0
|  |  | 160.2048 | 0 | 0
|  |  | 160.2816 | 0 | 0
|  |  | 160.3584 | 0 | 0
|  |  | 160.4352 | 0 | 0
|  |  | 160.512 | 0 | 0
|  |  | 160.5888 | 0 | 0
|  |  | 160.6656 | 0 | 0
|  |  | 160.7424 | 0 | 0
|  |  | 160.8192 | 0 | 0
|  |  | 160.896 | 0 | 0
|  |  | 160.9728 | 0 | 0
|  |  | 161.0496 | 0 | 0
|  |  | 161.1264 | 0 | 0
|  |  | 161.2032 | 0 | 0
|  |  | 161.28 | 0 | 0
|  |  | 161.3568 | 0 | 0
|  |  | 161.4336 | 0 | 0
|  |  | 161.5104 | 0 | 0
|  |  | 161.5872 | 0 | 0
|  |  | 161.664 | 0 | 0
|  |  | 161.7408 | 0 | 0
|  |  | 161.8176 | 0 | 0
|  |  | 161.8944 | 0 | 0
|  |  | 161.9712 | 0 | 0
|  |  | 162.048 | 0 | 0
|  |  | 162.1248 | 0 | 0
|  |  | 162.2016 | 0 | 0
|  |  | 162.2784 | 0 | 0
|  |  | 162.3552 | 0 | 0
|  |  | 162.432 | 0 | 0
|  |  | 162.5088 | 0 | 0
|  |  | 162.5856 | 0 | 0
|  |  | 162.6624 | 0 | 0
|  |  | 162.7392 | 0 | 0
|  |  | 162.816 | 0 | 0
|  |  | 162.8928 | 0 | 0
|  |  | 162.9696 | 0 | 0
|  |  | 163.0464 | 0 | 0
|  |  | 163.1232 | 0 | 0
|  |  | 163.2 | 0 | 0
|  |  | 163.2768 | 0 | 0
|  |  | 163.3536 | 0 | 0
|  |  | 163.4304 | 0 | 0
|  |  | 163.5072 | 0 | 0
|  |  | 163.584 | 0 | 0
|  |  | 163.6608 | 0 | 0
|  |  | 163.7376 | 0 | 0
|  |  | 163.8144 | 0 | 0
|  |  | 163.8912 | 0 | 0
|  |  | 163.968 | 0 | 0
|  |  | 164.0448 | 0 | 0
|  |  | 164.1216 | 0 | 0
|  |  | 164.1984 | 0 | 0
|  |  | 164.2752 | 0 | 0
|  |  | 164.352 | 0 | 0
|  |  | 164.4288 | 0 | 0
|  |  | 164.5056 | 0 | 0
|  |  | 164.5824 | 0 | 0
|  |  | 164.6592 | 0 | 0
|  |  | 164.736 | 0 | 0
|  |  | 164.8128 | 0 | 0
|  |  | 164.8896 | 0 | 0
|  |  | 164.9664 | 0 | 0
|  |  | 165.0432 | 0 | 0
|  |  | 165.12 | 0 | 0
|  |  | 165.1968 | 0 | 0
|  |  | 165.2736 | 0 | 0
|  |  | 165.3504 | 0 | 0
|  |  | 165.4272 | 0 | 0
|  |  | 165.504 | 0 | 0
|  |  | 165.5808 | 0 | 0
|  |  | 165.6576 | 0 | 0
|  |  | 165.7344 | 0 | 0
|  |  | 165.8112 | 0 | 0
|  |  | 165.888 | 0 | 0
|  |  | 165.9648 | 0 | 0
|  |  | 166.0416 | 41.325 | 0.65024
|  |  | 166.1184 | 43.92 | 0.77312
|  |  | 166.1952 | 45.79 | 0.9472
|  |  | 166.272 | 47.37 | 1.1418
|  |  | 166.3488 | 49.715 | 1.2237
|  |  | 166.4256 | 51.81 | 1.3363
|  |  | 166.5024 | 52.85 | 1.5821
|  |  | 166.5792 | 54.32 | 1.7408
|  |  | 166.656 | 55.79 | 1.8432
|  |  | 166.7328 | 56.985 | 1.9866
|  |  | 166.8096 | 58.025 | 2.1146
|  |  | 166.8864 | 58.98 | 2.2016
|  |  | 166.9632 | 60.305 | 2.1811
|  |  | 167.04 | 61.315 | 2.1606
|  |  | 167.1168 | 61.215 | 2.2835
|  |  | 167.1936 | 61.325 | 2.3757
|  |  | 167.2704 | 61.415 | 2.4474
|  |  | 167.3472 | 61.25 | 2.5395
|  |  | 167.424 | 61.39 | 2.5805
|  |  | 167.5008 | 62.31 | 2.4934
|  |  | 167.5776 | 61.13 | 2.9235
|  |  | 167.6544 | 62.86 | 2.5446
|  |  | 167.7312 | 61.775 | 2.9184
|  |  | 167.808 | 63.37 | 2.5856
|  |  | 167.8848 | 62.04 | 2.9696
|  |  | 167.9616 | 62.265 | 2.944
|  |  | 168.0384 | 62.9 | 2.775
|  |  | 168.1152 | 63.285 | 2.688
|  |  | 168.192 | 62.395 | 2.9389
|  |  | 168.2688 | 61.755 | 3.1488
|  |  | 168.3456 | 63.875 | 2.5907
|  |  | 168.4224 | 61.915 | 3.0464
|  |  | 168.4992 | 63.215 | 2.6726
|  |  | 168.576 | 61.53 | 3.0874
|  |  | 168.6528 | 62.955 | 2.6419
|  |  | 168.7296 | 60.47 | 3.2512
|  |  | 168.8064 | 62.11 | 2.6675
|  |  | 168.8832 | 60.7 | 2.9542
|  |  | 168.96 | 61.14 | 2.6317
|  |  | 169.0368 | 60.545 | 2.5651
|  |  | 169.1136 | 60.215 | 2.3706
|  |  | 169.1904 | 59.23 | 2.2374
|  |  | 169.2672 | 57.435 | 2.2016
|  |  | 169.344 | 55.71 | 2.1197
|  |  | 169.4208 | 54.275 | 1.9558
|  |  | 169.4976 | 53.015 | 1.8125
|  |  | 169.5744 | 52.22 | 1.6435
|  |  | 169.6512 | 50.685 | 1.6077
|  |  | 169.728 | 49.2 | 1.5258
|  |  | 169.8048 | 48.215 | 1.3056
|  |  | 169.8816 | 47.105 | 1.1315
|  |  | 169.9584 | 45.445 | 1.0342
|  |  | 170.0352 | 43.5 | 0.95232
|  |  | 170.112 | 41.665 | 0.8448
|  |  | 170.1888 | 39.925 | 0.6912
|  |  | 170.2656 | 0 | 0
|  |  | 170.3424 | 0 | 0
|  |  | 170.4192 | 0 | 0
|  |  | 170.496 | 0 | 0
|  |  | 170.5728 | 0 | 0
|  |  | 170.6496 | 0 | 0
|  |  | 170.7264 | 0 | 0
|  |  | 170.8032 | 0 | 0
|  |  | 170.88 | 0 | 0
|  |  | 170.9568 | 0 | 0
|  |  | 171.0336 | 0 | 0
|  |  | 171.1104 | 0 | 0
|  |  | 171.1872 | 0 | 0
|  |  | 171.264 | 0 | 0
|  |  | 171.3408 | 0 | 0
|  |  | 171.4176 | 0 | 0
|  |  | 171.4944 | 0 | 0
|  |  | 171.5712 | 0 | 0
|  |  | 171.648 | 0 | 0
|  |  | 171.7248 | 0 | 0
|  |  | 171.8016 | 0 | 0
|  |  | 171.8784 | 0 | 0
|  |  | 171.9552 | 0 | 0
|  |  | 172.032 | 0 | 0
|  |  | 172.1088 | 0 | 0
|  |  | 172.1856 | 0 | 0
|  |  | 172.2624 | 0 | 0
|  |  | 172.3392 | 0 | 0
|  |  | 172.416 | 0 | 0
|  |  | 172.4928 | 0 | 0
|  |  | 172.5696 | 0 | 0
|  |  | 172.6464 | 0 | 0
|  |  | 172.7232 | 0 | 0
|  |  | 172.8 | 0 | 0
|  |  | 172.8768 | 0 | 0
|  |  | 172.9536 | 0 | 0
|  |  | 173.0304 | 0 | 0
|  |  | 173.1072 | 0 | 0
|  |  | 173.184 | 0 | 0
|  |  | 173.2608 | 0 | 0
|  |  | 173.3376 | 0 | 0
|  |  | 173.4144 | 0 | 0
|  |  | 173.4912 | 0 | 0
|  |  | 173.568 | 0 | 0
|  |  | 173.6448 | 0 | 0
|  |  | 173.7216 | 0 | 0
|  |  | 173.7984 | 0 | 0
|  |  | 173.8752 | 0 | 0
|  |  | 173.952 | 0 | 0
|  |  | 174.0288 | 0 | 0
|  |  | 174.1056 | 0 | 0
|  |  | 174.1824 | 0 | 0
|  |  | 174.2592 | 0 | 0
|  |  | 174.336 | 0 | 0
|  |  | 174.4128 | 0 | 0
|  |  | 174.4896 | 0 | 0
|  |  | 174.5664 | 0 | 0
|  |  | 174.6432 | 0 | 0
|  |  | 174.72 | 0 | 0
|  |  | 174.7968 | 0 | 0
|  |  | 174.8736 | 0 | 0
|  |  | 174.9504 | 0 | 0
|  |  | 175.0272 | 0 | 0
|  |  | 175.104 | 0 | 0
|  |  | 175.1808 | 0 | 0
|  |  | 175.2576 | 0 | 0
|  |  | 175.3344 | 0 | 0
|  |  | 175.4112 | 0 | 0
|  |  | 175.488 | 0 | 0
|  |  | 175.5648 | 0 | 0
|  |  | 175.6416 | 0 | 0
|  |  | 175.7184 | 0 | 0
|  |  | 175.7952 | 0 | 0
|  |  | 175.872 | 0 | 0
|  |  | 175.9488 | 0 | 0
|  |  | 176.0256 | 0 | 0
|  |  | 176.1024 | 0 | 0
|  |  | 176.1792 | 0 | 0
|  |  | 176.256 | 0 | 0
|  |  | 176.3328 | 0 | 0
|  |  | 176.4096 | 0 | 0
|  |  | 176.4864 | 0 | 0
|  |  | 176.5632 | 0 | 0
|  |  | 176.64 | 0 | 0
|  |  | 176.7168 | 0 | 0
|  |  | 176.7936 | 0 | 0
|  |  | 176.8704 | 0 | 0
|  |  | 176.9472 | 0 | 0
|  |  | 177.024 | 0 | 0
|  |  | 177.1008 | 0 | 0
|  |  | 177.1776 | 0 | 0
|  |  | 177.2544 | 0 | 0
|  |  | 177.3312 | 0 | 0
|  |  | 177.408 | 0 | 0
|  |  | 177.4848 | 0 | 0
|  |  | 177.5616 | 0 | 0
|  |  | 177.6384 | 0 | 0
|  |  | 177.7152 | 0 | 0
|  |  | 177.792 | 0 | 0
|  |  | 177.8688 | 0 | 0
|  |  | 177.9456 | 0 | 0
|  |  | 178.0224 | 0 | 0
|  |  | 178.0992 | 0 | 0
|  |  | 178.176 | 0 | 0
|  |  | 178.2528 | 0 | 0
|  |  | 178.3296 | 0 | 0
|  |  | 178.4064 | 0 | 0
|  |  | 178.4832 | 0 | 0
|  |  | 178.56 | 0 | 0
|  |  | 178.6368 | 0 | 0
|  |  | 178.7136 | 0 | 0
|  |  | 178.7904 | 0 | 0
|  |  | 178.8672 | 0 | 0
|  |  | 178.944 | 0 | 0
|  |  | 179.0208 | 0 | 0
|  |  | 179.0976 | 0 | 0
|  |  | 179.1744 | 0 | 0
|  |  | 179.2512 | 0 | 0
|  |  | 179.328 | 0 | 0
|  |  | 179.4048 | 0 | 0
|  |  | 179.4816 | 0 | 0
|  |  | 179.5584 | 0 | 0
|  |  | 179.6352 | 0 | 0
|  |  | 179.712 | 0 | 0
|  |  | 179.7888 | 0 | 0
|  |  | 179.8656 | 0 | 0
|  |  | 179.9424 | 0 | 0
|  |  | 180.0192 | 0 | 0
|  |  | 180.096 | 0 | 0
|  |  | 180.1728 | 0 | 0
|  |  | 180.2496 | 0 | 0
|  |  | 180.3264 | 0 | 0
|  |  | 180.4032 | 0 | 0
|  |  | 180.48 | 0 | 0
|  |  | 180.5568 | 0 | 0
|  |  | 180.6336 | 0 | 0
|  |  | 180.7104 | 0 | 0
|  |  | 180.7872 | 0 | 0
|  |  | 180.864 | 0 | 0
|  |  | 180.9408 | 0 | 0
|  |  | 181.0176 | 0 | 0
|  |  | 181.0944 | 0 | 0
|  |  | 181.1712 | 0 | 0
|  |  | 181.248 | 0 | 0
|  |  | 181.3248 | 0 | 0
|  |  | 181.4016 | 0 | 0
|  |  | 181.4784 | 0 | 0
|  |  | 181.5552 | 0 | 0
|  |  | 181.632 | 0 | 0
|  |  | 181.7088 | 0 | 0
|  |  | 181.7856 | 0 | 0
|  |  | 181.8624 | 0 | 0
|  |  | 181.9392 | 0 | 0
|  |  | 182.016 | 0 | 0
|  |  | 182.0928 | 0 | 0
|  |  | 182.1696 | 0 | 0
|  |  | 182.2464 | 0 | 0
|  |  | 182.3232 | 0 | 0
|  |  | 182.4 | 0 | 0
|  |  | 182.4768 | 0 | 0
|  |  | 182.5536 | 0 | 0
|  |  | 182.6304 | 0 | 0
|  |  | 182.7072 | 39.73 | 0.88576
|  |  | 182.784 | 42.39 | 0.98304
|  |  | 182.8608 | 44.195 | 1.2186
|  |  | 182.9376 | 46.335 | 1.3414
|  |  | 183.0144 | 48.865 | 1.367
|  |  | 183.0912 | 51.375 | 1.3926
|  |  | 183.168 | 53.27 | 1.495
|  |  | 183.2448 | 54.655 | 1.664
|  |  | 183.3216 | 55.935 | 1.833
|  |  | 183.3984 | 57.295 | 1.9456
|  |  | 183.4752 | 58.535 | 2.0275
|  |  | 183.552 | 59.665 | 2.089
|  |  | 183.6288 | 60.185 | 2.2016
|  |  | 183.7056 | 61.09 | 2.1965
|  |  | 183.7824 | 62.01 | 2.1555
|  |  | 183.8592 | 60.27 | 2.6726
|  |  | 183.936 | 61.77 | 2.3603
|  |  | 184.0128 | 61.8 | 2.432
|  |  | 184.0896 | 62.52 | 2.3859
|  |  | 184.1664 | 61.145 | 2.816
|  |  | 184.2432 | 62.23 | 2.5958
|  |  | 184.32 | 61.805 | 2.8416
|  |  | 184.3968 | 63.25 | 2.5805
|  |  | 184.4736 | 62.15 | 2.9491
|  |  | 184.5504 | 61.68 | 3.1283
|  |  | 184.6272 | 63.525 | 2.6522
|  |  | 184.704 | 63.005 | 2.7904
|  |  | 184.7808 | 62.75 | 2.8826
|  |  | 184.8576 | 62.125 | 3.0771
|  |  | 184.9344 | 62.265 | 3.0413
|  |  | 185.0112 | 62.245 | 3.0618
|  |  | 185.088 | 62.065 | 3.0976
|  |  | 185.1648 | 61.915 | 3.0874
|  |  | 185.2416 | 63.125 | 2.6829
|  |  | 185.3184 | 61.955 | 2.944
|  |  | 185.3952 | 62.95 | 2.6266
|  |  | 185.472 | 61.27 | 2.944
|  |  | 185.5488 | 61.19 | 2.7802
|  |  | 185.6256 | 60.845 | 2.7136
|  |  | 185.7024 | 60.235 | 2.6829
|  |  | 185.7792 | 59.575 | 2.5651
|  |  | 185.856 | 58.31 | 2.4781
|  |  | 185.9328 | 57.495 | 2.263
|  |  | 186.0096 | 56.435 | 2.0685
|  |  | 186.0864 | 55.05 | 1.9405
|  |  | 186.1632 | 53.295 | 1.9046
|  |  | 186.24 | 51.91 | 1.8227
|  |  | 186.3168 | 50.48 | 1.7101
|  |  | 186.3936 | 48.75 | 1.6589
|  |  | 186.4704 | 47.43 | 1.495
|  |  | 186.5472 | 46.44 | 1.2851
|  |  | 186.624 | 45.95 | 1.0189
|  |  | 186.7008 | 44.35 | 0.89088
|  |  | 186.7776 | 42.435 | 0.79872
|  |  | 186.8544 | 40.79 | 0.67072
|  |  | 186.9312 | 0 | 0
|  |  | 187.008 | 0 | 0
|  |  | 187.0848 | 0 | 0
|  |  | 187.1616 | 0 | 0
|  |  | 187.2384 | 0 | 0
|  |  | 187.3152 | 0 | 0
|  |  | 187.392 | 0 | 0
|  |  | 187.4688 | 0 | 0
|  |  | 187.5456 | 0 | 0
|  |  | 187.6224 | 0 | 0
|  |  | 187.6992 | 0 | 0
|  |  | 187.776 | 0 | 0
|  |  | 187.8528 | 0 | 0
|  |  | 187.9296 | 0 | 0
|  |  | 188.0064 | 0 | 0
|  |  | 188.0832 | 0 | 0
|  |  | 188.16 | 0 | 0
|  |  | 188.2368 | 0 | 0
|  |  | 188.3136 | 0 | 0
|  |  | 188.3904 | 0 | 0
|  |  | 188.4672 | 0 | 0
|  |  | 188.544 | 0 | 0
|  |  | 188.6208 | 0 | 0
|  |  | 188.6976 | 0 | 0
|  |  | 188.7744 | 0 | 0
|  |  | 188.8512 | 0 | 0
|  |  | 188.928 | 0 | 0
|  |  | 189.0048 | 0 | 0
|  |  | 189.0816 | 0 | 0
|  |  | 189.1584 | 0 | 0
|  |  | 189.2352 | 0 | 0
|  |  | 189.312 | 0 | 0
|  |  | 189.3888 | 0 | 0
|  |  | 189.4656 | 0 | 0
|  |  | 189.5424 | 0 | 0
|  |  | 189.6192 | 0 | 0
|  |  | 189.696 | 0 | 0
|  |  | 189.7728 | 0 | 0
|  |  | 189.8496 | 0 | 0
|  |  | 189.9264 | 0 | 0
|  |  | 190.0032 | 0 | 0
|  |  | 190.08 | 0 | 0
|  |  | 190.1568 | 0 | 0
|  |  | 190.2336 | 0 | 0
|  |  | 190.3104 | 0 | 0
|  |  | 190.3872 | 0 | 0
|  |  | 190.464 | 0 | 0
|  |  | 190.5408 | 0 | 0
|  |  | 190.6176 | 0 | 0
|  |  | 190.6944 | 0 | 0
|  |  | 190.7712 | 0 | 0
|  |  | 190.848 | 0 | 0
|  |  | 190.9248 | 0 | 0
|  |  | 191.0016 | 0 | 0
|  |  | 191.0784 | 0 | 0
|  |  | 191.1552 | 0 | 0
|  |  | 191.232 | 0 | 0
|  |  | 191.3088 | 0 | 0
|  |  | 191.3856 | 0 | 0
|  |  | 191.4624 | 0 | 0
|  |  | 191.5392 | 0 | 0
|  |  | 191.616 | 0 | 0
|  |  | 191.6928 | 0 | 0
|  |  | 191.7696 | 0 | 0
|  |  | 191.8464 | 0 | 0
|  |  | 191.9232 | 0 | 0
|  |  | 192 | 0 | 0
|  |  | 192.0768 | 0 | 0
|  |  | 192.1536 | 0 | 0
|  |  | 192.2304 | 0 | 0
|  |  | 192.3072 | 0 | 0
|  |  | 192.384 | 0 | 0
|  |  | 192.4608 | 0 | 0
|  |  | 192.5376 | 0 | 0
|  |  | 192.6144 | 0 | 0
|  |  | 192.6912 | 0 | 0
|  |  | 192.768 | 0 | 0
|  |  | 192.8448 | 0 | 0
|  |  | 192.9216 | 0 | 0
|  |  | 192.9984 | 0 | 0
|  |  | 193.0752 | 0 | 0
|  |  | 193.152 | 0 | 0
|  |  | 193.2288 | 0 | 0
|  |  | 193.3056 | 0 | 0
|  |  | 193.3824 | 0 | 0
|  |  | 193.4592 | 0 | 0
|  |  | 193.536 | 0 | 0
|  |  | 193.6128 | 0 | 0
|  |  | 193.6896 | 0 | 0
|  |  | 193.7664 | 0 | 0
|  |  | 193.8432 | 0 | 0
|  |  | 193.92 | 0 | 0
|  |  | 193.9968 | 0 | 0
|  |  | 194.0736 | 0 | 0
|  |  | 194.1504 | 0 | 0
|  |  | 194.2272 | 0 | 0
|  |  | 194.304 | 0 | 0
|  |  | 194.3808 | 0 | 0
|  |  | 194.4576 | 0 | 0
|  |  | 194.5344 | 0 | 0
|  |  | 194.6112 | 0 | 0
|  |  | 194.688 | 0 | 0
|  |  | 194.7648 | 0 | 0
|  |  | 194.8416 | 0 | 0
|  |  | 194.9184 | 0 | 0
|  |  | 194.9952 | 0 | 0
|  |  | 195.072 | 0 | 0
|  |  | 195.1488 | 0 | 0
|  |  | 195.2256 | 0 | 0
|  |  | 195.3024 | 0 | 0
|  |  | 195.3792 | 0 | 0
|  |  | 195.456 | 0 | 0
|  |  | 195.5328 | 0 | 0
|  |  | 195.6096 | 0 | 0
|  |  | 195.6864 | 0 | 0
|  |  | 195.7632 | 0 | 0
|  |  | 195.84 | 0 | 0
|  |  | 195.9168 | 0 | 0
|  |  | 195.9936 | 0 | 0
|  |  | 196.0704 | 0 | 0
|  |  | 196.1472 | 0 | 0
|  |  | 196.224 | 0 | 0
|  |  | 196.3008 | 0 | 0
|  |  | 196.3776 | 0 | 0
|  |  | 196.4544 | 0 | 0
|  |  | 196.5312 | 0 | 0
|  |  | 196.608 | 0 | 0
|  |  | 196.6848 | 0 | 0
|  |  | 196.7616 | 0 | 0
|  |  | 196.8384 | 0 | 0
|  |  | 196.9152 | 0 | 0
|  |  | 196.992 | 0 | 0
|  |  | 197.0688 | 0 | 0
|  |  | 197.1456 | 0 | 0
|  |  | 197.2224 | 0 | 0
|  |  | 197.2992 | 0 | 0
|  |  | 197.376 | 0 | 0
|  |  | 197.4528 | 0 | 0
|  |  | 197.5296 | 0 | 0
|  |  | 197.6064 | 0 | 0
|  |  | 197.6832 | 0 | 0
|  |  | 197.76 | 0 | 0
|  |  | 197.8368 | 0 | 0
|  |  | 197.9136 | 0 | 0
|  |  | 197.9904 | 0 | 0
|  |  | 198.0672 | 0 | 0
|  |  | 198.144 | 0 | 0
|  |  | 198.2208 | 0 | 0
|  |  | 198.2976 | 0 | 0
|  |  | 198.3744 | 0 | 0
|  |  | 198.4512 | 0 | 0
|  |  | 198.528 | 0 | 0
|  |  | 198.6048 | 0 | 0
|  |  | 198.6816 | 0 | 0
|  |  | 198.7584 | 0 | 0
|  |  | 198.8352 | 0 | 0
|  |  | 198.912 | 0 | 0
|  |  | 198.9888 | 0 | 0
|  |  | 199.0656 | 0 | 0
|  |  | 199.1424 | 0 | 0
|  |  | 199.2192 | 0 | 0
|  |  | 199.296 | 0 | 0
|  |  | 199.3728 | 39.3 | 0.90624
|  |  | 199.4496 | 43.235 | 0.84992
|  |  | 199.5264 | 45.55 | 0.96256
|  |  | 199.6032 | 47.345 | 1.1366
|  |  | 199.68 | 49.515 | 1.2493
|  |  | 199.7568 | 51.155 | 1.4029
|  |  | 199.8336 | 52.31 | 1.6538
|  |  | 199.9104 | 53.935 | 1.7869
|  |  | 199.9872 | 55.735 | 1.8432
|  |  | 200.064 | 57.005 | 1.9712
|  |  | 200.1408 | 57.665 | 2.176
|  |  | 200.2176 | 58.6 | 2.2835
|  |  | 200.2944 | 59.72 | 2.3091
|  |  | 200.3712 | 60.38 | 2.3552
|  |  | 200.448 | 60.655 | 2.432
|  |  | 200.5248 | 60.73 | 2.4986
|  |  | 200.6016 | 61.02 | 2.5139
|  |  | 200.6784 | 61.925 | 2.4269
|  |  | 200.7552 | 62.28 | 2.432
|  |  | 200.832 | 61.615 | 2.7034
|  |  | 200.9088 | 62.68 | 2.5446
|  |  | 200.9856 | 61.89 | 2.8314
|  |  | 201.0624 | 63.49 | 2.5088
|  |  | 201.1392 | 62.44 | 2.8365
|  |  | 201.216 | 62.575 | 2.8365
|  |  | 201.2928 | 62.75 | 2.8211
|  |  | 201.3696 | 63.445 | 2.688
|  |  | 201.4464 | 63.165 | 2.7904
|  |  | 201.5232 | 62.675 | 2.9338
|  |  | 201.6 | 62.47 | 3.0003
|  |  | 201.6768 | 62.44 | 3.0208
|  |  | 201.7536 | 61.935 | 3.1488
|  |  | 201.8304 | 63.8 | 2.6266
|  |  | 201.9072 | 61.28 | 3.2512
|  |  | 201.984 | 63.44 | 2.6061
|  |  | 202.0608 | 61.455 | 3.031
|  |  | 202.1376 | 62.125 | 2.7238
|  |  | 202.2144 | 60.625 | 3.031
|  |  | 202.2912 | 61.295 | 2.6726
|  |  | 202.368 | 60.985 | 2.5856
|  |  | 202.4448 | 60.26 | 2.4934
|  |  | 202.5216 | 59.06 | 2.4013
|  |  | 202.5984 | 58.155 | 2.1965
|  |  | 202.6752 | 56.82 | 2.0736
|  |  | 202.752 | 54.965 | 2.0326
|  |  | 202.8288 | 53.495 | 1.92
|  |  | 202.9056 | 52.54 | 1.7357
|  |  | 202.9824 | 51.53 | 1.5718
|  |  | 203.0592 | 49.425 | 1.5821
|  |  | 203.136 | 47.825 | 1.4746
|  |  | 203.2128 | 46.96 | 1.2288
|  |  | 203.2896 | 46.25 | 1.0138
|  |  | 203.3664 | 44.55 | 0.91136
|  |  | 203.4432 | 42.22 | 0.8704
|  |  | 203.52 | 40.89 | 0.69632
|  |  | 203.5968 | 0 | 0
|  |  | 203.6736 | 0 | 0
|  |  | 203.7504 | 0 | 0
|  |  | 203.8272 | 0 | 0
|  |  | 203.904 | 0 | 0
|  |  | 203.9808 | 0 | 0
|  |  | 204.0576 | 0 | 0
|  |  | 204.1344 | 0 | 0
|  |  | 204.2112 | 0 | 0
|  |  | 204.288 | 0 | 0
|  |  | 204.3648 | 0 | 0
|  |  | 204.4416 | 0 | 0
|  |  | 204.5184 | 0 | 0
|  |  | 204.5952 | 0 | 0
|  |  | 204.672 | 0 | 0
|  |  | 204.7488 | 0 | 0
|  |  | 204.8256 | 0 | 0
|  |  | 204.9024 | 0 | 0
|  |  | 204.9792 | 0 | 0
|  |  | 205.056 | 0 | 0
|  |  | 205.1328 | 0 | 0
|  |  | 205.2096 | 0 | 0
|  |  | 205.2864 | 0 | 0
|  |  | 205.3632 | 0 | 0
|  |  | 205.44 | 0 | 0
|  |  | 205.5168 | 0 | 0
|  |  | 205.5936 | 0 | 0
|  |  | 205.6704 | 0 | 0
|  |  | 205.7472 | 0 | 0
|  |  | 205.824 | 0 | 0
|  |  | 205.9008 | 0 | 0
|  |  | 205.9776 | 0 | 0
|  |  | 206.0544 | 0 | 0
|  |  | 206.1312 | 0 | 0
|  |  | 206.208 | 0 | 0
|  |  | 206.2848 | 0 | 0
|  |  | 206.3616 | 0 | 0
|  |  | 206.4384 | 0 | 0
|  |  | 206.5152 | 0 | 0
|  |  | 206.592 | 0 | 0
|  |  | 206.6688 | 0 | 0
|  |  | 206.7456 | 0 | 0
|  |  | 206.8224 | 0 | 0
|  |  | 206.8992 | 0 | 0
|  |  | 206.976 | 0 | 0
|  |  | 207.0528 | 0 | 0
|  |  | 207.1296 | 0 | 0
|  |  | 207.2064 | 0 | 0
|  |  | 207.2832 | 0 | 0
|  |  | 207.36 | 0 | 0
|  |  | 207.4368 | 0 | 0
|  |  | 207.5136 | 0 | 0
|  |  | 207.5904 | 0 | 0
|  |  | 207.6672 | 0 | 0
|  |  | 207.744 | 0 | 0
|  |  | 207.8208 | 0 | 0
|  |  | 207.8976 | 0 | 0
|  |  | 207.9744 | 0 | 0
|  |  | 208.0512 | 0 | 0
|  |  | 208.128 | 0 | 0
|  |  | 208.2048 | 0 | 0
|  |  | 208.2816 | 0 | 0
|  |  | 208.3584 | 0 | 0
|  |  | 208.4352 | 0 | 0
|  |  | 208.512 | 0 | 0
|  |  | 208.5888 | 0 | 0
|  |  | 208.6656 | 0 | 0
|  |  | 208.7424 | 0 | 0
|  |  | 208.8192 | 0 | 0
|  |  | 208.896 | 0 | 0
|  |  | 208.9728 | 0 | 0
|  |  | 209.0496 | 0 | 0
|  |  | 209.1264 | 0 | 0
|  |  | 209.2032 | 0 | 0
|  |  | 209.28 | 0 | 0
|  |  | 209.3568 | 0 | 0
|  |  | 209.4336 | 0 | 0
|  |  | 209.5104 | 0 | 0
|  |  | 209.5872 | 0 | 0
|  |  | 209.664 | 0 | 0
|  |  | 209.7408 | 0 | 0
|  |  | 209.8176 | 0 | 0
|  |  | 209.8944 | 0 | 0
|  |  | 209.9712 | 0 | 0
|  |  | 210.048 | 0 | 0
|  |  | 210.1248 | 0 | 0
|  |  | 210.2016 | 0 | 0
|  |  | 210.2784 | 0 | 0
|  |  | 210.3552 | 0 | 0
|  |  | 210.432 | 0 | 0
|  |  | 210.5088 | 0 | 0
|  |  | 210.5856 | 0 | 0
|  |  | 210.6624 | 0 | 0
|  |  | 210.7392 | 0 | 0
|  |  | 210.816 | 0 | 0
|  |  | 210.8928 | 0 | 0
|  |  | 210.9696 | 0 | 0
|  |  | 211.0464 | 0 | 0
|  |  | 211.1232 | 0 | 0
|  |  | 211.2 | 0 | 0
|  |  | 211.2768 | 0 | 0
|  |  | 211.3536 | 0 | 0
|  |  | 211.4304 | 0 | 0
|  |  | 211.5072 | 0 | 0
|  |  | 211.584 | 0 | 0
|  |  | 211.6608 | 0 | 0
|  |  | 211.7376 | 0 | 0
|  |  | 211.8144 | 0 | 0
|  |  | 211.8912 | 0 | 0
|  |  | 211.968 | 0 | 0
|  |  | 212.0448 | 0 | 0
|  |  | 212.1216 | 0 | 0
|  |  | 212.1984 | 0 | 0
|  |  | 212.2752 | 0 | 0
|  |  | 212.352 | 0 | 0
|  |  | 212.4288 | 0 | 0
|  |  | 212.5056 | 0 | 0
|  |  | 212.5824 | 0 | 0
|  |  | 212.6592 | 0 | 0
|  |  | 212.736 | 0 | 0
|  |  | 212.8128 | 0 | 0
|  |  | 212.8896 | 0 | 0
|  |  | 212.9664 | 0 | 0
|  |  | 213.0432 | 0 | 0
|  |  | 213.12 | 0 | 0
|  |  | 213.1968 | 0 | 0
|  |  | 213.2736 | 0 | 0
|  |  | 213.3504 | 0 | 0
|  |  | 213.4272 | 0 | 0
|  |  | 213.504 | 0 | 0
|  |  | 213.5808 | 0 | 0
|  |  | 213.6576 | 0 | 0
|  |  | 213.7344 | 0 | 0
|  |  | 213.8112 | 0 | 0
|  |  | 213.888 | 0 | 0
|  |  | 213.9648 | 0 | 0
|  |  | 214.0416 | 0 | 0
|  |  | 214.1184 | 0 | 0
|  |  | 214.1952 | 0 | 0
|  |  | 214.272 | 0 | 0
|  |  | 214.3488 | 0 | 0
|  |  | 214.4256 | 0 | 0
|  |  | 214.5024 | 0 | 0
|  |  | 214.5792 | 0 | 0
|  |  | 214.656 | 0 | 0
|  |  | 214.7328 | 0 | 0
|  |  | 214.8096 | 0 | 0
|  |  | 214.8864 | 0 | 0
|  |  | 214.9632 | 0 | 0
|  |  | 215.04 | 0 | 0
|  |  | 215.1168 | 0 | 0
|  |  | 215.1936 | 0 | 0
|  |  | 215.2704 | 0 | 0
|  |  | 215.3472 | 0 | 0
|  |  | 215.424 | 0 | 0
|  |  | 215.5008 | 0 | 0
|  |  | 215.5776 | 0 | 0
|  |  | 215.6544 | 0 | 0
|  |  | 215.7312 | 0 | 0
|  |  | 215.808 | 0 | 0
|  |  | 215.8848 | 0 | 0
|  |  | 215.9616 | 0 | 0
|  |  | 216.0384 | 39.69 | 0.82432
|  |  | 216.1152 | 42.195 | 0.9472
|  |  | 216.192 | 44.435 | 1.0854
|  |  | 216.2688 | 46.82 | 1.1981
|  |  | 216.3456 | 48.835 | 1.3261
|  |  | 216.4224 | 50.815 | 1.449
|  |  | 216.4992 | 52.365 | 1.6128
|  |  | 216.576 | 54.165 | 1.7152
|  |  | 216.6528 | 55.865 | 1.8074
|  |  | 216.7296 | 57.355 | 1.9046
|  |  | 216.8064 | 58.585 | 2.0019
|  |  | 216.8832 | 59.32 | 2.135
|  |  | 216.96 | 59.945 | 2.2374
|  |  | 217.0368 | 60.365 | 2.3398
|  |  | 217.1136 | 61.035 | 2.3501
|  |  | 217.1904 | 61.28 | 2.3859
|  |  | 217.2672 | 61.645 | 2.4115
|  |  | 217.344 | 62.35 | 2.3654
|  |  | 217.4208 | 61.635 | 2.6266
|  |  | 217.4976 | 62.285 | 2.5549
|  |  | 217.5744 | 62.47 | 2.6317
|  |  | 217.6512 | 62.74 | 2.6522
|  |  | 217.728 | 62.5 | 2.7853
|  |  | 217.8048 | 62.2 | 2.9286
|  |  | 217.8816 | 62.47 | 2.8979
|  |  | 217.9584 | 62.715 | 2.8672
|  |  | 218.0352 | 63.185 | 2.729
|  |  | 218.112 | 62.685 | 2.8365
|  |  | 218.1888 | 62.44 | 2.944
|  |  | 218.2656 | 62.59 | 2.9747
|  |  | 218.3424 | 62.515 | 3.0106
|  |  | 218.4192 | 62.32 | 3.0515
|  |  | 218.496 | 62.46 | 2.9696
|  |  | 218.5728 | 62.685 | 2.8365
|  |  | 218.6496 | 62.78 | 2.7443
|  |  | 218.7264 | 61.805 | 2.9338
|  |  | 218.8032 | 62.75 | 2.6112
|  |  | 218.88 | 60.25 | 3.2051
|  |  | 218.9568 | 61.565 | 2.647
|  |  | 219.0336 | 61.05 | 2.5856
|  |  | 219.1104 | 60.555 | 2.473
|  |  | 219.1872 | 59.485 | 2.3706
|  |  | 219.264 | 58.375 | 2.2067
|  |  | 219.3408 | 56.91 | 2.0838
|  |  | 219.4176 | 54.655 | 2.1094
|  |  | 219.4944 | 53.535 | 1.9354
|  |  | 219.5712 | 52.605 | 1.751
|  |  | 219.648 | 51.28 | 1.6384
|  |  | 219.7248 | 49.675 | 1.5616
|  |  | 219.8016 | 48.275 | 1.4438
|  |  | 219.8784 | 47.36 | 1.2288
|  |  | 219.9552 | 46.425 | 1.0342
|  |  | 220.032 | 44.2 | 1.0086
|  |  | 220.1088 | 42.925 | 0.8192
|  |  | 220.1856 | 41.435 | 0.66048
|  |  | 220.2624 | 0 | 0
|  |  | 220.3392 | 0 | 0
|  |  | 220.416 | 0 | 0
|  |  | 220.4928 | 0 | 0
|  |  | 220.5696 | 0 | 0
|  |  | 220.6464 | 0 | 0
|  |  | 220.7232 | 0 | 0
|  |  | 220.8 | 0 | 0
|  |  | 220.8768 | 0 | 0
|  |  | 220.9536 | 0 | 0
|  |  | 221.0304 | 0 | 0
|  |  | 221.1072 | 0 | 0
|  |  | 221.184 | 0 | 0
|  |  | 221.2608 | 0 | 0
|  |  | 221.3376 | 0 | 0
|  |  | 221.4144 | 0 | 0
|  |  | 221.4912 | 0 | 0
|  |  | 221.568 | 0 | 0
|  |  | 221.6448 | 0 | 0
|  |  | 221.7216 | 0 | 0
|  |  | 221.7984 | 0 | 0
|  |  | 221.8752 | 0 | 0
|  |  | 221.952 | 0 | 0
|  |  | 222.0288 | 0 | 0
|  |  | 222.1056 | 0 | 0
|  |  | 222.1824 | 0 | 0
|  |  | 222.2592 | 0 | 0
|  |  | 222.336 | 0 | 0
|  |  | 222.4128 | 0 | 0
|  |  | 222.4896 | 0 | 0
|  |  | 222.5664 | 0 | 0
|  |  | 222.6432 | 0 | 0
|  |  | 222.72 | 0 | 0
|  |  | 222.7968 | 0 | 0
|  |  | 222.8736 | 0 | 0
|  |  | 222.9504 | 0 | 0
|  |  | 223.0272 | 0 | 0
|  |  | 223.104 | 0 | 0
|  |  | 223.1808 | 0 | 0
|  |  | 223.2576 | 0 | 0
|  |  | 223.3344 | 0 | 0
|  |  | 223.4112 | 0 | 0
|  |  | 223.488 | 0 | 0
|  |  | 223.5648 | 0 | 0
|  |  | 223.6416 | 0 | 0
|  |  | 223.7184 | 0 | 0
|  |  | 223.7952 | 0 | 0
|  |  | 223.872 | 0 | 0
|  |  | 223.9488 | 0 | 0
|  |  | 224.0256 | 0 | 0
|  |  | 224.1024 | 0 | 0
|  |  | 224.1792 | 0 | 0
|  |  | 224.256 | 0 | 0
|  |  | 224.3328 | 0 | 0
|  |  | 224.4096 | 0 | 0
|  |  | 224.4864 | 0 | 0
|  |  | 224.5632 | 0 | 0
|  |  | 224.64 | 0 | 0
|  |  | 224.7168 | 0 | 0
|  |  | 224.7936 | 0 | 0
|  |  | 224.8704 | 0 | 0
|  |  | 224.9472 | 0 | 0
|  |  | 225.024 | 0 | 0
|  |  | 225.1008 | 0 | 0
|  |  | 225.1776 | 0 | 0
|  |  | 225.2544 | 0 | 0
|  |  | 225.3312 | 0 | 0
|  |  | 225.408 | 0 | 0
|  |  | 225.4848 | 0 | 0
|  |  | 225.5616 | 0 | 0
|  |  | 225.6384 | 0 | 0
|  |  | 225.7152 | 0 | 0
|  |  | 225.792 | 0 | 0
|  |  | 225.8688 | 0 | 0
|  |  | 225.9456 | 0 | 0
|  |  | 226.0224 | 0 | 0
|  |  | 226.0992 | 0 | 0
|  |  | 226.176 | 0 | 0
|  |  | 226.2528 | 0 | 0
|  |  | 226.3296 | 0 | 0
|  |  | 226.4064 | 0 | 0
|  |  | 226.4832 | 0 | 0
|  |  | 226.56 | 0 | 0
|  |  | 226.6368 | 0 | 0
|  |  | 226.7136 | 0 | 0
|  |  | 226.7904 | 0 | 0
|  |  | 226.8672 | 0 | 0
|  |  | 226.944 | 0 | 0
|  |  | 227.0208 | 0 | 0
|  |  | 227.0976 | 0 | 0
|  |  | 227.1744 | 0 | 0
|  |  | 227.2512 | 0 | 0
|  |  | 227.328 | 0 | 0
|  |  | 227.4048 | 0 | 0
|  |  | 227.4816 | 0 | 0
|  |  | 227.5584 | 0 | 0
|  |  | 227.6352 | 0 | 0
|  |  | 227.712 | 0 | 0
|  |  | 227.7888 | 0 | 0
|  |  | 227.8656 | 0 | 0
|  |  | 227.9424 | 0 | 0
|  |  | 228.0192 | 0 | 0
|  |  | 228.096 | 0 | 0
|  |  | 228.1728 | 0 | 0
|  |  | 228.2496 | 0 | 0
|  |  | 228.3264 | 0 | 0
|  |  | 228.4032 | 0 | 0
|  |  | 228.48 | 0 | 0
|  |  | 228.5568 | 0 | 0
|  |  | 228.6336 | 0 | 0
|  |  | 228.7104 | 0 | 0
|  |  | 228.7872 | 0 | 0
|  |  | 228.864 | 0 | 0
|  |  | 228.9408 | 0 | 0
|  |  | 229.0176 | 0 | 0
|  |  | 229.0944 | 0 | 0
|  |  | 229.1712 | 0 | 0
|  |  | 229.248 | 0 | 0
|  |  | 229.3248 | 0 | 0
|  |  | 229.4016 | 0 | 0
|  |  | 229.4784 | 0 | 0
|  |  | 229.5552 | 0 | 0
|  |  | 229.632 | 0 | 0
|  |  | 229.7088 | 0 | 0
|  |  | 229.7856 | 0 | 0
|  |  | 229.8624 | 0 | 0
|  |  | 229.9392 | 0 | 0
|  |  | 230.016 | 0 | 0
|  |  | 230.0928 | 0 | 0
|  |  | 230.1696 | 0 | 0
|  |  | 230.2464 | 0 | 0
|  |  | 230.3232 | 0 | 0
|  |  | 230.4 | 0 | 0
|  |  | 230.4768 | 0 | 0
|  |  | 230.5536 | 0 | 0
|  |  | 230.6304 | 0 | 0
|  |  | 230.7072 | 0 | 0
|  |  | 230.784 | 0 | 0
|  |  | 230.8608 | 0 | 0
|  |  | 230.9376 | 0 | 0
|  |  | 231.0144 | 0 | 0
|  |  | 231.0912 | 0 | 0
|  |  | 231.168 | 0 | 0
|  |  | 231.2448 | 0 | 0
|  |  | 231.3216 | 0 | 0
|  |  | 231.3984 | 0 | 0
|  |  | 231.4752 | 0 | 0
|  |  | 231.552 | 0 | 0
|  |  | 231.6288 | 0 | 0
|  |  | 231.7056 | 0 | 0
|  |  | 231.7824 | 0 | 0
|  |  | 231.8592 | 0 | 0
|  |  | 231.936 | 0 | 0
|  |  | 232.0128 | 0 | 0
|  |  | 232.0896 | 0 | 0
|  |  | 232.1664 | 0 | 0
|  |  | 232.2432 | 0 | 0
|  |  | 232.32 | 0 | 0
|  |  | 232.3968 | 0 | 0
|  |  | 232.4736 | 0 | 0
|  |  | 232.5504 | 0 | 0
|  |  | 232.6272 | 0 | 0
|  |  | 232.704 | 38.89 | 0.86528
|  |  | 232.7808 | 41.36 | 0.9984
|  |  | 232.8576 | 43.76 | 1.1264
|  |  | 232.9344 | 46.555 | 1.1776
|  |  | 233.0112 | 48.4 | 1.3466
|  |  | 233.088 | 49.95 | 1.5667
|  |  | 233.1648 | 51.995 | 1.664
|  |  | 233.2416 | 53.745 | 1.7613
|  |  | 233.3184 | 55.545 | 1.833
|  |  | 233.3952 | 57.39 | 1.879
|  |  | 233.472 | 58.07 | 2.0685
|  |  | 233.5488 | 58.84 | 2.2016
|  |  | 233.6256 | 59.725 | 2.2579
|  |  | 233.7024 | 60.325 | 2.3194
|  |  | 233.7792 | 60.655 | 2.391
|  |  | 233.856 | 61.375 | 2.3757
|  |  | 233.9328 | 61.605 | 2.4269
|  |  | 234.0096 | 61.665 | 2.4832
|  |  | 234.0864 | 62.385 | 2.4269
|  |  | 234.1632 | 61.58 | 2.729
|  |  | 234.24 | 62.965 | 2.4678
|  |  | 234.3168 | 61.935 | 2.8262
|  |  | 234.3936 | 62.91 | 2.6573
|  |  | 234.4704 | 62.34 | 2.8826
|  |  | 234.5472 | 62.5 | 2.8979
|  |  | 234.624 | 62.855 | 2.816
|  |  | 234.7008 | 62.76 | 2.8826
|  |  | 234.7776 | 62.14 | 3.0771
|  |  | 234.8544 | 62.47 | 2.9798
|  |  | 234.9312 | 63.09 | 2.7904
|  |  | 235.008 | 63.2 | 2.7392
|  |  | 235.0848 | 63.035 | 2.7853
|  |  | 235.1616 | 62.64 | 2.8826
|  |  | 235.2384 | 61.785 | 3.0925
|  |  | 235.3152 | 63.38 | 2.6368
|  |  | 235.392 | 62.19 | 2.8518
|  |  | 235.4688 | 62.19 | 2.7392
|  |  | 235.5456 | 61.25 | 2.8979
|  |  | 235.6224 | 61.935 | 2.5907
|  |  | 235.6992 | 61.56 | 2.5139
|  |  | 235.776 | 60.73 | 2.4576
|  |  | 235.8528 | 59.87 | 2.345
|  |  | 235.9296 | 58.325 | 2.2784
|  |  | 236.0064 | 56.99 | 2.1197
|  |  | 236.0832 | 55.12 | 2.0531
|  |  | 236.16 | 53.71 | 1.9251
|  |  | 236.2368 | 52.63 | 1.7613
|  |  | 236.3136 | 51.125 | 1.705
|  |  | 236.3904 | 49.48 | 1.6435
|  |  | 236.4672 | 48.54 | 1.4234
|  |  | 236.544 | 47.405 | 1.2442
|  |  | 236.6208 | 46.14 | 1.0906
|  |  | 236.6976 | 45 | 0.91648
|  |  | 236.7744 | 44.195 | 0.71168
|  |  | 236.8512 | 43.33 | 0.52736
|  |  | 236.928 | 0 | 0
|  |  | 237.0048 | 0 | 0
|  |  | 237.0816 | 0 | 0
|  |  | 237.1584 | 0 | 0
|  |  | 237.2352 | 0 | 0
|  |  | 237.312 | 0 | 0
|  |  | 237.3888 | 0 | 0
|  |  | 237.4656 | 0 | 0
|  |  | 237.5424 | 0 | 0
|  |  | 237.6192 | 0 | 0
|  |  | 237.696 | 0 | 0
|  |  | 237.7728 | 0 | 0
|  |  | 237.8496 | 0 | 0
|  |  | 237.9264 | 0 | 0
|  |  | 238.0032 | 0 | 0
|  |  | 238.08 | 0 | 0
|  |  | 238.1568 | 0 | 0
|  |  | 238.2336 | 0 | 0
|  |  | 238.3104 | 0 | 0
|  |  | 238.3872 | 0 | 0
|  |  | 238.464 | 0 | 0
|  |  | 238.5408 | 0 | 0
|  |  | 238.6176 | 0 | 0
|  |  | 238.6944 | 0 | 0
|  |  | 238.7712 | 0 | 0
|  |  | 238.848 | 0 | 0
|  |  | 238.9248 | 0 | 0
|  |  | 239.0016 | 0 | 0
|  |  | 239.0784 | 0 | 0
|  |  | 239.1552 | 0 | 0
|  |  | 239.232 | 0 | 0
|  |  | 239.3088 | 0 | 0
|  |  | 239.3856 | 0 | 0
|  |  | 239.4624 | 0 | 0
|  |  | 239.5392 | 0 | 0
|  |  | 239.616 | 0 | 0
|  |  | 239.6928 | 0 | 0
|  |  | 239.7696 | 0 | 0
|  |  | 239.8464 | 0 | 0
|  |  | 239.9232 | 0 | 0
|  |  | 240 | 0 | 0
|  |  | 240.0768 | 0 | 0
|  |  | 240.1536 | 0 | 0
|  |  | 240.2304 | 0 | 0
|  |  | 240.3072 | 0 | 0
|  |  | 240.384 | 0 | 0
|  |  | 240.4608 | 0 | 0
|  |  | 240.5376 | 0 | 0
|  |  | 240.6144 | 0 | 0
|  |  | 240.6912 | 0 | 0
|  |  | 240.768 | 0 | 0
|  |  | 240.8448 | 0 | 0
|  |  | 240.9216 | 0 | 0
|  |  | 240.9984 | 0 | 0
|  |  | 241.0752 | 0 | 0
|  |  | 241.152 | 0 | 0
|  |  | 241.2288 | 0 | 0
|  |  | 241.3056 | 0 | 0
|  |  | 241.3824 | 0 | 0
|  |  | 241.4592 | 0 | 0
|  |  | 241.536 | 0 | 0
|  |  | 241.6128 | 0 | 0
|  |  | 241.6896 | 0 | 0
|  |  | 241.7664 | 0 | 0
|  |  | 241.8432 | 0 | 0
|  |  | 241.92 | 0 | 0
|  |  | 241.9968 | 0 | 0
|  |  | 242.0736 | 0 | 0
|  |  | 242.1504 | 0 | 0
|  |  | 242.2272 | 0 | 0
|  |  | 242.304 | 0 | 0
|  |  | 242.3808 | 0 | 0
|  |  | 242.4576 | 0 | 0
|  |  | 242.5344 | 0 | 0
|  |  | 242.6112 | 0 | 0
|  |  | 242.688 | 0 | 0
|  |  | 242.7648 | 0 | 0
|  |  | 242.8416 | 0 | 0
|  |  | 242.9184 | 0 | 0
|  |  | 242.9952 | 0 | 0
|  |  | 243.072 | 0 | 0
|  |  | 243.1488 | 0 | 0
|  |  | 243.2256 | 0 | 0
|  |  | 243.3024 | 0 | 0
|  |  | 243.3792 | 0 | 0
|  |  | 243.456 | 0 | 0
|  |  | 243.5328 | 0 | 0
|  |  | 243.6096 | 0 | 0
|  |  | 243.6864 | 0 | 0
|  |  | 243.7632 | 0 | 0
|  |  | 243.84 | 0 | 0
|  |  | 243.9168 | 0 | 0
|  |  | 243.9936 | 0 | 0
|  |  | 244.0704 | 0 | 0
|  |  | 244.1472 | 0 | 0
|  |  | 244.224 | 0 | 0
|  |  | 244.3008 | 0 | 0
|  |  | 244.3776 | 0 | 0
|  |  | 244.4544 | 0 | 0
|  |  | 244.5312 | 0 | 0
|  |  | 244.608 | 0 | 0
|  |  | 244.6848 | 0 | 0
|  |  | 244.7616 | 0 | 0
|  |  | 244.8384 | 0 | 0
|  |  | 244.9152 | 0 | 0
|  |  | 244.992 | 0 | 0
|  |  | 245.0688 | 0 | 0
|  |  | 245.1456 | 0 | 0
|  |  | 245.2224 | 0 | 0
|  |  | 245.2992 | 0 | 0
|  |  | 245.376 | 0 | 0
|  |  | 245.4528 | 0 | 0
|  |  | 245.5296 | 0 | 0
|  |  | 245.6064 | 0 | 0
|  |  | 245.6832 | 0 | 0
|  |  | 245.76 | 0 | 0
|  |  | 245.8368 | 0 | 0
|  |  | 245.9136 | 0 | 0
|  |  | 245.9904 | 0 | 0
|  |  | 246.0672 | 0 | 0
|  |  | 246.144 | 0 | 0
|  |  | 246.2208 | 0 | 0
|  |  | 246.2976 | 0 | 0
|  |  | 246.3744 | 0 | 0
|  |  | 246.4512 | 0 | 0
|  |  | 246.528 | 0 | 0
|  |  | 246.6048 | 0 | 0
|  |  | 246.6816 | 0 | 0
|  |  | 246.7584 | 0 | 0
|  |  | 246.8352 | 0 | 0
|  |  | 246.912 | 0 | 0
|  |  | 246.9888 | 0 | 0
|  |  | 247.0656 | 0 | 0
|  |  | 247.1424 | 0 | 0
|  |  | 247.2192 | 0 | 0
|  |  | 247.296 | 0 | 0
|  |  | 247.3728 | 0 | 0
|  |  | 247.4496 | 0 | 0
|  |  | 247.5264 | 0 | 0
|  |  | 247.6032 | 0 | 0
|  |  | 247.68 | 0 | 0
|  |  | 247.7568 | 0 | 0
|  |  | 247.8336 | 0 | 0
|  |  | 247.9104 | 0 | 0
|  |  | 247.9872 | 0 | 0
|  |  | 248.064 | 0 | 0
|  |  | 248.1408 | 0 | 0
|  |  | 248.2176 | 0 | 0
|  |  | 248.2944 | 0 | 0
|  |  | 248.3712 | 0 | 0
|  |  | 248.448 | 0 | 0
|  |  | 248.5248 | 0 | 0
|  |  | 248.6016 | 0 | 0
|  |  | 248.6784 | 0 | 0
|  |  | 248.7552 | 0 | 0
|  |  | 248.832 | 0 | 0
|  |  | 248.9088 | 0 | 0
|  |  | 248.9856 | 0 | 0
|  |  | 249.0624 | 0 | 0
|  |  | 249.1392 | 0 | 0
|  |  | 249.216 | 0 | 0
|  |  | 249.2928 | 0 | 0
|  |  | 249.3696 | 38.245 | 0.9472
|  |  | 249.4464 | 41.425 | 0.9472
|  |  | 249.5232 | 44.185 | 1.0189
|  |  | 249.6 | 46.325 | 1.152
|  |  | 249.6768 | 48.645 | 1.239
|  |  | 249.7536 | 50.325 | 1.4285
|  |  | 249.8304 | 52.52 | 1.5104
|  |  | 249.9072 | 54.56 | 1.5718
|  |  | 249.984 | 55.5 | 1.7869
|  |  | 250.0608 | 56.86 | 1.9251
|  |  | 250.1376 | 58.065 | 2.0275
|  |  | 250.2144 | 59.115 | 2.1299
|  |  | 250.2912 | 60.065 | 2.1811
|  |  | 250.368 | 61.175 | 2.176
|  |  | 250.4448 | 61.405 | 2.2682
|  |  | 250.5216 | 61.33 | 2.3706
|  |  | 250.5984 | 61.285 | 2.4576
|  |  | 250.6752 | 61.595 | 2.4627
|  |  | 250.752 | 62.195 | 2.4269
|  |  | 250.8288 | 61.645 | 2.6726
|  |  | 250.9056 | 62.935 | 2.4627
|  |  | 250.9824 | 61.955 | 2.8211
|  |  | 251.0592 | 63.515 | 2.519
|  |  | 251.136 | 61.755 | 3.0566
|  |  | 251.2128 | 64.06 | 2.5242
|  |  | 251.2896 | 62.725 | 2.857
|  |  | 251.3664 | 62.75 | 2.8774
|  |  | 251.4432 | 62.725 | 2.9082
|  |  | 251.52 | 62.6 | 2.9338
|  |  | 251.5968 | 62.65 | 2.944
|  |  | 251.6736 | 62.92 | 2.8518
|  |  | 251.7504 | 62.74 | 2.8723
|  |  | 251.8272 | 62.18 | 3.031
|  |  | 251.904 | 61.485 | 3.2154
|  |  | 251.9808 | 63.48 | 2.6266
|  |  | 252.0576 | 62.105 | 2.8826
|  |  | 252.1344 | 61.605 | 2.903
|  |  | 252.2112 | 62.285 | 2.6419
|  |  | 252.288 | 60.625 | 3.0003
|  |  | 252.3648 | 61.32 | 2.6112
|  |  | 252.4416 | 60.485 | 2.5754
|  |  | 252.5184 | 59.735 | 2.4166
|  |  | 252.5952 | 58.04 | 2.3654
|  |  | 252.672 | 56.365 | 2.2784
|  |  | 252.7488 | 55.3 | 2.0787
|  |  | 252.8256 | 53.62 | 2.0326
|  |  | 252.9024 | 51.87 | 1.9866
|  |  | 252.9792 | 51.18 | 1.7408
|  |  | 253.056 | 50.38 | 1.536
|  |  | 253.1328 | 48.74 | 1.449
|  |  | 253.2096 | 47.255 | 1.3107
|  |  | 253.2864 | 46.38 | 1.0906
|  |  | 253.3632 | 44.855 | 0.95232
|  |  | 253.44 | 42.385 | 0.9472
|  |  | 253.5168 | 40.325 | 0.88064
|  |  | 253.5936 | 39.49 | 0.63488
|  |  | 253.6704 | 0 | 0
|  |  | 253.7472 | 0 | 0
|  |  | 253.824 | 0 | 0
|  |  | 253.9008 | 0 | 0
|  |  | 253.9776 | 0 | 0
|  |  | 254.0544 | 0 | 0
|  |  | 254.1312 | 0 | 0
|  |  | 254.208 | 0 | 0
|  |  | 254.2848 | 0 | 0
|  |  | 254.3616 | 0 | 0
|  |  | 254.4384 | 0 | 0
|  |  | 254.5152 | 0 | 0
|  |  | 254.592 | 0 | 0
|  |  | 254.6688 | 0 | 0
|  |  | 254.7456 | 0 | 0
|  |  | 254.8224 | 0 | 0
|  |  | 254.8992 | 0 | 0
|  |  | 254.976 | 0 | 0
|  |  | 255.0528 | 0 | 0
|  |  | 255.1296 | 0 | 0
|  |  | 255.2064 | 0 | 0
|  |  | 255.2832 | 0 | 0
|  |  | 255.36 | 0 | 0
|  |  | 255.4368 | 0 | 0
|  |  | 255.5136 | 0 | 0
|  |  | 255.5904 | 0 | 0
|  |  | 255.6672 | 0 | 0
|  |  | 255.744 | 0 | 0
|  |  | 255.8208 | 0 | 0
|  |  | 255.8976 | 0 | 0
|  |  | 255.9744 | 0 | 0
|  |  | 256.0512 | 0 | 0
|  |  | 256.128 | 0 | 0
|  |  | 256.2048 | 0 | 0
|  |  | 256.2816 | 0 | 0
|  |  | 256.3584 | 0 | 0
|  |  | 256.4352 | 0 | 0
|  |  | 256.512 | 0 | 0
|  |  | 256.5888 | 0 | 0
|  |  | 256.6656 | 0 | 0
|  |  | 256.7424 | 0 | 0
|  |  | 256.8192 | 0 | 0
|  |  | 256.896 | 0 | 0
|  |  | 256.9728 | 0 | 0
|  |  | 257.0496 | 0 | 0
|  |  | 257.1264 | 0 | 0
|  |  | 257.2032 | 0 | 0
|  |  | 257.28 | 0 | 0
|  |  | 257.3568 | 0 | 0
|  |  | 257.4336 | 0 | 0
|  |  | 257.5104 | 0 | 0
|  |  | 257.5872 | 0 | 0
|  |  | 257.664 | 0 | 0
|  |  | 257.7408 | 0 | 0
|  |  | 257.8176 | 0 | 0
|  |  | 257.8944 | 0 | 0
|  |  | 257.9712 | 0 | 0
|  |  | 258.048 | 0 | 0
|  |  | 258.1248 | 0 | 0
|  |  | 258.2016 | 0 | 0
|  |  | 258.2784 | 0 | 0
|  |  | 258.3552 | 0 | 0
|  |  | 258.432 | 0 | 0
|  |  | 258.5088 | 0 | 0
|  |  | 258.5856 | 0 | 0
|  |  | 258.6624 | 0 | 0
|  |  | 258.7392 | 0 | 0
|  |  | 258.816 | 0 | 0
|  |  | 258.8928 | 0 | 0
|  |  | 258.9696 | 0 | 0
|  |  | 259.0464 | 0 | 0
|  |  | 259.1232 | 0 | 0
|  |  | 259.2 | 0 | 0
|  |  | 259.2768 | 0 | 0
|  |  | 259.3536 | 0 | 0
|  |  | 259.4304 | 0 | 0
|  |  | 259.5072 | 0 | 0
|  |  | 259.584 | 0 | 0
|  |  | 259.6608 | 0 | 0
|  |  | 259.7376 | 0 | 0
|  |  | 259.8144 | 0 | 0
|  |  | 259.8912 | 0 | 0
|  |  | 259.968 | 0 | 0
|  |  | 260.0448 | 0 | 0
|  |  | 260.1216 | 0 | 0
|  |  | 260.1984 | 0 | 0
|  |  | 260.2752 | 0 | 0
|  |  | 260.352 | 0 | 0
|  |  | 260.4288 | 0 | 0
|  |  | 260.5056 | 0 | 0
|  |  | 260.5824 | 0 | 0
|  |  | 260.6592 | 0 | 0
|  |  | 260.736 | 0 | 0
|  |  | 260.8128 | 0 | 0
|  |  | 260.8896 | 0 | 0
|  |  | 260.9664 | 0 | 0
|  |  | 261.0432 | 0 | 0
|  |  | 261.12 | 0 | 0
|  |  | 261.1968 | 0 | 0
|  |  | 261.2736 | 0 | 0
|  |  | 261.3504 | 0 | 0
|  |  | 261.4272 | 0 | 0
|  |  | 261.504 | 0 | 0
|  |  | 261.5808 | 0 | 0
|  |  | 261.6576 | 0 | 0
|  |  | 261.7344 | 0 | 0
|  |  | 261.8112 | 0 | 0
|  |  | 261.888 | 0 | 0
|  |  | 261.9648 | 0 | 0
|  |  | 262.0416 | 0 | 0
|  |  | 262.1184 | 0 | 0
|  |  | 262.1952 | 0 | 0
|  |  | 262.272 | 0 | 0
|  |  | 262.3488 | 0 | 0
|  |  | 262.4256 | 0 | 0
|  |  | 262.5024 | 0 | 0
|  |  | 262.5792 | 0 | 0
|  |  | 262.656 | 0 | 0
|  |  | 262.7328 | 0 | 0
|  |  | 262.8096 | 0 | 0
|  |  | 262.8864 | 0 | 0
|  |  | 262.9632 | 0 | 0
|  |  | 263.04 | 0 | 0
|  |  | 263.1168 | 0 | 0
|  |  | 263.1936 | 0 | 0
|  |  | 263.2704 | 0 | 0
|  |  | 263.3472 | 0 | 0
|  |  | 263.424 | 0 | 0
|  |  | 263.5008 | 0 | 0
|  |  | 263.5776 | 0 | 0
|  |  | 263.6544 | 0 | 0
|  |  | 263.7312 | 0 | 0
|  |  | 263.808 | 0 | 0
|  |  | 263.8848 | 0 | 0
|  |  | 263.9616 | 0 | 0
|  |  | 264.0384 | 0 | 0
|  |  | 264.1152 | 0 | 0
|  |  | 264.192 | 0 | 0
|  |  | 264.2688 | 0 | 0
|  |  | 264.3456 | 0 | 0
|  |  | 264.4224 | 0 | 0
|  |  | 264.4992 | 0 | 0
|  |  | 264.576 | 0 | 0
|  |  | 264.6528 | 0 | 0
|  |  | 264.7296 | 0 | 0
|  |  | 264.8064 | 0 | 0
|  |  | 264.8832 | 0 | 0
|  |  | 264.96 | 0 | 0
|  |  | 265.0368 | 0 | 0
|  |  | 265.1136 | 0 | 0
|  |  | 265.1904 | 0 | 0
|  |  | 265.2672 | 0 | 0
|  |  | 265.344 | 0 | 0
|  |  | 265.4208 | 0 | 0
|  |  | 265.4976 | 0 | 0
|  |  | 265.5744 | 0 | 0
|  |  | 265.6512 | 0 | 0
|  |  | 265.728 | 0 | 0
|  |  | 265.8048 | 0 | 0
|  |  | 265.8816 | 0 | 0
|  |  | 265.9584 | 0 | 0
|  |  | 266.0352 | 37.625 | 1.0342
|  |  | 266.112 | 41.615 | 0.94208
|  |  | 266.1888 | 44.655 | 0.96768
|  |  | 266.2656 | 46.67 | 1.1008
|  |  | 266.3424 | 48.815 | 1.2083
|  |  | 266.4192 | 50.505 | 1.3875
|  |  | 266.496 | 52.095 | 1.5462
|  |  | 266.5728 | 53.515 | 1.705
|  |  | 266.6496 | 55.025 | 1.833
|  |  | 266.7264 | 57.095 | 1.8534
|  |  | 266.8032 | 58.63 | 1.9149
|  |  | 266.88 | 59.32 | 2.0685
|  |  | 266.9568 | 59.59 | 2.2528
|  |  | 267.0336 | 59.725 | 2.4218
|  |  | 267.1104 | 60.94 | 2.345
|  |  | 267.1872 | 61.855 | 2.2784
|  |  | 267.264 | 61.945 | 2.3296
|  |  | 267.3408 | 61.39 | 2.4986
|  |  | 267.4176 | 61.95 | 2.4781
|  |  | 267.4944 | 62.35 | 2.473
|  |  | 267.5712 | 61.965 | 2.6573
|  |  | 267.648 | 62.84 | 2.56
|  |  | 267.7248 | 62.46 | 2.7597
|  |  | 267.8016 | 62.15 | 2.903
|  |  | 267.8784 | 61.99 | 3.0054
|  |  | 267.9552 | 63.265 | 2.7034
|  |  | 268.032 | 62.34 | 2.9952
|  |  | 268.1088 | 62.17 | 3.0618
|  |  | 268.1856 | 62.235 | 3.0618
|  |  | 268.2624 | 62.695 | 2.9133
|  |  | 268.3392 | 62.815 | 2.8928
|  |  | 268.416 | 62.665 | 2.9389
|  |  | 268.4928 | 62.395 | 2.9747
|  |  | 268.5696 | 62.245 | 2.9542
|  |  | 268.6464 | 62.17 | 2.9338
|  |  | 268.7232 | 62.065 | 2.903
|  |  | 268.8 | 61.84 | 2.8877
|  |  | 268.8768 | 62.76 | 2.5907
|  |  | 268.9536 | 60.8 | 2.985
|  |  | 269.0304 | 61.025 | 2.6778
|  |  | 269.1072 | 60.67 | 2.5395
|  |  | 269.184 | 59.735 | 2.4525
|  |  | 269.2608 | 58.515 | 2.3296
|  |  | 269.3376 | 56.905 | 2.2221
|  |  | 269.4144 | 55.505 | 2.0787
|  |  | 269.4912 | 54.11 | 1.9507
|  |  | 269.568 | 52.835 | 1.8227
|  |  | 269.6448 | 51.56 | 1.7101
|  |  | 269.7216 | 50.365 | 1.5821
|  |  | 269.7984 | 49.29 | 1.4131
|  |  | 269.8752 | 47.855 | 1.2646
|  |  | 269.952 | 46.905 | 1.0752
|  |  | 270.0288 | 45.47 | 0.93696
|  |  | 270.1056 | 43.79 | 0.82944
|  |  | 270.1824 | 41.64 | 0.75264
|  |  | 270.2592 | 39.255 | 0.68608
|  |  | 270.336 | 0 | 0
|  |  | 270.4128 | 0 | 0
|  |  | 270.4896 | 0 | 0
|  |  | 270.5664 | 0 | 0
|  |  | 270.6432 | 0 | 0
|  |  | 270.72 | 0 | 0
|  |  | 270.7968 | 0 | 0
|  |  | 270.8736 | 0 | 0
|  |  | 270.9504 | 0 | 0
|  |  | 271.0272 | 0 | 0
|  |  | 271.104 | 0 | 0
|  |  | 271.1808 | 0 | 0
|  |  | 271.2576 | 0 | 0
|  |  | 271.3344 | 0 | 0
|  |  | 271.4112 | 0 | 0
|  |  | 271.488 | 0 | 0
|  |  | 271.5648 | 0 | 0
|  |  | 271.6416 | 0 | 0
|  |  | 271.7184 | 0 | 0
|  |  | 271.7952 | 0 | 0
|  |  | 271.872 | 0 | 0
|  |  | 271.9488 | 0 | 0
|  |  | 272.0256 | 0 | 0
|  |  | 272.1024 | 0 | 0
|  |  | 272.1792 | 0 | 0
|  |  | 272.256 | 0 | 0
|  |  | 272.3328 | 0 | 0
|  |  | 272.4096 | 0 | 0
|  |  | 272.4864 | 0 | 0
|  |  | 272.5632 | 0 | 0
|  |  | 272.64 | 0 | 0
|  |  | 272.7168 | 0 | 0
|  |  | 272.7936 | 0 | 0
|  |  | 272.8704 | 0 | 0
|  |  | 272.9472 | 0 | 0
|  |  | 273.024 | 0 | 0
|  |  | 273.1008 | 0 | 0
|  |  | 273.1776 | 0 | 0
|  |  | 273.2544 | 0 | 0
|  |  | 273.3312 | 0 | 0
|  |  | 273.408 | 0 | 0
|  |  | 273.4848 | 0 | 0
|  |  | 273.5616 | 0 | 0
|  |  | 273.6384 | 0 | 0
|  |  | 273.7152 | 0 | 0
|  |  | 273.792 | 0 | 0
|  |  | 273.8688 | 0 | 0
|  |  | 273.9456 | 0 | 0
|  |  | 274.0224 | 0 | 0
|  |  | 274.0992 | 0 | 0
|  |  | 274.176 | 0 | 0
|  |  | 274.2528 | 0 | 0
|  |  | 274.3296 | 0 | 0
|  |  | 274.4064 | 0 | 0
|  |  | 274.4832 | 0 | 0
|  |  | 274.56 | 0 | 0
|  |  | 274.6368 | 0 | 0
|  |  | 274.7136 | 0 | 0
|  |  | 274.7904 | 0 | 0
|  |  | 274.8672 | 0 | 0
|  |  | 274.944 | 0 | 0
|  |  | 275.0208 | 0 | 0
|  |  | 275.0976 | 0 | 0
|  |  | 275.1744 | 0 | 0
|  |  | 275.2512 | 0 | 0
|  |  | 275.328 | 0 | 0
|  |  | 275.4048 | 0 | 0
|  |  | 275.4816 | 0 | 0
|  |  | 275.5584 | 0 | 0
|  |  | 275.6352 | 0 | 0
|  |  | 275.712 | 0 | 0
|  |  | 275.7888 | 0 | 0
|  |  | 275.8656 | 0 | 0
|  |  | 275.9424 | 0 | 0
|  |  | 276.0192 | 0 | 0
|  |  | 276.096 | 0 | 0
|  |  | 276.1728 | 0 | 0
|  |  | 276.2496 | 0 | 0
|  |  | 276.3264 | 0 | 0
|  |  | 276.4032 | 0 | 0
|  |  | 276.48 | 0 | 0
|  |  | 276.5568 | 0 | 0
|  |  | 276.6336 | 0 | 0
|  |  | 276.7104 | 0 | 0
|  |  | 276.7872 | 0 | 0
|  |  | 276.864 | 0 | 0
|  |  | 276.9408 | 0 | 0
|  |  | 277.0176 | 0 | 0
|  |  | 277.0944 | 0 | 0
|  |  | 277.1712 | 0 | 0
|  |  | 277.248 | 0 | 0
|  |  | 277.3248 | 0 | 0
|  |  | 277.4016 | 0 | 0
|  |  | 277.4784 | 0 | 0
|  |  | 277.5552 | 0 | 0
|  |  | 277.632 | 0 | 0
|  |  | 277.7088 | 0 | 0
|  |  | 277.7856 | 0 | 0
|  |  | 277.8624 | 0 | 0
|  |  | 277.9392 | 0 | 0
|  |  | 278.016 | 0 | 0
|  |  | 278.0928 | 0 | 0
|  |  | 278.1696 | 0 | 0
|  |  | 278.2464 | 0 | 0
|  |  | 278.3232 | 0 | 0
|  |  | 278.4 | 0 | 0
|  |  | 278.4768 | 0 | 0
|  |  | 278.5536 | 0 | 0
|  |  | 278.6304 | 0 | 0
|  |  | 278.7072 | 0 | 0
|  |  | 278.784 | 0 | 0
|  |  | 278.8608 | 0 | 0
|  |  | 278.9376 | 0 | 0
|  |  | 279.0144 | 0 | 0
|  |  | 279.0912 | 0 | 0
|  |  | 279.168 | 0 | 0
|  |  | 279.2448 | 0 | 0
|  |  | 279.3216 | 0 | 0
|  |  | 279.3984 | 0 | 0
|  |  | 279.4752 | 0 | 0
|  |  | 279.552 | 0 | 0
|  |  | 279.6288 | 0 | 0
|  |  | 279.7056 | 0 | 0
|  |  | 279.7824 | 0 | 0
|  |  | 279.8592 | 0 | 0
|  |  | 279.936 | 0 | 0
|  |  | 280.0128 | 0 | 0
|  |  | 280.0896 | 0 | 0
|  |  | 280.1664 | 0 | 0
|  |  | 280.2432 | 0 | 0
|  |  | 280.32 | 0 | 0
|  |  | 280.3968 | 0 | 0
|  |  | 280.4736 | 0 | 0
|  |  | 280.5504 | 0 | 0
|  |  | 280.6272 | 0 | 0
|  |  | 280.704 | 0 | 0
|  |  | 280.7808 | 0 | 0
|  |  | 280.8576 | 0 | 0
|  |  | 280.9344 | 0 | 0
|  |  | 281.0112 | 0 | 0
|  |  | 281.088 | 0 | 0
|  |  | 281.1648 | 0 | 0
|  |  | 281.2416 | 0 | 0
|  |  | 281.3184 | 0 | 0
|  |  | 281.3952 | 0 | 0
|  |  | 281.472 | 0 | 0
|  |  | 281.5488 | 0 | 0
|  |  | 281.6256 | 0 | 0
|  |  | 281.7024 | 0 | 0
|  |  | 281.7792 | 0 | 0
|  |  | 281.856 | 0 | 0
|  |  | 281.9328 | 0 | 0
|  |  | 282.0096 | 0 | 0
|  |  | 282.0864 | 0 | 0
|  |  | 282.1632 | 0 | 0
|  |  | 282.24 | 0 | 0
|  |  | 282.3168 | 0 | 0
|  |  | 282.3936 | 0 | 0
|  |  | 282.4704 | 0 | 0
|  |  | 282.5472 | 0 | 0
|  |  | 282.624 | 0 | 0
|  |  | 282.7008 | 39.78 | 0.67072
|  |  | 282.7776 | 41.82 | 0.85504
|  |  | 282.8544 | 44.17 | 0.9728
|  |  | 282.9312 | 47.495 | 0.95232
|  |  | 283.008 | 48.695 | 1.1981
|  |  | 283.0848 | 49.875 | 1.4541
|  |  | 283.1616 | 51.55 | 1.6026
|  |  | 283.2384 | 53.65 | 1.6589
|  |  | 283.3152 | 55.265 | 1.7818
|  |  | 283.392 | 56.505 | 1.9354
|  |  | 283.4688 | 58.165 | 1.9712
|  |  | 283.5456 | 59.225 | 2.0582
|  |  | 283.6224 | 59.705 | 2.1914
|  |  | 283.6992 | 60.34 | 2.2477
|  |  | 283.776 | 60.915 | 2.2886
|  |  | 283.8528 | 61.085 | 2.3706
|  |  | 283.9296 | 61.465 | 2.3962
|  |  | 284.0064 | 61.965 | 2.4013
|  |  | 284.0832 | 62.015 | 2.4832
|  |  | 284.16 | 61.87 | 2.6317
|  |  | 284.2368 | 62.47 | 2.5805
|  |  | 284.3136 | 62.29 | 2.7443
|  |  | 284.3904 | 62.16 | 2.8621
|  |  | 284.4672 | 61.785 | 3.0003
|  |  | 284.544 | 63.42 | 2.6266
|  |  | 284.6208 | 62.2 | 2.9952
|  |  | 284.6976 | 62.545 | 2.9491
|  |  | 284.7744 | 62.18 | 3.0925
|  |  | 284.8512 | 62.46 | 3.0106
|  |  | 284.928 | 63.125 | 2.8211
|  |  | 285.0048 | 62.44 | 3.0003
|  |  | 285.0816 | 62.17 | 3.0566
|  |  | 285.1584 | 61.955 | 3.0874
|  |  | 285.2352 | 63.68 | 2.6163
|  |  | 285.312 | 61.775 | 3.0669
|  |  | 285.3888 | 62.305 | 2.8109
|  |  | 285.4656 | 61.645 | 2.9389
|  |  | 285.5424 | 62.155 | 2.6982
|  |  | 285.6192 | 60.515 | 3.0669
|  |  | 285.696 | 61.985 | 2.5037
|  |  | 285.7728 | 60.82 | 2.5446
|  |  | 285.8496 | 60.29 | 2.3654
|  |  | 285.9264 | 59.035 | 2.2579
|  |  | 286.0032 | 56.905 | 2.2579
|  |  | 286.08 | 55.695 | 2.0787
|  |  | 286.1568 | 54.42 | 1.9354
|  |  | 286.2336 | 53.015 | 1.8227
|  |  | 286.3104 | 51.685 | 1.6998
|  |  | 286.3872 | 50 | 1.6486
|  |  | 286.464 | 48.18 | 1.6128
|  |  | 286.5408 | 47.54 | 1.3568
|  |  | 286.6176 | 46.45 | 1.1674
|  |  | 286.6944 | 45.29 | 0.97792
|  |  | 286.7712 | 44.15 | 0.80384
|  |  | 286.848 | 42.14 | 0.72704
|  |  | 286.9248 | 39.82 | 0.65536
|  |  | 287.0016 | 0 | 0
|  |  | 287.0784 | 0 | 0
|  |  | 287.1552 | 0 | 0
|  |  | 287.232 | 0 | 0
|  |  | 287.3088 | 0 | 0
|  |  | 287.3856 | 0 | 0
|  |  | 287.4624 | 0 | 0
|  |  | 287.5392 | 0 | 0
|  |  | 287.616 | 0 | 0
|  |  | 287.6928 | 0 | 0
|  |  | 287.7696 | 0 | 0
|  |  | 287.8464 | 0 | 0
|  |  | 287.9232 | 0 | 0
|  |  | 288 | 0 | 0
|  |  | 288.0768 | 0 | 0
|  |  | 288.1536 | 0 | 0
|  |  | 288.2304 | 0 | 0
|  |  | 288.3072 | 0 | 0
|  |  | 288.384 | 0 | 0
|  |  | 288.4608 | 0 | 0
|  |  | 288.5376 | 0 | 0
|  |  | 288.6144 | 0 | 0
|  |  | 288.6912 | 0 | 0
|  |  | 288.768 | 0 | 0
|  |  | 288.8448 | 0 | 0
|  |  | 288.9216 | 0 | 0
|  |  | 288.9984 | 0 | 0
|  |  | 289.0752 | 0 | 0
|  |  | 289.152 | 0 | 0
|  |  | 289.2288 | 0 | 0
|  |  | 289.3056 | 0 | 0
|  |  | 289.3824 | 0 | 0
|  |  | 289.4592 | 0 | 0
|  |  | 289.536 | 0 | 0
|  |  | 289.6128 | 0 | 0
|  |  | 289.6896 | 0 | 0
|  |  | 289.7664 | 0 | 0
|  |  | 289.8432 | 0 | 0
|  |  | 289.92 | 0 | 0
|  |  | 289.9968 | 0 | 0
|  |  | 290.0736 | 0 | 0
|  |  | 290.1504 | 0 | 0
|  |  | 290.2272 | 0 | 0
|  |  | 290.304 | 0 | 0
|  |  | 290.3808 | 0 | 0
|  |  | 290.4576 | 0 | 0
|  |  | 290.5344 | 0 | 0
|  |  | 290.6112 | 0 | 0
|  |  | 290.688 | 0 | 0
|  |  | 290.7648 | 0 | 0
|  |  | 290.8416 | 0 | 0
|  |  | 290.9184 | 0 | 0
|  |  | 290.9952 | 0 | 0
|  |  | 291.072 | 0 | 0
|  |  | 291.1488 | 0 | 0
|  |  | 291.2256 | 0 | 0
|  |  | 291.3024 | 0 | 0
|  |  | 291.3792 | 0 | 0
|  |  | 291.456 | 0 | 0
|  |  | 291.5328 | 0 | 0
|  |  | 291.6096 | 0 | 0
|  |  | 291.6864 | 0 | 0
|  |  | 291.7632 | 0 | 0
|  |  | 291.84 | 0 | 0
|  |  | 291.9168 | 0 | 0
|  |  | 291.9936 | 0 | 0
|  |  | 292.0704 | 0 | 0
|  |  | 292.1472 | 0 | 0
|  |  | 292.224 | 0 | 0
|  |  | 292.3008 | 0 | 0
|  |  | 292.3776 | 0 | 0
|  |  | 292.4544 | 0 | 0
|  |  | 292.5312 | 0 | 0
|  |  | 292.608 | 0 | 0
|  |  | 292.6848 | 0 | 0
|  |  | 292.7616 | 0 | 0
|  |  | 292.8384 | 0 | 0
|  |  | 292.9152 | 0 | 0
|  |  | 292.992 | 0 | 0
|  |  | 293.0688 | 0 | 0
|  |  | 293.1456 | 0 | 0
|  |  | 293.2224 | 0 | 0
|  |  | 293.2992 | 0 | 0
|  |  | 293.376 | 0 | 0
|  |  | 293.4528 | 0 | 0
|  |  | 293.5296 | 0 | 0
|  |  | 293.6064 | 0 | 0
|  |  | 293.6832 | 0 | 0
|  |  | 293.76 | 0 | 0
|  |  | 293.8368 | 0 | 0
|  |  | 293.9136 | 0 | 0
|  |  | 293.9904 | 0 | 0
|  |  | 294.0672 | 0 | 0
|  |  | 294.144 | 0 | 0
|  |  | 294.2208 | 0 | 0
|  |  | 294.2976 | 0 | 0
|  |  | 294.3744 | 0 | 0
|  |  | 294.4512 | 0 | 0
|  |  | 294.528 | 0 | 0
|  |  | 294.6048 | 0 | 0
|  |  | 294.6816 | 0 | 0
|  |  | 294.7584 | 0 | 0
|  |  | 294.8352 | 0 | 0
|  |  | 294.912 | 0 | 0
|  |  | 294.9888 | 0 | 0
|  |  | 295.0656 | 0 | 0
|  |  | 295.1424 | 0 | 0
|  |  | 295.2192 | 0 | 0
|  |  | 295.296 | 0 | 0
|  |  | 295.3728 | 0 | 0
|  |  | 295.4496 | 0 | 0
|  |  | 295.5264 | 0 | 0
|  |  | 295.6032 | 0 | 0
|  |  | 295.68 | 0 | 0
|  |  | 295.7568 | 0 | 0
|  |  | 295.8336 | 0 | 0
|  |  | 295.9104 | 0 | 0
|  |  | 295.9872 | 0 | 0
|  |  | 296.064 | 0 | 0
|  |  | 296.1408 | 0 | 0
|  |  | 296.2176 | 0 | 0
|  |  | 296.2944 | 0 | 0
|  |  | 296.3712 | 0 | 0
|  |  | 296.448 | 0 | 0
|  |  | 296.5248 | 0 | 0
|  |  | 296.6016 | 0 | 0
|  |  | 296.6784 | 0 | 0
|  |  | 296.7552 | 0 | 0
|  |  | 296.832 | 0 | 0
|  |  | 296.9088 | 0 | 0
|  |  | 296.9856 | 0 | 0
|  |  | 297.0624 | 0 | 0
|  |  | 297.1392 | 0 | 0
|  |  | 297.216 | 0 | 0
|  |  | 297.2928 | 0 | 0
|  |  | 297.3696 | 0 | 0
|  |  | 297.4464 | 0 | 0
|  |  | 297.5232 | 0 | 0
|  |  | 297.6 | 0 | 0
|  |  | 297.6768 | 0 | 0
|  |  | 297.7536 | 0 | 0
|  |  | 297.8304 | 0 | 0
|  |  | 297.9072 | 0 | 0
|  |  | 297.984 | 0 | 0
|  |  | 298.0608 | 0 | 0
|  |  | 298.1376 | 0 | 0
|  |  | 298.2144 | 0 | 0
|  |  | 298.2912 | 0 | 0
|  |  | 298.368 | 0 | 0
|  |  | 298.4448 | 0 | 0
|  |  | 298.5216 | 0 | 0
|  |  | 298.5984 | 0 | 0
|  |  | 298.6752 | 0 | 0
|  |  | 298.752 | 0 | 0
|  |  | 298.8288 | 0 | 0
|  |  | 298.9056 | 0 | 0
|  |  | 298.9824 | 0 | 0
|  |  | 299.0592 | 0 | 0
|  |  | 299.136 | 0 | 0
|  |  | 299.2128 | 0 | 0
|  |  | 299.2896 | 0 | 0
|  |  | 299.3664 | 0 | 0
|  |  | 299.4432 | 42.22 | 0.72704
|  |  | 299.52 | 44.2 | 0.90624
|  |  | 299.5968 | 46.29 | 1.0547
|  |  | 299.6736 | 48.18 | 1.1981
|  |  | 299.7504 | 50.01 | 1.3568
|  |  | 299.8272 | 52.01 | 1.4694
|  |  | 299.904 | 54.13 | 1.5514
|  |  | 299.9808 | 55.625 | 1.6691
|  |  | 300.0576 | 57.14 | 1.7766
|  |  | 300.1344 | 58.625 | 1.8739
|  |  | 300.2112 | 59.755 | 1.9763
|  |  | 300.288 | 59.965 | 2.1606
|  |  | 300.3648 | 60.11 | 2.3194
|  |  | 300.4416 | 60.615 | 2.3757
|  |  | 300.5184 | 61.955 | 2.2477
|  |  | 300.5952 | 62.11 | 2.2989
|  |  | 300.672 | 62.02 | 2.3808
|  |  | 300.7488 | 61.87 | 2.5037
|  |  | 300.8256 | 62.19 | 2.4986
|  |  | 300.9024 | 61.43 | 2.8109
|  |  | 300.9792 | 62.23 | 2.6931
|  |  | 301.056 | 61.935 | 2.8928
|  |  | 301.1328 | 63.785 | 2.5395
|  |  | 301.2096 | 62.48 | 2.9184
|  |  | 301.2864 | 63.005 | 2.8109
|  |  | 301.3632 | 63.285 | 2.7597
|  |  | 301.44 | 62.995 | 2.8416
|  |  | 301.5168 | 62.685 | 2.9235
|  |  | 301.5936 | 62.535 | 2.985
|  |  | 301.6704 | 63.135 | 2.8211
|  |  | 301.7472 | 63.605 | 2.6829
|  |  | 301.824 | 63.155 | 2.7648
|  |  | 301.9008 | 62.415 | 2.9286
|  |  | 301.9776 | 61.785 | 3.0669
|  |  | 302.0544 | 63.085 | 2.647
|  |  | 302.1312 | 61.41 | 3.031
|  |  | 302.208 | 62.655 | 2.5958
|  |  | 302.2848 | 61.195 | 2.7955
|  |  | 302.3616 | 61.07 | 2.4371
|  |  | 302.4384 | 59.96 | 2.4576
|  |  | 302.5152 | 59.38 | 2.4678
|  |  | 302.592 | 58.585 | 2.4115
|  |  | 302.6688 | 57.76 | 2.2528
|  |  | 302.7456 | 56.52 | 2.0787
|  |  | 302.8224 | 54.325 | 2.0685
|  |  | 302.8992 | 53.255 | 1.833
|  |  | 302.976 | 51.67 | 1.7101
|  |  | 303.0528 | 50.205 | 1.6026
|  |  | 303.1296 | 48.915 | 1.4694
|  |  | 303.2064 | 47.4 | 1.3773
|  |  | 303.2832 | 46.505 | 1.193
|  |  | 303.36 | 45.73 | 0.9728
|  |  | 303.4368 | 43.99 | 0.84992
|  |  | 303.5136 | 42.155 | 0.73216
|  |  | 303.5904 | 40.035 | 0.65536
|  |  | 303.6672 | 0 | 0
|  |  | 303.744 | 0 | 0
|  |  | 303.8208 | 0 | 0
|  |  | 303.8976 | 0 | 0
|  |  | 303.9744 | 0 | 0
|  |  | 304.0512 | 0 | 0
|  |  | 304.128 | 0 | 0
|  |  | 304.2048 | 0 | 0
|  |  | 304.2816 | 0 | 0
|  |  | 304.3584 | 0 | 0
|  |  | 304.4352 | 0 | 0
|  |  | 304.512 | 0 | 0
|  |  | 304.5888 | 0 | 0
|  |  | 304.6656 | 0 | 0
|  |  | 304.7424 | 0 | 0
|  |  | 304.8192 | 0 | 0
|  |  | 304.896 | 0 | 0
|  |  | 304.9728 | 0 | 0
|  |  | 305.0496 | 0 | 0
|  |  | 305.1264 | 0 | 0
|  |  | 305.2032 | 0 | 0
|  |  | 305.28 | 0 | 0
|  |  | 305.3568 | 0 | 0
|  |  | 305.4336 | 0 | 0
|  |  | 305.5104 | 0 | 0
|  |  | 305.5872 | 0 | 0
|  |  | 305.664 | 0 | 0
|  |  | 305.7408 | 0 | 0
|  |  | 305.8176 | 0 | 0
|  |  | 305.8944 | 0 | 0
|  |  | 305.9712 | 0 | 0
|  |  | 306.048 | 0 | 0
|  |  | 306.1248 | 0 | 0
|  |  | 306.2016 | 0 | 0
|  |  | 306.2784 | 0 | 0
|  |  | 306.3552 | 0 | 0
|  |  | 306.432 | 0 | 0
|  |  | 306.5088 | 0 | 0
|  |  | 306.5856 | 0 | 0
|  |  | 306.6624 | 0 | 0
|  |  | 306.7392 | 0 | 0
|  |  | 306.816 | 0 | 0
|  |  | 306.8928 | 0 | 0
|  |  | 306.9696 | 0 | 0
|  |  | 307.0464 | 0 | 0
|  |  | 307.1232 | 0 | 0
|  |  | 307.2 | 0 | 0
|  |  | 307.2768 | 0 | 0
|  |  | 307.3536 | 0 | 0
|  |  | 307.4304 | 0 | 0
|  |  | 307.5072 | 0 | 0
|  |  | 307.584 | 0 | 0
|  |  | 307.6608 | 0 | 0
|  |  | 307.7376 | 0 | 0
|  |  | 307.8144 | 0 | 0
|  |  | 307.8912 | 0 | 0
|  |  | 307.968 | 0 | 0
|  |  | 308.0448 | 0 | 0
|  |  | 308.1216 | 0 | 0
|  |  | 308.1984 | 0 | 0
|  |  | 308.2752 | 0 | 0
|  |  | 308.352 | 0 | 0
|  |  | 308.4288 | 0 | 0
|  |  | 308.5056 | 0 | 0
|  |  | 308.5824 | 0 | 0
|  |  | 308.6592 | 0 | 0
|  |  | 308.736 | 0 | 0
|  |  | 308.8128 | 0 | 0
|  |  | 308.8896 | 0 | 0
|  |  | 308.9664 | 0 | 0
|  |  | 309.0432 | 0 | 0
|  |  | 309.12 | 0 | 0
|  |  | 309.1968 | 0 | 0
|  |  | 309.2736 | 0 | 0
|  |  | 309.3504 | 0 | 0
|  |  | 309.4272 | 0 | 0
|  |  | 309.504 | 0 | 0
|  |  | 309.5808 | 0 | 0
|  |  | 309.6576 | 0 | 0
|  |  | 309.7344 | 0 | 0
|  |  | 309.8112 | 0 | 0
|  |  | 309.888 | 0 | 0
|  |  | 309.9648 | 0 | 0
|  |  | 310.0416 | 0 | 0
|  |  | 310.1184 | 0 | 0
|  |  | 310.1952 | 0 | 0
|  |  | 310.272 | 0 | 0
|  |  | 310.3488 | 0 | 0
|  |  | 310.4256 | 0 | 0
|  |  | 310.5024 | 0 | 0
|  |  | 310.5792 | 0 | 0
|  |  | 310.656 | 0 | 0
|  |  | 310.7328 | 0 | 0
|  |  | 310.8096 | 0 | 0
|  |  | 310.8864 | 0 | 0
|  |  | 310.9632 | 0 | 0
|  |  | 311.04 | 0 | 0
|  |  | 311.1168 | 0 | 0
|  |  | 311.1936 | 0 | 0
|  |  | 311.2704 | 0 | 0
|  |  | 311.3472 | 0 | 0
|  |  | 311.424 | 0 | 0
|  |  | 311.5008 | 0 | 0
|  |  | 311.5776 | 0 | 0
|  |  | 311.6544 | 0 | 0
|  |  | 311.7312 | 0 | 0
|  |  | 311.808 | 0 | 0
|  |  | 311.8848 | 0 | 0
|  |  | 311.9616 | 0 | 0
|  |  | 312.0384 | 0 | 0
|  |  | 312.1152 | 0 | 0
|  |  | 312.192 | 0 | 0
|  |  | 312.2688 | 0 | 0
|  |  | 312.3456 | 0 | 0
|  |  | 312.4224 | 0 | 0
|  |  | 312.4992 | 0 | 0
|  |  | 312.576 | 0 | 0
|  |  | 312.6528 | 0 | 0
|  |  | 312.7296 | 0 | 0
|  |  | 312.8064 | 0 | 0
|  |  | 312.8832 | 0 | 0
|  |  | 312.96 | 0 | 0
|  |  | 313.0368 | 0 | 0
|  |  | 313.1136 | 0 | 0
|  |  | 313.1904 | 0 | 0
|  |  | 313.2672 | 0 | 0
|  |  | 313.344 | 0 | 0
|  |  | 313.4208 | 0 | 0
|  |  | 313.4976 | 0 | 0
|  |  | 313.5744 | 0 | 0
|  |  | 313.6512 | 0 | 0
|  |  | 313.728 | 0 | 0
|  |  | 313.8048 | 0 | 0
|  |  | 313.8816 | 0 | 0
|  |  | 313.9584 | 0 | 0
|  |  | 314.0352 | 0 | 0
|  |  | 314.112 | 0 | 0
|  |  | 314.1888 | 0 | 0
|  |  | 314.2656 | 0 | 0
|  |  | 314.3424 | 0 | 0
|  |  | 314.4192 | 0 | 0
|  |  | 314.496 | 0 | 0
|  |  | 314.5728 | 0 | 0
|  |  | 314.6496 | 0 | 0
|  |  | 314.7264 | 0 | 0
|  |  | 314.8032 | 0 | 0
|  |  | 314.88 | 0 | 0
|  |  | 314.9568 | 0 | 0
|  |  | 315.0336 | 0 | 0
|  |  | 315.1104 | 0 | 0
|  |  | 315.1872 | 0 | 0
|  |  | 315.264 | 0 | 0
|  |  | 315.3408 | 0 | 0
|  |  | 315.4176 | 0 | 0
|  |  | 315.4944 | 0 | 0
|  |  | 315.5712 | 0 | 0
|  |  | 315.648 | 0 | 0
|  |  | 315.7248 | 0 | 0
|  |  | 315.8016 | 0 | 0
|  |  | 315.8784 | 0 | 0
|  |  | 315.9552 | 0 | 0
|  |  | 316.032 | 0 | 0
|  |  | 316.1088 | 40.475 | 0.94208
|  |  | 316.1856 | 43.795 | 0.93696
|  |  | 316.2624 | 45.375 | 1.1622
|  |  | 316.3392 | 47.585 | 1.28
|  |  | 316.416 | 50.435 | 1.2851
|  |  | 316.4928 | 52.595 | 1.3722
|  |  | 316.5696 | 53.725 | 1.5872
|  |  | 316.6464 | 55.16 | 1.7306
|  |  | 316.7232 | 56.325 | 1.9046
|  |  | 316.8 | 57.505 | 2.0429
|  |  | 316.8768 | 58.495 | 2.1606
|  |  | 316.9536 | 59.515 | 2.217
|  |  | 317.0304 | 60.13 | 2.2835
|  |  | 317.1072 | 60.52 | 2.3757
|  |  | 317.184 | 60.785 | 2.432
|  |  | 317.2608 | 61.485 | 2.3962
|  |  | 317.3376 | 61.59 | 2.4627
|  |  | 317.4144 | 61.635 | 2.5344
|  |  | 317.4912 | 62.435 | 2.4576
|  |  | 317.568 | 61.88 | 2.7034
|  |  | 317.6448 | 62.635 | 2.6112
|  |  | 317.7216 | 61.84 | 2.944
|  |  | 317.7984 | 63.885 | 2.5139
|  |  | 317.8752 | 62.275 | 2.8928
|  |  | 317.952 | 62.17 | 2.8211
|  |  | 318.0288 | 62.33 | 2.8262
|  |  | 318.1056 | 62.545 | 2.9286
|  |  | 318.1824 | 62.995 | 2.944
|  |  | 318.2592 | 63.305 | 2.9082
|  |  | 318.336 | 63.165 | 2.944
|  |  | 318.4128 | 62.31 | 3.1437
|  |  | 318.4896 | 61.73 | 3.2154
|  |  | 318.5664 | 63.255 | 2.7187
|  |  | 318.6432 | 61.475 | 3.1437
|  |  | 318.72 | 63.435 | 2.5702
|  |  | 318.7968 | 61.43 | 3.0208
|  |  | 318.8736 | 62.775 | 2.5958
|  |  | 318.9504 | 61.38 | 2.8467
|  |  | 319.0272 | 61.885 | 2.56
|  |  | 319.104 | 61.11 | 2.473
|  |  | 319.1808 | 59.67 | 2.4218
|  |  | 319.2576 | 58.16 | 2.4064
|  |  | 319.3344 | 57.11 | 2.2733
|  |  | 319.4112 | 56.02 | 2.0992
|  |  | 319.488 | 54.99 | 1.9302
|  |  | 319.5648 | 53.575 | 1.8381
|  |  | 319.6416 | 51.565 | 1.8227
|  |  | 319.7184 | 50.075 | 1.7152
|  |  | 319.7952 | 49.01 | 1.5155
|  |  | 319.872 | 47.65 | 1.3773
|  |  | 319.9488 | 45.7 | 1.3414
|  |  | 320.0256 | 43.965 | 1.2544
|  |  | 320.1024 | 42.825 | 1.0445
|  |  | 320.1792 | 41.725 | 0.83968
|  |  | 320.256 | 40.4 | 0.64512
|  |  | 320.3328 | 0 | 0
|  |  | 320.4096 | 0 | 0
|  |  | 320.4864 | 0 | 0
|  |  | 320.5632 | 0 | 0
|  |  | 320.64 | 0 | 0
|  |  | 320.7168 | 0 | 0
|  |  | 320.7936 | 0 | 0
|  |  | 320.8704 | 0 | 0
|  |  | 320.9472 | 0 | 0
|  |  | 321.024 | 0 | 0
|  |  | 321.1008 | 0 | 0
|  |  | 321.1776 | 0 | 0
|  |  | 321.2544 | 0 | 0
|  |  | 321.3312 | 0 | 0
|  |  | 321.408 | 0 | 0
|  |  | 321.4848 | 0 | 0
|  |  | 321.5616 | 0 | 0
|  |  | 321.6384 | 0 | 0
|  |  | 321.7152 | 0 | 0
|  |  | 321.792 | 0 | 0
|  |  | 321.8688 | 0 | 0
|  |  | 321.9456 | 0 | 0
|  |  | 322.0224 | 0 | 0
|  |  | 322.0992 | 0 | 0
|  |  | 322.176 | 0 | 0
|  |  | 322.2528 | 0 | 0
|  |  | 322.3296 | 0 | 0
|  |  | 322.4064 | 0 | 0
|  |  | 322.4832 | 0 | 0
|  |  | 322.56 | 0 | 0
|  |  | 322.6368 | 0 | 0
|  |  | 322.7136 | 0 | 0
|  |  | 322.7904 | 0 | 0
|  |  | 322.8672 | 0 | 0
|  |  | 322.944 | 0 | 0
|  |  | 323.0208 | 0 | 0
|  |  | 323.0976 | 0 | 0
|  |  | 323.1744 | 0 | 0
|  |  | 323.2512 | 0 | 0
|  |  | 323.328 | 0 | 0
|  |  | 323.4048 | 0 | 0
|  |  | 323.4816 | 0 | 0
|  |  | 323.5584 | 0 | 0
|  |  | 323.6352 | 0 | 0
|  |  | 323.712 | 0 | 0
|  |  | 323.7888 | 0 | 0
|  |  | 323.8656 | 0 | 0
|  |  | 323.9424 | 0 | 0
|  |  | 324.0192 | 0 | 0
|  |  | 324.096 | 0 | 0
|  |  | 324.1728 | 0 | 0
|  |  | 324.2496 | 0 | 0
|  |  | 324.3264 | 0 | 0
|  |  | 324.4032 | 0 | 0
|  |  | 324.48 | 0 | 0
|  |  | 324.5568 | 0 | 0
|  |  | 324.6336 | 0 | 0
|  |  | 324.7104 | 0 | 0
|  |  | 324.7872 | 0 | 0
|  |  | 324.864 | 0 | 0
|  |  | 324.9408 | 0 | 0
|  |  | 325.0176 | 0 | 0
|  |  | 325.0944 | 0 | 0
|  |  | 325.1712 | 0 | 0
|  |  | 325.248 | 0 | 0
|  |  | 325.3248 | 0 | 0
|  |  | 325.4016 | 0 | 0
|  |  | 325.4784 | 0 | 0
|  |  | 325.5552 | 0 | 0
|  |  | 325.632 | 0 | 0
|  |  | 325.7088 | 0 | 0
|  |  | 325.7856 | 0 | 0
|  |  | 325.8624 | 0 | 0
|  |  | 325.9392 | 0 | 0
|  |  | 326.016 | 0 | 0
|  |  | 326.0928 | 0 | 0
|  |  | 326.1696 | 0 | 0
|  |  | 326.2464 | 0 | 0
|  |  | 326.3232 | 0 | 0
|  |  | 326.4 | 0 | 0
|  |  | 326.4768 | 0 | 0
|  |  | 326.5536 | 0 | 0
|  |  | 326.6304 | 0 | 0
|  |  | 326.7072 | 0 | 0
|  |  | 326.784 | 0 | 0
|  |  | 326.8608 | 0 | 0
|  |  | 326.9376 | 0 | 0
|  |  | 327.0144 | 0 | 0
|  |  | 327.0912 | 0 | 0
|  |  | 327.168 | 0 | 0
|  |  | 327.2448 | 0 | 0
|  |  | 327.3216 | 0 | 0
|  |  | 327.3984 | 0 | 0
|  |  | 327.4752 | 0 | 0
|  |  | 327.552 | 0 | 0
|  |  | 327.6288 | 0 | 0
|  |  | 327.7056 | 0 | 0
|  |  | 327.7824 | 0 | 0
|  |  | 327.8592 | 0 | 0
|  |  | 327.936 | 0 | 0
|  |  | 328.0128 | 0 | 0
|  |  | 328.0896 | 0 | 0
|  |  | 328.1664 | 0 | 0
|  |  | 328.2432 | 0 | 0
|  |  | 328.32 | 0 | 0
|  |  | 328.3968 | 0 | 0
|  |  | 328.4736 | 0 | 0
|  |  | 328.5504 | 0 | 0
|  |  | 328.6272 | 0 | 0
|  |  | 328.704 | 0 | 0
|  |  | 328.7808 | 0 | 0
|  |  | 328.8576 | 0 | 0
|  |  | 328.9344 | 0 | 0
|  |  | 329.0112 | 0 | 0
|  |  | 329.088 | 0 | 0
|  |  | 329.1648 | 0 | 0
|  |  | 329.2416 | 0 | 0
|  |  | 329.3184 | 0 | 0
|  |  | 329.3952 | 0 | 0
|  |  | 329.472 | 0 | 0
|  |  | 329.5488 | 0 | 0
|  |  | 329.6256 | 0 | 0
|  |  | 329.7024 | 0 | 0
|  |  | 329.7792 | 0 | 0
|  |  | 329.856 | 0 | 0
|  |  | 329.9328 | 0 | 0
|  |  | 330.0096 | 0 | 0
|  |  | 330.0864 | 0 | 0
|  |  | 330.1632 | 0 | 0
|  |  | 330.24 | 0 | 0
|  |  | 330.3168 | 0 | 0
|  |  | 330.3936 | 0 | 0
|  |  | 330.4704 | 0 | 0
|  |  | 330.5472 | 0 | 0
|  |  | 330.624 | 0 | 0
|  |  | 330.7008 | 0 | 0
|  |  | 330.7776 | 0 | 0
|  |  | 330.8544 | 0 | 0
|  |  | 330.9312 | 0 | 0
|  |  | 331.008 | 0 | 0
|  |  | 331.0848 | 0 | 0
|  |  | 331.1616 | 0 | 0
|  |  | 331.2384 | 0 | 0
|  |  | 331.3152 | 0 | 0
|  |  | 331.392 | 0 | 0
|  |  | 331.4688 | 0 | 0
|  |  | 331.5456 | 0 | 0
|  |  | 331.6224 | 0 | 0
|  |  | 331.6992 | 0 | 0
|  |  | 331.776 | 0 | 0
|  |  | 331.8528 | 0 | 0
|  |  | 331.9296 | 0 | 0
|  |  | 332.0064 | 0 | 0
|  |  | 332.0832 | 0 | 0
|  |  | 332.16 | 0 | 0
|  |  | 332.2368 | 0 | 0
|  |  | 332.3136 | 0 | 0
|  |  | 332.3904 | 0 | 0
|  |  | 332.4672 | 0 | 0
|  |  | 332.544 | 0 | 0
|  |  | 332.6208 | 0 | 0
|  |  | 332.6976 | 0 | 0
|  |  | 332.7744 | 42.01 | 0.72704
|  |  | 332.8512 | 43.635 | 0.94208
|  |  | 332.928 | 45.345 | 1.1366
|  |  | 333.0048 | 47.48 | 1.239
|  |  | 333.0816 | 49.505 | 1.3517
|  |  | 333.1584 | 51.42 | 1.5002
|  |  | 333.2352 | 52.455 | 1.7562
|  |  | 333.312 | 54.095 | 1.8432
|  |  | 333.3888 | 55.93 | 1.8944
|  |  | 333.4656 | 57.085 | 2.0275
|  |  | 333.5424 | 58.195 | 2.1248
|  |  | 333.6192 | 59.01 | 2.2272
|  |  | 333.696 | 60.295 | 2.2374
|  |  | 333.7728 | 61.02 | 2.2784
|  |  | 333.8496 | 61.915 | 2.2477
|  |  | 333.9264 | 62.12 | 2.3142
|  |  | 334.0032 | 61.615 | 2.519
|  |  | 334.08 | 61.935 | 2.473
|  |  | 334.1568 | 62.16 | 2.4986
|  |  | 334.2336 | 61.795 | 2.7034
|  |  | 334.3104 | 63.07 | 2.5088
|  |  | 334.3872 | 63.035 | 2.6061
|  |  | 334.464 | 62.92 | 2.6931
|  |  | 334.5408 | 62.685 | 2.8211
|  |  | 334.6176 | 62.93 | 2.8416
|  |  | 334.6944 | 63.335 | 2.7597
|  |  | 334.7712 | 63.125 | 2.8058
|  |  | 334.848 | 63.005 | 2.857
|  |  | 334.9248 | 62.92 | 2.8979
|  |  | 335.0016 | 62.545 | 2.9901
|  |  | 335.0784 | 61.935 | 3.1693
|  |  | 335.1552 | 63.85 | 2.6522
|  |  | 335.232 | 62.245 | 3.0003
|  |  | 335.3088 | 62.055 | 2.9798
|  |  | 335.3856 | 61.665 | 3.0618
|  |  | 335.4624 | 63.15 | 2.6112
|  |  | 335.5392 | 61.515 | 2.9184
|  |  | 335.616 | 61.87 | 2.6778
|  |  | 335.6928 | 61.655 | 2.6061
|  |  | 335.7696 | 60.735 | 2.6266
|  |  | 335.8464 | 59.815 | 2.5651
|  |  | 335.9232 | 58.83 | 2.4064
|  |  | 336 | 57.55 | 2.2477
|  |  | 336.0768 | 55.385 | 2.263
|  |  | 336.1536 | 53.91 | 2.1402
|  |  | 336.2304 | 52.83 | 1.9712
|  |  | 336.3072 | 51.585 | 1.8381
|  |  | 336.384 | 50.435 | 1.6691
|  |  | 336.4608 | 49.065 | 1.5514
|  |  | 336.5376 | 47.98 | 1.367
|  |  | 336.6144 | 46.735 | 1.2032
|  |  | 336.6912 | 45.09 | 1.111
|  |  | 336.768 | 42.97 | 1.0496
|  |  | 336.8448 | 40.715 | 0.9984
|  |  | 336.9216 | 39.295 | 0.80896
|  |  | 336.9984 | 0 | 0
|  |  | 337.0752 | 0 | 0
|  |  | 337.152 | 0 | 0
|  |  | 337.2288 | 0 | 0
|  |  | 337.3056 | 0 | 0
|  |  | 337.3824 | 0 | 0
|  |  | 337.4592 | 0 | 0
|  |  | 337.536 | 0 | 0
|  |  | 337.6128 | 0 | 0
|  |  | 337.6896 | 0 | 0
|  |  | 337.7664 | 0 | 0
|  |  | 337.8432 | 0 | 0
|  |  | 337.92 | 0 | 0
|  |  | 337.9968 | 0 | 0
|  |  | 338.0736 | 0 | 0
|  |  | 338.1504 | 0 | 0
|  |  | 338.2272 | 0 | 0
|  |  | 338.304 | 0 | 0
|  |  | 338.3808 | 0 | 0
|  |  | 338.4576 | 0 | 0
|  |  | 338.5344 | 0 | 0
|  |  | 338.6112 | 0 | 0
|  |  | 338.688 | 0 | 0
|  |  | 338.7648 | 0 | 0
|  |  | 338.8416 | 0 | 0
|  |  | 338.9184 | 0 | 0
|  |  | 338.9952 | 0 | 0
|  |  | 339.072 | 0 | 0
|  |  | 339.1488 | 0 | 0
|  |  | 339.2256 | 0 | 0
|  |  | 339.3024 | 0 | 0
|  |  | 339.3792 | 0 | 0
|  |  | 339.456 | 0 | 0
|  |  | 339.5328 | 0 | 0
|  |  | 339.6096 | 0 | 0
|  |  | 339.6864 | 0 | 0
|  |  | 339.7632 | 0 | 0
|  |  | 339.84 | 0 | 0
|  |  | 339.9168 | 0 | 0
|  |  | 339.9936 | 0 | 0
|  |  | 340.0704 | 0 | 0
|  |  | 340.1472 | 0 | 0
|  |  | 340.224 | 0 | 0
|  |  | 340.3008 | 0 | 0
|  |  | 340.3776 | 0 | 0
|  |  | 340.4544 | 0 | 0
|  |  | 340.5312 | 0 | 0
|  |  | 340.608 | 0 | 0
|  |  | 340.6848 | 0 | 0
|  |  | 340.7616 | 0 | 0
|  |  | 340.8384 | 0 | 0
|  |  | 340.9152 | 0 | 0
|  |  | 340.992 | 0 | 0
|  |  | 341.0688 | 0 | 0
|  |  | 341.1456 | 0 | 0
|  |  | 341.2224 | 0 | 0
|  |  | 341.2992 | 0 | 0
|  |  | 341.376 | 0 | 0
|  |  | 341.4528 | 0 | 0
|  |  | 341.5296 | 0 | 0
|  |  | 341.6064 | 0 | 0
|  |  | 341.6832 | 0 | 0
|  |  | 341.76 | 0 | 0
|  |  | 341.8368 | 0 | 0
|  |  | 341.9136 | 0 | 0
|  |  | 341.9904 | 0 | 0
|  |  | 342.0672 | 0 | 0
|  |  | 342.144 | 0 | 0
|  |  | 342.2208 | 0 | 0
|  |  | 342.2976 | 0 | 0
|  |  | 342.3744 | 0 | 0
|  |  | 342.4512 | 0 | 0
|  |  | 342.528 | 0 | 0
|  |  | 342.6048 | 0 | 0
|  |  | 342.6816 | 0 | 0
|  |  | 342.7584 | 0 | 0
|  |  | 342.8352 | 0 | 0
|  |  | 342.912 | 0 | 0
|  |  | 342.9888 | 0 | 0
|  |  | 343.0656 | 0 | 0
|  |  | 343.1424 | 0 | 0
|  |  | 343.2192 | 0 | 0
|  |  | 343.296 | 0 | 0
|  |  | 343.3728 | 0 | 0
|  |  | 343.4496 | 0 | 0
|  |  | 343.5264 | 0 | 0
|  |  | 343.6032 | 0 | 0
|  |  | 343.68 | 0 | 0
|  |  | 343.7568 | 0 | 0
|  |  | 343.8336 | 0 | 0
|  |  | 343.9104 | 0 | 0
|  |  | 343.9872 | 0 | 0
|  |  | 344.064 | 0 | 0
|  |  | 344.1408 | 0 | 0
|  |  | 344.2176 | 0 | 0
|  |  | 344.2944 | 0 | 0
|  |  | 344.3712 | 0 | 0
|  |  | 344.448 | 0 | 0
|  |  | 344.5248 | 0 | 0
|  |  | 344.6016 | 0 | 0
|  |  | 344.6784 | 0 | 0
|  |  | 344.7552 | 0 | 0
|  |  | 344.832 | 0 | 0
|  |  | 344.9088 | 0 | 0
|  |  | 344.9856 | 0 | 0
|  |  | 345.0624 | 0 | 0
|  |  | 345.1392 | 0 | 0
|  |  | 345.216 | 0 | 0
|  |  | 345.2928 | 0 | 0
|  |  | 345.3696 | 0 | 0
|  |  | 345.4464 | 0 | 0
|  |  | 345.5232 | 0 | 0
|  |  | 345.6 | 0 | 0
|  |  | 345.6768 | 0 | 0
|  |  | 345.7536 | 0 | 0
|  |  | 345.8304 | 0 | 0
|  |  | 345.9072 | 0 | 0
|  |  | 345.984 | 0 | 0
|  |  | 346.0608 | 0 | 0
|  |  | 346.1376 | 0 | 0
|  |  | 346.2144 | 0 | 0
|  |  | 346.2912 | 0 | 0
|  |  | 346.368 | 0 | 0
|  |  | 346.4448 | 0 | 0
|  |  | 346.5216 | 0 | 0
|  |  | 346.5984 | 0 | 0
|  |  | 346.6752 | 0 | 0
|  |  | 346.752 | 0 | 0
|  |  | 346.8288 | 0 | 0
|  |  | 346.9056 | 0 | 0
|  |  | 346.9824 | 0 | 0
|  |  | 347.0592 | 0 | 0
|  |  | 347.136 | 0 | 0
|  |  | 347.2128 | 0 | 0
|  |  | 347.2896 | 0 | 0
|  |  | 347.3664 | 0 | 0
|  |  | 347.4432 | 0 | 0
|  |  | 347.52 | 0 | 0
|  |  | 347.5968 | 0 | 0
|  |  | 347.6736 | 0 | 0
|  |  | 347.7504 | 0 | 0
|  |  | 347.8272 | 0 | 0
|  |  | 347.904 | 0 | 0
|  |  | 347.9808 | 0 | 0
|  |  | 348.0576 | 0 | 0
|  |  | 348.1344 | 0 | 0
|  |  | 348.2112 | 0 | 0
|  |  | 348.288 | 0 | 0
|  |  | 348.3648 | 0 | 0
|  |  | 348.4416 | 0 | 0
|  |  | 348.5184 | 0 | 0
|  |  | 348.5952 | 0 | 0
|  |  | 348.672 | 0 | 0
|  |  | 348.7488 | 0 | 0
|  |  | 348.8256 | 0 | 0
|  |  | 348.9024 | 0 | 0
|  |  | 348.9792 | 0 | 0
|  |  | 349.056 | 0 | 0
|  |  | 349.1328 | 0 | 0
|  |  | 349.2096 | 0 | 0
|  |  | 349.2864 | 0 | 0
|  |  | 349.3632 | 0 | 0
|  |  | 349.44 | 41.425 | 0.7424
|  |  | 349.5168 | 44.375 | 0.80384
|  |  | 349.5936 | 46.175 | 0.9728
|  |  | 349.6704 | 47.88 | 1.1571
|  |  | 349.7472 | 49.535 | 1.3363
|  |  | 349.824 | 51.23 | 1.4899
|  |  | 349.9008 | 52.9 | 1.623
|  |  | 349.9776 | 54.615 | 1.7357
|  |  | 350.0544 | 55.85 | 1.9046
|  |  | 350.1312 | 57.34 | 2.0019
|  |  | 350.208 | 58.61 | 2.0941
|  |  | 350.2848 | 59.465 | 2.176
|  |  | 350.3616 | 59.985 | 2.2784
|  |  | 350.4384 | 60.785 | 2.2835
|  |  | 350.5152 | 61.265 | 2.3347
|  |  | 350.592 | 61.36 | 2.4013
|  |  | 350.6688 | 61.575 | 2.432
|  |  | 350.7456 | 61.96 | 2.4422
|  |  | 350.8224 | 62.61 | 2.4115
|  |  | 350.8992 | 61.905 | 2.6675
|  |  | 350.976 | 62.655 | 2.56
|  |  | 351.0528 | 62.845 | 2.6061
|  |  | 351.1296 | 63.1 | 2.6214
|  |  | 351.2064 | 62.77 | 2.7853
|  |  | 351.2832 | 62.705 | 2.8877
|  |  | 351.36 | 62.425 | 2.9901
|  |  | 351.4368 | 62.675 | 2.9338
|  |  | 351.5136 | 62.565 | 2.9901
|  |  | 351.5904 | 62.695 | 2.9491
|  |  | 351.6672 | 62.575 | 2.9696
|  |  | 351.744 | 62.405 | 3.0003
|  |  | 351.8208 | 62.33 | 3.0054
|  |  | 351.8976 | 62.33 | 2.9645
|  |  | 351.9744 | 62.59 | 2.8467
|  |  | 352.0512 | 62.265 | 2.8774
|  |  | 352.128 | 61.22 | 3.113
|  |  | 352.2048 | 62.42 | 2.6778
|  |  | 352.2816 | 60.98 | 2.9338
|  |  | 352.3584 | 60.87 | 2.7802
|  |  | 352.4352 | 61.105 | 2.5651
|  |  | 352.512 | 60.705 | 2.4115
|  |  | 352.5888 | 59.265 | 2.3501
|  |  | 352.6656 | 57.905 | 2.2067
|  |  | 352.7424 | 56.44 | 2.0685
|  |  | 352.8192 | 54.73 | 2.007
|  |  | 352.896 | 53.37 | 1.8995
|  |  | 352.9728 | 52.325 | 1.7408
|  |  | 353.0496 | 50.68 | 1.6794
|  |  | 353.1264 | 49.62 | 1.4899
|  |  | 353.2032 | 48.165 | 1.367
|  |  | 353.28 | 46.5 | 1.3005
|  |  | 353.3568 | 45.155 | 1.1469
|  |  | 353.4336 | 44.235 | 0.91136
|  |  | 353.5104 | 42.815 | 0.768
|  |  | 353.5872 | 41.03 | 0.65024
|  |  | 353.664 | 0 | 0
|  |  | 353.7408 | 0 | 0
|  |  | 353.8176 | 0 | 0
|  |  | 353.8944 | 0 | 0
|  |  | 353.9712 | 0 | 0
|  |  | 354.048 | 0 | 0
|  |  | 354.1248 | 0 | 0
|  |  | 354.2016 | 0 | 0
|  |  | 354.2784 | 0 | 0
|  |  | 354.3552 | 0 | 0
|  |  | 354.432 | 0 | 0
|  |  | 354.5088 | 0 | 0
|  |  | 354.5856 | 0 | 0
|  |  | 354.6624 | 0 | 0
|  |  | 354.7392 | 0 | 0
|  |  | 354.816 | 0 | 0
|  |  | 354.8928 | 0 | 0
|  |  | 354.9696 | 0 | 0
|  |  | 355.0464 | 0 | 0
|  |  | 355.1232 | 0 | 0
|  |  | 355.2 | 0 | 0
|  |  | 355.2768 | 0 | 0
|  |  | 355.3536 | 0 | 0
|  |  | 355.4304 | 0 | 0
|  |  | 355.5072 | 0 | 0
|  |  | 355.584 | 0 | 0
|  |  | 355.6608 | 0 | 0
|  |  | 355.7376 | 0 | 0
|  |  | 355.8144 | 0 | 0
|  |  | 355.8912 | 0 | 0
|  |  | 355.968 | 0 | 0
|  |  | 356.0448 | 0 | 0
|  |  | 356.1216 | 0 | 0
|  |  | 356.1984 | 0 | 0
|  |  | 356.2752 | 0 | 0
|  |  | 356.352 | 0 | 0
|  |  | 356.4288 | 0 | 0
|  |  | 356.5056 | 0 | 0
|  |  | 356.5824 | 0 | 0
|  |  | 356.6592 | 0 | 0
|  |  | 356.736 | 0 | 0
|  |  | 356.8128 | 0 | 0
|  |  | 356.8896 | 0 | 0
|  |  | 356.9664 | 0 | 0
|  |  | 357.0432 | 0 | 0
|  |  | 357.12 | 0 | 0
|  |  | 357.1968 | 0 | 0
|  |  | 357.2736 | 0 | 0
|  |  | 357.3504 | 0 | 0
|  |  | 357.4272 | 0 | 0
|  |  | 357.504 | 0 | 0
|  |  | 357.5808 | 0 | 0
|  |  | 357.6576 | 0 | 0
|  |  | 357.7344 | 0 | 0
|  |  | 357.8112 | 0 | 0
|  |  | 357.888 | 0 | 0
|  |  | 357.9648 | 0 | 0
|  |  | 358.0416 | 0 | 0
|  |  | 358.1184 | 0 | 0
|  |  | 358.1952 | 0 | 0
|  |  | 358.272 | 0 | 0
|  |  | 358.3488 | 0 | 0
|  |  | 358.4256 | 0 | 0
|  |  | 358.5024 | 0 | 0
|  |  | 358.5792 | 0 | 0
|  |  | 358.656 | 0 | 0
|  |  | 358.7328 | 0 | 0
|  |  | 358.8096 | 0 | 0
|  |  | 358.8864 | 0 | 0
|  |  | 358.9632 | 0 | 0
|  |  | 359.04 | 0 | 0
|  |  | 359.1168 | 0 | 0
|  |  | 359.1936 | 0 | 0
|  |  | 359.2704 | 0 | 0
|  |  | 359.3472 | 0 | 0
|  |  | 359.424 | 0 | 0
|  |  | 359.5008 | 0 | 0
|  |  | 359.5776 | 0 | 0
|  |  | 359.6544 | 0 | 0
|  |  | 359.7312 | 0 | 0
|  |  | 359.808 | 0 | 0
|  |  | 359.8848 | 0 | 0
|  |  | 359.9616 | 0 | 0
|  |  | 360.0384 | 0 | 0
|  |  | 360.1152 | 0 | 0
|  |  | 360.192 | 0 | 0
|  |  | 360.2688 | 0 | 0
|  |  | 360.3456 | 0 | 0
|  |  | 360.4224 | 0 | 0
|  |  | 360.4992 | 0 | 0
|  |  | 360.576 | 0 | 0
|  |  | 360.6528 | 0 | 0
|  |  | 360.7296 | 0 | 0
|  |  | 360.8064 | 0 | 0
|  |  | 360.8832 | 0 | 0
|  |  | 360.96 | 0 | 0
|  |  | 361.0368 | 0 | 0
|  |  | 361.1136 | 0 | 0
|  |  | 361.1904 | 0 | 0
|  |  | 361.2672 | 0 | 0
|  |  | 361.344 | 0 | 0
|  |  | 361.4208 | 0 | 0
|  |  | 361.4976 | 0 | 0
|  |  | 361.5744 | 0 | 0
|  |  | 361.6512 | 0 | 0
|  |  | 361.728 | 0 | 0
|  |  | 361.8048 | 0 | 0
|  |  | 361.8816 | 0 | 0
|  |  | 361.9584 | 0 | 0
|  |  | 362.0352 | 0 | 0
|  |  | 362.112 | 0 | 0
|  |  | 362.1888 | 0 | 0
|  |  | 362.2656 | 0 | 0
|  |  | 362.3424 | 0 | 0
|  |  | 362.4192 | 0 | 0
|  |  | 362.496 | 0 | 0
|  |  | 362.5728 | 0 | 0
|  |  | 362.6496 | 0 | 0
|  |  | 362.7264 | 0 | 0
|  |  | 362.8032 | 0 | 0
|  |  | 362.88 | 0 | 0
|  |  | 362.9568 | 0 | 0
|  |  | 363.0336 | 0 | 0
|  |  | 363.1104 | 0 | 0
|  |  | 363.1872 | 0 | 0
|  |  | 363.264 | 0 | 0
|  |  | 363.3408 | 0 | 0
|  |  | 363.4176 | 0 | 0
|  |  | 363.4944 | 0 | 0
|  |  | 363.5712 | 0 | 0
|  |  | 363.648 | 0 | 0
|  |  | 363.7248 | 0 | 0
|  |  | 363.8016 | 0 | 0
|  |  | 363.8784 | 0 | 0
|  |  | 363.9552 | 0 | 0
|  |  | 364.032 | 0 | 0
|  |  | 364.1088 | 0 | 0
|  |  | 364.1856 | 0 | 0
|  |  | 364.2624 | 0 | 0
|  |  | 364.3392 | 0 | 0
|  |  | 364.416 | 0 | 0
|  |  | 364.4928 | 0 | 0
|  |  | 364.5696 | 0 | 0
|  |  | 364.6464 | 0 | 0
|  |  | 364.7232 | 0 | 0
|  |  | 364.8 | 0 | 0
|  |  | 364.8768 | 0 | 0
|  |  | 364.9536 | 0 | 0
|  |  | 365.0304 | 0 | 0
|  |  | 365.1072 | 0 | 0
|  |  | 365.184 | 0 | 0
|  |  | 365.2608 | 0 | 0
|  |  | 365.3376 | 0 | 0
|  |  | 365.4144 | 0 | 0
|  |  | 365.4912 | 0 | 0
|  |  | 365.568 | 0 | 0
|  |  | 365.6448 | 0 | 0
|  |  | 365.7216 | 0 | 0
|  |  | 365.7984 | 0 | 0
|  |  | 365.8752 | 0 | 0
|  |  | 365.952 | 0 | 0
|  |  | 366.0288 | 0 | 0
|  |  | 366.1056 | 40.23 | 0.90624
|  |  | 366.1824 | 42.835 | 1.0035
|  |  | 366.2592 | 45.87 | 1.0189
|  |  | 366.336 | 48.07 | 1.1366
|  |  | 366.4128 | 49.02 | 1.4234
|  |  | 366.4896 | 50.965 | 1.5462
|  |  | 366.5664 | 52.975 | 1.6435
|  |  | 366.6432 | 54.38 | 1.7869
|  |  | 366.72 | 55.78 | 1.8739
|  |  | 366.7968 | 56.85 | 2.0275
|  |  | 366.8736 | 58.11 | 2.0992
|  |  | 366.9504 | 58.715 | 2.2323
|  |  | 367.0272 | 59.545 | 2.2938
|  |  | 367.104 | 60.295 | 2.3347
|  |  | 367.1808 | 61.025 | 2.3501
|  |  | 367.2576 | 62.255 | 2.2426
|  |  | 367.3344 | 60.305 | 2.8365
|  |  | 367.4112 | 61.58 | 2.5395
|  |  | 367.488 | 62.11 | 2.5395
|  |  | 367.5648 | 61.59 | 2.816
|  |  | 367.6416 | 63.07 | 2.5344
|  |  | 367.7184 | 62.49 | 2.7699
|  |  | 367.7952 | 62.59 | 2.8006
|  |  | 367.872 | 62.665 | 2.8262
|  |  | 367.9488 | 62.45 | 2.944
|  |  | 368.0256 | 62.62 | 2.9133
|  |  | 368.1024 | 62.265 | 3.0054
|  |  | 368.1792 | 62.47 | 2.985
|  |  | 368.256 | 63.08 | 2.8467
|  |  | 368.3328 | 63.185 | 2.8723
|  |  | 368.4096 | 62.995 | 2.9235
|  |  | 368.4864 | 62.575 | 3.0106
|  |  | 368.5632 | 62 | 3.1539
|  |  | 368.64 | 61.74 | 3.1642
|  |  | 368.7168 | 63.2 | 2.6778
|  |  | 368.7936 | 61.86 | 2.9491
|  |  | 368.8704 | 62.84 | 2.6112
|  |  | 368.9472 | 60.94 | 3.031
|  |  | 369.024 | 61.955 | 2.5958
|  |  | 369.1008 | 61.47 | 2.5498
|  |  | 369.1776 | 60.835 | 2.4678
|  |  | 369.2544 | 59.715 | 2.3859
|  |  | 369.3312 | 58.555 | 2.2067
|  |  | 369.408 | 56.595 | 2.135
|  |  | 369.4848 | 54.415 | 2.1197
|  |  | 369.5616 | 52.685 | 2.0531
|  |  | 369.6384 | 51.71 | 1.8637
|  |  | 369.7152 | 50.66 | 1.6947
|  |  | 369.792 | 49.37 | 1.5462
|  |  | 369.8688 | 48.095 | 1.4029
|  |  | 369.9456 | 47.245 | 1.193
|  |  | 370.0224 | 46.29 | 0.9984
|  |  | 370.0992 | 44.825 | 0.86016
|  |  | 370.176 | 43.66 | 0.68096
|  |  | 370.2528 | 41.555 | 0.60928
|  |  | 370.3296 | 0 | 0
|  |  | 370.4064 | 0 | 0
|  |  | 370.4832 | 0 | 0
|  |  | 370.56 | 0 | 0
|  |  | 370.6368 | 0 | 0
|  |  | 370.7136 | 0 | 0
|  |  | 370.7904 | 0 | 0
|  |  | 370.8672 | 0 | 0
|  |  | 370.944 | 0 | 0
|  |  | 371.0208 | 0 | 0
|  |  | 371.0976 | 0 | 0
|  |  | 371.1744 | 0 | 0
|  |  | 371.2512 | 0 | 0
|  |  | 371.328 | 0 | 0
|  |  | 371.4048 | 0 | 0
|  |  | 371.4816 | 0 | 0
|  |  | 371.5584 | 0 | 0
|  |  | 371.6352 | 0 | 0
|  |  | 371.712 | 0 | 0
|  |  | 371.7888 | 0 | 0
|  |  | 371.8656 | 0 | 0
|  |  | 371.9424 | 0 | 0
|  |  | 372.0192 | 0 | 0
|  |  | 372.096 | 0 | 0
|  |  | 372.1728 | 0 | 0
|  |  | 372.2496 | 0 | 0
|  |  | 372.3264 | 0 | 0
|  |  | 372.4032 | 0 | 0
|  |  | 372.48 | 0 | 0
|  |  | 372.5568 | 0 | 0
|  |  | 372.6336 | 0 | 0
|  |  | 372.7104 | 0 | 0
|  |  | 372.7872 | 0 | 0
|  |  | 372.864 | 0 | 0
|  |  | 372.9408 | 0 | 0
|  |  | 373.0176 | 0 | 0
|  |  | 373.0944 | 0 | 0
|  |  | 373.1712 | 0 | 0
|  |  | 373.248 | 0 | 0
|  |  | 373.3248 | 0 | 0
|  |  | 373.4016 | 0 | 0
|  |  | 373.4784 | 0 | 0
|  |  | 373.5552 | 0 | 0
|  |  | 373.632 | 0 | 0
|  |  | 373.7088 | 0 | 0
|  |  | 373.7856 | 0 | 0
|  |  | 373.8624 | 0 | 0
|  |  | 373.9392 | 0 | 0
|  |  | 374.016 | 0 | 0
|  |  | 374.0928 | 0 | 0
|  |  | 374.1696 | 0 | 0
|  |  | 374.2464 | 0 | 0
|  |  | 374.3232 | 0 | 0
|  |  | 374.4 | 0 | 0
|  |  | 374.4768 | 0 | 0
|  |  | 374.5536 | 0 | 0
|  |  | 374.6304 | 0 | 0
|  |  | 374.7072 | 0 | 0
|  |  | 374.784 | 0 | 0
|  |  | 374.8608 | 0 | 0
|  |  | 374.9376 | 0 | 0
|  |  | 375.0144 | 0 | 0
|  |  | 375.0912 | 0 | 0
|  |  | 375.168 | 0 | 0
|  |  | 375.2448 | 0 | 0
|  |  | 375.3216 | 0 | 0
|  |  | 375.3984 | 0 | 0
|  |  | 375.4752 | 0 | 0
|  |  | 375.552 | 0 | 0
|  |  | 375.6288 | 0 | 0
|  |  | 375.7056 | 0 | 0
|  |  | 375.7824 | 0 | 0
|  |  | 375.8592 | 0 | 0
|  |  | 375.936 | 0 | 0
|  |  | 376.0128 | 0 | 0
|  |  | 376.0896 | 0 | 0
|  |  | 376.1664 | 0 | 0
|  |  | 376.2432 | 0 | 0
|  |  | 376.32 | 0 | 0
|  |  | 376.3968 | 0 | 0
|  |  | 376.4736 | 0 | 0
|  |  | 376.5504 | 0 | 0
|  |  | 376.6272 | 0 | 0
|  |  | 376.704 | 0 | 0
|  |  | 376.7808 | 0 | 0
|  |  | 376.8576 | 0 | 0
|  |  | 376.9344 | 0 | 0
|  |  | 377.0112 | 0 | 0
|  |  | 377.088 | 0 | 0
|  |  | 377.1648 | 0 | 0
|  |  | 377.2416 | 0 | 0
|  |  | 377.3184 | 0 | 0
|  |  | 377.3952 | 0 | 0
|  |  | 377.472 | 0 | 0
|  |  | 377.5488 | 0 | 0
|  |  | 377.6256 | 0 | 0
|  |  | 377.7024 | 0 | 0
|  |  | 377.7792 | 0 | 0
|  |  | 377.856 | 0 | 0
|  |  | 377.9328 | 0 | 0
|  |  | 378.0096 | 0 | 0
|  |  | 378.0864 | 0 | 0
|  |  | 378.1632 | 0 | 0
|  |  | 378.24 | 0 | 0
|  |  | 378.3168 | 0 | 0
|  |  | 378.3936 | 0 | 0
|  |  | 378.4704 | 0 | 0
|  |  | 378.5472 | 0 | 0
|  |  | 378.624 | 0 | 0
|  |  | 378.7008 | 0 | 0
|  |  | 378.7776 | 0 | 0
|  |  | 378.8544 | 0 | 0
|  |  | 378.9312 | 0 | 0
|  |  | 379.008 | 0 | 0
|  |  | 379.0848 | 0 | 0
|  |  | 379.1616 | 0 | 0
|  |  | 379.2384 | 0 | 0
|  |  | 379.3152 | 0 | 0
|  |  | 379.392 | 0 | 0
|  |  | 379.4688 | 0 | 0
|  |  | 379.5456 | 0 | 0
|  |  | 379.6224 | 0 | 0
|  |  | 379.6992 | 0 | 0
|  |  | 379.776 | 0 | 0
|  |  | 379.8528 | 0 | 0
|  |  | 379.9296 | 0 | 0
|  |  | 380.0064 | 0 | 0
|  |  | 380.0832 | 0 | 0
|  |  | 380.16 | 0 | 0
|  |  | 380.2368 | 0 | 0
|  |  | 380.3136 | 0 | 0
|  |  | 380.3904 | 0 | 0
|  |  | 380.4672 | 0 | 0
|  |  | 380.544 | 0 | 0
|  |  | 380.6208 | 0 | 0
|  |  | 380.6976 | 0 | 0
|  |  | 380.7744 | 0 | 0
|  |  | 380.8512 | 0 | 0
|  |  | 380.928 | 0 | 0
|  |  | 381.0048 | 0 | 0
|  |  | 381.0816 | 0 | 0
|  |  | 381.1584 | 0 | 0
|  |  | 381.2352 | 0 | 0
|  |  | 381.312 | 0 | 0
|  |  | 381.3888 | 0 | 0
|  |  | 381.4656 | 0 | 0
|  |  | 381.5424 | 0 | 0
|  |  | 381.6192 | 0 | 0
|  |  | 381.696 | 0 | 0
|  |  | 381.7728 | 0 | 0
|  |  | 381.8496 | 0 | 0
|  |  | 381.9264 | 0 | 0
|  |  | 382.0032 | 0 | 0
|  |  | 382.08 | 0 | 0
|  |  | 382.1568 | 0 | 0
|  |  | 382.2336 | 0 | 0
|  |  | 382.3104 | 0 | 0
|  |  | 382.3872 | 0 | 0
|  |  | 382.464 | 0 | 0
|  |  | 382.5408 | 0 | 0
|  |  | 382.6176 | 0 | 0
|  |  | 382.6944 | 0 | 0
|  |  | 382.7712 | 40.22 | 0.8448
|  |  | 382.848 | 43.565 | 0.84992
|  |  | 382.9248 | 45.205 | 1.0394
|  |  | 383.0016 | 46.8 | 1.2749
|  |  | 383.0784 | 49.045 | 1.3824
|  |  | 383.1552 | 50.785 | 1.5309
|  |  | 383.232 | 52.56 | 1.6589
|  |  | 383.3088 | 54.225 | 1.7818
|  |  | 383.3856 | 55.73 | 1.8739
|  |  | 383.4624 | 56.695 | 2.0531
|  |  | 383.5392 | 57.855 | 2.1453
|  |  | 383.616 | 59.16 | 2.176
|  |  | 383.6928 | 59.965 | 2.2374
|  |  | 383.7696 | 60.8 | 2.2682
|  |  | 383.8464 | 61.2 | 2.3347
|  |  | 383.9232 | 61.565 | 2.3706
|  |  | 384 | 62.065 | 2.3603
|  |  | 384.0768 | 61.11 | 2.6982
|  |  | 384.1536 | 62.125 | 2.5293
|  |  | 384.2304 | 62.245 | 2.601
|  |  | 384.3072 | 61.905 | 2.8006
|  |  | 384.384 | 63.305 | 2.5395
|  |  | 384.4608 | 62.545 | 2.8006
|  |  | 384.5376 | 63.035 | 2.7494
|  |  | 384.6144 | 63.1 | 2.7904
|  |  | 384.6912 | 62.685 | 2.9389
|  |  | 384.768 | 62.275 | 3.072
|  |  | 384.8448 | 62.15 | 3.1027
|  |  | 384.9216 | 62.415 | 3.0413
|  |  | 384.9984 | 62.215 | 3.1078
|  |  | 385.0752 | 63.135 | 2.8314
|  |  | 385.152 | 63.575 | 2.7136
|  |  | 385.2288 | 62.78 | 2.9082
|  |  | 385.3056 | 62.35 | 2.9747
|  |  | 385.3824 | 62.01 | 3.0003
|  |  | 385.4592 | 61.485 | 3.0566
|  |  | 385.536 | 62.655 | 2.647
|  |  | 385.6128 | 60.96 | 2.985
|  |  | 385.6896 | 61.3 | 2.729
|  |  | 385.7664 | 60.83 | 2.6829
|  |  | 385.8432 | 60.09 | 2.6112
|  |  | 385.92 | 58.83 | 2.5293
|  |  | 385.9968 | 57.39 | 2.391
|  |  | 386.0736 | 56.295 | 2.1811
|  |  | 386.1504 | 55.175 | 1.9866
|  |  | 386.2272 | 54.02 | 1.8227
|  |  | 386.304 | 53.025 | 1.6589
|  |  | 386.3808 | 51.61 | 1.5514
|  |  | 386.4576 | 50.23 | 1.4387
|  |  | 386.5344 | 48.42 | 1.3824
|  |  | 386.6112 | 46.785 | 1.2902
|  |  | 386.688 | 45.3 | 1.1776
|  |  | 386.7648 | 43.78 | 1.0496
|  |  | 386.8416 | 42.16 | 0.88576
|  |  | 386.9184 | 40.15 | 0.77824
|  |  | 386.9952 | 0 | 0
|  |  | 387.072 | 0 | 0
|  |  | 387.1488 | 0 | 0
|  |  | 387.2256 | 0 | 0

## Settings
XLPRO ™ Version 4.06 |  |  | BW-2 | DEFAULT
Operation Settings Area |  |  | 3500 | 3500 | Max. Freq. (Hz)
|  |  | 20 | 20 | Initial Delay (ms)
Units |  |  | 7 | 7 | End Delay (ms)
Radiation: |  | R | 100 | 100 | Width (ms)
/Time: |  | /min | 7 | 7 | Pulses in window
Scale: |  |  | 3 | 3 | Initial Interval (ms)
Pulse Rate: |  | 15-9-6-5 | 20 | 20 | Noise threshold (kV)
Mode |  |  | 0.9 | 0.9 | Upper threshold
Units Display: |  | Rescale | 0.7 | 0.7 | Lower threshold
Cell Update: |  | Incremental | 3 | 3 | FT-resolution control
Auto Zero: |  | Disabled |  | INSTRUMENT PARAMETERS
Capture Mode: |  | Auto | 0.09838 | 0.09838 | 9000 Sample Interval (ms)
Screen Tracking: |  | Yes | 38.08 | 38.08 | 9000 Input Filter (us) |  |  | 4081/4082
kVp |  |  | 37.3 | 37.3 | 37.3 | 37.3 | Sample Interval (us)
Config: | kVSDmA | Dose+Time+kV | 2.9764 | 2.9764 | 2.9764 | 2.9764 | Input Filter -3dB (kHz)
Scale: |  | 9095 | 16.2 | 16.2 | 16.2 | 16.2 | Output Filter (us)
Start delay (ms): |  | 0 | 4 | 4 | 2.6 | 2.6 | Averaging Intervals
Width (ms): |  | 5000 | 0.75 | 0.75 | Width fraction of kVp
Filtration (mm-Al): |  | 2.5 | 0.17 | 0.17 | Duplicate width (ms)
Dental threshold: |  | 0 | 0.0768 | 0.0768 | 9095 Sample Interval (ms)
Device Index |  | 0
Monitor Configuration |  |  |  |  | COEFFICIENTS FOR PRACTICAL kV
Serial No: |  | N/A | Power | Values
ROM Ver: |  | N/A | Min. kV | 20 | 36 | 150 | 161
Detector: |  | N/A | 0 | -18.6452 | -0.017472 | -0.017472 | 0
Chamb Kq: |  | N/A | 1 | 0.85068 | 1.0308e-05 | 1.0308e-05 | 0
Atm Corr: |  | Automatic | 2 | -0.019681 | 2.3082e-05 | 2.3082e-05 | 0
Temp (°C): |  | 22 | 3 | 0.00025663 | -1.662e-07 | -1.662e-07 | 0
Press (kPa): |  | 101.3 | 4 | -2.1424e-06 | 4.3106e-10 | 4.3106e-10 | 0
|  |  | 5 | 0 | 0 | 0 | 0
|  |  | EXP? | 1 | 0 | 0 | 0
NON-LINEAR CORRECTION FACTORS
|  |  | 1030 | 2700 | 0.12
|  |  | 0.343 | -0.569 | 0.2809 | -0.0331
|  |  | 2.9804 | -5.4582 | 4.6956 | -1.4557