## Instructions
This spreadsheet is provided to assist you in collecting data for quality control testing
of a mammographic x-ray unit.
The tab following this Instruction sheet is the sheet on which you will import the
generator test data from your Accu-Gold system. To simplify the analysis, it is
recommended that the template as presented by followed as closely as possible. If
a generator does not have one of the indicated settings and that row is not used,
do not delete the row from the spreadsheet; rather, just clear the cells of that row
of any entries.
The cells on the Data sheet are color coded as follows: |  | Sheet
Green | cells are for data that is collected from AG2 software. | AG2 Data
Yellow | cells are where you enter information specific to your testing session | AG2 Data
The Report sheet is calculated and completed automatically. Except for pasting the
waveform, your intervention should not be required.
The Windows snipping tool can be used to copy a screen shot of the waveform.
Prior to measuring set the unit of measure.
This can be done by accessing the "unit" settings on the "List" tab.
Getting started
Open Accu-Gold 2 and activate "Companion Mode"
If the software is already open you can just click the "Toggle" button
Accu-Gold 2 will automatically re-size the template and application so they can be viewed
simultaneously.
The starting cell "E3" will automatically be selected.
On Accu-Gold 2 with Accu-Gold format selected click the "Get Cell" button
and check the auto-send box.
You can press the "Start" button and begin measuring when the software
is in the "Ready" state.
Exposures can be redone by selecting any cell in column "E" of the exposure line then using
the "Get Cell" button on AccuGold 2.
Please Note ---
These templates are provided with the compliments of Radcal Corporation as samples and
are available for modification by the user. Certain cells are locked and protected but may be
unlocked without a password. (Use Tools/Protection/Unprotect Workbook/[enter])
Radcal assumes no responsibility or liability for any unintentional errors.
It remains the user's responsibility to validate the calculations and to confirm suitability for the particular purpose.

## AG2 Data
Machine Settings |  |  |  |  | Measured Data
Set kV | mAs | Target/Filter | Focal Spot | Index | Date | Time |  | kVp |  | Dose |  | Dose Rate |  | HVL | [14 empty] | Average kV | [11 empty] | Filtration
24 | 100 | Mo/Mo | L | 1 | 2022-04-25 15:56 | 1.0709 | s | 23.8 | kV | 653.6 | mR | 610.4 | mR/s | 0.36 | mm | 1 | USB | 1 | 46.7 | Hz | Standard | Multi-Sensor | Diagnostic | General | 3.227 | mm | 120.2 | kV | 119.8 | kV | 1 | 0.03113 | s | 120.2 | kV | 43-0001 |  |  |  |  | 3.227 | mm | 32.6 | kV
25 | 100 | Mo/Mo | L | 2 | 2022-04-25 15:57 | 0.9994 | s | 24.9 | kV | 764.9 | mR | 765.4 | mR/s | 0.37 | mm | 1 | USB | 1 | 7.8 | Hz | Standard | Multi-Sensor | Diagnostic | General | 3.227 | mm | 120.2 | kV | 119.8 | kV | 1 | 0.1385 | s | 120.2 | kV | 43-0001 | 2572 | mR/s | 4.698 | mm | 3.227 | mm | 119.8 | kV
26 | 100 | Mo/Mo | L | 3 | 2022-04-25 15:57 | 0.9992 | s | 26 | kV | 878.4 | mR | 879.1 | mR/s | 0.38 | mm | 1 | USB | 1 | 6.7 | Hz | Standard | Multi-Sensor | Diagnostic | General | 3.227 | mm | 120.2 | kV | 119.8 | kV | 1 | 0.1591 | s | 120.2 | kV | 43-0001 | 2572 | mR/s | 4.698 | mm | 3.227 | mm | 119.8 | kV
27 | 100 | Mo/Mo | L | 4 | 2022-04-25 15:57 | 0.9993 | s | 27 | kV | 995.3 | mR | 996 | mR/s | 0.39 | mm | 1 | USB | 1 | 6.1 | Hz | Standard | Multi-Sensor | Diagnostic | General | 3.227 | mm | 120.2 | kV | 119.8 | kV | 1 | 0.1739 | s | 120.2 | kV | 43-0001 | 2572 | mR/s | 4.698 | mm | 3.227 | mm | 119.8 | kV
28 | 100 | Mo/Mo | L | 5 | 2022-04-25 15:57 | 0.9987 | s | 28.1 | kV | 1116 | mR | 1117 | mR/s | 0.4 | mm | 1 | USB | 1 | 12.8 | Hz | Standard | Multi-Sensor | Diagnostic | General | 3.227 | mm | 120.2 | kV | 119.8 | kV | 1 | 0.08817 | s | 120.2 | kV | 43-0001 |  |  |  |  | 3.227 | mm
28 | 100 | Mo/Mo | L | 6 | 2022-04-25 15:57 | 0.9988 | s | 28.1 | kV | 1116 | mR | 1117 | mR/s | 0.3 | mm | 1 | USB | 1 | 8.1 | Hz | Standard | Multi-Sensor | Diagnostic | General | 3.227 | mm | 120.2 | kV | 119.8 | kV | 1 | 0.1327 | s | 120.2 | kV | 43-0001 | 2572 | mR/s | 4.698 | mm | 3.227 | mm | 119.8 | kV
28 | 100 | Mo/Mo | L | 7 | 2022-04-25 15:57 | 0.9989 | s | 28.1 | kV | 1115 | mR | 1117 | mR/s | 0.3 | mm | 1 | USB | 1 | 6 | Hz | Standard | Multi-Sensor | Diagnostic | General | 3.227 | mm | 120.2 | kV | 119.8 | kV | 1 | 0.1761 | s | 120.2 | kV | 43-0001 | 2572 | mR/s | 4.698 | mm | 3.227 | mm | 119.8 | kV
28 | 100 | Mo/Mo | L | 8 | 2022-04-25 15:57 | 0.9985 | s | 28.1 | kV | 1116 | mR | 1118 | mR/s | 0.3 | mm | 1 | USB | 1 | 7.5 | Hz | Standard | Multi-Sensor | Diagnostic | General | 3.227 | mm | 120.2 | kV | 119.8 | kV | 1 | 0.1437 | s | 120.2 | kV | 43-0001 | 2572 | mR/s | 4.698 | mm | 3.227 | mm | 119.8 | kV
29 | 100 | Mo/Mo | L | 9 | 2022-04-25 15:57 | 0.9976 | s | 29.3 | kV | 1240 | mR | 1243 | mR/s | 0.31 | mm | 1 | USB | 1 | 5.7 | Hz | Standard | Multi-Sensor | Diagnostic | General | 3.227 | mm | 120.2 | kV | 119.8 | kV | 1 | 0.1843 | s | 120.2 | kV | 43-0001 | 2572 | mR/s | 4.698 | mm | 3.227 | mm | 119.8 | kV
30 | 100 | Mo/Mo | L | 10 | 2022-04-25 15:57 | 0.998 | s | 30.4 | kV | 1369 | mR | 1371 | mR/s | 0.32 | mm | 1 | USB | 1 | 6.5 | Hz | Standard | Multi-Sensor | Diagnostic | General | 3.227 | mm | 120.2 | kV | 119.8 | kV | 1 | 0.1628 | s | 120.2 | kV | 43-0001 | 2572 | mR/s | 4.698 | mm | 3.227 | mm | 119.8 | kV
32 | 100 | Mo/Mo | L | 11 | 2022-04-25 15:57 | 1.0647 | s | 32.4 | kV | 1643 | mR | 1543 | mR/s | 0.33 | mm | 1 | USB | 1 | 6.8 | Hz | Standard | Multi-Sensor | Diagnostic | General | 3.227 | mm | 120.2 | kV | 119.8 | kV | 1 | 0.1574 | s | 120.2 | kV | 43-0001 | 2572 | mR/s | 4.698 | mm | 3.227 | mm | 119.8 | kV
28 | 25 | Mo/Mo | L | 12 | 2022-04-25 15:57 | 0.2489 | s | 28 | kV | 274.2 | mR | 1102 | mR/s | 0.29 | mm | 1 | USB | 1 | 5.2 | Hz | Standard | Multi-Sensor | Diagnostic | General | 3.227 | mm | 120.2 | kV | 119.8 | kV | 1 | 0.2007 | s | 120.2 | kV | 43-0001 | 2572 | mR/s | 4.698 | mm | 3.227 | mm | 119.8 | kV
28 | 50 | Mo/Mo | L | 13 | 2022-04-25 15:57 | 0.4987 | s | 28.1 | kV | 554.4 | mR | 1112 | mR/s | 0.3 | mm | 1 | USB | 1 | 6.7 | Hz | Standard | Multi-Sensor | Diagnostic | General | 3.227 | mm | 120.2 | kV | 119.8 | kV | 1 | 0.1589 | s | 120.2 | kV | 43-0001 | 2572 | mR/s | 4.698 | mm | 3.227 | mm | 119.8 | kV
28 | 200 | Mo/Mo | L | 14 | 2022-04-25 15:58 | 1.9986 | s | 28.1 | kV | 2239 | mR | 1120 | mR/s | 0.3 | mm | 1 | USB | 1 | 6 | Hz | Standard | Multi-Sensor | Diagnostic | General | 3.227 | mm | 120.2 | kV | 119.8 | kV | 1 | 0.1775 | s | 120.2 | kV | 43-0001 | 2572 | mR/s | 4.698 | mm | 3.227 | mm | 119.8 | kV
28 | 320 | Mo/Mo | L | 15 | 2022-04-25 15:58 | 3.1985 | s | 28.1 | kV | 3587 | mR | 1122 | mR/s | 0.3 | mm | 1 | USB | 1 | 6.5 | Hz | Standard | Multi-Sensor | Diagnostic | General | 3.227 | mm | 120.2 | kV | 119.8 | kV | 1 | 0.1642 | s | 120.2 | kV | 43-0001 | 2572 | mR/s | 4.698 | mm | 3.227 | mm | 119.8 | kV
28 | 100 | Mo/Rh | L | 16 | 2022-04-25 15:58 | 0.9988 | s | 28.1 | kV | 915.6 | mR | 916.7 | mR/s | 0.36 | mm | 1 | USB | 1 | 6.5 | Hz | Standard | Multi-Sensor | Diagnostic | General | 3.227 | mm | 120.2 | kV | 119.8 | kV | 1 | 0.1637 | s | 120.2 | kV | 43-0001 | 2572 | mR/s | 4.698 | mm | 3.227 | mm | 119.8 | kV
28 | 100 | Rh/Rh | L | 17 | 2022-04-25 15:58 | 1.6117 | s | 27.8 | kV | 850.4 | mR | 527.6 | mR/s | 0.37 | mm | 1 | USB | 1 | 6.9 | Hz | Standard | Multi-Sensor | Diagnostic | General | 3.227 | mm | 120.2 | kV | 119.8 | kV | 1 | 0.1553 | s | 120.2 | kV | 43-0001 | 2572 | mR/s | 4.698 | mm | 3.227 | mm | 119.8 | kV
28 | 100 | Rh/Al | L | 18 | 2022-04-25 15:58 | 0.1321 | s | 28.1 | kV | 338.8 | mR | 2572 | mR/s | 0.36 | mm | 1 | USB | 1 | 8.2 | Hz | Standard | Multi-Sensor | Diagnostic | General | 3.227 | mm | 120.2 | kV | 119.8 | kV | 1 | 0.1317 | s | 120.2 | kV | 43-0001 | 2572 | mR/s | 4.698 | mm | 3.227 | mm | 119.8 | kV
28 | 100 | W/Rh | L | 19 | 2022-04-25 15:58 | 0.1451 | s | 28.1 | kV | 372.6 | mR | 2572 | mR/s | 0.36 | mm | 1 | USB | 1 | 7.4 | Hz | Standard | Multi-Sensor | Diagnostic | General | 3.227 | mm | 120.2 | kV | 119.8 | kV | 1 | 0.1447 | s | 120.2 | kV | 43-0001 | 2572 | mR/s | 4.698 | mm | 3.227 | mm | 119.8 | kV
28 | 100 | W/Ag | L | 20 | 2022-04-25 15:58 | 0.1642 | s | 28.1 | kV | 421.8 | mR | 2572 | mR/s | 0.36 | mm | 1 | USB | 1 | 6.5 | Hz | Standard | Multi-Sensor | Diagnostic | General | 3.227 | mm | 120.2 | kV | 119.8 | kV | 1 | 0.1639 | s | 120.2 | kV | 43-0001 | 2572 | mR/s | 4.698 | mm | 3.227 | mm | 119.8 | kV
28 | 25 | Mo/Mo | S | 21 | 2022-04-25 15:58 | 0.2734 | s | 28.1 | kV | 702 | mR | 2572 | mR/s | 0.36 | mm | 1 | USB | 1 | 3.8 | Hz | Standard | Multi-Sensor | Diagnostic | General | 3.227 | mm | 120.2 | kV | 119.8 | kV | 1 | 0.2729 | s | 120.2 | kV | 43-0001 | 2572 | mR/s | 4.698 | mm | 3.227 | mm | 119.8 | kV
28 | 25 | Mo/Rh | S | 22 | 2022-04-25 15:58 | 0.167 | s | 28.1 | kV | 428.8 | mR | 2572 | mR/s | 0.36 | mm | 1 | USB | 1 | 6.4 | Hz | Standard | Multi-Sensor | Diagnostic | General | 3.227 | mm | 120.2 | kV | 119.8 | kV | 1 | 0.1667 | s | 120.2 | kV | 43-0001 | 2572 | mR/s | 4.698 | mm | 3.227 | mm | 119.8 | kV
28 | 25 | Rh/Rh | S | 23 | 2022-04-25 15:58 | 0.1522 | s | 28.1 | kV | 390.7 | mR | 2572 | mR/s | 0.36 | mm | 1 | USB | 1 | 7 | Hz | Standard | Multi-Sensor | Diagnostic | General | 3.227 | mm | 120.2 | kV | 119.8 | kV | 1 | 0.1519 | s | 120.2 | kV | 43-0001 | 2572 | mR/s | 4.698 | mm | 3.227 | mm | 119.8 | kV
28 | 25 | Rh/Al | S | 24 | 2022-04-25 15:58 | 0.1714 | s | 28.1 | kV | 440 | mR | 2572 | mR/s | 0.36 | mm | 1 | USB | 1 | 6.2 | Hz | Standard | Multi-Sensor | Diagnostic | General | 3.227 | mm | 120.2 | kV | 119.8 | kV | 1 | 0.171 | s | 120.2 | kV | 43-0001 | 2572 | mR/s | 4.698 | mm | 3.227 | mm | 119.8 | kV
28 | 25 | W/Rh | S | 25 | 2022-04-25 15:58 | 0.1569 | s | 28.1 | kV | 402.6 | mR | 2572 | mR/s | 0.36 | mm | 1 | USB | 1 | 6.8 | Hz | Standard | Multi-Sensor | Diagnostic | General | 3.227 | mm | 120.2 | kV | 119.8 | kV | 1 | 0.1565 | s | 120.2 | kV | 43-0001 | 2572 | mR/s | 4.698 | mm | 3.227 | mm | 119.8 | kV
28 | 25 | W/Ag | S | 26 | 2022-04-25 15:58 | 0.1598 | s | 28.1 | kV | 409.9 | mR | 2572 | mR/s | 0.36 | mm | 1 | USB | 1 | 6.7 | Hz | Standard | Multi-Sensor | Diagnostic | General | 3.227 | mm | 120.2 | kV | 119.8 | kV | 1 | 0.1592 | s | 120.2 | kV | 43-0001 | 2572 | mR/s | 4.698 | mm | 3.226 | mm | 119.8 | kV

## Data (hidden)
Location: | Outpatient Center
Registration ID: | 42
Registrant/Owner: | Community Hospital |  |  |  |  | Yellow cells are to be entered manually.
Facility Name: | Main Campus
Address: | 1000 Wormwood Drive |  |  |  |  | Blue cells are calculated automatically.
City, State, ZIP: | Somewhere, Else 00001
Contact: | Dr. Smith
Telephone: | (888) 555-1212
1 | 2022-04-26 11:39 | 0.07741 | s | 32.7 | kV | 290.3 | mR | 3765 | mR/s | 0.3538 | mm | 1 |  | USB | 1 | 14.8 | Hz | Standard | Multi-Sensor | Mo/Mo | Mammo | 0.01977 | mm | 32.7 | kV | 32.6 | kV | 1 | 0.077 | s | 32.7 | kV | 43-0001
Testing by: | John Doe
Test date: | 2011-03-09
Rapidose data file:
Unit: | Mammography |  |  |  | All regulatory criteria are typical values and should be edited to match local regulatory requirements.
Tube ID: | 1
System ID: | 24
Film-screen or digital? | d | (f/d)
GEN mfg
model
serial
date
TUBE mfg
model
serial
date
***********************************************************************************************************
RADIATION OUTPUT MEASUREMENT:
|  | System SID = | 66 | cm
|  | Detector to breast support surface distance = | 1.6 | cm
|  | Breast support surface to Rapidose distance = | 0.5 | cm
|  | Point above breast support surface to calculate dose = | 4.2 | cm
|  | Inverse square correction = | 1.1267
Set kVp | mAs | Target/Filter | Focal Spot | Measured mR | Measured Time (msec) | Measured kVp | kVp error | Corrected mR @ 4.2 cm | mR/mAs @ 4.2 cm | HVL (mm Al) | T/F code |  |  | kVp | mR/mAs @ 4.2 cm | ln(kVp) | ln(mR/mAs)
24 | 100 | Mo/Mo | L | 653.6 | 1070.9 | 23.8 | 0.8333% | 736.4119 | 7.3641 | 0.36 | 1 |  |  | 24 | 7.3641 | 3.1781 | 1.9966
25 | 100 | Mo/Mo | L | 764.9 | 999.4 | 24.9 | 0.4% | 861.8137 | 8.6181 | 0.37 | 1 |  |  | 25 | 8.6181 | 3.2189 | 2.1539
26 | 100 | Mo/Mo | L | 878.4 | 999.2 | 26 | 0% | 989.6943 | 9.8969 | 0.38 | 1 |  |  | 26 | 9.8969 | 3.2581 | 2.2922
27 | 100 | Mo/Mo | L | 995.3 | 999.3 | 27 | 0% | 1121.4056 | 11.2141 | 0.39 | 1 |  |  | 27 | 11.2141 | 3.2958 | 2.4172
28 | 100 | Mo/Mo | L | 1116 | 998.7 | 28.1 | 0.3571% | 1257.3985 | 12.574 | 0.4 | 1 |  |  | 28 | 12.5712 | 3.3322 | 2.5314
28 | 100 | Mo/Mo | L | 1116 | 998.8 | 28.1 | 0.3571% | 1257.3985 | 12.574 | 0.3 | 1 |  |  | 29 | 13.9711 | 3.3673 | 2.637
28 | 100 | Mo/Mo | L | 1115 | 998.9 | 28.1 | 0.3571% | 1256.2718 | 12.5627 | 0.3 | 1 |  |  | 30 | 15.4245 | 3.4012 | 2.736
28 | 100 | Mo/Mo | L | 1116 | 998.5 | 28.1 | 0.3571% | 1257.3985 | 12.574 | 0.3 | 1 |  |  | 32 | 18.5117 | 3.4657 | 2.9184
29 | 100 | Mo/Mo | L | 1240 | 997.6 | 29.3 | 1.034% | 1397.1094 | 13.9711 | 0.31 | 1
30 | 100 | Mo/Mo | L | 1369 | 998 | 30.4 | 1.333% | 1542.4539 | 15.4245 | 0.32 | 1 |  |  |  |  | b
32 | 100 | Mo/Mo | L | 1643 | 1064.7 | 32.4 | 1.25% | 1851.17 | 18.5117 | 0.33 | 1 |  |  |  |  | 3.1896 | -8.112
28 | 25 | Mo/Mo | L | 274.2 | 248.9 | 28 | 0% | 308.9415 | 12.3577 | 0.29 | 1 |  |  |  |  | 0.00029991
28 | 50 | Mo/Mo | L | 554.4 | 498.7 | 28.1 | 0.3571% | 624.6431 | 12.4929 | 0.3 | 1 |  |  |  |  | a
28 | 200 | Mo/Mo | L | 2239 | 1998.6 | 28.1 | 0.3571% | 2522.6839 | 12.6134 | 0.3 | 1
28 | 320 | Mo/Mo | L | 3587 | 3198.5 | 28.1 | 0.3571% | 4041.477 | 12.6296 | 0.3 | 1
28 | 100 | Mo/Rh | L | 915.6 | 998.8 | 28.1 | 0.3571% | 1031.6076 | 10.3161 | 0.36 | 2
28 | 100 | Rh/Rh | L | 850.4 | 1611.7 | 27.8 | 0.7143% | 958.1466 | 9.5815 | 0.37 | 3
28 | 100 | Rh/Al | L | 338.8 | 152.2 | 28.1 | 0.3571% | 381.7263 | 3.8173 | 0.36 | 4
28 | 100 | W/Rh | L | 372.6 | 171.4 | 28.1 | 0.3571% | 419.8088 | 4.1981 | 0.36 | 4
28 | 100 | W/Ag | L | 421.8 | 156.9 | 28.1 | 0.3571% | 475.2425 | 4.7524 | 0.36 | 4
28 | 25 | Mo/Mo | S | 702 | 159.8 | 28.1 | 0.3571% | 790.9442 | 31.6378 | 0.36 | 1
28 | 25 | Mo/Rh | S | 428.8 | 0 | 28.1 | 0.3571% | 483.1294 | 19.3252 | 0.36 | 2
28 | 25 | Rh/Rh | S | 390.7 | 0 | 28.1 | 0.3571% | 440.2021 | 17.6081 | 0.36 | 3
28 | 25 | Rh/Al | S | 440 | 0 | 28.1 | 0.3571% | 495.7485 | 19.8299 | 0.36 | 4
28 | 25 | W/Rh | S | 402.6 | 0 | 28.1 | 0.3571% | 453.6099 | 18.1444 | 0.36 | 4
28 | 25 | W/Ag | S | 409.9 | 0 | 28.1 | 0.3571% | 461.8348 | 18.4734 | 0.36 | 4
******************************
MA(S) LINEARITY
mA(s) station | mR/mAs | linearity
25 | 12.3577
50 | 12.4929 | 0.0054407
100 | 12.5712 | 0.0031242
200 | 12.6134 | 0.0016777
320 | 12.6296 | 0.00064162 |  |  | Tolerance = | 0.1 | linearity coefficient
| Pass | [5 empty] | Mark to include on Summary page
Comment: | Radiation output linearity between different stations meets standards.
*******************************
EXPOSURE REPRODUCIBILITY
Coefficient of Variation from above data:
CV = | 0.00044813 |  |  |  | Tolerance = | 5% | coefficient of variation
| Pass | [5 empty] | Mark to include on Summary page
Comment: | Exposure reproducibility is acceptable.
*******************************
kVp accuracy and HVL
kVp accuracy standard | 5%
Set kVp | Target/Filter | Measured kVp | kVp error | Measured HVL | HVL lower limit | HVL upper limit | kVp pass? | HVL pass? |  |  | HVL C-factor
24 | Mo/Mo | 23.8 | -0.8333% | 0.36 | 0.27 | 0.36 | TRUE | TRUE |  | Mo/Mo | 0.12
25 | Mo/Mo | 24.9 | -0.4% | 0.37 | 0.28 | 0.37 | TRUE | TRUE |  | Mo/Rh | 0.19
26 | Mo/Mo | 26 | 0% | 0.38 | 0.29 | 0.38 | TRUE | TRUE |  | Rh/Rh | 0.22
27 | Mo/Mo | 27 | 0% | 0.39 | 0.3 | 0.39 | TRUE | TRUE
28 | Mo/Mo | 28.1 | 0.3571% | 0.4 | 0.31 | 0.4 | TRUE | TRUE
28 | Mo/Mo | 28.1 | 0.3571% | 0.31 | 0.31 | 0.4 | TRUE | TRUE
28 | Mo/Mo | 28.1 | 0.3571% | 0.32 | 0.31 | 0.4 | TRUE | TRUE
28 | Mo/Mo | 28.1 | 0.3571% | 0.33 | 0.31 | 0.4 | TRUE | TRUE
29 | Mo/Rh | 29.3 | 1.034% | 0.36 | 0.32 | 0.41 | TRUE | TRUE
30 | Rh/Rh | 30.4 | 1.333% | 0.37 | 0.33 | 0.42 | TRUE | TRUE
32 | Rh/Al | 32.4 | 1.25% | 0.36 | 0.35 | 0.44 | TRUE | TRUE
28 | W/Rh | 28 | 0% | 0.36 | 0.31 | 0.4 | TRUE | TRUE
28 | W/Ag | 28.1 | 0.3571% | 0.36 | 0.31 | 0.4 | TRUE | TRUE
28 | Mo/Mo | 28.1 | 0.3571% | 0.36 | 0.31 | 0.4 | TRUE | TRUE
28 | Mo/Rh | 28.1 | 0.3571% | 0.36 | 0.31 | 0.4 | TRUE | TRUE
28 | Rh/Rh | 28.1 | 0.3571% | 0.36 | 0.31 | 0.47 | TRUE | TRUE
28 | Rh/Al | 27.8 | -0.7143% | 0.36 | 0.31 | 0.5 | TRUE | TRUE
28 | W/Rh | 28.1 | 0.3571% | 0.36 | 0.31 |  | TRUE | TRUE
28 | W/Ag | 28.1 | 0.3571% | 0.36 | 0.31 |  | TRUE | TRUE
kVp accuracy | Pass | [5 empty] | Mark to include on Summary page
Comment: | kVp accuracy meets applicable standards.
HVL | Pass | [5 empty] | Mark to include on Summary page
Comment: | X-ray beam filtration is adequate.
***********************************
RADIATION OUTPUT
kVp | mR/mAs @ 4.2 cm
24 | 7.3641
25 | 8.6181
26 | 9.8969
27 | 11.2141
28 | 12.5712
29 | 13.9711
30 | 15.4245
32 | 18.5117
Output acceptable? | Pass
[5 empty] | No standard |  | Mark to include on Summary page
Comment: | Radiation output is typical.
***********************************
RADIATION AND KV WAVEFORM
Capture the waveform image and paste it in the appropriate section of the Report.
Waveform acceptable? | Pass
[5 empty] | No standard |  | Mark to include on Summary page
Comment: | Waveform is typical.
*******************************
General Comments
Mark to include on Summary page
| Unit meets regulatory requirements. | [5 empty] | x
*******************************
Summary Comments
Unit meets regulatory requirements.
Unit meets regulatory requirements.
*****************************
Note on G10: Look for this symbol for additional comments. Hover your mouse pointer over it to see the comment.
Note on B39: mA values must be edited to meet the requirements of the unit being tested. Erase, do not delete, any lines not used. Keep the boxed area as the reproducibility evaluation exposures.
Note on E39: All generator values are brought over automatically from the "Rapidose import" page.

## Report
Mammographic Unit Performance Report
| Facility Identification |  |  |  | System Identification
| Outpatient Center |  |  |  | Mammography
| Registration ID 42 |  |  |  | 1
| Community Hospital |  |  |  | Unique ID |  |  | 24
| Main Campus |  |  |  | Control Manufacturer |  |  | 0
| 1000 Wormwood Drive |  |  |  | Control Model |  |  | 0
| Somewhere, Else 00001 |  |  |  | Control S/N |  |  | 0
| Dr. Smith |  |  |  | Control Date |  |  | 0
| (888) 555-1212 |  |  |  | Tube Manufacturer |  |  | 0
[5 empty] | Tube Model |  |  | 0
[5 empty] | Tube S/N |  |  | 0
[5 empty] | Tube Date |  |  | 0
Test Date | 40611
Tested by | John Doe
Signature
Summary of the results of this performance survey
(N/A = not applicable, N/M = not measured)
| Pass | Radiation output
| Pass | Beam quality
| Pass | Reproducibility of radiation exposures
| Pass | kVp accuracy
| Pass | mAs radiation output linearity
General Comments:
Unit meets regulatory requirements.
X-Ray Machine Performance
Pass | Radiation Output
|  |  |  | Measurement distance = | 0.5 above breast support surface
| Set kVp | mR/mAs @ 4.2 cm
| 24 | 7.3641
| 25 | 8.6181
| 26 | 9.8969
| 27 | 11.2141
| 28 | 12.5712
| 29 | 13.9711
| 30 | 15.4245
| 32 | 18.5117
| Comment: | Radiation output is typical.
Pass | kVp Accuracy
| Set kVp | Measured kVp | kVp error
| 24 | 23.8 | -0.8333% |  | 20 | 5%
| 25 | 24.9 | -0.4% |  | 40 | 5%
| 26 | 26 | 0% |  | 20 | -5%
| 27 | 27 | 0% |  | 40 | -5%
| 28 | 28.1 | 0.3571%
| 28 | 28.1 | 0.3571%
| 28 | 28.1 | 0.3571%
| 28 | 28.1 | 0.3571%
The kVp error limit is 5%.
| Comment: | kVp accuracy meets applicable standards.
Pass | Beam Quality
| Set kVp | Target / Filter | Measured HVL | HVL lower limit | HVL upper limit
| 24 | Mo/Mo | 0.36 | 0.27 | 0.36
| 25 | Mo/Mo | 0.37 | 0.28 | 0.37
| 26 | Mo/Mo | 0.38 | 0.29 | 0.38
| 27 | Mo/Mo | 0.39 | 0.3 | 0.39
| 28 | Mo/Mo | 0.4 | 0.31 | 0.4
| 28 | Mo/Mo | 0.31 | 0.31 | 0.4
| 28 | Mo/Mo | 0.32 | 0.31 | 0.4
| 28 | Mo/Mo | 0.33 | 0.31 | 0.4
| 29 | Mo/Rh | 0.36 | 0.32 | 0.41
| 30 | Rh/Rh | 0.37 | 0.33 | 0.42
| 32 | Rh/Al | 0.36 | 0.35 | 0.44
| 28 | W/Rh | 0.36 | 0.31 | 0.4
| 28 | W/Ag | 0.36 | 0.31 | 0.4
| 28 | Mo/Mo | 0.36 | 0.31 | 0.4
| 28 | Mo/Rh | 0.36 | 0.31 | 0.4
| 28 | Rh/Rh | 0.36 | 0.31 | 0.47
| 28 | Rh/Al | 0.36 | 0.31 | 0.5
| 28 | W/Rh | 0.36 | 0.31
| 28 | W/Ag | 0.36 | 0.31
Measured HVL must exceed the stated limit.
| Comment: | X-ray beam filtration is adequate.
Pass | Exposure Reproducibility
| Exposure | Measured mR
| #1 | 1116
| #2 | 1116
| #3 | 1115
| #4 | 1116
| Average | 1115.75
| Std. Dev | 0.5
| CV | 0.04481%
CV must be less than 5%.
| Comment: | Exposure reproducibility is acceptable.
Pass | mAs Linearity
| mAs | mR/mAs | Linearity Coeff.
| 25 | 12.3577
| 50 | 12.4929 | 0.0054407 |  |  | 0 | 0.1
| 100 | 12.5712 | 0.0031242 |  |  | 1000 | 0.1
| 200 | 12.6134 | 0.0016777
| 320 | 12.6296 | 0.00064162
Linearity coefficient must be less than 0.10.
| Comment: | Radiation output linearity between different stations meets standards.
Pass | Radiation and kV Waveforms
| Comment: | Waveform is typical.
Note on B133: Paste waveform in this area.